42022702 |  | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure, the jaws of the pk dissecting forceps instrument would not close. No patient harm, adverse outcome or injury was reported, and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering discovered a broken pitch cable and a broken wire at the yaw pulley exit, where a piece of material is also missing. Engineering also observed deep scratches on a large area around the main tube, where material has been removed. The scratches extend approximately 1. 750" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. Due to the damage observed on the main tube, it's possible that the wire and pitch cable breakages may have also been a result of instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1331514 | 1172294 | 1123585 | 2955842200801287 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-18 | 2008-09-19 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-08-26 | Yes | No | Reuse | 2008
42011005 | 2008-08-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the tips of the precise bipolar forceps instrument broke off and fell into the patient. The tips were retrieved, and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed only one broken grip, with the breakage located at approximately. 110" from its base. There appears to be no indication of high force applied to the broken grip or it's surrounding components, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the broken tip was successfully retrieved from the patient. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1331745 | 1172297 | 1123588 | 2955842200801283 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-09-19 | 2008-09-19 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that a connector on the bipolar maryland forceps instrument was loose. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found both pins to be properly installed and not loose. Engineering also found the distal end of the main tube to have a 4" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1331752 | 1172299 | 1123590 | 2955842200801276 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-19 | 2008-09-19 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-08-21 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the monopolar curved scissors instrument did not cut anymore. No patient harm, adverse outcome or injury was reported, and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted, and the instrument was not able to cleanly cut through a. 006" piece of latex. Engineering observed that the instrument blades are not damaged, however, the blade edges are slightly rounded at the tips, which negatively affects cutting performance. Engineering determined that decreased cable tension and/or belleville spring force may have also contributed to poor cutting performance. Engineering also observed deep scratches on the clevis end of the main tube with material removed. The damage is located approximately. 750" from the proximal clevis and spans an area of 1. 120" x. 250" long. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1331458 | 1172303 | 1123594 | 2955842200801277 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-19 | 2008-09-19 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument did not work properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to move intuitively with a full range of motion in all directions. The grips opened and closed properly, and passing both recognition and engagement testing. Engineering observed deep scratches on the main tube with material removed. The scratches are approximately 1. 35 inches from the proximal clevis to main tube extension. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1331528 | 1172308 | 1123599 | 2955842200801281 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-19 | 2008-09-19 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-08-29 | Other | Other | It was reported that during a da vinci s hysterectomy procedure the customer experienced difficulty opening and closing the jaws of an instrument installed on a patient side manipulator (psm) arm. With the assistance of an isi technical support engineer the customer re-seated the sterile adapter, re-draped the psm, and hit the fault over ride button, however, the engagement issue continued to occur. The customer placed an instrument of a different type on the psm and stated it appeared to be responding accurately. The customer also reported hearing a "clicking" noise. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering was unable to confirm the non intuitive instrument engagement issue reported by the customer. As a precautionary measure, the suspected patient side manipulator (psm) arm was replaced. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. Engineering also replaced the camera head and camera cable to address the customer reported "clicking" noise. The psm and camera head were returned to isi for failure analysis investigation. Engineering evaluation found the psm to have low cable tension on two axis' with one axis wrist drive shaft shifted to the left by approximately 40 degrees. The camera head showed indication of a physical drop causing a screw to loosen inside of the camera, thus generating the "clicking" noise experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1344940 | 1172778 | 1124049 | 2955842200801289 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-09-09 | 2008-09-19 | Yes | Yes | Health Professional | IS2000 A51 P4 | Yes | No | 2008-08-29 | Yes | No | Reuse | 2008
400180 | 2008-08-12 | Injury | HospitalizationOther | It was reported that during a da vinci surgical procedure, while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing and a burn on the patient's bowel. The procedure was converted to a mini laparatomy to repair the bowel. The patient was administered antibiotics and required additional hospitalization, due to the bowel injury. No additional patient harm, adverse outcome or injury was reported. Multiple attempts have been made to contact the attending physician present for this procedure to obtain additional information, however, no response has been received at this time. | The accessory and instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1188700 | 1161192 | 1112823 | 2955842200801265 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-09-12 | 2008-09-12 | Yes | Yes | Health Professional | 400180 | No | No | 2008-08-14 | Device Not Retu | Yes | Unkown | 2008
40020705 | 2008-08-19 | Malfunction | Other | It was reported that approximately 20 mins into a da vinci myomectomy procedure, while the surgeon was removing a fibroid, the tennaculum forceps instrument clevis broke and a piece fell into the pt. The broken piece was immediately retrieved and the procedure continued with a replacement tennaculum forceps instrument. While using the second tennaculum forceps instrument the clevis broke and a piece fell into the pt. The broken piece was retrieved and the surgical procedure continued with a third replacement tennaculum forceps instrument. The planned surgical procedure was successfully completed and not significantly lengthened. No pt harm, adverse outcome or injury was reported. | The instrument has been returned and evaluated. Per the customer reported complaint, engineering found the instrument is broken at the proximal clevis to main tube interface with the clevis dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces. Based on the engineering evaluation, a likely failure mode was overloading of the instrument at the tip, causing the breakage. Per the case notes. The broken instrument components were successfully retrieved from the pt. The following medwatch report listed below was also sent to the fda for the first tennaculum forceps instrument used during the surgical procedure. Mfr report #2955842-2008-01273. | TENNACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1330197 | 1170799 | 1122179 | 2955842200801274 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-09-12 | 2008-09-12 | Yes | Yes | Health Professional | 40020705 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42014204 |  | Malfunction |  | It was reported that shaft of the 5 mm monopolar cautery instrument was peeling. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube insulation is damaged in areas toward the distal end. The insulation is damaged in areas toward the distal end. The insulation has small cuts grouped in twos, 90 degrees apart, that expose the bare metal of the tube. The cuts have created a flap of insulation material roughly equivalent to the size of the exposed section. The damage to the tube insulation was most likely caused by instrument collisions and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Such as disconnection of the instrument tip. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1334070 | 1170800 | 1122170 | 2955842200801275 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-12 | 2008-09-12 | No | Yes | Health Professional | 42014204 | Yes | No | 2008-08-20 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument was not recognized by the da vinci s system. No additional info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to be successfully recognized by a da vinci s system. The pogo pins do not stick and are not contaminated. Engineering could not replicate the customer reported recognition issue. Engineering also observed the distal end of the main tube to have various scratch marks concentrated in a 1. 6" long section on one side of the tube. A few scratches are deep enough to have removed tube material. Based on the location and appearance, the damage was most likely caused by instrument collisions and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1330479 | 1170804 | 1122174 | 2955842200801267 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-12 | 2008-09-12 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-08-18 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the large needle driver instrument was not recognized by the system. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a test system and the instrument was intermittently recognized. It was determined that the recognition issue experienced by the customer was due to the programming error or a faulty rti board. Engineering also observed the distal end of the main tube has a couple of sections with material removed due to deep scratches. The scratches are under. 250" long and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions and/or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRVIER ISNTRUMENT | ENDOSCOPIC INSTRUMENT | 1330492 | 1170806 | 1122176 | 2955842200801269 | 1 | GCJ | Manufacturer | DistributorOther | Other | Initial | 2008-09-12 | 2008-09-12 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the monopolar curved scissors instrument was not recognized by the system. Additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a test system and the instrument was successfully recognized. Engineering observed distal end of the main tube had a couple of sections. , approximately 180 degrees apart, with light material removed. The damaged areas are 1. 1" and 1. 25" long and parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1330506 | 1170807 | 1122177 | 2955842200801270 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-09-12 | 2008-09-12 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument was not working properly. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the yaw pulley on one grip has the crimp pocket feature broken off, allowing the cable crimp to slip out. The grip motion in non-intuitive as a result. The yaw pulley breakage is most likely due to rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, suck as disconnection of the instrument tip. Engineering also found the distal end of the main tube has a 2. 25" long section with material removed all around the tube. The damage is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1334337 | 1170808 | 1122178 | 2955842200801266 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-12 | 2008-09-12 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-08-18 | Yes | No | Reuse | 2008
40020705 | 2008-08-19 | Malfunction | Other | It was reported that approx 20 mins into a da vinci myomectomy procedure, while the surgeon was removing a fibroid, the tennaculum forceps instrument clevis broke and a piece fell into the pt. The broken piece was immediately retrieved and the procedure continued with a replacement tennaculum forceps instrument. While using the second tennaculum forceps instrument, the clevis broke and a piece fell into the pt. The broken piece was retrieved and the surgical procedure continued with a third replacement tennaculum forceps instrument. The planned surgical procedure was successfully completed and not significantly lengthened. No pt harm, adverse outcome or injury was reported. | The instrument has been returned and evaluated. Per the customer reported complaint, engineering found the instrument is broken at the proximal clevis to main tube interface, with the clevis dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces. Engineering also observed that one grip close cable is broken near the proximal pulleys and proximal clevis cable hole causing the pulleys to spin freely. Other cables at the wrist are loose, but not damaged. Based on the engineering eval, a likely failure mode was overloading of the instrument at the tip and high grasping force, causing the breakage. Per the same notes, the broken instrument components were successfully retrieved from the pt. The following medwatch report listed below was also sent to the fda for the second tennaculum forceps instrument used during the surgical procedure. Mfr report: 2955842-2008-01274. | TENNACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1327934 | 1170809 | 1122179 | 2955842200801273 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-09-12 | 2008-09-12 | Yes | Yes | Health Professional | 40020705 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument did not work and nothing fell in the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable broken near the proximal pulleys and proximal clevis cable hole, allowing the pulley to spin freely. The proximal clevis cable hole exhibits wear on one edge and the other cables at the wrist are derailed, but undamaged. It was concluded that cable wear on the pulleys/clevis combined with high grasping force likely contributed to the cable breakage and may have caused the derailment of the other cables. Engineering also observed multiple deep scratches on the main tube with material removed. The scratches are 1. 25",. 5" and. 4" long and each located consecutively at. 5",. 750" and 1. 5" from the intersection of the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences. Such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1328032 | 1170810 | 1122180 | 2955842200801268 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-12 | 2008-09-12 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-08-18 | Yes | No | Reuse | 2008
IS1200 A43P1 | 2008-08-19 | Other | Other | It was reported that prior to starting a da vinci myomectomy procedure the customer experienced system error code #20008. With the assistance of an isi technical support engineer (tse), the site exercised the grip on the right master tool manipulator (mtm), however, the faults continued to occur. The tse instructed the site re-boot the system and during homing the error recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Results and conclusions - the investigation conducted by field service engineering concluded that system error code #20008 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #20008 system error code appeared when the da vinci safety system determined the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering evaluation found that encoder sensors had malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1342030 | 1172649 | 1123920 | 2955842200801271 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-09-12 | 2008-09-12 | Yes | Yes | Health Professional | IS1200 A43P1 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
IS2000 A51P4 |  | Other | Other | It was reported that while setting up the da vinci s surgical system for a prostatectomy procedure, the site experienced system error code #23023. With the assistance of an intuitive surgical technical support engineer, the site recycled power to the system but the system error code remained. The patient was under anesthesia when the site decided to abort the procedure. No patient harm was reported. | Conclusion - the investigation conducted by isi field service engineering concluded that system error code #23023 experienced by the customer was associated with a patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart, the provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. System error code #23023 appears when the davinci safety system determines that the power sense signal on the cannula is not as expected. Signaling that the cannula sensor signal cannot be trusted. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put davinci in a "recoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering discovered that the embedded serializer patient manipulator (espm) board, which monitors the status of logic signals from the external push button switches and the cannula id sensors, had malfunctioned and caused the cannula sensors to work incorrectly, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1344501 | 1172653 | 1123924 | 2955842200801272 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-09-12 | 2008-09-12 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument did not work properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension to be dislodged from the main tube. The entire tube extension wall is circumferentially broken at the base of the axial keys, resulting in the tube extension rotating 360 degrees. No pieces are missing. Engineering concluded that the damage may be due to excessive side loading. Engineering also observed the main tube to have a 2. 25" long section 1" below the midpoint with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1332512 | 1156504 | 1108653 | 2955842200801262 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-16 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering conducted manual performance tests, and found the instrument to fail a cut test, as latex gets pinched at the tips. The blade edges are rounded at the tips, negatively affecting cut performance. Engineering also observed the distal end of main tube to have a 2" long section with material removed on one side of the tube. The damage area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1332538 | 1156505 | 1108654 | 2955842200801221 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-19 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument was not working properly. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on a test system and the grips did not initially move. A popping sound was then heard coming from the instrument, before motion was translated to the grips. It was determined that the popping sound was likely the hypotubes which were stuck together and later became dislodged, which then allowed grip motion. Bio debris was found towards the distal end of the flush tube, suggesting that the hypotube may have also had dried bio debris, which caused the hypotube to stick. The instrument moved appropriately in all directions, once the grips became unstuck and free to move. Engineering also observed, that one side of the distal end of the instrument main tube, has a 2" long section with material removed. The damage appears to have been caused by a combination of instrument collisions and a cannula accessory. There are also various scratches concentrated in a. 5" long section which may have been caused by instrument collisions or rough handling. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1332548 | 1156506 | 1108655 | 2955842200801263 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-08-16 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument were not coagulating. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the conductor wires for damage as this may cause cautery issues, however, the wires were found to be intact and electrical continuity passed. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. The customer reported failure mode cannot be confirmed. Engineering also observed that one side of the distal end of the instrument main tube has a 2. 5" long section with light material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1297239 | 1156507 | 1108656 | 2955842200801259 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-08-14 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the "tip cover" of the precise bipolar forceps instrument was peeling. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering could not confirm the "tip cover peeling", as the instrument does not have a tip cover. Engineering observed the distal end of the main tube has a couple of sections, 90 degrees apart, with material removed. Sections are approx 2" long and. 108 to. 190" wide. Based on the location and appearance, the damage areas were most likely caused by a combination of a cannula accessory and instrument collisions, and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1332546 | 1156508 | 1108657 | 2955842200801257 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-08-12 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument is defective. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was able to open and close the grips, and move the wrist properly, when manually rotating the input discs. Engineering also observed that the distal end of the main tube has three adjacent deep scratch marks, on one side of the tube with material removed. The scratches are spaced close together, and are not aligned with the tube axis. Engineering concluded that the scratches are most likely from instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1332544 | 1156509 | 1108658 | 2955842200801260 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-08-15 | Yes | No | Reuse | 2008
40016906 | 2008-08-08 | Malfunction | Other | It was reported that during a da vinci s hysterectomy surgical procedure, the surgeon inserted the harmonic curved shears instrument into the robotic trocar and the inactive "blade" of the harmonic curved shears insert fell into the pt's abdomen. The "blade" was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument and insert accessory were not returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. A f/u mdr will be submitted if the instrument and/or insert accessory are returned (post engineering eval) or if add'l info is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1332542 | 1156510 | 1108659 | 2955842200801255 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-09-05 | 2008-09-05 | Yes | Yes | Health Professional | 40016906 | No | No | 2008-08-11 | Device Not Retu | Yes | Unkown | 2008
400180 | 2008-08-05 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing was observed. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No pt harm, adverse outcome or injury was reported. | The mcs tip cover accessory has not been returned for eval, therefore, reported failure could not be confirmed nor denied. A f/u mdr will be submitted, if the tip cover is returned for eval or if add'l info is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1332540 | 1156511 | 1108660 | 2955842200801256 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-09-05 | 2008-09-05 | Yes | Yes | Health Professional | 400180 | No | No | 2008-08-11 | Device Not Retu | Yes | Unkown | 2008
42017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a conductor wire broken at the yaw pulley exit. The yaw pulley shows no sign of arcing. Electrical continuity failed. Engineering also found the distal end of the main tube has a 1 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1297235 | 1156512 | 1108661 | 2955842200801264 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-05 | 2008-09-05 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-08-16 | Yes | No | Reuse | 2008
420214 | 2008-06-10 | Injury | Other | On august 15, 2008, isi rec'd medwatch uf/importer report with the following event description: pt undergoing robot-assisted laparoscopic-assisted vaginal hysterectomy. During the procedure, the pk dissecting forceps were placed through the trocar and was connected to the cautery. The pedal on the cautery was depressed several times, yet the forceps did not respond. When the forceps were removed from the trocar, a broken wire was detected. Post-op the pt developed a purple mark around the site of the trocar. Initially it was thought to be a bruise, after further examination, it was thought to be a burn. | The instrument used during the surgical procedure has not been returned for eval. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the pt side manipulator arms to ground through the cannula accessory and to the pt if the pt is not properly grounded, however, the root cause of the customer reported event cannot be definitively determined. A f/u mdr will be submitted if the instrument is returned (post engineering evaluation) or if add'l info is rec'd. As of september 4, 2008, there have been no reported recurrences of the issue at this hosp. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1173997 | 1156543 | 1083474 | 2955842200801261 | 1 | GEI | Manufacturer | Other | NOT APPLICABLE | Initial | 2008-09-05 | 2008-09-05 | Yes | Yes | Health Professional | 420214 | No | No | 2008-08-15 | Device Not Retu | No | Unkown | 2008
IS1200 A43P9 | 2008-08-13 | Other | Other | It was reported that during the beginning of a da vinci hysterectomy procedure, the connector on the right camera controller unit (ccu) broke off, resulting in the site being unable to connect to the camera. The patient was under anesthesia and port incisions had been created before the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome of injury was reported. | An investigation conducted by the field service engineer confirmed the connector on the right camera control unit (ccu) to be broken. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1294317 | 1156692 | 1108743 | 2955842200801258 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2008-09-05 | 2008-09-05 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-08-13 | Yes | No | Reuse | 2008
42000604 | 2008-07-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, black splinters from the large needle driver instrument were found inside of the patient. All fragments were removed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed there to be scratch marks at the distal end of the main tube with a small amount of material removed. The change in surface finish over the scratched area is under. 200" long and less than. 040" wide. Scratches may have been caused by instrument collisions and/or rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handling instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1290448 | 1148952 | 1101443 | 2955842200801246 | 1 | GCJ | Manufacturer | User facilityOther | Unknown | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-08-04 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument was splintered. This was not discovered during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the presence of various deep scratches concentrated in a 2. 5" long section on one side of the distal end of the main tube. The scratches have removed tube material. The scratches are not quite aligned with the tube axis. Their narrow width and close spacing suggests they may have been caused by instrument collisions or rough handling in central processing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1290447 | 1148953 | 1101444 | 2955842200801242 | 1 | GCJ | Manufacturer | User facilityOther | Unknown | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-08-02 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft on the monopolar curved scissors instrument was splintered. This was not discovered during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the presence of various deep scratches concentrated in a 1. 25" long section on one side of the distal end of the main tube. The scratches have removed tube material and are not quite aligned with the tube axis. Their narrow width and close spacing suggests they may have been caused by instrument collisions or rough handling in central processing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or another unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1276919 | 1148954 | 1101445 | 2955842200801243 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-02 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument is defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument moved intuitively with a full range of motion in all directions. Recognition and engagement passed. Engineering also observed that the distal end of the main tube has various scratch marks concentrated in a 1" long section on one side of the tube. A couple of the scratches are deep enough to have removed tube material. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1276941 | 1148955 | 1101446 | 2955842200801254 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42000605 | Yes | Yes | 2008-08-08 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the tube extension is dislodged from the main tube. The entire tube extension wall is circumstantially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees with no pieces missing. Engineering also observed that the distal end of the main tube has a 2. 25" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1276952 | 1148956 | 1101447 | 2955842200801245 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-08 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the monopolar curved scissors instrument was dull. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and the instrument was not able to cleanly cut through a. 006" of latex, with the latex getting pinched at the scissor tips. Engineering observed that the instrument blades are not damaged, however, the blade edges are slightly rounded at the tips which negatively affects cutting performance. Engineering determined that decreased cable tension and/or belleville spring force may have also contributed to poor cutting performance. Engineering also observed that one side of the distal end of the main tube has a 1. 75" long section with material removed. The damaged area is located 1" below the midpoint of the main tube, is parallel to the tube axis and has a rouge surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1276978 | 1148957 | 1101448 | 2955842200801247 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-05 | Yes | No | Reuse | 2008
400180 | 2008-07-28 | Malfunction |  | It was reported that approximately 1. 5 hours into a da vinci s hysterectomy procedure, burn holes were seen on the monopolar curved scissors (mcs) instrument tip cover while the surgeon was taking down the round ligament. The surgeon immediately removed the mcs instrument and tip cover. The planned surgical procedure was completed with a replacement mcs instrument and tip cover and without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was discarded by the customer and will not be returned for evaluation, however, the customer did return the mcs instrument used during the procedure. Per the customer reported complaint, engineering inspected the instrument under magnification and found no char marks on the metal wrist components. None of the components had any burrs or abnormally sharp edges that may have damaged the silicone portion of the tip cover. Due to the tip cover accessory not returning, the root cause of the customer reported failure mode cannot be determined or confirmed. However, engineering concluded that arcing was likely due to insufficient silicone dielectric strength with strength weakened due to high generator settings by the customer, which may have contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1288128 | 1148958 | 1101449 | 2955842200801253 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-07 | 2008-08-29 | No | Yes | Health Professional | 400180 | No | No | 2008-08-07 | Yes | Yes | Initial | 2008
400156 | 2008-07-31 | Malfunction |  | It was reported that halfway through a da vinci s hysterectomy procedure, while the surgeon was taking down the vaginal cuff, the cautery hook tip of the monopolar cautery instrument arced and then fell off inside of the patient. The surgeon immediately removed the cautery hook tip from the patient and replaced the monopolar cautery instrument. The planned surgical procedure was completed with a replacement instrument and hook tip and without significant delay. No patient harm, adverse outcome or injury was reported. | The cautery hook tip accessory was discarded by the customer and will not be returned for evaluation. However, the customer stated they will be returning the 5mm monopolar cautery instrument. The instrument has not been returned for evaluation at the time of this report, therefore, the root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | 5MM CAUTERY HOOK TIP ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1288152 | 1148959 | 1101450 | 2955842200801252 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 400156 | No | No | 2008-08-07 | Device Not Retu | Yes | Initial | 2008
40018005 | 2008-07-30 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, the surgeon noticed a spark from the tip of the msc instrument. The surgeon immediately removed the mcs instrument and replaced the tip cover accessory. The planned surgical procedure was completed with the replacement tip cover. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, however, the customer provided pictures of the accessory for further analysis. Per the customer reported complaint, engineering viewed a hole in the silicone, approximately. 080 inches above the silicone to sleeve interface. The hole appears to be. 025 to. 030 inches in diameter with localized melting all around the hole that is indicative or arcing. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1288135 | 1148960 | 1101451 | 2955842200801250 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-08-29 | 2008-08-29 | No | Yes | Health Professional | 40018005 | No | No | 2008-08-05 | Device Not Retu | Yes | Initial | 2008
400179 | 2008-08-05 | Malfunction | Other | It was reported that approximately one hour into a da vinci adrendectomy procedure, the surgeon noticed a spark from the tip of the monopolar curved scissors (mcs) instrument while dissecting the adrenal gland. The mcs instrument did have a tip cover accessory installed. The planned surgical procedure was completed without significant delay. After the procedure, the surgical staff observed a burn on the patient's skin surrounding the port incision where the cannula accessory and mcs instrument were inserted. No additional patient harm, adverse outcome or injury was reported. | The instrument used during the surgical procedure has not been returned for evaluation. Previous investigations have shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of 2008, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1288144 | 1148961 | 1101452 | 2955842200801249 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-08-29 | 2008-08-29 | Yes | Yes | Health Professional | 400179 | No | No | 2008-08-05 | Device Not Retu | No | Unkown | 2008
40018005 | 2008-07-30 | Other | Required Intervention | It was reported that upon removal of the monopolar curved scissor (mcs) instrument during a vinci s total hysterectomy procedure, the mcs tip cover fell off into the patient. The tip cover was retrieved and the planned procedure was completed without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover to be intact and not broken. Visual inspection under a microscope did not reveal any tearing of silicone or cracking of the sleeve. A fit test was performed using the mcs instrument used during this procedure and the tip cover fit snugly without coming off easily. The inner diameter of the sleeve is properly lined with silicone, which creates a tight fit with the mcs tube extension, even if both surfaces are wet. Engineering was unable to replicate the complaint of the tip cover falling off. No physical damage found. The mcs instrument was also returned. Engineering found no physical damage that may have contributed to the reported problem. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1284021 | 1149359 | 1101850 | 2955842200801235 | 1 | GEI | Manufacturer | Company Representative | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-07-30 | Yes | No | Unkown | 2008
IS1200 A43P9 | 2008-08-06 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure, the left eye in the high resolution stereo viewer (hrsv) was displaying "squiggly" lines as well as a loss of video from the camera head. With the assistance of an intuitive surgical representative, the site swapped vision cables, cycled power on the system and unplugged the system several times, however, the problem was not resolved. The patient was under anesthesia before the surgical staff made the decision to abort the procedure. No port incisions had been created. No patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer concluded the vision issues experienced by the customer was associated with the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected hrsv. The high resolution stereo viewer (hrsv) is located on the surgeon console and provides the video image for the surgeon console operator. The hrsv was returned to isi for evaluation. Engineering discovered that the left side monitor had malfunctioned, thus causing the reported vision issues. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1284060 | 1149360 | 1101855 | 2955842200801251 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-08-06 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-08-05 | Other | Other | It was reported that at the beginning of a da vinci prostatectomy procedure, the surgeon side console switched to backup battery power and the "overvault" led indicator was lit on the console power supply. With the assistance of an intuitive surgical technical support engineer, the site was unsuccessful in troubleshooting the problem. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that he customer reported complaint was related to the primary power supply (pps). The pps is a +48v power supply with battery backup and is mounted outside of the system control electronics chassis. The system was repaired by replacing the pps. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1283923 | 1149364 | 1101855 | 2955842200801248 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-29 | 2008-08-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-08-05 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument is rough and scratched. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a 1. 5 inch and 1. 75 inch long sections 180 degrees apart with material removed. The damaged areas are parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVE INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1328320 | 1133414 | 1086644 | 2955842200801232 | 1 | GCJ | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported tha during a da vinci s prostatectomy procedure, the wrist cable of the maryland bipolar forceps instrument was broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one conductor wire is broken at the yaw pulley exit and the yaw pulley shows no sign of arcing. Electrical continuity failed. Engineering also observed the distal end of the main tube has a 3" long section just below the midpoint, with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1279697 | 1133415 | 1086645 | 2955842200801224 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-31 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that the permanent cautery hook instrument was not working properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the pitch cable is broken at the distal clevis hub and most likely broke under tensile loading. Engineering also observed that there are various deep scratches with light material removed in a concentrated area on one side of the distal end of the main tube. Scratches are not aligned with tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1287797 | 1133416 | 1086646 | 2955842200801225 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-08-02 | Yes | No | Reuse | 2008
42019006 |  | Malfunction |  | It was reported that during a da vinci surgical procedure, the customer observed that there were "metal shavings" coming off of the cobra grasper instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering checked metal wrist components for damage and missing material. All of the components were intact with no metal removed. Engineering also observed that the distal end of the main tube has two narrow, approx. 5" long sections with light material removed. Sections are separated by approx. 035" and are parallel to the tube axis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1279719 | 1133417 | 1086647 | 2955842200801226 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42019006 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the precise bipolar cautery instrument is defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument moved intuitively with full range of motion in all directions. The system successfully recognized the instrument and the grips opened and closed properly. Engineering also observed the distal end of the main tube has a few deep scratches with material removed. There is a. 250" long scratch across the tube and a section 90 degrees apart with a few scratches that are parallel to the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1281458 | 1133418 | 1086648 | 2955842200801223 | 1 | NAY | Manufacturer | ForeignOther | Other | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-07-31 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that after a da vinci s total hysterectomy surgical procedure, the monopolar curved scissors instrument snapped, however, nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the instrument main tube, with the thin walled section of the tube extension broken circumferentially at the interface. No pieces are missing. Engineering also observed charring on the tube extension pieces, indicating that an arcing event occurred. Engineering also discovered a charred. 120" x. 060" oblong section on the distal end of the main tube, located. 5" above the tube extension. The charring indicates that a second arcing event occurred. The instrument was disassembled and a small char was noticed on the tube extension seal. Engineering also found a crack in the tube that starts at one slot and ends at the charred section. It was concluded that the tube extension breakage may have propagated the tube crack at the slot which likely created a path for arcing to occur and high generator settings may have also contributed to the damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1267849 | 1133425 | 1086655 | 2955842200801239 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-01 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s urological surgical procedure, the monopolar curved scissors instrument was dull. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and the instrument was not able to cleanly cut through a. 006" of latex, with the latex getting snagged at the scissor tips. Engineering observed that the instrument blades are not damaged, however, the blade edges are slightly rounded at the tips which negatively affects cutting performance. Engineering also observed light material removed all around the midpoint to the distal end of the instrument main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1266817 | 1133457 | 1086688 | 2955842200801238 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-01 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument is rough and scratched. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a 1. 8 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1266800 | 1133458 | 1086689 | 2955842200801228 | 1 | GEI | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
40018005 | 2008-07-10 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon saw a spark from the tip of the monopolar curved scissors (msc) instrument with a mcs tip cover accessory installed. The surgeon immediately removed the mcs instrument and replaced the tip cover accessory. The planned surgical procedure was completed with the replacement tip cover. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, however, the customer provided pictures of the accessory for further analysis. Per the customer reported complaint, engineering viewed two holes in the silicone, between the parting lines, approximately. 075 inches and. 175 inches above the silicone to sleeve interface. The holes appears to be. 020 to. 030 inches in diameter with localized melting all around the hole that is indicative of arcing. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1266785 | 1133459 | 1086690 | 2955842200801237 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 40018005 | No | No | 2008-07-31 | Device Not Retu | Yes | Initial | 2008
40016906 | 2008-07-30 | Malfunction | Other | It was reported that during a da vinci s ovarian cystectomy procedure, when the surgeon was taking out a specimen of cyst using the tip of the harmonic curved shears instrument, the "blade" came off and fell into the pt. The "blade" was retrieved and the planned procedure was completed. Incident lengthened procedure about 2 hours. No pt harm adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering found the curved blade is intact, however, the clamp arm is missing from the insert. One side of the clamp arm attachment hinge feature (inner and outer tubes) is bent, suggesting the clamp arm was overloaded. It had been concluded that the "blade" reference by the customer is actually the clamp arm of the harmonic curved shears insert. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" reference by the customer is actually the clamp arm of the harmonic curved shears insert. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1269143 | 1133557 | 1086788 | 2955842200801227 | 1 | LFL | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-08-22 | 2008-08-22 | Yes | Yes | Health Professional | 40016906 | Yes | No | 2008-07-30 | Yes | Yes | Initial | 2008
42017206 |  | Malfunction |  | It was reported that the shaft of the maryland bipolar forceps instrument is rough and scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has multiple sections, 180 degrees apart, with varying degrees of material removed. One area is 1. 7 inches long and parallel to the tube axis with light material removed. Based on the location and appearance, the damage was most likely caused by a cannula accessory. The other damaged area appears to be a combination of tube abrasions and scratch marks, most likely from instrument collisions. This section has a rougher surface finish with heavy material removed. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do no use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1267041 | 1133562 | 1086793 | 2955842200801229 | 1 | GEI | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the shaft of the endowrist prograsp instrument is rough and scratched. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the main tube has multiple sections, 180 degrees apart, with material removal from scrapes and scratches. The scrape marks are parallel to the tube axis and based on the location and appearance, the damage was most likely caused by a cannula accessory. The scratches are not aligned with the tube axis and were most likely from instrument collisions and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1266919 | 1133563 | 1086794 | 2955842200801230 | 1 | GCJ | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the shaft of the endowrist prograsp instrument is rough and scratched. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a scratched section with light material removed. The scratches are not aligned with the tube axis and were most likely from instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1266997 | 1133564 | 1086795 | 2955842200801231 | 1 | NAY | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument is rough and scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a 1 inch long section with material removed on one side of the tube. The marks are parallel to the tube axis and based on the location and appearance, the damage was most likely caused by a cannula accessory. Also observed were random scratch marks that are not aligned with the tube axis and were most likely from instrument collisions and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1267037 | 1133565 | 1086796 | 2955842200801233 | 1 | GCJ | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
42022702 | 2008-07-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon saw shavings from the shaft of the pk dissecting forceps instrument fall into the pt. The surgeon immediately removed the pk dissecting forceps instrument, irrigated the area, and removed all visible shaving from the pt's anatomy. The planned surgical procedure was completed with a different instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a couple of sections with material removed on one side. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1269297 | 1133566 | 1086797 | 2955842200801234 | 1 | GEI | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-07-28 | Yes | No | Reuse | 2008
420183 | 2008-07-29 | Malfunction |  | It was reported that during a da vinci s surgical procedure, at the start of the dissection, the permanent cautery hook instrument began to arc from the neck of the instrument. The surgeon immediately removed the permanent cautery hook instrument and replaced it with another instrument of the same type. The planned surgical procedure was completed with replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1269339 | 1133567 | 1086798 | 2955842200801236 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-08-22 | 2008-08-22 | No | Yes | Health Professional | 420183 | No | Yes | 2008-08-01 | Device Not Retu | No | Unkown | 2008
IS1200 A43P9 | 2008-05-14 | Other | Other |  |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1160067 | 1145703 | 1035156 | 2955842200801244 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-08-22 | 2008-08-22 | Yes | Yes | Invalid Data | IS1200 A43P9 | Device Returned To Manufacturer | No | 2008-05-14 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-08-01 | Other | Other | It was reported that during a da vinci s surgical prostatectomy procedure the customer experienced a non recoverable #25515 system error. An isi technical support engineer (tse) had the customer power cycle the system, however, they then received a #30253 system error. Customer then power cycled and emergency power off (epo'd) the patient side cart (psc) which powered up normally. The tse asked customer to call back at the end of procedure to review logs. The customer called back immediately with a reoccurrence of the #25515 system error code. The customer was then instructed to disable arm2 and continue procedure with arm1 and arm3. The customer called back stating the #30253 system error code is unrecoverable, the system would fault out as soon as customer attempted to override. The tse had them power the system down, epo, and reset the circuit breaker on both the psc and surgeon console. When the site powered back on the fault returned immediately. The surgeon decided to convert to an open surgical technique to complete the planned procedure. It was noted there was a 40 minute delay. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #25515 and #30253 experienced by the customer were associated with the remote arm controller (rac), the patient cart controller (pcc) and the low voltage differential signal (lvds) cable. The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The pcc is the board embedded in the patient cart that provides all user interface, communications control, the patient cart motor control for the patient side system. The lvds cable connects the rac to the pcc. The system was repaired by replacing the affected rac, pcc and lvds cable. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #25515 indicates a communication problem on the lvds cable between the rac and pcc. System error code #30253 is a sympathetic error which occurs when a fault is detected during power up. The pcc and rac were returned to isi for failure analysis investigation. Engineering found no issues with the rac, thus it was determined that a communication failure occurred due to a damaged connector the lvds cable and/or the pcc. As of august 19,(b) (6)hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1299875 | 1148468 | 1100959 | 2955842200801240 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-08-22 | 2008-08-22 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No | 2008-08-01 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-06-25 | Other | Other | It was reported that during a da vinci s total hysterectomy surgical procedure, the customer experienced nonrecoverable system error code #25580. The site decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that system error code #25580 experienced by the customer was associated with the remote arm controller (rac) board and the switcher power supply. The rac consists of 5 printed circuit assembly boards which operate together to provide control of the system arms. The switcher power supply is an electrical transformer which converts the ac voltage from the wall outlet to 48 v dc for use by the circuit boards and motors on the pt side cart. The system was repaired by replacing the affected rac and switcher power supply. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #25580 appears when the da vinci s safety determines a hardware parameter (like voltage or temperature) on the rac was exceeded a threshold of concern. Upon determining this condition, the safety system puts da vinci s in a "recoverable safe state". The rac and switcher power supply were returned to isi. In-house eval found no issues with the rac. The switcher power supply has been forwarded to the supplier for investigation. A f/u mdr will be submitted when the supplier investigation results are received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1290438 | 1148471 | 1100962 | 2955842200801214 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-22 | 2008-08-22 | Yes | No | Health Professional | IS2000 A5 1P4 | Yes | No | 2008-07-23 | Yes | No | Reuse | 2008
ISI200 A43P1 | 2008-08-01 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced nonrecoverable system error code #10025. The isi rep in attendance, at the procedure contacted an isi technical support engineer, who instructed the site to power cycle the system, however, the error continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering concluded, that system error code #10025 experienced by the customer, was associated with a pt side manipulator (psm) arm and/or a multiple servo driver (msd) pca board. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system contains five multiple servo driver (msd) boards, which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected psm and msd pca board. The system alarm (system generated fault code) functioned as designed, and there was no injury to the pt. The #10025 system error code appears when the da vinci safety system detects a failure to enable amps. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The system error code #10025 experienced by the customer, indicated that another fault occurred during the system start up, not allowing the amps to enable. The psm and msd pca were returned for eval, however, at this time the msd pca failure analysis has not been completed. No issues were found with psm, however, engineering found system error codes #212 and #21008 in the system error code logs. The psm affected potentiometer and motor were replaced as a precautionary measure. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1296947 | 1148473 | 1100964 | 2955842200801241 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-08-22 | 2008-08-22 | Yes | Yes | Health Professional | ISI200 A43P1 | Yes | No | 2008-08-01 | Yes | No | Reuse | 2008
ISI200 A43PL | 2008-07-22 | Other | Other | It was reported that during a da vinci surgical procedure, the customer experienced unexpected set up joint movement on the endoscopic camera manipulator (ecm) arm. An isi technical support engineer attempted to assist the customer, however, the surgeon decided to convert the procedure to traditional open surgical technique to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering confirmed that unexpected set up joint movement experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The field service engineer performed a clutch move of the ecm set up joint which cleared the system error message. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1264126 | 1123768 | 1077799 | 2955842200801213 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-08-15 | 2008-08-18 | Yes | Yes | Health Professional | ISI200 A43PL | Yes | Yes | 2008-07-22 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-07-21 | Other | Other | It was reported that during a da vinci s cardiac revascularization procedure, the customer experienced instrument engagement issues on a pt side manipulator (psm) arm. The customer tried multiple instruments and re-seated the sterile adapter prior to calling an isi technical support engineer (tse). The isi tse had the customer power off the system and hit fault over ride, however, the psm intermittently responded. The surgeon decided to convert the procedure to a small thoracotomy to complete the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering discovered system error code #23015 associated with a pt side manipulator (psm) arm. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #20015 system error code appeared when the da vinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state. " the psm was returned to isi for failure analysis investigation. Engineering eval found that a encoder had malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1264152 | 1123769 | 1077800 | 2955842200801222 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-08-15 | 2008-08-18 | Yes | Yes | Health Professional | IS2000 A51 P4 | Yes | No | 2008-07-21 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the tips of the precise bipolar pyramid tip instrument were crossed and no pieces fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side-to-side misalignment of the grips and a. 075" offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Engineering evaluation also observed that one side of the distal end of the main tube has a 2. 2" long section with material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1342081 | 1124391 | 1078402 | 2955842200801216 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-18 | 2008-08-18 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-07-25 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one of the two release levers is missing. A housing pin and two of the adjacent snap tabs are broken off, as well as a small piece of the housing next to the release lever slot. It was determined that the housing most likely popped off during mishandling which caused the housing, pin, and tabs to break. The breakage allowed the lever to fall out and it appears that the housing was then loosely put back on. Engineering evaluation also observed a deep scratch with material removed, located on the instrument main tube and approximately. 750" from the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1346193 | 1124393 | 1078404 | 2955842200801211 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-08-18 | 2008-08-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-22 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to pass both recognition and engagement testing with the grips opening and closing properly. The distal end of the main tube has a 1. 25" long section with material removed on one side of the tube. The damaged area is located 4. 5" above the proximal clevis, parallel to the tube axis with a rough surface finish. The tube also has a. 150" long deep scratch located 1. 7" above the proximal clevis with a small amount of material removed. Based on the location and appearance, the damage was most likely caused by a cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1342969 | 1119266 | 1073557 | 2955842200801215 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-07-24 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument appears scratched, but is smooth and is without splinters. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the instrument presents multiple small-deep scratches. The scratched area is within 2" from the interface of the main tube and the proximal clevis where material has been removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1342964 | 1119267 | 1073558 | 2955842200801218 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-07-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument appears scratched, but is smooth and is without splinters. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of the main tube has a 6. 75" long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1342944 | 1119268 | 1073559 | 2955842200801219 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-26 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument appears scratched, but is smooth and is without splinters. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument to have deep scratches on the main tube. Material has been removed in an area approximately. 725" by. 300". The damage is located approximately. 875" from the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. Electrical continuity passed. No ther damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1298885 | 1119333 | 1073624 | 2955842200801220 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-07-26 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument appears scratched, but is smooth and is without splinters. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the instrument presents multiple small-deep scratches. The scratched area is within 5" from the interface of the main tube and the proximal clevis where material has been removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1275480 | 1119334 | 1073625 | 2955842200801217 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-07-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to move intuitively with a full range of motion in all directions. The instrument passed recognition and engagement testing. The scissors do not cut cleanly through. 006" latex. The latex gets snagged at the scissor tips. The blades of the scissors are not damaged but the blade edges are slightly rounded at the tips, negatively affecting the cut performance. Engineering also observed the distal end of the main tube to have a 1. 75" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1275497 | 1119343 | 1073634 | 2955842200801212 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-22 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the tips of the maryland bipolar forceps instrument are not aligned. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the teeth do not properly mesh due to one grip being. 012" longer than the other. The side to side grip alignment is also slightly off by roughly. 010". The yaw pulley and conductor cap of the longer grip are separated at the glue joint. Engineering also observed the distal end of the main tube to have a. 250" long deep scratch with material removed. The scratch mark is not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1275767 | 1119363 | 1073654 | 2955842200801210 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-08-15 | 2008-08-15 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-22 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | It was reported that the cannula accessory was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering observed a visible dent at the distal tip and a. 342 gage pin does not fit through the distal tip inner diameter. The appearance of the dent leads engineering to conclude that the damage was likely caused by user mishandling. It was determined that the deformation of the tip has the potential to cause scraping of the instrument main tube in certain loading conditions. No other damage found. The cannula and obturator instructions for use (ifu) specifically states: cannula inspection instructions - warning: do not use a cannula with imperfections on its distal end. Examples of defects include a rough edge, dents, or an out-of-round shape. Use of a defective cannula may result in instrument damage, including scraping of the instrument shaft and particulate falling into the patient. Inspect the cannula prior to use for any defects. Warning: cannula defects can be caused by dropping the cannula or handling it roughly. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 1348243 | 1119744 | 1074035 | 2955842200801158 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 42000202 | Yes | No | 2008-05-16 | Yes | No | Unkown | 2008
42000604 |  | Malfunction |  | The large needle driver instrument was returned with no information provided. No patient harm, adverse outcome or injury was reported. | The instrument was evaluated. Engineering conducted performance testing and recognition and engagement had passed. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. Engineering observed dried bio debris surrounding 3 of 4 pogo pins which created some friction during actuation, however, recognition was not affected. Engineering also observed a 1. 5 inch long section with light material removal all around the distal end of the instrument main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1348279 | 1119745 | 1074036 | 2955842200801168 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-05-21 | Yes | No | Reuse | 2008
42020702 |  | Malfunction |  | It was reported that the tenaculum forceps instrument had a frayed fiber and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable frayed at the grip hub, however, the grips can still open and close. It appears that the cable may have been nicked by another instrument, causing the fraying. Engineering also observed that the distal end of the main tube has various deep scratch marks with material removed. The scratches are not aligned with the tube axis, suggesting that they were most likely caused by instrument collisions. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1348259 | 1119746 | 1074037 | 2955842200801167 | 1 | GCJ | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 42020702 | Yes | No | 2008-05-16 | Yes | No | Reuse | 2008
40003622 |  | Malfunction |  | It was reported that during a da vinci surgical procedure, shavings from the shaft of the debakey forceps instrument fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned to isi for evaluation, therefore, the root cause of the customer reported failure mode could not be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) and/or if additional information is received. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1261457 | 1119747 | 1074038 | 2955842200801171 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 40003622 | No | No | 2008-06-05 | Device Not Retu | No | Unkown | 2008
42022702 |  | Malfunction |  | It was reported that the cable of the pk dissecting forceps instrument is broken and pieces did not fall into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip open cable is frayed at the yaw pulley. The instrument grips can still open and close with no loss in functionality. The conductor wire has a small indentation in the insulation at approximately the same location on the yaw pulley as the fray. Engineering concluded that the cable and wire may have been damaged from an instrument collision or mishandling. Engineering evaluation also observed that one side of the distal end of the main tube has a few deep scratches with material removed. The scratches are not aligned to the tube axis, suggesting that they were most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions or warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1257951 | 1119748 | 1074039 | 2955842200801170 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-06-11 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering conducted a cut test and the scissors performed as intended. The instrument blades were not damaged and electrical continuity passed. The customer alleged failure mode could not be confirmed. Engineering also observed that one side of the distal end of the main tube has a 1. 75 inch long section with material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1257978 | 1119763 | 1074054 | 2955842200801169 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-08-14 | 2008-08-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-11 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-05-09 | Other | Other | It was reported that during a da vinci s partial nephrectomy surgical procedure, the customer experienced nonrecoverable system error code #212. An isi technical support engineer attempted to troubleshoot the issue via telephone, however, the issue could not be resolved and the planned surgical procedure was aborted and rescheduled for a later date. No pt harm was reported. | The investigation conducted by field service engineering concluded that system error code #212 experienced by the customer was associated with the left master tool manipulator (mtml) and/or a multiple servo driver (msd) pca board. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to their right (mtmr). The system contains five multiple servo driver (msd) boards which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected mtml and msd pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #212 system error code appears when the davinci s safety system detects an error in the expected motor current and voltage values. Upon determining this condition, the safety systems put davinci s in a "recoverable safe state". The mtml and msd pca were returned for eval. No issues were found with the msd pca, however, engineering was able to replicate system error code #212 during testing of the mtml. Further eval of the mtml revealed that an electrical cable/harness had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrence of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1297405 | 1120674 | 1074948 | 2955842200801166 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-08-14 | 2008-08-14 | Yes | Yes | Health Professional | IS2000 A5 1P4 | Yes | Yes | 2008-05-09 | Yes | No | Reuse | 2008
| 2008-07-30 | Other |  | Surgical procedure: robotic hysterectomy. During surgical procedure, while using monopolar scissors with tip cover in place and settings at 38w, instrument had unintended arcing (spark). No pt injury or harm resulted. Tip cover removed and cover inspected for holes. Photos taken of tip cover and sent to mfr, intuitive. |  | DAVINCI SURGICAL SYSTEM INTUITIVE | MONOPOLARCURVED SCISSORSTIP CO | 2633717 | 2602389 | 2499076 | 2602389 | 1 | NEY | User Facility | Unknown | RISK MANAGER | Initial | 2008-08-06 | 2008-08-13 | No | No | Service Personnel |  | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was not recognized by the system. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, recognition testing was performed and the instrument was successfully recognized by the system. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. The customer reported failure mode could not be duplicated. Engineering also observed that one side of the distal end of the main tube has a couple of sections, 180 degrees apart, with light material removed. The damaged areas are 1" and 1. 5" long, both parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. See scanned page. | CMARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1354441 | 1107057 | 1063208 | 2955842200801206 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-11 | 2008-08-11 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-14 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument was not recognized by the system. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, recognition testing was performed. The instrument was successfully recognized by the system and moved intuitively with full range of motion in all directions. The customer reported failure mode could not be duplicated. Engineering also observed that the distal end of the main tube has a couple of deep scratches with material removed. The scratches are not aligned with the tube axis, suggesting that they may have been caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1354544 | 1107058 | 1063209 | 2955842200801209 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-11 | 2008-08-11 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-07-17 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-07-17 | Injury | Required InterventionLife Threatening | This report is being filed based an isi investigation concluding 08/08/2008. (b)(4) was received on 08/01/2008 from the (b)(6). It was reported that during da vinci s bilateral internal mammary arteries revascularization procedure, the customer experienced a system error code #23. With the assistance of an isi (b)(4), the site powered down the system to clear the fault. The site continued with the procedure, however, the system error reoccurred. The site disabled the endoscopic camera manipulator (ecm) to continue the case. The site then elected to manually manipulate the camera and endoscope for approximately 5 to 6 hours when a loss of carbon dioxide insufflation occurred resulting in the heart pushing up into the endoscope two times causing lacerations to the patient's right ventricle. The procedure was converted to a thoracotomy to performed a non robotic repair of the damaged ventricle with several stitches and to complete the planned procedure. After the 14 hour procedure, the patient could not be extubated necessitating a tracheostomy. As of (b)(6), the patient remained under care at (b)(6). | The investigation conducted by (b)(4) concluded that the system error code #23 experienced by the customer was associated with a configured embedded sterilizer setup joint (cfg, essj) and remote arm controller (rac). The embedded sterilizer for setup-joint is the printed circuit assembly (pca) inside a system arm that monitors the potentiometer for each of four joints and their associated backup potentiometers. The rac consists of five printed circuit assembly boards which operate together to provide control of the system arms. The system was repaired by replacing the affected cfg, essj and rac. System error code #23 is reported by software to denote that the hardware wheel "wdog" has tripped on one of the digital communication links in the system. This means that the system cannot reliably communicate over the digital link and therefore cannot continue normal operation. Communication faults occasionally occur due to typically either faulty electronics or poor connections in the communication link. The error 23 fault indicated by the system in this case pointed to a communication error involving the ecm's rac and essj modules. The system was repaired as the (b)(4) both removed the electronics that could have experienced an intermittent failure and re-secured all of the connections involved. Field experience has shown that these measures are effective in resolving this type of system communication issue. As of (b)(6) 2007, the site has continued to use the system and has not reported recurrences of this issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2264513 | 2240665 | 2137581 | 2955842200801205 | 1 | NAY | Manufacturer | User facilityCompany Representative | Other | Initial | 2008-08-11 | 2008-08-11 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-07-17 | No | No | Unkown | 2008
42017908 |  | Malfunction |  | It was reported that a cable on the monopolar curved scissors instrument was found to be broken. This was not discovered during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable to be broken at the distal idlers resulting in the pulley spinning freely. The cable broke under tensile loading. The other cables at the wrist are not damaged. Engineering concluded that the breakage is likely due to the cable wear on the pulleys combined with the high drive torques. Engineering also observed the distal end of the main tube to have various deep scratches concentrated on one side of the tube with material removed. Scratch marks are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345064 | 1107029 | 1063180 | 2955842200801204 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-17 | Yes | No | Reuse | 2008
42017204 |  | Malfunction |  | It was reported that a plug on the maryland bipolar forceps instrument had come out. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed the plug riser insert to be missing from the back end of the instrument. The pins and conductor wire are sticking out of the chassis. The insert may have not been fully seated in the chassis or may be pulled out when the bipolar cord was removed. Some customer supplied bipolar cords are a tight fit on the bipolar pins and may require a high removal force, potentially causing the insert to pull out. Engineering also observed the distal end of the main tube to have a 1. 75" long section with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345104 | 1107030 | 1063181 | 2955842200801203 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 42017204 | Yes | No | 2008-07-14 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument has a loose piece. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer complaint, engineering observed that the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Engineering believes the damage is most likely due to excessive side loading. Engineering also observed the distal end of the main tube has a 3" long section with material removed on one side of the tube. The damage area is parallel to the tube axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345100 | 1107031 | 1063182 | 2955842200801200 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-14 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the jaws of the monopolar curved scissors instrument are not working. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer complaint, engineering observed that the scissor blades open and close when the input discs are manually rotated. The tips stick slightly during opening of the blades. Engineering also observed the distal end of the main tube has a 3" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely cause by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345075 | 1107032 | 1063183 | 2955842200801199 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-14 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument "tip" is broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer complaint, engineering observed the plug riser snap tabs to be broken, enabling the instrument housing to be easily removed. Engineering also observed the distal end of the main tube to have various deep scratches concentrated on one side of the tube with material removed. Scratch marks are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345060 | 1107033 | 1063184 | 2955842200801201 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-07-19 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument stopped working and has a loose wire. No additional information was provided. No patient harm, adverse outcome for injury was reported. | The instrument was returned and evaluated. Per the customer complaint, engineering observed the plug riser snap tabs to be broken, enabling the instrument housing to be easily removed. Engineering also observed the distal end of the main tube to have various deep scratched concentrated on one side of the tube with material removed. Scratch marks are not aligned with the tube axis. Based on the location and appearance, the damage was most likely cause by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1243057 | 1107034 | 1063185 | 2955842200801198 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-07-14 | Yes | No | Reuse | 2008
40018005 | 2008-07-09 | Malfunction |  | It was reported that approximately 2 hours into a da vinci s hysterectomy procedure, the surgeon noticed holes in the monopolar curved scissors (msc) instrument tip cover accessory. The surgeon immediately removed the mcs instrument and replaced the tip cover accessory. The planned surgical procedure was completed with the replacement tip cover. No patient harm, adverse outcome or injury reported. | The tip cover accessory has not been returned for evaluation, however, the customer provided pictures of the accessory for further analysis. Per the customer reported complaint, engineering viewed a hole in the silicone, approximately. 100 to. 120 inches above the silicone to sleeve interface. The hole appears to be. 020 to. 030 inches in diameter with localized melting all around the hole that is indicative of arcing. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1345080 | 1107035 | 1063186 | 2955842200801194 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-08-08 | 2008-08-08 | No | Yes | Health Professional | 40018005 | No | No | 2008-07-10 | Device Not Retu | Yes | Initial | 2008
420174 | 2008-07-14 | Malfunction | Other | It was reported that during a da vinci s nissen fundoplication procedure, a "blade" of the harmonic curved shears instrument broke and fell into the patient. The "blade" was retrieved and the planned procedure was completed without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. There are no components in the harmonic curved shears instrument or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. | HARMONIC CURVED SHEARS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1345043 | 1107036 | 1063187 | 2955842200801202 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-08-08 | 2008-08-08 | Yes | Yes | Health Professional | 420174 | No | No | 2008-07-14 | Device Not Retu | No | Unkown | 2008
IS2000 A51P4 | 2008-04-29 | Injury | Required InterventionHospitalization | On 7/11/08, isi received medwatch report, which indicates that an adverse event occurred in '08, during the use of the da vinci s surgical system and a gynecare x-tract tissue morcellator in a surgical procedure. The report was submitted to the fda by the pt, however, a description of the event was not provided. | On 7/30/08, additional info regarding the event was provided by the pt's sibling. It was reported that at the end of the da vinci s hysterectomy surgical procedure, the surgical staff inserted a gynecare x-tract morcellator into the pt, however, it was not securely attached. When the surgeon tried to attach the morcellator intraabdominally, the surgeon lost control of the morcellator and injured the pt's aorta, inferior vena cava and common iliac artery (cia). As a result, the pt needed a vein graft for the cia repair. On 8/7/08, additional info was provided by an isi rep who reported that during the time of the event, a system arm was undocked in order for the surgical staff to insert the morcellator through that port site and only the system camera was being used when the injury occurred. Based on the info provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the pt injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1029947 | 1118753 | 1017587 | 2955842200801195 | 1 | NAY | Manufacturer | Other | Invalid Data | Initial | 2008-08-08 | 2008-08-08 | Yes | No | Health Professional | IS2000 A51P4 | No | No Answer Provided | 2008-07-11 | No | No | Reuse | 2008
| 2008-07-10 | Malfunction |  | Surgical procedure: robot-assisted total laparoscopic hysterectomy, bilateral salpingo-oophorectomy, removal of 6x5 cm right ovarian mass. During surgical procedure, spark noted. Tip cover changed. No injury to pt occurred. Tip cover preserved for later inspection. Photos taken of tip cover and sent to intuitive. Tip cover inspected for holes. |  | DAVINCI SURGICAL SYSTEM INTUITIVE | MONOPOLARCURVED SCISSORSTIP CO | 1209558 | 1172804 | 1124094 | 1172804 | 1 | NEY | User Facility |  | RISK MANAGER | Initial | 2008-07-17 | 2008-08-07 | No | Yes | Health Professional |  | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2008
| 2008-07-09 | Malfunction |  | Surgical procedure: total laparoscopic hysterectomy, bilateral salpingo-oophorectomy and washings via robotic assistance. Bilateral pelvic lymphadenectomy and aortic lymph node sampling via robotic assistance. Reduction and repair of umbilical hernia. Approximately almost 2 hours into case, surgeon noted hole in tip cover and immediately removed tip and replaced with new tip cover. No spark or injury to pt occurred. Tip cover preserved for later inspection. Photos taken of tip cover and sent to intuitive. |  | DAVINCI SURGICAL SYSTEM INTUITIVE | TIP COVER FOR MONOPOLAR CURVED | 1209540 | 1172823 | 1124094 | 1172823 | 1 | NEY | User Facility |  | Other | Initial | 2008-07-17 | 2008-08-07 | No | Yes | Service Personnel |  |  | Yes |  | No Answer Provi |  | Invalid Data | 2008
42017206 |  | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure, the system was not recognizing the maryland bipolar forceps instrument, and nothing fell into patient. The planned surgical procedure was completed with a replacement instrument of the same type. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering conducted performance tests and found that the system intermittently failed to recognize the instrument. There were two pogo pins that were roughly. 007" shorter than the other pins. Repeated mechanical actuation of the short pins caused them to rise, therefore, intermittently engaging the instrument. Engineering believes the cleaning process may have affected the height of pins and cause a slight sticking. Engineering also observed that the main tube has two sections, 180 degrees apart, with material removed on one side of the tube. One section is 1. 75" long and located just below the midpoint of the tube. The other section is 1. 5" long and located 2" above the proximal clevis. The damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damages were most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1354917 | 1099833 | 1056704 | 2955842200801186 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42020704 | 2008-07-01 | Malfunction | Other | It was reported that during a da vinci s surgical procedure, a wire on the tenaculum forceps instrument broke and fell inside of the patient. The wire was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument has a broken grip close cable just past the distal idler pulley, allowing the pulley to spin freely. The cable has different fracture strand lengths, however, the cable does not appear to be missing any pieces. Engineering concluded that the damage was likely caused by high grasping force and wear against the pulley causing the breakage. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1355126 | 1099834 | 1056705 | 2955842200801182 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-07-09 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tips of the monopolar curved scissors instrument were burnt. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the tube extension has a. 250 x. 110" charred section directly below the flush hole on the proximal clevis. The damage is indicative that an arcing event occurred. The charring is located below the silicone portion of the tip cover accessory, assuming that the tip cover accessory was installed (a tip cover accessory was not returned with the instrument). The tube extension also has an indentation and crack at one of the keys that mates with the proximal clevis. It was determined that the arcing event may have been caused by insulation failure in the ultem tip cover sleeve, which likely created a path for arcing to occur. High generator settings may have also contributed to the arcing. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355066 | 1099835 | 1056706 | 2955842200801191 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the fiberglass towards the end of the monopolar curved scissors instrument is frayed. The instrument damage was not identified prior or during a surgical procedure. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one side of the distal end of the main tube has a 1. 5 inch long section with material removed. The damaged area is located directly above the tube extension, parallel to the tube axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1342371 | 1099836 | 1056707 | 2955842200801192 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-08 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was defective. No additional information provided. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Engineering observed the scissor blade edges are slightly rounded at the tips, negatively affecting cutting performance. The scissors do not cut cleanly through. 006 inch latex as the latex gets pinched at the scissor tips. Engineering also found that the distal end of the main tube has a 2. 5 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355123 | 1099843 | 1056714 | 2955842200801181 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-12 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the system was not recognizing the permanent cautery spatula instrument, and no pieces fell into patient. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Engineering conducted performance test and found they could not replicate the customer reported recognition issue. One ear of the distal clevis was broken off and the conductor cap was missing. The size of the missing ear piece is roughly. 280" by. 215". The ceramic sleeve is cracked, but does not have any missing pieces. The location and appearance of the damage indicates that it was most likely due to mishandling or misuse. Engineering also observed that the distal end of the main tube has a few deep scratches with material removed. The location and appearance of the damage indicates that it was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355062 | 1099844 | 1056715 | 2955842200801188 | 1 | GEI | Manufacturer | DistributorForeign | Other | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure, the system was not recognizing the maryland bipolar forceps instrument, and nothing fell into patient. The planned surgical procedure was completed with a replacement instrument of the same type. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering conducted performance tests and found that they could not replicate the customer reported recognition issue. Engineering also observed that the distal end of the main tube has a 1. 5" long section with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355118 | 1099845 | 1056716 | 2955842200801187 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument was not recognized on the system. The procedure was successfully completed using another instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to successfully be recognized by a system. The pogo pins on the instrument were not found to stick or to be contaminated. Engineering was unable to replicate the recognition issue. Engineering also observed the distal end of the instrument main tube to have a 1. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1355105 | 1099846 | 1056717 | 2955842200801179 | 1 | GCJ | Manufacturer | DistributorOther | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that the permanent cautery hook instrument was not working. No pieces were reported as falling into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the ceramic sleeve to be missing a. 085" x. 120" piece on the large diameter section which is no longer glued to the yaw pulley. The sleeve can move up and down the hook shank as a result. Engineering concluded that the damaged ceramic sleeve may be due to mishandling and misuse. Engineering also found the main tube to have a 3. 5" long section just below the midpoint with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355081 | 1099847 | 1056718 | 2955842200801197 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-07-12 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar forceps instrument was splintering. This was not noticed intraoperatively. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering confirmed that the distal end of the main tube has a 2" long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355099 | 1099848 | 1056719 | 2955842200801190 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that the permanent cautery hook instrument was not working. No pieces were reported as falling into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the cautery hook to be broken off at the base of the yaw pulley. The hook and ceramic sleeve were not returned. A yaw pulley section beneath the ceramic sleeve was also found to have broken off. The hook fractured through one of the holes machined in the shank. Engineering concluded that the damage may be due to the excessive forces and user abuse. Engineering also found the main tube to have a 1. 5" long section located 1. 5" below the midpoint with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355088 | 1099849 | 1056720 | 2955842200801196 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-07-12 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that during cleaning the fenestrated bipolar forceps, instrument was found to have a splintered shaft. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section of the main tube with light material removed. The damaged area is 1. 5 inches long and parallel to the tube axis, with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355132 | 1099850 | 1056721 | 2955842200801180 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-08-06 | 2008-08-06 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
IS1200 A4 3P9 | 2008-07-08 | Other | Other | It was reported that during a da vinci hysterectomy surgical procedure, the customer experienced multiple occurences of system error code #20013. The procedure was in process for 2. 5 hours when the site converted to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering found system error code #20013 that was experienced by the customer as well as system error code #21013. The field service engineer concluded that these system error codes were associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #20013 and #21013 system error code appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety system put da vinci in a "recoverable safe state". The mtml was returned to isi for failure analysis investigation. Engineering evaluation found that a potentiometer had malfunctioned, thus generating the system errors experienced by the customer. The mtml's motor assembly was also replaced as an additional precaution. As of 2008, there have been no reported recurrences of the issue at this hospital. | DE VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1245922 | 1099513 | 1056464 | 2955842200801193 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-08-01 | 2008-08-04 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2008-07-08 | Yes | No | Reuse | 2008
ISI200 A43 P1 | 2008-07-11 | Other | Other | It was reported that prior to starting a da vinci hysterectomy procedure, the robotic coordinator was unable to white balance the right camera controller unit (ccu). With the assistance of an isi technical support engineer (tse) the site verified the ccu settings and replaced the camera cable. The site continued with preparation for the procedure. The site called back stating that the right eye appeared washed out and white. The isi tse had the site redo the black and white balance as well as check, adjust and change additional ccu settings. The endoscope was changed and different inputs on the camera head were enabled, however, the right eye color continued to appear unbalanced. The surgeon decided to convert to traditional open surgical techniques to finish the planned surgery after approximately 90 minutes of troubleshooting. No patient harm was reported. | The investigation conducted by isi field service engineering (fse) discovered that the issue experienced by the customer was associated with a defective camera controller unit (ccu). The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1346783 | 1099516 | 1056467 | 2955842200801185 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-08-01 | 2008-08-04 | Yes | Yes | Health Professional | ISI200 A43 P1 | No | No | 2008-07-11 | Yes | No | Reuse | 2008
IS1200 A4 3P9 | 2008-07-07 | Other | Other | It was reported that during a da vinci mitral valve repair procedure, one of the pt side manipulator. (psm) arms was not responsive and the surgical staff could not get the arm into following mode. The pt was under anesthesia for approx five hours, with a 30 minute delay due to the system problem, when the procedure was converted to traditional open surgical techniques. The planned surgery was completed and no pt harm was reported. | The investigation conducted by the field service engineer concluded that a pt side manipulator (psm) was having intermittent issues. The system was repaired by replacing the psm and ria/lia board. The pt side manipulator is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. Engineering conducted tests for the psm and confirmed the customer reported problem. The psm was repaired by replacing the motor/encoder on axis-6 and all 4 potentiometers. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1354239 | 1099519 | 1056470 | 2955842200801184 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-08-01 | 2008-08-04 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2008-07-07 | Yes | No | Reuse | 2008
ISI200 A4 3P9 | 2008-07-07 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced recurring system error #20008. With the assistance of an isi representative the customer exercised the affected patient side manipulator (psm) arm several times, changed instruments, and sterile adapter, however, the system error #20008 recurred and could not be overridden. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. A delay of approximately 15 minutes was noted. No patient harm was reported. | The investigation conducted by field service engineering could not reproduce the faults experienced by the customer. The system was repaired by replacing the suspected patient side manipulator (psm), multiple servo driver pca board, the psm cable. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #20008 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder), and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state. " engineering concluded the psm arm generated the system error when attempting to run internal system sensor checks. Engineering concluded tests for the psm, cable, and pca board, and could not confirm the customer reported problem. The psm was repaired by replacing the motor/encoder on axis-6 as a precautionary measure. As of july 31, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1346741 | 1099525 | 1056476 | 2955842200801183 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-08-01 | 2008-08-04 | Yes | Yes | Health Professional | ISI200 A4 3P9 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-07-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a non recoverable system error code #23013. With the assistance fo an isi technical support engineer, the customer cycled the system power. The system powered back up and homed with no issues, therefore, the customer continued with the case. The customer called back stating the system error code #23013 recurred and the system could not recover. The error continued to occur with any attempt to move the endoscopic camera manipulator (ecm) arm. The patient was under anesthesia for approximately 60 minutes when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #23013 experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #23013 system error code appears when the da vinci s safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, the axis 1 and axis 2 encoder and motor assembly and potentiometer were replaced. As of 2008, ther have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1352574 | 1099530 | 1056481 | 2955842200801189 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-08-04 | 2008-08-04 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-07-07 | Device Not Retu | No | Reuse | 2008
IS2000 A51P4 | 2008-04-01 | Injury | Required Intervention | In 2008, it was reported that the site experienced a system fault while setting up for a da vinci s prostatectomy surgical procedure. No pt issues were reported at that time. On 7/3/08, additional info was received from the attending surgeon who indicated that continuous system faults occurred after the pt ports were placed, resulting in a 1 hour delay. During the delay, the pt's abdomen became deflated and the pt bucked vertically. With the assistance of an isi technical support engineer via telephone, the system faults were resolved by cleaning the fiberoptic cable and the planned surgical procedure was completed. The pt returned to the operating room three days after the procedure was performed with the system and a small bowel laceration was discovered and repaired. It is believed that the small bowel laceration was caused at some point in the procedure on 4/1/08, when the injured area of the bowel was not in view. This could have occurred during initial placement of the surgical ports, when the pt bucked on the table and the trocars were still installed during the system delay, during manipulation of the bowel, or during an unobserved insertion of a da vinci s or assistant instrument during the completed procedure. | The investigation conducted by isi field svc engineering (fse) concluded that the system faults experienced by the customer were associated with the cable connecting the pt side cart (psc) ans surgeon side console (ssc) and the connectors on the psc and ssc. The isi fse observed that the cable and connectors had major contamination, which may have caused system communication issues, thus resulting in the system faults experienced by the customer. The system was repaired by cleaning the cable and connectors. As of 7/14/08, it was reported that the pt was recovering well. As of 7/24/08, there have been no reported recurrences of the issue at this hosp. The hosp also submitted medwatch report regarding the system malfunction, however, no info was provided on the report indicating that the pt was impacted in any way. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1025862 | 1108804 | 1016724 | 2955842200801176 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-08-01 | 2008-08-01 | Yes | Yes | Health Professional | IS2000 A51P4 | No | Yes | 2008-07-03 | Yes | No | Reuse | 2008
IS1200 A43 P9 | 2008-06-30 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the customer experienced multiple recurrences of system error codes #20008. The surgical staff was able to clear the fault by restarting the system, however, the fault recurred. The pt was under anesthesia for approximately 1. 5 hours when the site converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that system error codes #20008 and #21008 were experienced by the customer and were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system error codes #20008 and #21008 functioned as designed and there was no injury to the pt. The davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci in a "recoverable safe state. " the ecm were returned to isi for failure analysis investigation. Engineering eval found that an axis motor encoder assembly had malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrence of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1231338 | 1085160 | 1043033 | 2955842200801174 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-07-30 | 2008-07-31 | Yes | Yes | Health Professional | IS1200 A43 P9 | Yes | Yes | 2008-06-30 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument could not cut. The instrument was examined and it was noticed that the wires were broken. The planned surgical procedure was completed with another instrument. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable frayed at the scissor grip and distal idlers. The scissors can still open and close, but the cable tension is below average, which affects cutting performance. The grip hub has a peak across the cable groove and cable wear against the grip and pulleys most likely contributed to the fraying. Engineering also observed that the distal end of the main tube has a 1 inch long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1228521 | 1085048 | 1042921 | 2955842200801177 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-30 | 2008-07-30 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-07-05 | Yes | No | Reuse | 2008
| 2008-07-17 | Injury |  | While using the intuitive surgical inc. 's davinci s robotic system in an elective coronary bypass surgical procedure, the robotic camera remote control function failed. The remainder of this surgical procedure was successfully completed using the robot, with manipulation of the camera arm performed by surgical staff. |  | DAVINCI S SURGICAL SYSTEM | ROBOTIC SYSTEM | 1210500 | 1180743 | 1131474 | 1180743 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2008-07-30 | 2008-07-30 | Yes | Yes | Invalid Data |  | No | No Answer Provided |  |  |  |  | 2008
42019006 |  | Malfunction |  | It was reported that the cobra grasper instrument has a broken cable. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable to be broken at the distal idler pulley and the adjacent grip open cable was derailed. The grip open cable derailment may have contributed to the grip close cable breakage. Engineering also observed the distal end of the main tube to have a 2. 75 inch long section with material removed on one side of the tube. The damaged area is parallel to tube axis and had a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will submitted if additional information is received. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1357494 | 1084414 | 1042280 | 2955842200801165 | 1 | GCJ | Manufacturer | Company RepresentativeOther | UNKNOWN | Initial | 2008-07-25 | 2008-07-28 | No | Yes | Health Professional | 42019006 | Yes | No | 2008-07-05 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a wire on the endowrist prograsp forceps instrument broke. The procedure was successfully completed using another instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the pitch cables at the wrist are loose and the wrist does not respond to movement of pitch input. The housing was removed to find the top pitch cable detached from the clamping pulley and the cable crimp to have slipped off the cable, which may have been a result of incomplete crimping. Engineering also observed the distal end of the main tube to have a 2. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1358690 | 1084415 | 1042281 | 2955842200801175 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-28 | 2008-07-28 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-07-02 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument would not work. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a broken conductor wire at the yaw pulley exit. The jaw pulley is missing a. 150" x. 060" piece opposite of the conductor break, allowing the grip to move within the pulley and have a larger range of jaw motion. The cause of the breakage cannot be determined, however, the added range of grip motion likely created excess tension on the wire, causing it to break. The wire may not have been securely attached to the grip and pulled out during articulation of the wrist due to poor strain relief. Electrical continuity failed. Engineering evaluation also observed that the distal end of the main tube has various deep scratches concentrated on one side of the tube. Tube material has been removed as a result of the scratches. The scratches are not aligned to the tube axis, suggesting that they were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1358679 | 1084416 | 1042282 | 2955842200801178 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-28 | 2008-07-28 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-07-05 | Yes | No | Reuse | 2008
420179 |  | Other | Other | It was reported that during a da vinci s total hysterectomy surgical procedure, the pt had giant fibroids and the surgeon experienced difficulty in moving the uterus. The monopolar curved scissors instrument broke due to the amount of torque being applied and the surgeon decided to convert to traditional open surgical techniques to finish the planned surgery to prevent any additional instrument breakage. No pt harm was reported. | The instrument has not been returned to isi for eval. Based on the info provided, it was determined that the difficulty with the pt's anatomy was a contributing factor in the decision to convert the planned surgical procedure. A follow up mdr will be submitted if the instrument is returned or if additional info is received. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. As of 2008, there have been no reported recurrences of the issue at this hosp. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1311139 | 1081621 | 1039491 | 2955842200801173 | 1 | GEI | Manufacturer | User facility | Physician | Initial | 2008-07-25 | 2008-07-25 | Yes | Yes | Health Professional | 420179 | No | Yes | 2008-06-27 | Yes | No | Reuse | 2008
42000202 | 2008-06-03 | Malfunction |  | The end of trocar appeared to be slightly bent at the insertion sleeve, scraping the instrument sliding through it. This left metal shavings in the patient's abdomen. |  | DA VINCI | TROCAR | 1076879 | 1100443 | 1057234 | 1100443 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2008-07-25 | 2008-07-25 | No | Yes | Physician | 42000202 | Yes | No Answer Provided |  |  |  |  | 2008
| 2008-06-10 | Malfunction |  | Patient undergoing robot-assisted laparoscopic-assisted vaginal hysterectomy. During the procedure the pk dissecting forceps was placed through the trocar and was connected to the cautery. The pedal on the cautery was depressed several times yet the forceps did not respond. When the forceps was removed from the trocar a broken wire was detected. Post-op the patient developed a purple mark around the site of the trocar. Initially it was thought to be a bruise, after further examination, it was thought to be a burn. |  | PK DISSECTING FORCEPS | FORCEPS ROBOTIC ACCESSORY | 1124039 | 1129663 | 1083474 | 1129663 | 1 | NAY | User Facility |  | Nurse | Initial | 2008-07-15 | 2008-07-25 | No | Yes | Physician |  | No | Yes |  |  |  |  | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering conducted a cut test and the scissors performed as intended. The instrument blades were not damaged. The customer alleged failure mode could not be confirmed. Engineering also observed that one side of the distal end of the main tube as a 1. 5 inch long section with material removed. The damaged area is parallel to the tube axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCIESSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1364646 | 1083233 | 1041080 | 2955842200801157 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-07-24 | 2008-07-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-24 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-06-26 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure, the customer experienced multiple recurrences of system error codes #23013 and #23008. The surgical staff was unable to clear the faults. The patient was under anesthesia for 2 hours before the surgical staff made the decision to reschedule the procedure to a later date. No port incisions had been created. No patient harm, adverse outcome or injury was reported. | Conclusion: the investigation conducted by field service engineering concluded that system error code #23013 and #23008 experienced by the customer were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system error codes #23013 and #23008 functioned as designed and there was no injury to the patient. The davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". The ecm were returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, the encoder assembly and a motor assembly were replaced as a precautionary measure. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1211760 | 1080669 | 1038580 | 2955842200801162 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-07-21 | 2008-07-21 | Yes | Yes | Health Professional | IS2000 A51 P4 | Yes | No | 2008-06-26 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-23 | Other | Other | It was reported that while setting up the da vinci s surgical system for a hysterectomy procedure, the endoscopic camera manipulator (ecm) cannula camera mount broke off. The pt was under anesthesia for approximately 30 minutes and the port incisions had been created when the site decided to convert to laparoscopic surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by an isi representative confirmed that the endoscopic camera manipulator (ecm) cannula mount was broken. The ecm cannula mount holds the cannula securely in place while the camera is delivered through the cannula. The cannula provides the port through the body wall of the pt, through which the endoscope is inserted. The system was repaired by replacing the affected ecm cannula camera mount. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1212257 | 1080670 | 1038581 | 2955842200801153 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-07-18 | 2008-07-21 | Yes | Yes | Health Professional | IS2000 A51P4 | No | Yes | 2008-06-23 | Yes | No | Reuse | 2008
IS1200 A4 3P1 | 2008-06-27 | Other | Other | It was reported that during a da vinci surgical procedure, the customer was unable to focus the camera from the surgeon side console (ssc) or the vision side cart (vsc). With the assistance of an isi technical support engineer and the onsite representative, the customer re-seated the focus cable and replaced the camera head cable, however, the focus problem continued to recur. The pt was under anesthesia for approximately one hour when the site decided to convert to laparoscopic surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by an isi field service engineer found that the 3d camera head was defective. The camera head consists of two video cameras, one camera is used for each optical path (right and left). The camera head is connected to a focus controller, which allows the surgeon to focus the camera from the surgeon side console (ssc) using a foot pedal. The system was repaired by replacing the affected camera head. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1212239 | 1080671 | 1038582 | 2955842200801155 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-07-18 | 2008-07-21 | Yes | Yes | Health Professional | IS1200 A4 3P1 | Device Returned To Manufacturer | No | 2008-06-27 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-25 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced problems with one of the pt side manipulator (psm) arms. An isi technical support engineer confirmed that the psm set up joint was not moving and found a system error code. The pt was under anesthesia when the site decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the remote arm controller (rac) board. The system was repaired by replacing the affected rac. The rac is an electronic module comprised of four printed circuit assemblies and cooling fans. System error code appears when the da vinci s safety system determines during the power up self test that the remote arm controller board (rac) power failed the set up joint reinforce test. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The rac was returned to isi for failure analysis investigation. Inspection was performed and two low voltage differential signaling connectors for the rac malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1212200 | 1080672 | 1038583 | 2955842200801154 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2008-07-18 | 2008-07-21 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-06-25 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2080-06-20 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure, the customer experienced a non recoverable system error code #25004. With the assistance of an isi technical support representative the customer powered down and restarted the system however, system error code #23008 occurred. Additional troubleshooting steps were taken and eventually the third system arm was disabled to successfully complete the case. The patient was under anesthesia for an additional 18 minutes while trying to resolve the system issue. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code #25004 and #23008 experienced by the customer were associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system error codes #25004 and #23008 functioned as designed and there was no injury to the patient. The davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering evaluation determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrence of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1198520 | 1079083 | 1037014 | 2955842200801150 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-18 | 2008-07-18 | Yes | Yes | Health Professional | IS2000 A51 P4 | Yes | No | 2008-06-20 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the maryland bipolar forceps instrument was broken. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering conducted performance tests and found the instrument moved intuitively with a full range of motion. Electrical continuity passed and the conductor wires are not damaged at the wrist or at the back end of the instrument. Engineering also observed that the distal end of the main tube has a 1. 5 inch long section with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337281 | 1079628 | 1037559 | 2955842200801163 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-06-28 | Yes | No | Reuse | 2008
40017905 |  | Malfunction |  | It was reported that a burning smell was noticed coming from the wires of the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a. 165" x 100" oblong hole burned through it. The hole is located roughly. 450" above the interface with the tube extension. The tube exhibits localized charring around the hole, indicating an arcing event occurred. The instrument was disassembled to find charring on the tube extension and hypotubes. The main tube has a crack parallel to the tube axis directly below the hole that runs all the way down the tube. The source of the crack is unclear, however, it was determined that the crack in the tube most likely created a path for arcing to occur and high generator settings may have also contributed to the damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337217 | 1079629 | 1037560 | 2955842200801156 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-07-21 | 2008-07-18 | No | Yes | Health Professional | 40017905 | Yes | No | 2008-06-24 | Yes | No | Reuse | 2008
42014304 |  | Malfunction |  | It was reported that during a da vinci s thyroidectomy procedure, damages on the 5mm maryland dissector instrument were observed. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed that the main tube has various damaged sections, some with material removed on one side of the tube. The damaged area is located roughly. 200 inch, 1. 5 inches, and 3 inches above the snake wrist. The tube insulation appears to have scratches and gouge marks. The location and appearance of the damage indicates that it was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337140 | 1079630 | 1037561 | 2955842200801160 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-06-25 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument would not grip tissue. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering was able to open and close the instrument grips by manually rotating the inputs. System tests could not be performed since the instrument was returned with no uses remaining. The cable tension was slight low, which is not unusual for an instrument with no uses remaining, however, the stretched cables may have affected the grasping force. Results and conclusions: engineering also observed that one side of the distal end of the instrument main tube has a 3. 5 inch long section with material removed. The damaged area is located directly above the proximal clevis, parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337191 | 1079631 | 1037562 | 2955842200801152 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-06-23 | Yes | No | Reuse | 2008
42014204 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure damages on the 5mm monopolar cautery instrument were observed. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed that the main tube has a. 650 inch long damaged section with material removed on one side of the tube. The damaged area is located roughly 2. 7 inches above the snake wrist. The tube insulation appears to have scratches and gouge marks. The location and appearance of the damage indicates that it was most likely caused by instrument collisions. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337101 | 1079632 | 1037563 | 2955842200801159 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 42014204 | Yes | No | 2008-06-25 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 275" x. 150" oblong hole burnt through the main tube. The hole center is located roughly. 500" about the tube extension. There is a ring of tube material removed around the perimeter of the hole, indicating an arcing event occurred. The main tube has a longitudinal crack that bisects the hole and is located approximately halfway between two slot features. It appears that the crack may have been created from rough user handling. One of the slots is damaged, with pieces coming off. The instrument was disassembled and engineering discovered that the hypotubes and tube extension seal were charred in the same location as the hole. It was determined that the crack in the tube most likely created a path for arcing to occur and high generator settings may have also contributed to the damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1337057 | 1079633 | 1037564 | 2955842200801164 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-06-28 | Yes | No | Reuse | 2008
42014304 |  | Malfunction |  | It was reported that during a da vinci s thyroidectomy procedure damages on the 5mm maryland dissector instrument were observed. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering confirmed that the main tube has a couple of different damaged sections with tube insulation removed. The damaged areas are located roughly 0. 7 inches and 1. 7 inches above the snake wrist. The tube insulation appears to have scratches and gouge marks. The location and appearance of the damage indicates that it was most likely caused by instrument collisions and removal from the cannula accessory may have exacerbated the damage. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments; do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2882151 | 2853175 | 2749711 | 2955842200801161 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2008-07-18 | 2008-07-18 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-06-25 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-06-16 | Other | Other | It was reported that prior to starting the da vinci s surgical procedure, the customer experienced a system error code #23. The patient has been under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule it to a later date. No pt harm, was reported. | The investigation conducted by field service engineering concluded that the system error code #23 experienced by the customer was associated with a configured embedded serializer setup joint (cfg, essj). The embedded serializer for setup-joint is the printed circuit assembly (pca) inside a system arm that monitors the potentiometer for each of four joints and their associated backup potentiometers. The system was repaired by replacing the affected cfg essj. System error code #23 is reported by software to denote that the hardware wheel "wdog" has tripped on one of the digital communication links in the system. This means that the system cannot reliably communicate over the digital link and therefore cannot continue normal operation. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was aborted. As of 2007, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1213921 | 1081625 | 1039495 | 2955842200801142 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-16 | 2008-07-16 | Yes | Yes | Health Professional | IS2000 A5 1P4 | Yes | No | 2008-06-17 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-16 | Injury | Required InterventionLife Threatening | It was reported that during a da vinci s beating heart double vessel coronary artery bypass graft procedure at the hospital, there was an unexplained movement on the system arm which had the endowrist stabilizer instrument attached to it. The unexpected movement caused the feet at the distal end of the endowrist stabilizer instrument to tip downward resulting in damage to the myocardium of the patient's left ventricle. The surgical team immediately converted to an open sternotomy to successfully repair the damaged ventricle with a single medtronic u-clip s50. The patient was reported as recovering well with no post-operative complications. | The investigation conducted by an isu field service engineer found the system to successfully completed all verification tests and to function as designed. No system errors related to the cause of this event were found. Additional investigations conducted by clinical and software engineering were inconclusive as a root cause could not be determined based on the event descriptions reported by several eye witnesses and the review of the system's event logs. The logs did show that near the time of the event, the arm (holding the endowrist stabilizer instrument) was "unlocked", indicating that the instrument was ready for surgeon control. The particular log used to determine if the instrument was controlled and actively in use at the time of the event had not retained the information as each log records info in a circular buffer of limited lengths. This means a finite number of the most recent events are recorded in the log. The hosp has continued to use the da vinci s system to perform surgery on patients. As of 2008, no adverse events or similar instances of this event have been reported. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1045565 | 1077464 | 1035415 | 2955842200801144 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-07-15 | 2008-07-15 | Yes | No | Health Professional | IS2000 A51P4 | Yes | No | 2008-06-17 | Yes | No | Reuse | 2008
40000620 | 2008-06-19 | Malfunction |  | It was reported that during da vinci s prostatectomy procedure, a piece of the large needle driver instrument broke off and fell into the pt. The piece could not be located and the surgical staff made the decision to leave the piece inside the pt. The procedure was completed without significant delay. The instrument was inspected and appeared complete. No pt harm, adverse outcome or injury was reported. | The instrument has returned and evaluated. Per the customer reported complaint, engineering confirmed that the carbide insert was missing from one grip. The brazing surface of the grip shows very little traces of filler metal used to join insert and grip. The insert detached from the grip due to poor adhesion between grip and insert. Materials used in the intuitive surgical needle driver instruments have been tested for biocompatibility. All materials were found to be appropriately biocompatible for their intended use. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1196080 | 1079147 | 1037078 | 2955842200801147 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-07-15 | 2008-07-15 | No | Yes | Health Professional | 40000620 | Yes | No | 2008-06-19 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on an in-house test system and successfully recognized. The pogo pins did not stick and were not contaminated. Engineering was unable to confirm the customer alleged failure mode. Engineering also observed that the distal end of the main tube has a couple of deep scratches with material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1220230 | 1075728 | 1033691 | 2955842200801151 | 1 | GCJ | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2008-07-14 | 2008-07-14 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42020504 | 2008-06-20 | Malfunction |  | It was reported that during a da vinci s nissen fundoplication procedure, shavings were observed coming off of the shaft of the fenestrated bipolar forceps instrument, a tiny piece fell into the pt and was retrieved. The planned surgical procedure was completed with a different type of instrument. No additional info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering confirmed that the distal end of the main tube has two sections with material removed. The sections are 180 degrees apart and each are roughly 2 inches long. The damaged areas are mostly parallel to the tube axis, however, the appearance of the damage indicates that it was most likely caused by a combination of instrument collisions and a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instrument. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1190718 | 1076732 | 1034688 | 2955842200801148 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-06-20 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument are not lining up. No additional info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a. 110 inch offset at the instrument tips. The bent grip has a separation of the yaw pulley and conductor cap at the glue joint with the yaw pulley cracked at the base of the bent grip. Engineering concluded the damage was likely caused by overloading at the tip. Engineering also observed the distal end of the main tube has a 1 inch long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis with a wear patter consistent with cannula related tube abrasions. A deep scratch below the scraped section was likely caused by an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequence. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1198478 | 1076764 | 1034720 | 2955842200801135 | 1 | GEI | Manufacturer | User facilityOther | Physician | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42022702 | Yes | Yes | 2008-05-13 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | It was reported that the cannula accessory had light resistance. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering observed a visible dent at the distal tip and a. 342 gage pin does not fit through the distal tip inner diameter. The appearance of the dent leads engineering to conclude that the damage was likely caused by user mishandling. It was determined that the deformation of the tip has the potential to cause scraping of the instrument main tube in certain loading conditions. No other damage found. The cannula and obturator instructions for use (ifu) specifically states: cannula inspection instructions warning: do not use a cannula with imperfections on its distal end. Examples of defects include a rough edge, dents, or an out-of-round shape. Use of a defective cannula may results in the instrument damage, including scraping of the instrument shaft and particulate falling into the pt. Inspect the cannula prior to use for any defects. Warnings: cannula defects can be caused by dropping the cannula or handling it roughly. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 1197917 | 1076769 | 1034725 | 2955842200801145 | 1 | GCJ | Manufacturer | User facility | Nurse | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42000202 | Yes | Yes | 2008-06-19 | Yes | No | Unkown | 2008
42020503 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument is not working properly and that the tip is loose. No additional info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips open and close properly when manually rotated. The instrument was returned missing both housing levers with one rear housing snap tab and three pins broken off. Engineering also observed the distal end of the main tube has a 2 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis with a wear patter consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1198400 | 1076770 | 1034726 | 2955842200801136 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-06-18 | Yes | No | Reuse | 2008
42017908 | 2008-06-18 | Malfunction |  | It was reported that when performing the initial prostate dissection during a da vinci s prostatectomy surgical procedure the monopolar curved scissors (mcs) instrument was in use for approx 2 minutes when a burning smell was noticed. The procedure was paused and the source of the smell was found at the tip of the instrument which connects to the monopolar cautery cord. The mcs instrument and cautery cord were removed and replaced to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 130 inch by. 070 inch oblong hole burnt through the distal end of the main tube. The hole center is located roughly. 500 inches above the tube extension. Localized material removed around the hole indicates that an arcing event occurred. The instrument was disassembled and engineering discovered that the tube extension and the hypotubes were charred in the same location as the hole. A crack in the distal end of the tube was noticed, which begins at one of the slot features and runs along the side of the hole. The source of the crack is unclear, however, it was determined that the crack in the tube most likely created a path for arcing to occur and high generator settings may have also contributed to the damage. No evidence of arcing at the instrument back end was observed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1193582 | 1076771 | 1034727 | 2955842200801146 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-19 | Yes | No | Reuse | 2008
42009305 | 2008-06-20 | Malfunction |  | It was reported that during a da vinci s nissen fundoplication procedure, the shaft of the endowrist prograsp forceps instrument had shavings coming off however no pieces fell into patient. No additional info provided. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the instrument was returned and evaluated. Per the isi rep reported complaint, engineering confirmed that the distal end of main tube has a 2. 5 inch long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish, however, the appearance of the damage indicates that it was most likely caused by a combination of instrument collisions and a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Conclusions - the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1193738 | 1076772 | 1034728 | 2955842200801149 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-07-11 | 2008-07-11 | No | Yes | Health Professional | 42009305 | Yes | No | 2008-06-20 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-06-19 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced a recoverable system error code #23005. The surgeon decided to convert to traditional open surgical techniques to finish the planned surgery prior to speaking to an isi technical support rep. No pt harm was reported. | The investigation and eval conducted by technical support engineering found that the system error code #23005 experienced by the customer was likely due to the pt port placement. The system error code #23005 appears when the da vinci s safety system determines a pt side manipulator (psm) arm has collided with another arm or something else in the environment. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1199897 | 1077081 | 1035037 | 2955842200801137 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-07-11 | 2008-07-11 | Yes | Yes | Health Professional | IS2000 A5 1P4 | No | No | 2008-06-19 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-20 | Other | Other | It was reported that near the end of a da vinci s hysterectomy procedure, the customer experienced a non recoverable system error code #1. With the assistance of an isi technical support engineer, the customer power cycled the surgeon side console and pt side cart multiple times, however, the system error code #1 recurred. The pt was under anesthesia for approx four hours when the site converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by the technical support engineering concluded that the system error code #1 experienced by the customer was associated with a pt side manipulator (psm) and remote arm controller (rac) board. The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm and rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #1 appeared when the da vinci s safety system detected a hardware power supply voltage was out of range. Upon determining these conditions, the safety system put the da vinci s system in a "non-recoverable safe state". The psm and rac were returned to isi for failure analysis investigation. Engineering concluded that the loose wrist drive cables on axis 6 of the psm generating the system errors experienced by the customer. The psm was repaired by replacing the insertion drive cables. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1199919 | 1077082 | 1035038 | 2955842200801138 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-11 | 2008-07-11 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-06-20 | Yes | No | Reuse | 2008
ISI200 A43P9 | 2008-06-16 | Other | Other | It was reported that prior to starting the da vinci s prostatectomy surgical procedure, the right "eye" of the high resolution stereo viewer was black, and no text or icons were visible. An isi technical support engineer attempted to troubleshoot the issue via telephone, however, the issue could not be resolved. The pt was under anesthesia for approximately 30 minutes, and port incisions were made when the surgeon decided to abort and reschedule the planned surgical procedure to a later date. No pt harm was reported. | The investigation conducted by isi field service engineering concluded that the vision loss experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. The hrsv was returned and evaluated. Per the customer reported complaint, engineering discovered that the right side hrsv monitor had malfunctioned, thus causing the right "eye" vision loss experienced by the customer. As of date, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1202039 | 1077083 | 1035039 | 2955842200801143 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-07-11 | 2008-07-11 | Yes | Yes | Health Professional | ISI200 A43P9 | Yes | No | 2008-06-16 | Yes | No | Reuse | 2008
IS12000 | 2007-01-29 | Death | Death | On june 10, 2008, intuitive surgical was notified of a pt death per litigation documentation which reports the following: in 2007, five days after undergoing a da vinci radical prostatectomy procedure, a pt underwent emergency exploratory laparotomy and inter alia. Bowel necrosis with septicemia was observed and said to be the result of an interoperative mesentery tear. Six centimeters of continuous stitch was applied to repair the affected area, however, the pt expired. | Intuitive surgical requested add'l info related to this event. A follow-up mdr will be submitted if add'l info becomes available. The hospital had continued to use the da vinci surgical system and since december of 2007 has used the da vinci s surgical system to perform surgery on pts. As of july 9, 2008, no similar instances of this event have been reported to isi. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1040455 | 1071654 | 1029665 | 2955842200801133 | 1 | NAY | Manufacturer | Other | NOT APPLICABLE | Initial | 2008-06-09 | 2008-07-09 | Yes | No | Health Professional | IS12000 | No | No | 2008-06-10 | No | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the casing of the large needle driver instrument is busted and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that three of four snap tabs were broken off the instrument housing and two housing pins were broken off. As a result of the breakage, one side of the housing can be lifted up. Engineering also observed that the distal end of the instrument main tube has two sections, 180 degrees apart, with light material removed. The damaged areas are 1 inch and 3 inches long, parallel to the tube axis and have a rough surface finish. Based on the location and the appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1048799 | 1071708 | 1029719 | 2955842200801141 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-09 | 2008-07-09 | No | Yes | Health Professional | 42000605 | Yes | Yes | 2008-06-10 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument had a broken needle and nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable broken near the proximal pulleys and proximal clevis cable hole, allowing the pulley to spin freely. It was determined that cable wear on the pulleys/clevis combined with high grasping force likely contributed to the cable breakage. Engineering also observed that one side of the distal end of the instrument main tube has a 1. 25 inch long section with light material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1176657 | 1071709 | 1029720 | 2955842200801139 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-09 | 2008-07-09 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-10 | Other | Other | It was reported that during a da vinci s surgical procedure the customer experienced a system error and powered down the system. On power up, the led of the endoscopic camera manipulator (ecm) arm was red and system error code # 903 appeared. The customer powered down the system for a second time and recycled the power on the surgeon side console (ssc) and pt side cart (psc). The system was then restarted and system error code # 40006 appeared. The system was shutdown and restarted for a third time, however, the system would not complete start up and system error code # 903 recurred. The pt had been under anesthesia for approx 3 hours when the surgeon decided to convert to an open surgical technique to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering concluded that system error codes # 903 and #40006 experienced by the customer were associated with the remote arm controller (rac) for the ecm and/or the cable connecting the ssc to the psc. The ecm is the system camera arm located on the psc which provides the sterile interface for the 3d endoscope. The rac consists of five pcas which operate together to provide control of the system arms. The system was repaired by replacing the affected rac and cable. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #903 is an indication that the error log is full and system error code # 40006 appears when an attempt is made to insert an element into a full queue. The vision cable and rac were returned to isi for failure analysis investigation. Engineering found no issues with the rac, however, it was observed that a vision cable fiber contact had major contamination, which may have caused a system communication fault thus leading to the error log/queue overflows that were indicated by the error codes. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1349697 | 1071753 | 1029764 | 2955842200801140 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-09 | 2008-07-09 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No | 2008-06-10 | Device Not Retu | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the back of the tenaculum forceps instrument broke off. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the luer plate dislodged from the chassis and pushed into the back end of the instrument. The chassis feature that retains that luer plate is undamaged , however , all of the housing pins adjacent to the snap tabs are broken off. Engineering concluded that the instrument was likely mishandled, causing the housing to come off and the luer plate to dislodge. Engineering also found the distal end of main tube has a couple of deep scratches on one side of the tube with material removed. The scratch marks are directly above the proximal clevis and are not aligned with the tube axis. Engineering concluded that the scratches were most likely caused by instrument collisions or rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tips. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1176984 | 1071663 | 1029674 | 2955842200801124 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-08 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-06-09 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that during a da vinci s lobectomy procedure, the permanent cautery hook instrument did not react to movements of articulation. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one pitch up cable is broken at the distal clevis hub. Engineering concluded that the cable likely broke from overload in tension. Engineering also observed the distal end of the main tube has various deep scratches with material removed. The scratches are narrow and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1176758 | 1071664 | 1029675 | 2955842200801121 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-07-03 | 2008-07-08 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-06-12 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one grip close cable is frayed at the scissor grip. The scissors can still open and close. The grip hub has a small peak of material across the groove that likely contributed to fraying. Engineering also observed the main tube has a large section with material removed all around the tube. The damaged area starts 1 inch below the tube midpoint and extends down to the interface with the tube extension. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1176959 | 1071665 | 1029676 | 2955842200801123 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-03 | 2008-07-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-09 | Yes | No | Reuse | 2008
42017908 | 2008-05-27 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a cable on the monopolar curved scissors instrument broke. The procedure was completed with another instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is broken at the distal idlers causing the pulley to spin freely. Cable wear on the pulleys combined with high drive torques likely contributed to the cable breakage. Engineering also observed the main tube has a sections with material removed. One section has material removal all around the tube between the midpoint and the tube extension. The damaged areas are parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. The other damaged section is above the midpoint of the tube, out of range of the distal tip of a cannula. Engineering concluded that this tube damage may be from cleaning or pressure against the internal bowl and tube transition of the cannula. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1176559 | 1071666 | 1029677 | 2955842200801122 | 1 | GEI | Manufacturer | User facilityOther | Nurse | Initial | 2008-07-03 | 2008-07-08 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-06-11 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-06-11 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon experienced a dark image and could not continue. With the assistance of an isi technical support engineer, the site inspected and changed the light guide cable and endoscope, however, the image continued to appear dark. The camera control unit and touch screen advanced image settings were confirmed correct with the image recurring dark. The surgeon decided to converted to traditional open surgical techniques to finish the planned surgery after 30 mins of troubleshooting. No pt harm was reported. | The investigation conducted by isi field service engineering (fse) discovered that the issue experienced by the customer was associated with a defective camera head. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The defective camera head has not been returned to isi for evaluation. If the camera head is returned for evaluation a follow-up mdr will be submitted. As of 2008, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1353381 | 1071757 | 1029768 | 2955842200801127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-03 | 2008-07-08 | Yes | Yes | Health Professional | IS2000 A51 P4 | No | No |  | Device Not Retu | No | Reuse | 2008
IS2000 A51P4 | 2008-06-10 | Other | Other | It was reported that during a da vinci s thymectomy procedure, the customer experienced multiple occurrences of system error code #22001. With the assistance of an isi rep, the customer rebooted the system and moved the endoscopic camera manipulator (ecm) along its range of motion. The customer proceeded with procedure, however, a non recoverable system error #22001 recurred. At this time the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. A delay of 30 mins was noted. No pt harm was reported. | The investigation conducted by field service engineering found that the system error code # 22001 was experienced by the customer as well as system error code # 22003. The field service engineer conclude that these system error codes were associated with the potentiometer on the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm potentiometer. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #22001 and #22003 system error code appeared when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". The ecm was repaired by replacing the affected potentiometer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1353342 | 1071758 | 1029769 | 2955842200801118 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-03 | 2008-07-08 | Yes | Yes | Service Personnel | IS2000 A51P4 | Yes | Yes | 2008-06-10 | Device Not Retu | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument is not holding the needle. No add'l info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a decrease in the cable tension that could reduce the grasping force. The carbide teeth are not damaged. Engineering also found the distal end of the main tube has an area on one side of the tube with various light scratches and one deep scratch with material removed. The deep scratch is located 1. 5 inches above the proximal clevis and is not aligned to the tube axis. Based on the location and appearance, the damage was most likely caused by a instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings; handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1350132 | 1071376 | 1029367 | 2955842200801116 | 1 | GCJ | Manufacturer | User facilityOther | Nurse | Initial | 2008-07-07 | 2008-07-07 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-06-10 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer's reported complaint, engineering conducted performance testing and found the instrument to pass recognition and engagement tests. The instrument moved intuitively with a full range of motion in all directions with the grips opened and closed properly. Engineering also found the distal end of the main tube to have a 3. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1193717 | 1071377 | 1029368 | 2955842200801128 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-07 | 2008-07-07 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-06-13 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument is not holding the needle. No add'l info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument passed the recognition and engagement testing. The instrument had a full range of motion in all directions with grips opening and closing properly. A needle was grasped and suturing was simulated without any issues. Engineering also found the distal end of the main tube has a 1. 4 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | LARGER NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1350210 | 1071378 | 1029369 | 2955842200801115 | 1 | GCJ | Manufacturer | User facilityOther | Nurse | Initial | 2008-07-07 | 2008-07-07 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-06-10 | Yes | No | Reuse | 2008
42020501 |  | Malfunction |  | It was reported that the tip of the fenestrated bipolar forceps instrument does not close well. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to not grasp well due to one grip being slightly bent, creating side to side misalignment of grips. There is a. 030" offset at the tips. The bent grip has a slight separation of the yaw pulley and the conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed the distal end of the main tube with material removed. The scratches are narrow, under. 250" long, and not aligned with the tube axis, suggesting they may have been caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1350139 | 1071379 | 1029370 | 2955842200801134 | 1 | GEI | Manufacturer | OtherForeign | Other | Initial | 2008-07-07 | 2008-07-07 | No | Yes | Health Professional | 42020501 | Yes | No | 2008-06-12 | Yes | No | Reuse | 2008
40007702 |  | Other | Other | It was reported that upon completion of a da vinci s prostatectomy procedure, the inner, clear seal was missing from the cannula seal accessory. An inspection of the inside of the pt was performed and the missing portion was not located. No pt harm was reported. | The cannula seal accessory was not returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info becomes available a follow-up mdr will be submitted. | CANNULA SEAL ACCESSORY | ENDOSCOPIC ACCESSORY | 1177260 | 1071380 | 1029371 | 2955842200801132 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-07-07 | 2008-07-07 | Yes | Yes | Health Professional | 40007702 | No | Yes | 2008-06-17 | Device Not Retu | No | Unkown | 2008
40018005 |  | Malfunction |  | The customer reported the tip cover accessory has holes on the side. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda relating to this event: 2955842-2008-01126. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found multiple holes in the silicone with localized melting around all holes that is indicative or arcing. All of the damage is contained within one half of the silicone, between the two parting lines. The holes are all under. 030 inches with varying shapes, round and oblong, and are located from. 275 inches to. 530 inches above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has various mechanical tears in the general vicinity of the holes, likely due to instrument collisions or repeated instrument wrist articulations. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1349327 | 1070036 | 1028067 | 2955842200801125 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-06-06 | Yes | Yes | Initial | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument jaws will not close. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument moved intuitively with full range of motion and the grips opened and closed properly. Engineering also found the main tube has a 1. 75 inch long section located 1. 5 inches below the midpoint with light material removed all around the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1349278 | 1070037 | 1028068 | 2955842200801119 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-06-06 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument has wires sticking out, and there was no patient involvement. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed both grip cables on one side of the proximal clevis are derailed at the distal idlers and the grip close cable is frayed. Engineering concluded that the cable may have derailed onto the pulley edge during articulation of the wrist, and became frayed. Engineering also observed the distal end of the main tube has material removed due to a combination of scratches and scraping. Tube has a 2 inch long scrape mark parallel to the tube axis. Based on the location and appearance of the scraping, the damage was most likely caused by a cannula accessory. There are also various deep scratches embedded within the scraped section. These scratches are not aligned with the tube axis. Based on the location and appearance of these deep scratches, the damage was most likely caused by instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1357815 | 1070038 | 1028069 | 2955842200801110 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument has "metal slivers" coming off after sterilization process. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of the main tube has various deep scratches concentrated on one side of the tube with material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1357817 | 1070039 | 1028070 | 2955842200801111 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
40018005 |  | Malfunction |  | The customer reported the tip cover accessory has holes on the side. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda relating to this event: 2955842-2008-01125. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a hole in the silicone with localized melting that is indicative of arcing. The hole center is located. 165 inches above the silicone to sleeve interface and is approximately. 030 inches in diameter. The hole is bisected by one of the parting lines. Engineering concluded that insufficient silicone dielectric strength and repeated instrument wrist articulations likely caused arcing. High generator settings may have also contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1357821 | 1070040 | 1028071 | 2955842200801126 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-06-06 | Yes | Yes | Initial | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument is not working. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument moved intuitively with full range of motion and the grips open and closed properly. Engineering also found the distal end of main tube has a section located 2 inches above the proximal clevis with various deep scratches concentrated on one side with material removed. Engineering concluded that the scratches were most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1210018 | 1070071 | 1028102 | 2955842200801120 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-06-04 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tip monopolar curved scissors instrument is broken, and there was no patient involvement. No additional info provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the tube extension is broken and missing a. 250 inch by. 285 inch piece at the proximal clevis interface. The clevis is dislodged from the tube extension. Engineering concluded the tube extension was most likely broken due to excessive side loading or other mishandling or misuse. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a patient during a surgical procedure. Engineering also observed the distal end of the main tube has a 3. 5 inch long section with material removed all around the tube. The damaged section is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1210077 | 1070072 | 1028103 | 2955842200801112 | 1 | GEI | Manufacturer | OtherUser facility | Unknown | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that a wire at the wrist of the monopolar curved scissors instrument had broken. The instrument was replaced with an instrument of the same type to successfully complete the procedure. No pieces of the instrument had fallen into the patient. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable to be broken at the scissor grip and distal idlers. The grip hub has a peak across the cable groove that should have been machined off by the supplier. The cables wear against this peak likely contributed to the breakage. The supplier has been informed of this issue. Engineering also observed the main tube to have a 2" long section near the midpoint with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Part of the damaged area is above the midpoint, which is outside the range of motion of the distal tip of the cannula. The scraping may be due to tube pressure against the internal bowl/tube interface of the cannula or from rough handling during cleaning process. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress than can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1346744 | 1070073 | 1028104 | 2955842200801131 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the shaft of the precise bipolar pyramid tip instrument has splinters and there was no patient involvement. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the main tube has a 3 inch long section located 1. 5 inches below the midpoint with material removed all around the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1355970 | 1070075 | 1028106 | 2955842200801113 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the fenestrated maryland bipolar instrument was not recognized by the system. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and found the instrument to successfully be recognized on the system. The instrument moved intuitively with a full range of motion in all directions with grips opening and closing properly. The pogo pins do not stick and are not contaminated. Engineering was unable to replicate the recognition issue. Engineering also observed the main tube to have a deep v-shaped scratch with material removed located. 250" above the proximal clevis. Engineering concluded that the scratch was most likely caused by instrument collisions or other mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1356047 | 1070076 | 1028107 | 2955842200801130 | 1 | GEI | Manufacturer | User facilityForeign | UNKNOWN | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-06-04 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | It was reported that the permanent cautery spatula instrument was not recognized by the system. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and found the instrument to successfully be recognized on the system. The instrument moved intuitively with a full range of motion in all directions with grips opening and closing properly. The pogo pins do not stick and are not contaminated. Engineering was unable to replicate the recognition issue. Engineering also found the distal end of the main tube to have a couple of deep scratches with material removed. The scratches are not aligned with tube axis. Engineering concluded that the scratch was most likely caused by instrument collisions or other mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1356061 | 1070077 | 1028108 | 2955842200801129 | 1 | GEI | Manufacturer | User facilityForeign | UNKNOWN | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-06-04 | Yes | No | Reuse | 2008
40000621 |  | Malfunction |  | It was reported that the insulation is scratched on the large needle driver instrument and there was no patient involvement. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of main tube was various deep scratches concentrated on one side of the tube with material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. -do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTURMENT | ENDOSCOPIC INSTUMENT | 1178163 | 1070711 | 1028739 | 2955842200801114 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-07-03 | 2008-07-03 | No | Yes | Health Professional | 40000621 | Yes | No | 2008-06-05 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-06-04 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced a non recoverable system error #20008. With the assistance of an isi representative, the customer power cycled the system and placed the affected patient side manipulator (psm) arm out of use position. The customer proceeded with procedure, however, the non-recoverable system error #20008 recurred. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. A delay of 30 minutes was noted. No patient harm was reported. | The investigation conducted by the field service engineer concluded the patient side manipulator (psm) was having intermittent issues. The system was repaired by replacing the affected psm. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #20008 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder), and the secondary sensor (potentiometer) were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". Engineering concluded the psm arm generated the system error when attempting to run internal system sensor checks. Engineering conducted tests for the psm and confirmed the customer reported problem. The system was repaired by replacing the motor/encoder on axis-3. As of july 01, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1178165 | 1070712 | 1028740 | 2955842200801117 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-07-03 | 2008-07-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-06-04 | Yes | No | Reuse | 2008
8MM REF 420179 | 2008-05-21 | Injury | Hospitalization | Surgical procedure: total laparoscopic hysterectomy, bilateral salpingo-oophorectomy, washings via robotic assistance. Laparoscopic bilateral pelvic and aortic lymph node dissection via robotic assistance. Incident occurred during procedure. Complete lymphadenectomy was performed along the external, removing all lymph nodes. The obturator nerves was being dissected. In dissecting out the obturator nodes, it was noted that there was a spark from the proximal portion of the moving portion of the monopolar shears. The instrument was quickly removed and the protective coating (sheath/tip cover accessory) of the shear removed and inspected. Two 2 small defects (holes) were noted in this in addition to multiple small brown markings. This was preserved for later inspection. It was noted at this time that there was what appeared to be a thermal burn to the external iliac artery. A 5 mm rounded area of injury with a zone of blanching, measuring about 8mm proximal to this. Site marked with clips. |  | DAVINCI SURGICAL SYSTEM INTUITIVE | MONOPOLARCURVED SCISSORSTIP CO | 1209637 | 1172824 | 1124095 | 1172824 | 1 | NEY | User Facility |  | RISK MANAGER | Initial | 2008-05-30 | 2008-07-03 | Yes | Yes | Health Professional | 8MM REF 420179 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2008
STANDARD | 2008-05-14 | Malfunction |  | Davinci robot kept sounding an alarm and the surgeon would override the alarm. It continued to alarm frequently and was delaying progress of the case. We undocked the robot and restarted the system but the system continue to sound the alarm. |  | DANINCI CONSOLE | ROBOTIC SURGICAL SYSTEM | 1043749 | 1077205 | 1035156 | 1077205 | 1 | NAY | User Facility |  |  | Initial | 2008-05-14 | 2008-07-01 | No | Yes | Physician | STANDARD |  | No Answer Provided |  |  |  |  | 2008
420208 | 2008-06-04 | Malfunction | Other | It was reported that towards the end of a da vinci s hysterectomy procedure, the fenestrated bipolar forceps instrument cable broke and the carbon sheath fractured while inside the patient. The yellow led on the patient side manipulator (psm) arm came on when the surgical staff tried to remove the instrument. With assistance from isi technical support, the cannula accessory and instrument were undocked from the patient and removed. The planned surgical procedure was completed with minimal delay. The instrument was inspected and appeared complete. No patient harm, adverse outcome or injury was reported. | Conclusion: the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1330891 | 1066913 | 1024964 | 2955842200801097 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-06-27 | 2008-06-27 | Yes | Yes | Health Professional | 420208 | No | No | 2008-06-04 | Device Not Retu | No | Unkown | 2008
42017908 | 2008-06-03 | Malfunction | Other | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument shaft was "shaved" off and pieces fell into the patient's abdomen. The shaft fragments were immediately removed from the patient and the case was completed with another instrument and cannula accessory of the same type. No patient harm, adverse outcome or injury was reported. | Results and conclusion: the instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube has a two inch long section with material removed on one side of the tube. The damaged area is located one inch above the tube extension and is parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1330886 | 1066914 | 1024965 | 2955842200801096 | 1 | NAY | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-06-27 | 2008-06-27 | Yes | Yes | Health Professional | 42017908 | Yes | No | 2008-06-03 | Yes | No | Initial | 2008
IS1200 A43P9 | 2008-06-03 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgeon stated that a pt side manipulator (psm) arm seemed "very sluggish". The pt was under anesthesia for approx one hour when the procedure was converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by the field service engineer concluded that a pt side manipulator (psm) was having intermittent issues. The system was repaired by replacing the psm and testing system. The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. Engineering conducted performance tests for the pt side manipulator (psm) arm and the customer reported problem could not be duplicated. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1149338 | 1067018 | 1025069 | 2955842200801107 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-27 | 2008-06-27 | No | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-06-03 | Yes | No | Reuse | 2008
40009308 | 2008-05-30 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the prograsp forceps instrument failed to insert into the instrument cannula. The site removed the instrument and replaced it with another instrument of the same type. The surgeon was then unable to move instruments on either patient side manipulator (psm). The power was cycled and the instrument re-inserted to successfully complete the procedure. The procedure was lengthened by approximately 45 minutes. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to have zero lives remaining. Dallas chip verification did not produce error codes to indicate a instrument failure. No physical defects could be found on the instrument. Additional investigation is currently being conducted to determine the cause of the customer reported problem. If additional information is obtained, a follow-up report will be submitted. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1175399 | 1067019 | 1025070 | 2955842200801109 | 1 | GCJ | Manufacturer | ForeignDistributor | Other | Initial | 2008-06-27 | 2008-06-27 | Yes | No | Health Professional | 40009308 | Yes | No | 2008-06-06 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-06-02 | Other | Other | It was reported that while setting up the da vinci s surgical system for a hysterectomy procedure the surgeon console's left eye view was black. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the problem without success. The patient was under anesthesia for approximately 30 minutes and the port incisions had been created when the site decided to convert to laparoscopic surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the black left eye view experienced by the customer was associated with configure system video processor. The device is primarily responsible for combining the incoming streams with graphics to create the outgoing streams. The system was repaired by replacing the affected device. The device was returned to isi for evaluation. Functional and system tests were performed and engineering was able to reproduce the customer reported failure mode. The device was repaired after modification of settings. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1175398 | 1067020 | 1025071 | 2955842200801108 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-06-27 | 2008-06-27 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-06-02 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument was not working. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Results - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube is cracked/bent at a 10 degree angle, roughly 3. 2 inches above the proximal clevis. The main tube has a section with severe material removal stretching from the bend point to the proximal clevis. The round tube is locally flat at the damaged section. It was determined that the damage was most likely due to mishandling and the tube may have been pulled/dragged against a sharp edge with high pressure, causing the material removal. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1159477 | 1065428 | 1023499 | 2955842200801106 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-25 | 2008-06-25 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42017205 |  | Malfunction |  | It was reported that the electrocautery function of the maryland bipolar forceps instrument was broken. No pt harm, adverse out come or injury was reported. | Results & conclusions - the instrument was returned and evaluated by engineering. Per the customer reported complaint, electrical testing was performed and electrical continuity failed. The instrument housing was removed and engineering discovered a detached wire from the top bipolar pin, thus causing the failure mode experienced by the customer. Engineering also observed that the main tube has a 1. 5 inch long section below the midpoint with light material removed all around the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. The tube also has a couple of scratches directly above the proximal clevis which appear to have been caused by instrument collisions. No other damage found. The endowrist instruments instructions for (ifu) especially sites. General precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1157197 | 1065429 | 1023500 | 2955842200801105 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-25 | 2008-06-25 | No | Yes | Health Professional | 42017205 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is defective. No add'l info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one grip close cable to be broken at the scissor grip. The grip hub has a peak across the cable groove causing the cable to wear against the peak which likely contributed to the breakage. Engineering also found the distal end of the main tube has two sections on the same side of the tube with deep scratches. The scratches are not aligned with the tube axis and were likely caused by an instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1165207 | 1065543 | 1023614 | 2955842200801093 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is broken on the end and nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tube extension is dislodged from the main tube and the entire tube extension wall is broken at the base of the axil keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing and it appears that the damage may have been caused by excessive side loading. Engineering also observed a 2. 75 inch long section with light material removed all around the distal end of the main tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1165205 | 1065545 | 1023616 | 2955842200801103 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-24 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s pyeloplasty surgical procedure, a broken cable was observed on the monopolar curved scissors instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a broken grip close cable at the distal idlers, allowing the pulley to spin freely. It was determined that cable wear on the pulleys, combined with high drive torques, most likely contributed to the cable breakage. Engineering also observed that one side of the distal end of the main tube has a 1. 25 inch section with material removed. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted in add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453507 | 1065547 | 1023618 | 2955842200801100 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-21 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the grips of the endowrist prograsp instrument are broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grips are not broken and the instrument moves intuitively when the inputs are manually rotated. The instrument is missing one release lever. The housing is intact, but likely popped off due to mishandling, thereby causing the lever to fall out. Engineering also observed that the distal end of the maint tube has three adjacent deep scratches with material removed. The scratches are all roughly. 015 inches long and not aligned with the tube axis, suggesting that they were most likely caused by instrument collisions. The endo wrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1453518 | 1065549 | 1023620 | 2955842200801104 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-24 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the endowrist prograsp forceps instrument would not grasp tissue or hold tissue in place. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to pass both recognition and engagement testing. The instrument was found to move intuitively with a full range of motion in all directions with the grips opening and closing properly. The grasping force was also found to be normal. Engineering also observed the distal end of the main tube to have various scratches on one side of the main tube, some of which have light material removed. The scratches are not aligned with the tube axis, suggesting possible instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1175386 | 1065571 | 1023642 | 2955842200801099 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to pass recognition and engagement testing. The instrument had a full range of motion in all directions with grips opening and closing properly. The instrument also passed a latex cut test and an electrical continuity test. Engineering also found the main tube to have a 3. 5" long section located 1" below the midpoint with material removed all around the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1175385 | 1065573 | 1023644 | 2955842200801098 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is defective. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the distal end of the main tube to have multiple deep scratches concentrated on one side of the tube, with material removed. Some of the scratches are not aligned with the tube axis and were likely caused by an instrument collisions. A few of the scratches are parallel to the tube axis, approximately. 040 inches wide, and were likely caused by rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1175383 | 1065577 | 1023648 | 2955842200801094 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that the tip of the fenestrated bipolar forceps instrument is not aligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of the grips. There is also a. 075 inch offset at the tips. The bent grip has separation of the yaw pulley and the conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed that one side of the distal end of the main tube has a 2 inch long section with material removed. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damaged found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1175357 | 1065578 | 1023649 | 2955842200801101 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-05-24 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument is defective. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the distal end of the main tube to have deep scratched and scrapes with light material removed. The scratches are not aligned with the tube axis and are likely to result of instrument collisions. The main tube also has a one inch long by. 110 inch wide scraped section that is parallel to the tube axis with a wear pattern consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1330078 | 1065579 | 1023650 | 2955842200801095 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-26 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the tips of the endowrist prograsp instrument do not open or close properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, testing was performed by engineering and the instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. Engineering evaluation also observed that the distal end of the main tube has a. 250 inch long deep scratch with material removed. The scratch is perpendicular to the tube axis, suggesting that it was most likely caused by an instrument collision. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This amy result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1331705 | 1065580 | 1023651 | 2955842200801102 | 1 | GCJ | Manufacturer | Health ProfessionalOther | UNKNOWN | Initial | 2008-06-24 | 2008-06-24 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-22 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-05-27 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon heard a "popping" sound from the high resolution stereo viewer (hrsv), then vision loss in the right "eye" of the hrsv. The right "eye" image became solid black and no test or icons were visible. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | An investigation conducted by the field service engineer concluded the vision loss experienced by the customer was associated with the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected hrsv. The high resolution stereo viewer (hrsv) is located on the surgeon console and provides the video image for the surgeon console operator. The hrsv was returned to isi for eval. Engineering discovered that the right side monitor had malfunctioned, thus causing the vision loss. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1146007 | 1064811 | 1022882 | 2955842200801085 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-23 | 2008-06-23 | No | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-05-27 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-05-29 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, after the instruments were inserted in the pt, the right "eye" of the high resolution stereo viewer became black and no text or icons were visible. An isi technical support engineer attempted to troubleshoot the issue via telephone, however, the issue could not be resolved. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering concluded that the vision loss experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1146003 | 1064812 | 1022883 | 2955842200801088 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-06-23 | 2008-06-23 | Yes | Yes | Health Professional | IS2000 A5 1P4 | Yes | No | 2008-05-29 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-05-22 | Other | Other | It was reported while setting up the da vinci s surgical system for a hysterectomy procedure, the illuminator would not powered on. With the assistance of an isi technical support engineer, the site reseated the power connections and lamp modules, however the illuminator would not turn on. The pt was under anesthesia for approx one hour and the port incisions had been created, when the site decided to convert to laparoscopic surgical techniques to finish the planned surgery. No pt harm was reported. | The illuminator has been returned to isi for eval and forwarded to the original equipment mfr (oem) for analysis. A follow-up mdr will be submitted if the illuminator is evaluated and additional info is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1146004 | 1064813 | 1022884 | 2955842200801087 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-23 | 2008-06-23 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No |  | Yes | No | Reuse | 2008
IS1200 A43 P8 | 2008-05-30 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced multiple occurrences of system error code #20008. With the assistance of an isi technical support representative, the customer powered down and restarted the system then reseated the instrument, however, the error continued to recur. The customer decided to complete the procedure with their second da vinci s system. The patient was under anesthesia for an additional 30 minutes while the da vinci s system was being prepared. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code #20008 experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generate error code) functioned as designed and there was no injury to the patient. System error code #20008 appears when the da vinci s safety systems determined that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering evaluation determined that a potentiometer assembly within the mtmr had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1175388 | 1064814 | 1022885 | 2955842200801084 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-06-20 | 2008-06-23 | Yes | Yes | Health Professional | IS1200 A43 P8 | Yes | No | 2008-05-30 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | It was reported that the tip of the cannula accessory was damaged. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering observed a visible dent at the distal tip. The appearance of the dent leads engineering to conclude that the damage was likely caused by user mishandling. It was determined that the deformation of the tip has the potential to cause scraping of the instrument main tube in certain loading conditions. No other damage found. The cannula and obturator instructions for use (ifu) specifically states: cannula inspection instructions warning: do not use a cannula with imperfections on its distal end. Examples of defects include a rough edge, dents, or an out-of-round shape. Use of a defective cannula may result in instrument damage, including scraping of the instrument shaft and particulate falling into the pt. Inspect the cannula prior to use for any defects. Warning: cannula defects can be caused by dropping the cannula or handling it roughly. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 1331563 | 1065865 | 1023916 | 2955842200801089 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-06-23 | 2008-06-23 | No | Yes | Health Professional | 42000202 | Yes | No | 2008-05-17 | Yes | No | Unkown | 2008
42000202 |  | Malfunction |  | It was reported that the tip of the cannula accessory was damaged. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering observed a visible dent at the distal tip. The appearance of the dent leads engineering to conclude that the damage was likely caused by user mishandling. It was determined that the deformation of the tip has the potential to cause scraping of the instrument main tube in certain loading conditions. No other damage found. The cannula and obturator instructions for use (ifu) specifically states: cannula inspection instructions warning: do not use a cannula with imperfections on its distal end. Examples of defects include a rough edge, dents, or an out-of-round shape. Use of a defective cannula may result in instrument damage, including scraping of the instrument shaft and particulate falling into the pt. Inspect the cannula prior to use for any defects. Warning: cannula defects can be caused by dropping the cannula or handling it roughly. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 1331672 | 1065869 | 1023920 | 2955842200801090 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2008-06-23 | 2008-06-23 | No | Yes | Health Professional | 42000202 | Yes | No | 2008-05-17 | Yes | No | Unkown | 2008
IS2000 A51P4 | 2008-05-21 | Injury | OtherRequired Intervention | It was reported that towards the end of a da vinci s gynecologic diagnostic laparoscopy procedure, the assistant surgeon placed the monopolar curved scissors (mcs) instrument inside of the pt's body and was moving the instrument outside of the field of view. The console surgeon took control of the instrument without having a visual of where the instrument was located and activated the instrument tip, which created a transection cut to the pt's right internal iliac artery and right ureter. The procedure was converted to open surgical techniques and a vascular surgeon and urologist were called in to repair the pt's right internal iliac artery and stent the ureter. A blood transfusion was required during the repair surgery. It was reported that the pt is recovering as planned post-operative day 27. | An isi rep present during the procedure did not observe the sys or instrument malfunction in a way that would cause or contribute to the event. It was determined that the pt injuries were caused by the instrument when the surgeon inadvertently moved the instrument tip from their field of view. The da vinci s user manual specifically states: caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the pt. Caution: removal of instruments during a procedure should be done very carefully and only with the knowledge of and full view of the surgeon console operator. Warning: instruments must be in the surgeon's field of view when matching grips. Matching grips for an instrument that is outside the surgeon's field of view may cause serious harm to the pt. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1031580 | 1063449 | 1021500 | 2955842200801076 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-06-20 | 2008-06-20 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No | 2008-05-21 | Device Not Retu | No | Reuse | 2008
42017206 | 2008-05-08 | Malfunction |  | The fenestrated maryland bipolar instrument was returned with an unknown problem. No additional info was provided. No pt. Harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering conducted tests and found the instrument to move intuitively in all directions with the grips opening and closing properly. The instrument also passed recognition and engagement testing. The grasping force is normal despite cables having been stretched due to normal wear and tear. Electrical continuity passed. Engineering also observed the main tube to have a 3" long section below the midpoint with material removed all around the tube. The damaged area is parallel to tube axis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow- up mdr will be submitted if additional info is received. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1129738 | 1063553 | 1021604 | 2955842200801082 | 1 | GEI | Manufacturer | ForeignUser facility | Unknown | Initial | 2008-05-24 | 2008-06-20 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-24 | Yes | No | Reuse | 2008
42017908 | 2008-05-08 | Malfunction |  | The monopolar curved scissors instrument was returned with an unknown problem. No additional info was provided. No pt. Harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the tube extension to be dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys resulting in the tube extension having the ability to rotate 360 degrees. No pieces are missing. Engineering has concluded that the damage may be due to excessive side loading. Engineering also observed the main tube to have light material removed all around the tube from the midpoint to the distal end. The damage area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow- up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1129818 | 1063554 | 1021605 | 2955842200801081 | 1 | GEI | Manufacturer | ForeignUser facility | Unknown | Initial | 2008-05-24 | 2008-06-20 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-24 | Yes | No | Reuse | 2008
42020503 | 2008-05-27 | Malfunction |  | It was reported during a da vinci s myomectomy procedure metal shards were observed on the fenestrated bipolar forceps instrument. When the site attempted to remove the instrument from the trocar, it would not come out, therefore, the instrument and trocar were removed concurrently from the pt. The planned surgical procedure was delayed, however, it was completed with a replacement. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned to isi for evaluation. A follow-up mdr will be submitted if the instrument is returned or if additional info is rec'd. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1129677 | 1063556 | 1021607 | 2955842200801086 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-06-20 | 2008-06-20 | No | Yes | Health Professional | 42020503 | No | No |  | Device Not Retu | No | Unkown | 2008
42018904 |  | Malfunction |  | It was reported that a cable on the fenestrated grasper instrument had become loose. No additional information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch cables to be loose at the wrist. The pitch input rotates 45 degrees before movement is translated to wrist. The housing was removed to find that both screws on the bottom pitch clamping pulley had become loose. The clamping pulley appears to have rotated slightly on the input shaft, creating decreased cable tension. Engineering concluded that the screws may not have been fully tightened during assembly or may have became loose from vibrations in ultrasonic bath. Engineering also found the distal end of the main tube to have a. 150" long deep scratch with material removed. The scratch is located 2. 15" above the proximal clevis and is perpendicular to the tube axis, possibly caused by an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1319166 | 1063557 | 1021608 | 2955842200801080 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-20 | 2008-06-20 | No | Yes | Service Personnel | 42018904 | Yes | No | 2008-05-24 | No Answer Provi | No | Initial | 2008
42009307 |  | Malfunction |  | It was reported that fiberglass was coming off the endowrist prograsp instrument. Nothing was observed falling into the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed that the distal end of the main tube has various scratch marks concentrated in a 1. 5 inch long area, directly above the proximal clevis. A couple of the scratch marks were deep and had material removed. The scratches are randomly oriented relative to the tube axis, suggesting that they were most likely caused by instrument collisions. Conclusions:the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1328002 | 1063558 | 1021609 | 2955842200801092 | 1 | GCJ | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2008-06-20 | 2008-06-20 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the cables were sticking out and fiberglass was coming off the maryland bipolar forceps instrument. Nothing was observed falling into the pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the bipolar insert was missing from the instrument back end. As a result, the bipolar pins were no longer retained and the conductor wires stick out of the chassis. Engineering also observed that one side of the distal end of the main tube has various deep scratches with material removed. Based on the location and appearance, engineering determined that the scratches were most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1320138 | 1063559 | 1021610 | 2955842200801091 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-06-20 | 2008-06-20 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not opening and closing. It was also reported that the shaft was scratched and that no pieces of the instrument fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to pas recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Engineering also found the distal end of the main tube to have a. 250" long deep scratch above the proximal clevis with material removed. The scratch is not oriented parallel the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1296350 | 1061496 | 1019547 | 2955842200801083 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-16 | 2008-06-16 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
40018005 | 2008-05-22 | Other |  | It was reported that during a da vinci s total abdominal hysterectomy procedure while dissecting the obturator nerve, the surgeon noticed a spark from the proximal portion of the monopolar curved scissors (mcs) instrument. A thermal burn was discovered on the pt's external iliac artery. The surgeon immediately removed the mcs instrument and a stent was placed in the pt's iliac artery. The planned surgical procedure was completed. No add'l pt harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for eval, however, pictures were provided by the customer. Based on the pictures provided engineering found two holes burned through the silicone, indicative of arcing. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. The root cause of the customer reported failure mode cannot be confirmed based on the photographs. If the tip cover or mcs instrument is returned for eval, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1306468 | 1061709 | 1019760 | 2955842200801079 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-06-13 | 2008-06-13 | Yes | Yes | Health Professional | 40018005 | No | No | 2008-05-23 | Device Not Retu | Yes | Initial | 2008
IS2000 A5 1P4 | 2008-05-22 | Other | Other | It was reported while setting up the da vinci s surgical system for a prostatectomy procedure the customer experienced system error code #297. The customer restarted the system, however, system error code #297 recurred upon power up. The pt was under anesthesia and the port incisions had been created when the site decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system error code #297 experienced by the customer was associated with a configured embedded serializer setup joint (cfg, essj). The embedded serializer for setup-joint is the printed circuit assembly (pca) inside a system arm that monitors the potentiometer for each of the four joints and their associated backup potentiometers. The system was repaired by replacing the affected cfg essj and reconnecting the low voltage differential signal (lvds) cable tp the pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23 indicates a communication failure between the surgeon console and the pt side cart and this communication failure caused an incomplete pca software download which created the persistent fault condition error code of #297. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1136340 | 1061710 | 1019761 | 2955842200801077 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-13 | 2008-06-13 | Yes | Yes | Health Professional | IS2000 A5 1P4 | Yes | No | 2008-05-22 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-05-19 | Other | Other | It was reported that during a da vinci s beating heart tecab procedure, approx five hours into the procedure, the surgeon stated that a pt side manipulator (psm) arm felt "jerky". The site converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. The customer contacted isi technical support after the procedure was complete. | The investigation conducted by field service engineering found no errors in the da vinci s system log. Performance tests for the specific pt side manipulator (psm) arm were conducted and the customer reported problem could not be duplicated. The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. Engineering concluded that the instrument inserted in the psm was most likely at the end of the cannula accessory and the "jerky" motion was due to this poor positioning. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1306457 | 1061711 | 1019762 | 2955842200801073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-13 | 2008-06-13 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that before a procedure the da vinci s surgical system was not recognizing the endowrist prograsp forceps instrument and nothing fell into the pt. The planned surgical procedure was completed with another instrument of the same type. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering unable to replicate the issue when conducting tests that found the instrument was successfully recognized by the system. The pogo pins do not stick and are not contaminated. The instrument was able to move intuitively with full range of motion in all directions and the grips opened and closed properly. Engineering also observed that distal end of the main tube has two deep scratch marks and a scrape mark with material removed. Scratches are not aligned with tube axis. Based on the location and appearance, the scratches are most likely caused by instrument collisions. The scrape mark is 180 degrees from the scratches. The damaged area is 1. 2 inches long and parallel to the tube axis. Based on the location and appearance, the scrape was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1132839 | 1061783 | 1019833 | 2955842200801069 | 1 | GCJ | Manufacturer | DistributorOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not recognized by the da vinci s system. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that recognition passed. Engineering also found the distal end of the main tube has a. 75 inch long section below the mid point with various scratches and light material removed. Based on the location and appearance, the scratch is most likely due to instrument collisions. A scrape mark below the mid point of the distal end of the main tube was also observed, most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1314257 | 1061784 | 1019834 | 2955842200801067 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that the cable of the fenestrated bipolar forceps instrument "would not engage in arm". No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that the da vinci s system intermittently recognized the instrument. Engineering determined a mechanical interaction between the chassis, input discs, and sterile adapter is resulting in slight positioning variations of the pogo pins and causing the intermittent recognition issues. Engineering also found the distal end of the main tube has deep scratches and scrape marks, approx 180 degrees apart, with material removed. The scrape mark is 1. 25 inches long and parallel to the tube axis. Based on the location and appearance the scrape mark most likely caused by a cannula accessory. The other side of the tube has three adjacent deep scratches perpendicular to the tube axis, likely caused by instrument collisions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1135857 | 1061785 | 1019835 | 2955842200801065 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument was not recognized by the da vinci s system. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that recognition passed. Engineering also observed the distal end of the main tube has a couple of sections with deep scratches. The scratches are 180 degrees apart and have material removed. Based on the location and appearance, the scratch is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as the disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1314437 | 1061786 | 1019836 | 2955842200801070 | 1 | GCJ | Manufacturer | OtherForeign | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42019006 |  | Malfunction |  | It was reported that the jaws of the cobra grasper instrument stopped opening and closing, however, this did not cause any delays in the procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Engineering also observed that the distal end of the main tube has various scratches. A couple of the scratches are deep enough to have material removed. The scratches are short, narrow and not aligned with the tube axis. Based on the location and appearance, the damage was most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1316177 | 1061788 | 1019838 | 2955842200801063 | 1 | GCJ | Manufacturer | OtherCompany Representative | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42019006 | Yes | No | 2008-05-17 | Yes | No | Initial | 2008
42022702 |  | Malfunction |  | It was reported that the cadiere forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions with the grips opening and closing properly. Engineering also found the distal end of the main tube to have a sections, 180 degrees apart, with material removed on one side of the tube. The damaged areas are 2" and 2. 5" long, parallel to tube axis and have a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted fi additional information is received. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1316459 | 1061789 | 1019839 | 2955842200801074 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-16 | 2008-06-13 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-20 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the wrist of the permanent cautery hook instrument did not move wrist and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the wrist functioned properly. The instrument was also successfully recognized on the test system and moved intuitively with full range of motion in all directions. Engineering also found the distal end of the main tube has a deep scratch with material removed. The damaged area is located. 150 inches above the proximal clevis and is not aligned with the tube axis. Based on the location and appearance, the scratch is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1316480 | 1061790 | 1019840 | 2955842200801071 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that prior to starting a da vinci prostatectomy procedure, system was not recognizing the bipolar maryland forceps instrument and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that the instrument was successfully recognized the system. The pogo pins do not stick and are not contaminated. The instrument was able to move intuitively with a full range of motion in all directions and the grips opened and closed properly. Engineering also observed that the distal end of the main tube was a 1. 5 inch long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1316485 | 1061791 | 1019841 | 2955842200801072 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument has scratches on the shaft and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube has a couple of sections with light material removed roughly parallel to the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions or rough handling during the cleaning process. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as a disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1316498 | 1061792 | 1019842 | 2955842200801066 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-17 | Yes | No | Initial | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the cable of the monopolar curved scissors instrument broke, however, nothing fell into the patient. The planned surgical procedure was completed with a replacement instrument of the same type. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is broken at the scissor grip. The grip hub has a peak across the cable groove and the wear against this peak most likely contributed to the breakage. Engineering also observed that the main tube has various sections above and below midpoint with light material removed. The damaged areas are parallel to the tube axis. Based on the location and appearance, these scratches are most likely caused by a cannula accessory. The tube damage above the midpoint may have been caused by the cleaning process. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1302522 | 1061798 | 1019848 | 2955842200801064 | 1 | GEI | Manufacturer | OtherDistributor | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that prior to starting a da vinci s prostatectomy procedure, the system was not recognizing the maryland bipolar forceps instrument and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument was successfully recognized by the system, however, a couple of the pogo pins have friction when actuated and have debris/coating at the base of the pin. Which may have caused the pogo pins to stick, resulting in the recognition issues at the site. Engineering also observed that distal end of the main tube has a 1 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Below this section there is a section with various deep scratches concentrated on one side of the tube. Based on the location and appearance, the damaged area was most likely caused by instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1302541 | 1061799 | 1019849 | 2955842200801078 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-23 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that prior to a procedure the da vinci s surgical system was not recognizing the monopolar curved scissors instrument. The planned surgical procedure was completed with another instrument of the same type. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument was successfully recognized the system. The pogo pins do not stick and are not contaminated. The instrument was able to move intuitively with full range of motion in all directions. Engineering also observed that distal end of the main tube has a 1. 5 inch long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECOTROSURGICAL INSTRUMENT | 1302549 | 1061800 | 1019850 | 2955842200801068 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-19 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that a pk dissecting forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip to be severely bent, causing side to side misalignment of grips. There is a. 200" offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. The severity of the bend suggests some kind of mishandling/misuse had occurred. The distal end of the main tube has material removed in the form of a combination of deep scratches superimposed over a scrape mark. The scraping may have been caused by cleaning/handling. Scratches are likely from an instrument collision. The endowrist instruments instructions for use (ifu) specifically states" general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1302578 | 1061801 | 1019851 | 2955842200801075 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-06-13 | 2008-06-13 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-20 | Yes | No | Reuse | 2008
42022702 | 2008-07-14 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the surgeon was taking down the endopelvic fascia when shavings from the pk dissecting forceps instrument were noted inside of the pt. The procedure was delayed by 30 mins to remove a majority of the shavings and the planned surgical procedure was completed with a replacement instrument and cannula accessory. No pt harm, adverse outcome or injury was reported. | The instrument and cannula accessory have not yet been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A f/u mdr will be submitted if the instrument and cannula accessory are returned (upon completion of failure analysis) or if add'l info is rec'd. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1247879 | 1106896 | 1063026 | 2955842200801208 | 1 | GEI | Manufacturer | User facility | Other | Initial |  | 2008-06-11 | Yes | Yes | Health Professional | 42022702 | No | No | 2008-07-14 | Device Not Retu | No | Unkown | 2008
40015201 |  | Other | Other | It was reported that while prepping the instruments to be used during a da vinci s beating heart cabg surgical procedure utilizing the endowrist stabilizer, the scrub nurse observed that the 15 degree blue blade accessory fit loosely into the snap-fit instrument. Another blade accessory of the same type was opened and attached to the instrument, however, it also had a loose fit. The surgical staff decided to move forward with the procedure using the accessory as is. The arteriotomy was performed with the blade accessory in the instrument, however, the blade became dislodged and fell into the pt during instrument removal. Since the heart was ischemic at the time of the dislodgement, the blade was left to be retrieved at a later time. During retrieval, the blade accessory was not readily visible and a two hour search was performed before the blade accessory was found and removed. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | A blade accessory from the same lot used during the event was returned and evaluated. Per the customer reported complaint, engineering observed that the retention legs of the blade accessory were not bent to specification, thereby, causing the loose fit experienced by the customer and allowing the blade to detach from the instrument. | 15 DEGREE BLUE BLADE | ENDOSCOPIC INSTRUMENT ACCESSOR | 1254698 | 1106898 | 1063028 | 2955842200801207 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial |  | 2008-06-11 | Yes | Yes | Health Professional | 40015201 | No | No | 2008-07-14 | No | Yes | Initial | 2008
IS2000 A51P4 | 2008-05-13 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, while starting up the system, the customer experienced recurring system error code #25517. An isi technical support engineer had the customer turn off the system and reset the circuit breaker on the pt side cart (psc), however, when the system powered back up the system error code #25517 reappeared. The pt was under anesthesia for approx 30 minutes when the site decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system error code #25517 experienced by the customer was associated with a pt side manipulator (psm) embedded serializer set up joint (essj). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The essj was returned to isi for failure analysis investigation. System error code # 25517 appeared when the da vicini safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety system put davinci in a "recoverable safe state". Engineering concluded that the two low voltage differential signaling connectors for the essj malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1149457 | 1059611 | 1017702 | 2955842200801058 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-10 | 2008-06-10 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-05-15 | Other | Other | It was reported that prior to starting a da vinci s cardiac surgical procedure, while starting up the sys, the customer experienced recurring sys error code #25589. An isi technical support engineer had the customer turn off the sys and reset the circuit breaker on the pt side cart (psc), however, when the sys powered back up the sys error code #25589 reappeared. The pt was under anesthesia for approx one hr when the site decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | Results, conclusions: the investigation conducted by field svc engineering concluded that the sys fault experienced by the customer was associated with the remote arm controller (rac) board. The sys was repaired by replacing the affected rac. The rac is an electronic module comprised of four printed circuit assemblies and cooling fans. Sys error code #25589 appears when the da vinci s safety sys determines during the power up self test that the remote arm controller board (rac) brakes failed the brake voltage test. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The rac was returned to isi for failure analysis investigation. Inspection was performed and two low voltage differential signaling connectors for the rac malfunctioned, thus generating the sys errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1297477 | 1059612 | 1017703 | 2955842200801059 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-06-10 | 2008-06-10 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-05-15 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has fiber glass shavings "peeling" off of the shaft. This was not identified during a procedure, however, a surgical staff member ran her hand down the shaft and received a splinter of the shaft material in her finger. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has a 1. 5 inch long section with light material removed on one side. The damage area is located 3 inches above the proximal clevis and is parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1291641 | 1059011 | 1017102 | 2955842200801056 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-06 | 2008-06-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-01 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the cable of the pk dissecting forceps instrument is broken and the tip will not close. There was no pt impact, nothing fell into pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that none of the mechanical drive cables are broken. One conductor wire is broken at the yaw pulley exit and the yaw pulley is missing a piece opposite of the conductor break which is allowing the grip to move within the pulley and have a larger range of motion in the yaw. Engineering concluded that the added range of motion of the grip likely created excess tension on the wire causing it to break. The wire was likely not attached securely to the grip, had poor strain relief and pulled out during articulation of the wrist. Engineering also found the distal end of the main tube has a 1. 5 inch section with light material removed on one side of the tube. The damage area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1291557 | 1059012 | 1017103 | 2955842200801054 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-09 | 2008-06-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-14 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is not recognized by the da vinci s system. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a system and recognition passed. Engineering was unable to replicate customer reported recognition problem. Engineering also observed the distal end of main tube has various scratches and one small scrape mark with material removed. The scratches are randomly oriented relative to the tube axis, suggesting they were caused by instrument collisions. The scrape mark is. 850 inches long and. 100 inches wide and is parallel to the tube axis. Based on the location and appearance, the scrape mark damage was most likely caused by a cannula accessory. No other damage found. The endowrist instruments instructions for use (ifu) specifically states; general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1283543 | 1059013 | 1017104 | 2955842200801053 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-06 | 2008-06-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-12 | Yes | No | Reuse | 2008
42017908 | 2008-01-03 | Malfunction |  | It was reported that during a low anterior resection da vinci s procedure, the monopolar curved scissors instrument wrist cable broke. The instrument was immediately replaced with another instrument of the same type. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is frayed at the scissor grip. Engineering concluded that the frayed cable was caused by a peak across the cable groove at the grip hub. Engineering also observed the distal end of main tube has a 1. 5 inch long section with light material removed on one side. The damaged area is located 3 inches above the proximal clevis and is parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1283517 | 1059014 | 1017105 | 2955842200801049 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-06-06 | 2008-06-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-17 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument does not work. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire is broken at the yaw pulley exit. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. The added range of motion of the grip likely created excess tension on the wire, causing it to break. Engineering concluded that the wire was not securely attached to the grip causing poor strain relief and pulled out during articulation of the wrist. Electrical continuity failed. Engineering also observed the distal end of main tube has a couple of sections with material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damages found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1283579 | 1059015 | 1017106 | 2955842200801048 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-06 | 2008-06-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument is scratched near the instrument tip. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube has a 1. 9 inch long by. 260 inch wide section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1283601 | 1059016 | 1017107 | 2955842200801050 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-06-09 | 2008-06-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
42020906 |  | Malfunction |  | It was reported that 1. 5 hours into a da vinci s surgical procedure, two small wires from the suturecut needle driver instrument fell into the pt. Both wires were removed from the pt, and the instrument was removed and replaced with an instrument of the same type. The planned surgical procedure was completed without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the broken instrument components were successfully retrieved from the pt. A follow-up mdr will be submitted if the instrument is returned or if add'l info is received. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1283558 | 1059017 | 1017108 | 2955842200801055 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-06-09 | 2008-06-09 | No | Yes | Health Professional | 42020906 | No | No |  | Device Not Retu | No | Unkown | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is dull. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and confirmed the scissors do not cut cleanly through. 006 inches of latex. The latex gets snagged at the midpoint of the blade. Engineering observed that one blade edge has a small indentation in curvature at the midpoint which was negatively affecting performance. Engineering also found the distal end of the main tube has a 1. 5 inch section with light material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1283613 | 1059018 | 1017109 | 2955842200801051 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-06-09 | 2008-06-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument does not grasp well. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the grip open angle was larger than usual and the grasping force was slightly below average. Engineering concluded that the cables may be stretched, reducing the grasping force. Engineering also found the distal end of the main tube has a 1. 5 inch section with light material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1283625 | 1059019 | 1017110 | 2955842200801052 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2008-06-09 | 2008-06-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-05-13 | Yes | No | Reuse | 2008
42017904 | 2008-05-27 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument would not close completely. The surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The instrument has not been returned to isi for evaluation. A follow up mdr will be submitted if the instrument is returned or if additional info is received. As of may 27, 2008, there have been no reported recurrences of the issue at this hosp. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1094222 | 1059606 | 1017697 | 2955842200801060 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-06-09 | 2008-06-09 | Yes | Yes | Health Professional | 42017904 | No | No |  | Device Not Retu | No | Unkown | 2008
IS1200 A43P9 | 2008-05-16 | Other | Other | It was reported that during a da vinci hysterectomy surgical procedure the surgeon was unable to move the one of the instrument arms. An isi field service engineer attempted to troubleshoot the issue via the telephone, however, the surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No patient harm or injury was reported. | The investigation conducted by isi field service engineer discovered that the patient side manipulator (psm) issue (instrument arm) experienced by the customer was caused by the psm being positioned to close to the end of range motion. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. As of may 16, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1094221 | 1059607 | 1017698 | 2955842200801061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-06-09 | 2008-06-09 | Yes | Yes | Health Professional | IS1200 A43P9 | No | No |  | Yes | No | Reuse | 2008
IS1200 A43P8 |  | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure, the grip spring of the left master tool manipulator (mtml) broke and fell out of grip assembly. The mtml grip assembly was repaired by the site's biomedical engineer with the assistance of an isi technical support engineer and the procedure was continued. Later in the procedure, the grip assembly again came apart and the surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | Results & conclusions - per the customer reported complaint, the investigation conducted by the field svc engineer confirmed that the grip assembly of the left mater tool manipulator (mtml) was broken. The master tool manipulator refers to the master manipulators at the surgeon's console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The sys was repaired by replacing the affected mtml. The (mtml) was returned to isi for eval and engineering observed that the grip springs had fallen out of the grip assembly. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1288700 | 1059608 | 1017699 | 2955842200801057 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-06-09 | 2008-06-09 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2008-06-03 | Yes | No | Reuse | 2008
40016906 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the harmonic curved shears insert accessory broke and a components fell into the patient. The broken pieces were retrieved and the planned surgical procedure was completed with a 5-10 minute delay. No patient harm, adverse outcome or injury was reported. | The instrument accessory has been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (upon completion of failure analysis) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1278933 | 1058115 | 1016226 | 2955842200801062 | 1 | LFL | Manufacturer | User facility | UNKNOWN | Initial | 2008-06-06 | 2008-06-06 | No | Yes | Service Personnel | 40016906 | No | No | 2008-05-02 | Device Not Retu | Yes | Initial | 2008
37076104 | 2008-04-01 | No Answer Provided |  | In attempting to dock the intuitive robot for a robotic-assisted laparoscopic prostatectomy, the patient cart would go forward a few inches, then alarm, and the screen would read "fault. " it was determined that the fault was due to either a short in the red cable between the cart and the console or by dust in the cable connections. All connections were cleaned and the robot worked without fault. |  | DA VINCI | ROBOTIC SURGICAL SYSTEM | 1025862 | 1058634 | 1016724 | 1058634 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2008-06-03 | 2008-06-03 | No | No | Nurse | 37076104 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
| 2008-03-05 | Injury | HospitalizationOtherLife ThreateningRequired Intervention | Failure of davinci robot prior to surgery in '08 and during surgery in the next day, requiring surgeon to abandon procedure and perform sternotomy to complete coronary artery bypass grafting (cabg). Dates of use: 2008. Diagnosis: cabg. Event abated after use stopped? yes. |  | DAVINCI | DAVINCI SURGICAL ROBOTIC SYSTE | 1025950 | 1055496 | 1013627 | MW5007100 | 1 | NAY | Voluntary |  | Other | Initial |  | 2008-06-02 | Yes | Yes | Health Professional |  | Yes | No |  |  |  |  | 2008
42019006 |  | Malfunction |  | It was reported that the shaft of the cobra grasper instrument appears to be scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has various deep scratches with material removed. The scratch marks are all under. 300" long and are randomly oriented with respect to the tube axis. Engineering concluded that the scratches were most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1050184 | 1054815 | 1012965 | 2955842200801044 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42019006 | Yes | No | 2008-05-08 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the shaft of the tenaculum forceps instrument appears to be scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube has a couple of light scratches and a few small sections with material removed. The damaged areas are roughly. 5 inches long or less, with a rough surface finish, but are more narrow than cannula related tube abrasions. Based on the location and appearance, the scratches are most likely due to instrument collisions. Engineering also observed that one pitch up cable is broken at the distal clevis. Engineering concluded that the breakage may be due to lifting of heavy object. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1050162 | 1054816 | 1012966 | 2955842200801043 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-05-08 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | It was reported that during a da vinci s prostectomy procedure, shavings from the cannula fell into the pt. All fragments were suctioned out of the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The accessory was returned and evaluated. Per the customer reported complaint, engineering found that the cannula is not bent at distal tip. A. 342 gage pin fits through the inner diameter and does not have any burrs. The tube and base do not separate at the weld. Engineering concluded that the material that fell into the pt is most likely from an instrument main tube abrasion and not from the cannula. High loading on the instrument can cause tube pressure against the inner edge of the cannula resulting in scraping. No other damage found. | 8MM CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 1049928 | 1054817 | 1012967 | 2955842200801040 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42000202 | Yes | No | 2008-05-15 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar forceps instrument appears to be scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has two scratch marks with a small amount of material removed. The scratches are both under. 225 inches long and perpendicular to the tube's longitudinal axis. Engineering concluded that the scratches were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1050156 | 1054819 | 1012968 | 2955842200801042 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-05-08 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the fenestrated maryland bipolar instrument is broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed and the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed the distal end of the main tube has a single deep scratch with material removed. The scratch is. 150 inches long and perpendicular to tube axis. Based on the location and appearance, the scratch is most likely due to an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1049931 | 1054820 | 1012969 | 2955842200801038 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-07 | Yes | No | Reuse | 2008
42017908 | 2008-05-06 | Malfunction |  | It was reported that during a da vinci s surgical procedure a burned area on the "shaft" of the monopolar curved scissors (mcs) instrument was found. The mcs instrument was immediately removed and replaced with another instrument of the same type to complete the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 070 inch by. 100 inch oblong hole burned through the main tube. The hole is located roughly. 375 inches above the tube extension. A ring of tube material is removed around the perimeter of the hole, indicative of arcing, and the main tube has a longitudinal crack that bisects the hole. The instrument was disassembled to find the hypotubes charred in the same location as the hole. The tube crack starts at one of the slots at the distal end of the tube. Engineering has determined that the crack in the tube and high generator settings likely created a path for arcing to occur. The source of the crack is unclear. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1049936 | 1054821 | 1012970 | 2955842200801046 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-05-06 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported during cleaning the tip of the fenestrated bipolar maryland instrument was found to be charred. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found both yaw pulleys have charring at the grip base and the inner mating surfaces above the cable crimp. Material has been removed from yaw pulleys due to burning. Engineering concluded the damage is likely caused by repeated or prolonged activation of cautery without tissue between grips. High generator settings may also contribute to melting. Engineering also observed that the distal end of main tube has a 2. 75 inch long section directly above the proximal clevis with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | FENESTRATED BIPOLAR MARYLAND INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1049935 | 1054822 | 1012971 | 2955842200801039 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-05-08 | Yes | No | Reuse | 2008
40020705 | 2008-05-06 | Malfunction |  | It was reported that during a da vinci myomectomy procedure two "green pea size" pieces from the tenaculum forceps instrument shaft fell into the pt. Both broken instrument pieces were found removed from the pt, causing approx a 5 minute delay, and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces. The clevis is missing 3/4 of the section that mates with the main tube. The size of the missing section suggests the clevis broke in more than one piece, likely in two pieces, referred to by the customer as "green pea sized" roughly. 200 inches by. 160 inches. Based on engineering's eval, a likely failure mode was overloading of the instrument at the tip, causing the breakage. Per the case notes, the broken instrument components were successfully retrieved from the pt. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1049930 | 1054823 | 1012972 | 2955842200801047 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 40020705 | Yes | No | 2008-05-06 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar forceps instrument appears to be scratched. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has a couple of sections, 180 degrees apart, 2 inches and 3 inches long, with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELETROSURGICAL INSTRUMENT | 1050150 | 1054824 | 1012973 | 2955842200801041 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-05-08 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument has a scratched shaft. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube has various deep scratches with material removed on one side of the tube. The scratches are all under. 300 inches in length and are randomly oriented to the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnections of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1050148 | 1054825 | 1012974 | 2955842200801045 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-30 | 2008-05-30 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-05-07 | Yes | No | Initial | 2008
42020704 |  | Malfunction |  | It was reported that wires were coming out of the tenaculum forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is broken at the distal idlers allowing the pulley to spin freely. The other cables at the wrist are not damaged. Engineering concluded the cable broke under tensile loading and that the cable wear on pulleys combined with the high grasping force most likely contributed to the breakage. Engineering also observed the distal end of main tube has a 2 inch long section directly above the proximal clevis with material removed on one side of the tube. The damaged is parallel to the tube axis with a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1077817 | 1052757 | 1010908 | 2955842200801025 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-04-28 | Yes | No | Reuse | 2008
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a crack on the base. It was also reported that no instrument components fell inside of the pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated, per the customer reported complaint, engineering found one ear of the proximal clevis is missing and exposing the proximal pulleys. The size of the missing piece is approximately. 180 inches by. 210 inches. The clevis is not dislodged from the main tube and the interface between the two components is undamaged. Engineering concluded that the ear may have broke due to high impact force on the outside of the ear, possibly from an instrument collision. Engineering also observed the distal end of main tube has material removed from a combination of scratches and scrape marks parallel to the tube axis. Based on the location and appearance the scratches are likely from instrument collisions while the scrape marks are consistent with cannula accessory tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1050025 | 1052758 | 1010909 | 2955842200801030 | 1 | GEI | Manufacturer | Company RepresentativeOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42021403 | Yes | No | 2008-03-13 | Yes | No | Reuse | 2008
40017208 |  | Malfunction |  | It was reported that the tip of the fenestrated maryland bipolst instrument is broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal tip of the instrument is undamaged. There is one bipolar pin and plug riser insert that are missing from the back end of the instrument. The pin was likely pulled off during removal of a tight fitting bipolar cord. The other pin is properly crimped. Engineering also observed the distal end of the main tube has a couple of deep scratches with light material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instruction for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1050145 | 1052759 | 1010910 | 2955842200801024 | 1 | GEI | Manufacturer | Other | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 40017208 | Yes | No | 2008-04-28 | Yes | No | Reuse | 2008
40017908 | 2008-04-21 | Malfunction |  | It was reported that during a da vinci prostectomy procedure, the tips of the monopolar curved scissors instrument were dull. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument scissors cut cleanly through. 006" latex. Blades are not damaged. Electrical continuity passed. Engineering also observed the distal end of the main tube has a 3. 5 inch long section with material removed on one side of the tube. The damaged area starts 1. 25 inches above the tube extension and is parallel to the tube axis with a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1065789 | 1052760 | 1010911 | 2955842200801027 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial |  | 2008-05-27 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-04-28 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the housing of the endowrist prograsp forceps instrument is broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument was returned with the housing and levers missing. Engineering concluded that the housing and levers likely came off due to mishandling of instrument. Engineering also observed that distal end of main tube has a 2. 5 inch long by. 125 inch wide section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1050245 | 1052761 | 1010912 | 2955842200801026 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-05-01 | Yes | No | Initial | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a broken wire. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found none of the cables or wires are broken, however, one grip close cable is frayed near the distal idlers. The grips can still open and close. The cable may have been nicked by another instrument, causing the breakage. Engineering also observed that the distal end of main tube has a 5 inch long section directly above the proximal clevis with light material removed all around the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1360142 | 1052762 | 1010913 | 2955842200801033 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-04-10 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | The maryland bipolar forceps instrument was returned with no information. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley shows no signs of arcing. The yaw pulley is missing a piece opposite the conductor break which is allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that the added range of motion of the grip likely created excess tension on the wire causing it to break. Engineering also observed that the distal end of main tube has two sections, 180 degrees apart, with material removed all around the tube. The damaged areas are 1. 25 inches and 2 inches long and are parallel to tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1359965 | 1052763 | 1010914 | 2955842200801034 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-30 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is broken at the distal idlers and the pulley spins freely. Engineering concluded that the cable broke under tensile loading. Other cables at the wrist are not damaged. The cable wear on the pulleys with high grasping force likely contributed to breakage. Engineering also observed that the distal end of main tube has a 1. 75 inch long section with material removed. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1359910 | 1052764 | 1010915 | 2955842200801035 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-04-30 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument has shavings coming off. This was not identified during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found distal end of main tube has a 1 inch long by. 065 inch wide section with material removed. The damaged area is located. 850 inches above the tube extension and is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Engineering also observed that the scissor crossover at the tips is less than usual, but the blades are not damaged. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1358844 | 1052765 | 1010916 | 2955842200801029 | 1 | GEI | Manufacturer | OtherHealth Professional | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-05-01 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a broken wire. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the pitch up cable is broken at the distal clevis. The other cables at the wrist are not damaged. Engineering concluded that the cable may have broke due to repeated lifting of heavy tissue mass. Engineering also found the distal end of the main tube has various deep scratches with light material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1364689 | 1052766 | 1010917 | 2955842200801028 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-23 | 2008-05-27 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-04-28 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-04-29 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure the patient side manipulator (psm) was not moving correctly. An isi representative was present for the case and began to troubleshoot. The system was restarted however, the psm fault recurred. The customer decided to complete the planned surgical procedure laparoscopically. An additional incision was made for the atrial retractor and the patient was under anesthesia for approximately an additional 20 minutes. No patient harm was reported. | The investigation conducted by field service engineer concluded that fault experienced by the customer was associated with a patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generate error code) functioned as designed and there was no injury to the patient. The psm was returned to isi for failure analysis investigation. Engineering evaluation found two loose wrist drive cables on axis 4 and determined that they were the cause of the system error experienced by the customer. The psm was repaired by replacing the wrist drive cables on axis 4 as well as the wrist drive cables on axis 5, 6 and 7 as a precautionary measure. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1048360 | 1053081 | 1011232 | 2955842200801036 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-05-23 | 2008-05-27 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-04-29 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-02-26 | Other | Other | It was reported that prior to beginning a da vinci s surgical procedure, the system surgeon side console (ssc) view was unstable and 3d calibration could not be performed. The patient was under anesthesia, however, no surgical tasks had been performed when the surgeon decided to postpone the procedure. The procedure was completed the following day with the da vinci s surgical system. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found a small cut on one of the leads of the camera cable and therefore concluded that the unstable view experienced by the customer was associated with the camera cable. The camera cable attaches the camera to the camera control unit (ccu), focus control unit (fcu), and the vision cart. The system was repaired by replacing the affected camera cable. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1048684 | 1053086 | 1011237 | 2955842200801002 | 1 | NAY | Manufacturer | User facilityForeign | Other | Initial | 2008-05-16 | 2008-05-27 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No | 2008-02-26 | Device Not Retu | No | Reuse | 2008
IS1200 A43P9 | 2008-05-02 | Other | Other | It was reported that during a da vinci surgical procedure the customer experienced persistant #20008 system error codes. An isi technical support engineer (tse) determined that the fault was related to a pt side manipulator (psm) arm and had the customer examine the psm. The customer found a loose cable. The pt was under anesthesia and port incisions were made when the surgeon decided to abort and reschedule the planned surgical procedure to a later date. | Conclusions - the investigation conducted by field svc engineering concluded that the sys fault experienced by the customer was associated with a pt side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The sys was repaired by replacing the affected psm. Results - the sys alarm (sys generated fault code) functioned as designed and there was no injury to the pt. The #20008 sys error codes appeared when the davinci safety sys determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a "recoverable safe state. " the psm was returned and evaluated. Engineering found the cable intact, however, it derailed at the carriage end of the axis. Engineering determined that drape material may gotten caught in between the pulley and the cable, causing the cable to derail from the carriage pulley when the drape was removed. The psm was repaired by replacing the two axis cables. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1050475 | 1053087 | 1011238 | 2955842200801037 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-23 | 2008-05-27 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-05-02 | Yes | No | Reuse | 2008
40019302 | 2008-04-24 | Other | Other | It was reported that during a da vinci s single vessel small thoracotomy procedure, the da vinci balloon port accessory "broke" and co2 began to leak from the cannula. The surgeon decided not to replace the balloon port accessory and complete the planned surgical procedure. The procedure was lengthened by approximately 30 minutes due to the co2 leakage. No patient harm, adverse outcome or injury was reported. | The balloon port accessory was discarded by the customer and therefore will not be returned for evaluation. The root cause of the customer reported failure cannot be determined. The da vinci balloon port instructions for use (ifu) specifically states: general precautions and warnings. Caution: over inflation of the balloon may cause it to rupture. Note that 4cc of fluid will completely inflate the balloon and further inflation is unnecessary and will increase the risk of balloon rupture. Insertion and removal of instruments through the cannula should be done carefully to prevent unintentional damage to the internal seals, which may compromise insufflation. | DA VINCI BALLOON PORT 12MM ECM | ENDOSCOPIC INSTRUMENT ACCESSOR | 1048681 | 1053088 | 1011243 | 2955842200801031 | 1 | GCJ | Manufacturer | Health ProfessionalCompany Representativ | Nurse | Initial | 2008-05-23 | 2008-05-27 | Yes | Yes | Health Professional | 40019302 | No | Yes | 2008-04-28 | Device Not Retu | Yes | Initial | 2008
40019302 | 2008-04-21 | Other | Other | It was reported that during a da vinci s single vessel small thoracotomy procedure, the da vinci balloon port accessory "broke" and co2 began to leak from the cannula. The surgeon decided not to replace the balloon port accessory and complete the planned surgical procedure. The procedure was lengthened by approximately 30 minutes due to the co2 leakage. No patient harm, adverse outcome or injury was reported. | The balloon port accessory was discarded by the customer and therefore will not be returned for evaluation. The root cause of the customer reported failure cannot be determined. The da vinci balloon port instructions for use (ifu) specifically states: general precautions and warnings. Caution: over inflation of the balloon may cause it to rupture. Note that 4cc of fluid will completely inflate the balloon and further inflation is unnecessary and will increase the risk of balloon rupture. Insertion and removal of instruments through the cannula should be done carefully to prevent unintentional damage to the internal seals, which may compromise insufflation. | DA VINCI BALLOON PORT 12MM ECM | ENDOSCOPIC INSTRUMENT ACCESSOR | 1048678 | 1053092 | 1011243 | 2955842200801032 | 1 | GCJ | Manufacturer | Health ProfessionalCompany Representativ | Nurse | Initial | 2008-05-23 | 2008-05-27 | Yes | Yes | Health Professional | 40019302 | No | Yes | 2008-04-28 | Device Not Retu | Yes | Initial | 2008
42000604 |  | Malfunction |  | It was reported that after a da vinci s surgical procedure, the surgical staff found a grey piece of the large needle driver instrument was broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument is missing both release levers and the two rear snap tabs and the adjacent pins are broken off, indicating instrument was mishandled. Housing likely popped off causing levers to fall out, and then the housing was then put back on without levers. Engineering also observed the distal end of the main tube has a 2 inch long section with material removed. The damaged area is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1045322 | 1046596 | 1004786 | 2955842200801012 | 1 | GCJ | Manufacturer | OtherUser facility | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42000604 | Yes | Yes | 2008-04-25 | Yes | No | Reuse | 2008
40018005 | 2008-04-21 | Malfunction |  | It was reported that during a da vinci surgical procedure while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing. The surgeon immediately stopped using and the mcs instrument and replaced the mcs tip cover accessory to complete the planned procedure. No pt harm, adverse outcome or injury was reported. The following med watch reports were sent to the fda regarding this event: 2955872-2008-01018, 2955872-2008-01019, 2955872-2008-01021. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a. 030 inch diameter hole in the silicone. The hole center is located. 220 inches above the silione to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and eval found char marks in approx the same location as the holes in the tip cover. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1045702 | 1046597 | 1004789 | 2955842200801020 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-04-21 | Yes | Yes | Initial | 2008
42017206 | 2008-03-12 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the fenestrated maryland bipolar instrument stopped working. The surgeon completed the procedure with another instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument moves intuitively in all directions. The cables feel slightly stretched due to heavy usage. Electrical continuity passed. The chassis is missing a piece below one lever, most likely from mishandling. Engineering also observed the distal end of main tube has a 6 inch long section directly above the proximal clevis with material removed. The damaged area is parallel to tube axis and has a rough surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1046789 | 1046598 | 1004788 | 2955842200801010 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-25 | Yes | No | Reuse | 2008
40018005 | 2008-04-21 | Malfunction |  | It was reported that during a da vinci surgical procedure while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing. The surgeon immediately stopped using and the mcs instrument and replaced the mcs tip cover accessory to complete the planned procedure. No pt harm, adverse outcome or injury was reported. The following med watch reports were sent to the fda regarding this event: 2955872-2008-01018, 2955872-2008-01020, 2955872-2008-01021. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a. 020 inch by. 40 inch oval shaped hole in the silicone. The hole center is located. 070 inches above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and eval found char marks in approx the same location as the holes in the tip cover. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1045661 | 1046599 | 1004789 | 2955842200801019 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-04-21 | Yes | Yes | Initial | 2008
42000604 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the surgical staff felt the cables of the large needle driver instrument were broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers and the pulley spins freely. Engineering concluded the cable broke under tensile loading. The other cables at wrist are not damaged. Cable wear on the pulleys combined with high grasping force likely contributed to breakage. Engineering also observed the distal end of the main tube has a 2 inch long section with material removed on the side of the tube. The damaged area is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted. If additional info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1046713 | 1046600 | 1004790 | 2955842200801013 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42000604 | Yes | Yes | 2008-04-25 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that after cleaning the monopolar curved scissors instrument frayed fiber glass was found at the distal end. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has a 3 inch long section with material removed. The damaged area covers the width of the tube and is parallel to the tub axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1045686 | 1046601 | 1004791 | 2955842200801015 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-24 | Yes | No | Reuse | 2008
40018005 |  | Malfunction |  | The monopolar curved scissors tip cover accessory was returned without any info. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering found various mechanical tears and holes burned through the silicone. The damage is contained within one half of the tip cover silicone, between the two parting lines. The holes are. 010 inches in size with varying shapes and are located. 200 inches to. 350 inches above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. Various mechanical tears were found in the general vicinity of the holes. Engineering concluded tht arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively,. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1046761 | 1046602 | 1004792 | 2955842200801001 | 1 | GEI | Manufacturer | Other | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-04-21 | Yes | Yes | Initial | 2008
40018005 |  | Malfunction |  | The monopolar curved scissors tip cover accessory was returned without any information. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was evaluated. Engineering found a. 050 inch long mechanical tear in the silicone. The silicone exhibits localized melting at the tear, suggesting an arcing event occurred. The tear is located roughly halfway between the parting lines and starts. 160 inches above the silicone to sleeve interface. Engineering concluded that the initial damage to the silicone may have been caused by a collision, which later became a path for energy to escape through. Arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1042378 | 1046616 | 1004806 | 2955842200801000 | 1 | GEI | Manufacturer | Other | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-04-21 | Yes | Yes | Initial | 2008
42009307 |  | Malfunction |  | An isi rep reported that the tip of the prograsp forceps instrument would not open or close. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted testing and found that the grips opened and closed properly. The instrument moved intuitively with full range of motion, in all directions. One of the rear housing snap tabs is broken off and the luer plate is dislodged from the chassis. Engineering concluded the housing damage was most likely due to mishandling. Engineering also observed the distal end of main tube has a couple of deep scratches with material removed. The scratches are not aligned with tube axis and were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1042040 | 1046617 | 1004807 | 2955842200801006 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-04-20 | Yes | No | Reuse | 2008
42017908 | 2008-03-24 | Malfunction |  | It was reported that during a prostatectomy procedure, the da vinci s system did not recognize the monopolar curved scissors instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the system failed to recognize the instrument. The "instrument not recognized" error message appeared on the screen during multiple instrument installation attempts. A couple of the pogo pins were stiff and required high friction to be actuated. This was most likely causing an intermittent electrical contact with the sterile adapter, resulting in the recognition issues. Engineering also observed the distal end of the main tube has a 3. 5 inch long section with material removed all around the tube. The damaged area is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042074 | 1046618 | 1004808 | 2955842200801009 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-22 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | An isi rep reported that the tip of the monopolar curved scissors instrument would not open or close. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips open and close properly. One grip is bent, causing side to side misalignment of grips. There is a. 050 inch offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, most likely indicating overloading at the tip. Engineering conducted testing and the recognition and engagement passed. Electrical continuity passed. Engineering also observe the distal end of main tube has a 1. 3 inch section with material removed on one side of the tube. The damaged area is 2 inches above the proximal clevis, parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042058 | 1046619 | 1004809 | 2955842200801005 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-20 | Yes | No | Reuse | 2008
42009307 | 2008-03-12 | Malfunction |  | It was reported that during cleaning the housing of the endowrist prograsp forceps instrument broke. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the luer plate is not dislodged from the chassis, however, a section of the outer wall on the chassis feature is broken. Engineering concluded the damage is most likely due to mishandling during cleaning and possibly from applying a non-axial force at the flush ports. Engineering also observed the distal end of the main tube has various sections with material removal. One. 75 inch long by. 135 inch wide section has created a localized flat spot in the round tube. The damaged area is parallel to tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. There is also a deep scratch with material removed located 3. 33 inches above the proximal clevis. The scratch is likely due to an instrument collision. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1045815 | 1046620 | 1004810 | 2955842200801011 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2008-04-25 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tip of the monopolar curved scissors instrument will not close. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that the scissors do not close smoothly due to damaged blades. There are momentary pauses in closing when blades hit a damaged section. One blade edge has various pitting related to chips that affect closing, however, both of the blades exhibit pitting. , engineering also observed the distal end of main tube has a 3. 5 inch long section with material removed. The damaged area is located. 750 of an inch above tube extension. It is parallel with the tube's longitudinal axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042064 | 1046621 | 1004811 | 2955842200801014 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-04-24 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the prograsp forceps instrument has a frayed cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one pitch cable frayed at the distal clevis. The fraying may be due to repeatedly grabbing and lifting large objects. Engineering also found the distal end of the main tube has a couple of deep scratches with material removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damge to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1042387 | 1046622 | 1004812 | 2955842200801016 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-04-24 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | An isi rep reported that the tip of the monopolar curved scissors instrument would not open or close. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted testing and found the grips opened and closed properly. The instrument moved intuitively with a full range of motion and in all directions. Engineering also observed the distal end of main tube has a 2 inch section just below the midpoint with material removed. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042050 | 1046623 | 1004813 | 2955842200801007 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-20 | Yes | No | Reuse | 2008
40018005 | 2008-04-21 | Malfunction |  | It was reported that during a da vinci surgical procedure while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing. The surgeon immediately stopped using the mcs instrument and replaced the mcs tip cover accessory to complete the planned procedure. No pt harm, adverse outcome or injury was reported. The following med watch reports were sent to the fda regarding this event: 2955872-2008-01018, -01019, -01020. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found two adjacent. 010 inch by. 020 inch oblong holes in the silicone. The hole centers are located. 320 and. 360 inches above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and eval found char marks in approx the same location as the holes in the tip cover. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1045820 | 1046624 | 1004789 | 2955842200801021 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-04-21 | Yes | Yes | Initial | 2008
42017206 |  | Malfunction |  | An isi rep reported that the connector on the fenestrated maryland bipolar instrument was broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the top plug riser pin is pushed into back end and the bottom pin is bent. A bipolar cord cannot be installed due to damaged. The chassis feature that retains bipolar insert and pins are broken. Engineering concluded the instrument was most likely mishandled, breaking the chassis feature. Engineering also observed the distal end of main tube has a 2 inch long by. 200 of an inch wide section with material removed. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042080 | 1046625 | 1004815 | 2955842200801008 | 1 | GEI | Manufacturer | Company RepresentativeOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-20 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that one of the jaws on the monopolar curved scissors instrument was not moving and that a cable snapped. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grip close cable is broken at the distal idlers with 1/3 of the cable broken at the scissor grip causing the pulley to spin freely. The cable groove of the grip has a sharp peak. Damage may be due to tensile loading and cable wear on the pulleys and grip, combined with high grasping force. Engineering also observed the distal end of the main tube has a 2 inch long section with material removed on one side. The damaged area is 2 inches above the tube extension, parallel to the axis, and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Additional investigation is under-way regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042373 | 1046626 | 1004816 | 2955842200801017 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-08 | Yes | No | Reuse | 2008
40017907 | 2008-04-21 | Malfunction |  | It was reported that during a da vinci surgical procedure while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing. The surgeon immediately stopped using and the mcs instrument and replaced the mcs tip cover accessory to complete the planned procedure. No pt harm, adverse outcome or injury was reported. The following med watch reports were sent to the fda regarding this event: 2955872-2008-01019, 2955872-2008-01020, 2955872-2008-01021. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found three areas with char marks: the proximal clevis body, one proximal clevis ear, and the edge of the distal clevis. The char marks are indicative of multiple arcing events. Three tip cover accessories were returned with instrument. All tip covers were found to have arced holes through the silicone. The holes in the tip covers align with the different char marks on the instrument. Instrument does not have any abnormally sharp edges. Engineering also observed the distal end of the main tube has a 1. 25 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1364636 | 1046929 | 1005119 | 2955842200801018 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-16 | 2008-05-16 | No | Yes | Health Professional | 40017907 | Yes | No | 2008-04-21 | Yes | No | Reuse | 2008
40018005 | 2008-04-22 | Other | Other | It was reported that during a da vinci s hysterectomy procedure while using the monopolar curved scissors instrument ( mcs) with the tip cover accessory installed, the surgeon noticed arcing and a burn on the back side of the pt's uterus. The surgeon immediately stopped using the mcs instrument and replaced the mcs tip cover accessory, however, a few moments later the surgeon noticed a similar arcing event. The planned procedure was completed with approx a 90 min delay. No additional pt harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda regarding this event. 2955842-2008-01022. The endowrist instruments instructions for use ( ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and evaluation found a char mark in approx the same location as the tears in the tip cover. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a. 015 inches by. 030 inch oval shaped hole in the silicone. The hole is located. 160 inches above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. The tip cover also has various mechanical tears, likely due to instrument collisions, roughly 90 degrees from the hole. The endowrist instruments instructions for use ( ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and evaluation found a char mark in approx the same location as the tears in the tip cover. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1044637 | 1047898 | 1006089 | 2955842200801023 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-05-16 | 2008-05-16 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-04-23 | Yes | No | Reuse | 2008
40018005 | 2008-04-22 | Other | Other | It was reported that during a da vinci s hysterectomy procedure while using the monopolar curved scissors instrument ( mcs) with the tip cover accessory installed, the surgeon noticed arcing and a burn on the back side of the pt's uterus. The surgeon immediately stopped using the mcs instrument and replaced the mcs tip cover accessory, however, a few moments later the surgeon noticed a similar arcing event. The planned procedure was completed with approx a 90 min delay. No additional pt harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda regarding this event. 2955842-2008-01023. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found three adjacent holes burned through the silicone, ranging in size from. 015 inches to. 030 inches. The holes are located from. 140 inches to. 220 inches above the silicone to sleeve interface and the silicone exhibits localized melting around the hole, indicative of arcing. There is a smaller arced hole with an adjacent mechanical tear located 60 degrees from the holes and. 385 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears near the tip, likely due to instrument collisions, roughly 90 degrees from the hole. The endowrist instruments instructions for use ( ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The mcs instrument allegedly involved was returned and evaluation found a char mark in approx the same location as the tears in the tip cover. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1044626 | 1047901 | 1006089 | 2955842200801022 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-05-16 | 2008-05-16 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-04-23 | Yes | No | Reuse | 2008
40018005 | 2008-03-20 | Other | Other | It as reported that during a da vinci nephrectomy procedure while cauterizing tissue, the surgeon noticed sparks coming from the monopolar curved scissors (mcs) instrument. Afterwards a burn mark was observed on the pt's kidney. The surgeon immediately removed the mcs instrument and replaced the tip cover accessory. The planned surgical procedure was completed with the replacement tip cover, with approx a four minute delay. No add'l pt harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the tip cover is returned for eval a f/u mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1041114 | 1047902 | 1006090 | 2955842200801003 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-05-16 | 2008-05-16 | Yes | Yes | Health Professional | 40018005 | No | No | 2008-04-24 | Device Not Retu | Yes | Initial | 2008
IS1200 A43P9 | 2008-04-21 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the system surgeon side console (ssc) view was cloudy due to too much condensation. The customer tried multiple endoscopes to resolve the issue, however, the cloudy image remained. An isi technical support engineer suggested that the customer clean the camera head lense, however, the procedure was converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by isi field service engineering (fse) found a faulty bifurcated light guide cable. The bifurcated light guide cable is connected to the illuminator via a single connection allowing one cable to support two light ports on the endoscope. The system was repaired by replacing the affected bifurcated light guide cable. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1041108 | 1047906 | 1006094 | 2955842200801004 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-05-16 | 2008-05-16 | Yes | Yes | Health Professional | IS1200 A43P9 | No | No | 2008-04-21 | Device Not Retu | No | Reuse | 2008
IS1200 A43 P8 |  | Other | Other | It was reported that during a da vinci esophagactomy procedure, the customer experienced multiple occurrences of system error code #20008. The customer powered down and restarted the system. After restarting the system, the error continued to recur. The customer decided to convert to a laparoscopic technique to complete the planned surgery. The pt was under anesthesia for an add'l 40 minutes while the operating room was prepped for the laparoscopic procedure. No pt harm was reported. | The investigation conducted by field svc engineering concluded that system error code #20008 experienced by the customer was associated with the right master tool manipulator (mtmr). The sys was repaired by replacing the affected mtmr. The system alarm (system generate error code) functioned as designed and there was no injury to the pt. System error code #20008 appears when the da vinci s safety systems determined that the angular position of one or more robotic joint's on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering eval determined that a potentiometer assembly within the mtmr had malfunctioned, thus generating the sys error code experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1048249 | 1047874 | 1006062 | 2955842200800261 | 1 | GCJ | Manufacturer | ForeignDistributor | Other | Initial | 2008-05-09 | 2008-05-12 | Yes | Yes | Health Professional | IS1200 A43 P8 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
ISI200 A43P8 |  | Other | Other | It was reported that prior to beginning a da vinci prostatectomy surgical procedure, the system surgeon console left eye view was black. The pt was under anesthesia, however, no surgical tasks had been performed with the system when the surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by isi field service engineering concluded that the black left eye view experienced by the customer was associated with the system camera. The system was repaired by replacing the affected camera. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1043001 | 1047875 | 1006063 | 2955842200800255 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-09 | 2008-05-12 | Yes | Yes | Health Professional | ISI200 A43P8 | No | No |  | Device Not Retu | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that a monopolar curved scissors instrument was dull. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and observed that the instrument cut cleanly and as intended. The blades are not damaged and electrical continuity passed. Engineering also observed the distal end of the main tube has two distinct sections located above and below the midpoint with material removed. Damaged areas are both roughly 2. 5" long, parallel to tube axis with a rough surface finish. Based on the location and appearance, the damage area below the midpoint may have been caused by a cannula accessory. The damage above the midpoint area appears to have been caused by high pressure against the cannula bowl/tube transition. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1043450 | 1041285 | 999536 | 2955842200800273 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-20 | Yes | No | Reuse | 2008
42017206 | 2008-04-17 | Malfunction |  | It was reported that the monopolar curved scissors instrument has shavings coming off of the fiberglass shaft. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube has a 2. 5 inch long section just below the mid point with light material removed. The damaged area is parallel to the tube axis and has a rough surface finish. The wear pattern is consistent with cannula related tube abrasions, however, the position of the damaged area is slightly higher on the tube. The scraping may have been caused by high pressure against the internal bowl/tube transition of the cannula accessory. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1043444 | 1041286 | 999537 | 2955842200800267 | 1 | GEI | Manufacturer | User facilityOther | Nurse | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2008-04-17 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument had excessive wear of the shaft, and no pieces fell into the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has a 5. 5 inch long section with material removed all around the tube. Damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely cause by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Engineering also observed that one grip close cable is derailed at the distal idlers. No other damge found. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1043439 | 1041287 | 999538 | 2955842200800258 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-04-15 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument had excessive wear of the shaft, and no pieces fell into the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reportd. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube has a section ranging from 2 to 3 inches in length with light material removed all around the tube. The damaged areas are both parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damge was most likely cause by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Engineering also found the entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. There are no pieces missing. The damage is most likely due to excessive side loading. Electrical continuity passed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1043441 | 1041288 | 999539 | 2955842200800257 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-15 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, a broken cable was found on a maryland bipolar forceps instrument. The planned surgical procedure was completed. No add'l info was provided. No pt harm adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no cables or wires broken or damaged at the wrist. One grip is bent, creating a. 040" offset at the tips. Yaw pulley does not have separation at glue joint, however, the most likely cause of bending is overloading the tip. Engineering also observed the distal end of main tube has a 2. 0" long section with material removed on one side of the tube. Damaged area is located directly above the proximal clevis, parallel to tube axis, and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042840 | 1041289 | 999540 | 2955842200800265 | 1 | NAY | Manufacturer | DistributorOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-17 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, a maryland bipolar forceps instrument was not recognized by the system. The planned surgical procedure was completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a da vinci s test system to check for recognition. System successfully recognized the instrument. Pogo pins did not stick, but have some contamination (bio debris) on the sides of the pin. This contamination may have caused intermittent recognition issues at the site. Engineering also observed the distal end of main tube has a 1. 5" long section with material removed. Damaged area is located 4" above the proximal clevis, parallel to tube axis with a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr wil be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040602 | 1041290 | 999541 | 2955842200800268 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-17 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, a broken cable was found on a maryland bipolar forceps instrument. The planned surgical procedure was completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no cables or wires broken or damaged at the wrist. One grip is bent, creating a. 030" offset at the tips. Yaw pulley does not have separation at glue joint, however the most likely cause of bending is overloading the tip. Engineering also observed the distal end of main tube, tube has a 2. 8" long section with material removed on one side of the tube. Damaged area is located. 5" above the proximal clevis, parallel to tube axis, and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Additionally, several deep scratches adjacent to the scraped tube area were found. Their location, orientation and appearance suggest instrument collision may have occurred. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040615 | 1041291 | 999542 | 2955842200800266 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-17 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, an error was generated when the monopolar curved scissors instrument was inserted for use on the system. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted testing and found that the "instrument not recognized" error message occurred when the instrument was attached to multiple system arms. Engineering observed two of the pogo pins stick when depressed, which is most likely caused the recognition issue. Engineering also discovered that the scissor blades do not fully close due to blade damage. One blade edge has a deformation near the tip, and the other blade appears to have below average curvature. This combination results in a mechanical stop below the tips. Engineering also observed the distal end of main tube has a 2. 5 inch long section with light material removed all around the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1043468 | 1041292 | 999543 | 2955842200800252 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the connector prongs are loose and bent on a maryland bipolar forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found both bipolar pins are bent in towards each other, preventing installation of a bipolar cord. Damage to pins is likely due to mishandling. Engineering also observed the distal end of the main tube has a 2. 5" long section with light material removed. Damaged area is located 2. 75" above the proximal clevis, parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040631 | 1041293 | 999544 | 2955842200800271 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-18 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | It was reported that the permanent cautery spatula instrument was broken, however, the patient was not impacted. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one of the distal clevis "ears" and the conductor cap were missing. The ear was broken off at base and it appears that the breakage was most likely due to overloading of the tip with the wrist pitched. Engineering also observed the distal end of main tube has a section directly above the proximal clevis with material removed on one side. Based on the location and appearance, the damage was most likely caused by a cannula accessory and instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress and can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040638 | 1041294 | 999545 | 2955842200800243 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
42014504 | 2008-03-27 | Malfunction |  | It was reported that the debakey forceps instrument shaft is scratched. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the main tube has three sections with missing tube insulation. The missing sections range in size from. 110 inches by. 90 inches to. 375 inches by. 100 inches. The damaged areas are located 3. 6 inches above the snake wrist all the way up to the midpoint of the tube. No other damage found. Based on the location and appearance, the scratches are most likely due to instrument collisions or a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1040651 | 1041295 | 999546 | 2955842200800253 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42014504 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
42017908 | 2008-04-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while the surgeon was attempting to cut tissue with the monopolar curved scissors (mcs) instrument, one half of the tip broke off and fell into the patient. The fragment was immediately removed and the procedure was completed with another instrument of the same type, with approximately a 10 minute delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one scissor blade broke off at the cross section of the grip cable crimp. The broken blade was returned. The intact blade shows signs of light pitting on the non-cutting edge, indicating possible corrosion which may have caused corrosive cracking, causing the blade failure. Laboratory analysis discovered chlorine present on the fracture surface, indicating that the tip breakage was due to stress corrosion cracking. The endowrist instruments instructions for use (ifu) specifically states: general instructions, cleaning the instruments. Do not expose instruments to hydrogen peroxide (h202), bleach or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040664 | 1041296 | 999547 | 2955842200800256 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
42022702 | 2008-04-18 | Malfunction | Other | It was reported that during a da vinci s sacrocolpopexy procedure, while using the pk dissecting forceps instrument, the surgical staff observed charring of the patient's skin surrounding the port incision where the cannula accessory and instrument were inserted. The surgeon immediately replaced the pk dissecting forceps instrument with another instrument of the same type to complete the planned surgical procedure. The "burned" skin was removed and sutured close. The case was not significantly lengthened due to this event. No additional patient harm, adverse outcome or injury was reported. | The instrument used during the surgical procedure has not been returned for evaluation. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of 2008, there have been no reported recurrences of the issue at this hospital. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040671 | 1041297 | 999548 | 2955842200800272 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-05-09 | 2008-05-09 | Yes | Yes | Health Professional | 42022702 | No | No | 2008-04-18 | Device Not Retu | No | Unkown | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument cable is broken and no pieces fell in the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grip close cable is broken at the scissor grip, with one strand broken at the distal idlers. The grip hub has a peak across the cable groove that most likely contributed to breakage. The instrument has no uses left, indicating that the cable most likely broke on its' last life. Engineering also observed the distal end of main tube has a 1. 5 inch section with material removed all around the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040875 | 1041298 | 999549 | 2955842200800263 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2008-04-17 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the sys was not recognizing the precise bipolar forceps instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted testing and found that instrument recognition was successful. Engineering was unable to replicate the site's instrument recognition issue. Engineering also observed one conductor wire is broken at the yaw pulley exit. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of yaw motion. The added range of motion of the grip most likely created excess tension on the wire, causing it to break. Engineering also observed the distal end of the main tube has a 2. 25 inch long section with material removed all around the tube. The damaged area is located 4 inches above proximal clevis, parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely cause by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042576 | 1041299 | 999550 | 2955842200800264 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-04-17 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure the bipolar maryland forceps instrument had a broken cable. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that there were no cables broken, but one conductor wire is broken at the yaw pulley exit. The wire insulation does not appear cut or melted. The wire most likely was not securely attached to the grip, had poor strain relief, and pulled out during the articulation of the wrist. Engineering also observed the distal end of the main tube has two distinct sections with light material removal. The damaged areas are 1 inch and 1. 6 inchs long, both based on the location and appearance, the damge was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1043437 | 1041300 | 999551 | 2955842200800269 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-18 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure the endowrist prograsp forceps instrument had a broken cable. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip open cable is frayed at the force amp pulley. The cable may have been nicked by another instrument. Engineering also observed the distal end of the main tube has two distinct sections, 1. 3 inch and 2 inchs long, with material removed. The 2 inch long section starts. 625 of an inch above the proximal clevis and is roughly parallel to the tube axis. The 1. 3 inch long section starts 4. 2 inches above the proximal clevis and is parallel to the tube axis. Both of the damaged areas have rough surface finishes. Based on the location and appearance, the damge was most likely cause by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1043472 | 1041301 | 999552 | 2955842200800270 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-04-18 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, a broken cable was found on a monopolar curved scissors instrument. The planned surgical procedure was completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a broken grip close cable at the distal idlers, allowing the pulley to spin freely. It was determined that the cable broken under tensile loading and cable wear on the pulleys combined with drive torques likely contributed to breakage. Engineering also observed the proximal clevis has an. 085" diameter char mark above the flush hole. The char mark includes part of the silicone overmold window, which has melted/material removal, suggesting an arcing event occurred. The site did not return any tip cover accessory, however, a sample tip cover accessory was installed. The position of the char mark lines up with the silicone to sleeve interface of the tip cover accessory. Additionally, the main tube has a 4. 5" long section directly above the tube extension with material removed all around the tube. Damage area is parallel to the tube axis, with rough surface finish. Based on location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040624 | 1041302 | 999553 | 2955842200800262 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-17 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a large needle driver instrument was not holding the needle. The planned surgical procedure was completed, and no pieces fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument had zero uses remaining, therefore, was unable to conduct performance tests on a system. Manually rotating the input discs confirmed the instrument moves intuitively with full range of motion, and the grips open and close properly. Engineering also observed the distal end of main tube has various deep scratches with material removed. Scratches are not aligned with the tube axis, suggesting they were caused by instrument collisions. Instrument is fully functional otherwise. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1045963 | 1041303 | 999554 | 2955842200800259 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-09 | 2008-05-09 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-04-15 | Yes | No | Reuse | 2008
42017908 | 2008-03-31 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the surgeon noticed arcing and a burn on the pt's bowel. The surgeon immediately stopped using the mcs instrument and replaced the mcs tip cover accessory. The planned procedure was completed with the replacement tip cover accessory and with approx. A 30 minute delay. No add'l pt harm, adverse outcome or injury was reported. Multiple attempts have been made to contact the attending physician present for this procedure, to obtain add'l info, however, no response has been received at this time. | The monopolar curved scissors instrument (mcs) was returned with the tip cover accessory installed. The instrument does not show signs of possible arcing such as char marks. Engineering observed the distal end of the main tube has a 2 inch section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The tip cover accessory used in conjunction with this instrument during the reported case has been returned and evaluated, please see mdr 2955842-2008-00219. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1046874 | 1047805 | 1005993 | 2955842200800274 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-05-09 | 2008-05-09 | Yes | Yes | Health Professional | 42017908 | Yes | No | 2008-04-16 | Yes | Yes | Reuse | 2008
IS2000 A51P4 | 2008-04-16 | Other | Other | It was reported that 30 minutes into a da vinci s prostatectomy procedure, the pt side cart (psc) switched to battery power. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the site try multiple power outlets and powered off the psc, however, the issue could not be resolved. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | On site investigation conducted by isi field service engineering (fse) discovered that the power supply was defective. The system was repaired by replacing the power supply. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1046355 | 1047807 | 1005995 | 2955842200800260 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-05-09 | 2008-05-09 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-04-16 | Yes | No | Reuse | 2008
| 2008-04-29 | Malfunction |  | During a laparoscopic robotic assisted procedure, the vcare broke down in parts into the patient's uterus. The physician scrubbed in, removed the uterus and all of the remaining parts of the vcare (uterine manipulator) through the vagina. No additional intervention was needed at the time. |  | VCARE | UTERINE MANIPULATORRETRACTOR | 1007210 | 1039273 | 997626 | 1039273 | 1 | LKF | User Facility |  |  | Initial | 2008-04-29 | 2008-05-05 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2008
42009307 |  | Malfunction |  | It was reported that the tip of an endowrist prograsp forceps instrument was broken. No instrument fragments fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned an devaluated. Per the customer reported complaint, engineering inspected the instrument and found that the tip was not broken. All components at the wrist were undamaged. Instrument recognition and engagement were confirmed. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. Instrument is fully functional. The complaint was not confirmed. Engineering also observed the distal end of main tube has a 1. 25" long,. 100" wide section with material removed. Damaged area is located 1. 25" above proximal clevis, parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Additionally, a couple of light scratch marks below the damaged area were found. Their location, orientation and appearance suggests instrument collision may have occurred. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1040740 | 1037945 | 996306 | 2955842200800235 | 1 | GCJ | Manufacturer | User facilityOther | Physician | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2008-04-07 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the tip of a endowrist prograsp forceps instrument was broken. No instrument fragments fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the instrument and found frayed cables on the side of the grip open and grip close assembly at the force amp pulley. Grips can still open and close and motion is intuitive. Location of fray is nearly identical for both cables. Cables may have been nicked by another instrument, causing fraying. Engineering also observed the distal end of the main tube has some scratch marks, including one deep scratch. 300" above the proximal clevis with material removed. Scratch is curved and not aligned with the tube axis, suggesting it was caused by an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1041487 | 1037948 | 996309 | 2955842200800234 | 1 | GCJ | Manufacturer | User facilityOther | Physician | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2008-04-07 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that a wire on a pk dissecting forceps instrument was loose. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one conductor wire is broken at the yaw pulley exit. Wire insulation does not appear cut or melted. The wire may not have been securely attached to the grip, had poor strain relief, and pulled out during articulation of the wrist. Electrical continuity failed. Engineering also observed the distal end of main tube has a 2. 5" long section directly above the proximal clevis with material removed. The distal end also has other sections around its circumference with light material removal damaged areas are parallel to tube axis and have a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1042424 | 1037949 | 996310 | 2955842200800245 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-25 | 2008-05-02 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-04-09 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that prior to starting a da vinci s nissen fundoplication surgical procedure, a broken cable was found on a large needle driver instrument. The planned surgical procedure was completed. No additional information was provided. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable broken at the distal idlers. Pulley spins freely. Other cables at wrist were not damaged. Cable wear on pulleys combined with high grasping force may have contributed to the cable breakage. Engineering also observed the distal end of main tube had material removed from a combination of scratches and scrape marks. Scratches are randomly oriented relative to tube axis, suggesting instrument collision had occurred. The source of the scraping cannot be determined. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1040665 | 1037950 | 996311 | 2955842200800238 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-04-08 | Yes | No | Reuse | 2008
42021408 |  | Malfunction |  | It was reported that the shaft of the pk dissecting forceps instrument is scratched and no pieces fell into the patient. No further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of the main tube has a 3. 25 inch long section with material removed all around the tube. Damaged area is located 1. 5 inches above tube extension, parallel to tube axis, with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. The instrument is fully functional. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040711 | 1037952 | 996313 | 2955842200800246 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42021408 | Yes | No | 2008-04-09 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that shavings were observed coming off the shaft of the monopolar curved scissors instrument during processing. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the main tube has a 3 inch long section with material removed. Damaged area is located 2. 5 inches above the tube extension, parallel to the tube's longitudinal axis, and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1038563 | 1037953 | 996314 | 2955842200800242 | 1 | GEI | Manufacturer | User facilityOther | Nurse | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-18 | Yes | No | Reuse | 2008
42017907 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not opening. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed and the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed the distal end of the main tube has a couple of deep scratches 1 inch above the tube extension with material removed. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040686 | 1037955 | 996316 | 2955842200800240 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017907 | Yes | No | 2008-04-03 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was broken at the shaft. It was also reported that no instrument components fell inside of the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the transition between the thin wall and the. 330 inch diameter ring. As a result, the tube extension can be rotated 360 degrees. Tube extension is missing a. 150 inch by. 220 inch section as a result of breakage. Damage may be due to excessive side loading. Lack of a radius at the transition may have also contributed to breakage. Engineering also observed the distal end of the main tube has a couple of scratches with material removed, both within. 500 inches of the tube extension. Based on the location and appearance, the scratches were most likely caused by instrument collisions or some sort of user mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040720 | 1037956 | 996317 | 2955842200800236 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, a broken cable was found on the monopolar curved scissors instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is frayed at the scissor grip. The scissors can still open and close. The grip hub has a sharp peak across the cable groove located near the fraying. Repeated cable wearing across the peak may have contributed to the cable fraying. Engineering also observed the main tube has two distinct sections with material removed. One section is 2. 75 inches long and is located below the midpoint, and the other section is 2. 5 inches long and located above the mid point. Both damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1038566 | 1037958 | 996319 | 2955842200800249 | 1 | GEI | Manufacturer | ForeignUser facility | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-10 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that prior to starting a da vinci s nissen fundoplication surgical procedure, a fenestrated bipolar foreceps instrument was not recognized by the system. The procedure was completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed instrument recognition and engagement. Pogo pins do not stick and are not contaminated. Instrument moves intuitively when driven on the da vinci s test system. The complaint was not confirmed. Engineering also observed the distal end of main tube has two distinct sections, 180 degrees apart, with material removed. Sections are 4" and 2" long, parallel to tube axis, and have a rough surface finish. Based on the location and appearance, the damage may have been caused by cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040674 | 1037959 | 996320 | 2955842200800239 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-04-08 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that approximately 30 minutes into a da vinci s prostatectomy procedure, the surgeon noticed a hole in the "shaft" of the monopolar curved scissors (mcs) instrument. The surgeon immediately removed the instrument and replaced it with another instrument of the same type to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 270 inch by. 075 inch oblong hole burned through the main tube. The hole is located roughly 1. 25 inches above the tube extension. A ring of tube material is removed around the perimeter of the hole, indicative of arcing, and the main tube has a longitudinal crack that bisects the hole. The instrument was disassembled to find the hypotubes and cap coupling tube charred in the same location as the hole. The tube crack that bisects the hole, starts at one of the slots at the distal end of the tube. There is also a second longitudinal crack at the distal end located 90 degrees radially from the crack at the slot. The second crack also extends past the location of the hole. Engineering has determined that the crack in the tube and high generator settings likely created a path for arcing to occur. The source of the crack is unclear. Engineering also observed the distal end of the main tube has a 3. 5 inch long section directly above the tube extension, with material removed. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040702 | 1037960 | 996321 | 2955842200800244 | 1 | GEI | Manufacturer | Health ProfessionalOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-04-08 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument does not open. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed and the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. One pitch cable is frayed slightly at the distal clevis; however, the other cables are not damaged. The latex cut test and electric continuity passed. The instrument is fully functional. Engineering also observed the distal end of the main tube has various deep scratches with material removed. Based on the location and appearance, the scratches are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040694 | 1037961 | 996322 | 2955842200800250 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-04-11 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that shavings were observed coming off the shaft of a maryland bipolar forceps instrument during processing. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of main tube had various deep scratches with material removed. Scratches were mostly on one side of the tube, under. 250" in length and not aligned with the tube axis. Their location, orientation and appearance suggest instrument collision may have occurred. The device was confirmed to be fully functional, with no other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1040679 | 1037963 | 996324 | 2955842200800254 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-05-02 | 2008-05-02 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2008-04-14 | Yes | No | Reuse | 2008
IS1200 A43P8 | 2008-04-08 | Other | Other | It was reported that during a da vinci nephrectomy procedure, the customer experienced system error code #212. Due to pt bleeding the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No pt harm was reported. | The investigation conducted by isi field svc engineering determined that sys error code #212 experienced by the customer, which was associated with a pt side manipulator (psm). The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The sys was repaired by replacing the left psm. The sys alarm (sys generated fault code) functioned as designed and there was no injury to the pt. Sys error code #212 appears when the da vinci safety sys determines a voltage tracking fault reported by the digital signal processor (dsp). Upon determining these conditions, the safety sys put da vinci in a "no-recoverable safe state". The psm was returned to isi for failure analysis investigation. Based on the sys error logs, an axis motor and encoder assembly were placed. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1035380 | 1038874 | 997223 | 2955842200800247 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2008-04-09 | Yes | No | Reuse | 2008
420179 | 2008-04-03 | Other | Other | It was reported that during a da vinci hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument, a burning smell was noticed by the surgical staff. The instrument was removed and observed to be cracked and bent roughly 2 inches from the tip of the instrument. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. On april 18, 2008, the coordinator spoke to an isi representative and stated that a possible burn may have occurred as the pt is feeling discomfort in her abdomen. No further info was provided. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. The site indicated that they will not be returning the instrument to isi for add'l eval. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1035370 | 1038883 | 997232 | 2955842200800237 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | 420179 | No | No | 2008-04-07 | Device Not Retu | No | Unkown | 2008
IS1200 A43P9 | 2008-04-11 | Other | Other | It was reported that during a da vinci surgical procedure, the customer experienced intermittent sys faults (error codes). Due to the pt anatomy and recurring sys faults, the site converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field svc engineering concluded that the sys error code #2009 experienced by the customer was associated with a pt side manipulator (psm. The pt side manipulator is an instrument arm located on the pt side cart, that provides sterile interface for the endowrist instruments. The sys was repaired by replacing the affected psm. The sys alarm (sys generated fault code) functioned as designed and there was no injury to the pt. The psm was returned to isi for failure analysis investigation. Sys error code # 20009 appeared when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety sys put davinci in a "recoverable safe state". Engineering concluded that the motor encoder for a psm axis has malfunctioned, thus generating the sys errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1035387 | 1038889 | 997238 | 2955842200800251 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-04-14 | Yes | No | Reuse | 2008
40017908 | 2008-03-16 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, the cable to the monopolar curved scissors instrument broke. The surgeon continued with another instrument, however, the procedure was delayed by 1. 5 hrs. No pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one grip close cable is broken at the distal idlers, allowing the pulley to spin freely. The other cables at wrist are not damaged. Cable wear on the pulleys, combined with high blade closing force, most likely contributed to breakage. Engineering also observed the distal end of main tube has a 3. 5 inch long section directly above the tube extension with light material removed. The damaged area is parallel to the tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. The tube extension is missing half of the orange coating, possibly due to exposure to harsh cleaning agents. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1035393 | 1038892 | 997241 | 2955842200800231 | 1 | GEI | Manufacturer | Health ProfessionalOther | Other | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | 40017908 | Yes | No |  | Yes | No | Reuse | 2008
IS2000 A43P8 | 2008-04-08 | Other | Other | It was reported that while starting up the da vinci surgical system for a procedure, the customer experienced non- recoverable system error code #253. The pt had been under anesthesia for less than an hour and no ports had been placed when the customer contact isi technical support. No pt harm was reported. The surgeon decided to complete the planned surgery using laparoscopic technique. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the remote interface adapter (ria) pca. The system was repaired by replacing the ria pca. System error code #253 appears when the da vinci safety system determines an mnet mode has failed to communicate with the master system controller (msc). Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state. " as of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1034710 | 1038949 | 997298 | 2955842200800241 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | IS2000 A43P8 | Yes | No |  | Yes | No | Reuse | 2008
IS2000 AS 1P4 | 2008-04-09 | Other | Other | It was reported that prior to a da vinci s total hysterectomy surgical procedure, while starting up the system, the customer experienced system error code #25589. The pt was not in the room and did not have ports placed, however the pt had been given anesthesia 30 minutes prior to the site calling isi technical support. No pt harm was reported. The surgeon decided to reschedule the surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the remote arm controller board (rac). The system was repaired by replacing the affected rac. The remote arm controller (rac) is an electronic module comprised of four printed circuit assemblies (pcas) and cooling fans. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #25589 appears when the da vinci s safety system determines during the power up self test that the remote arm controller board (rac) brakes failed the brake voltage test. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state. " the rac was returned to isi for failure analysis investigation. It was determined that faulty chips most likely contributed to the system error experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1034695 | 1038951 | 997300 | 2955842200800248 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-05-02 | 2008-05-02 | Yes | Yes | Health Professional | IS2000 AS 1P4 | Yes | No |  | Yes | No | Reuse | 2008
40018005 | 2008-03-31 | No Answer Provided | Other | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curves scissors instrument ( mcs) with the tip cover accessory installed, the surgeon noticed arcing and a burn on the pt's bowel. The surgeon immediately stopped using and the mcs instrument and replaced the mcs tip cover accessory. The planned procedure was completed with the replacement tip cover accessory and with approx a 30 min delay. No additional pt harm, adverse outcome or injury was reported. Multiple attempts have been made to contact the attending physician present for this procedure, to obtain additional info, however, no response has been received at this time. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has a. 090 inch long mechanical tear through the entire thickness of the silicone. The tear is along one of the parting lines, located roughly. 450 inches above the silicone to sleeve interface. There are no signs of melting around this tear that are indicative or arcing. There are a couple of other mechanical tears in the silicone near the tip, however, there are no obvious signs of arching. The silicone stretches when placed on the mcs instrument, creating separation at the tear that exposes electrically active components. The instrument may have arced through the separation created by the tear. It was determined that mechanical tearing in the silicone is likely due to instrument collisions. Collisions may have damaged the silicone, locally reducing dielectric strength and creating a path for arcing. High generator setting may have also contributed to arcing. The monopolar curved scissors instrument ( mcs) was returned with the tip cover. The instrument does not show signs of possible arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1036332 | 1036846 | 995217 | 2955842200800219 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-04-29 | 2008-04-28 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-03-31 | Yes | Yes | Initial | 2008
IS2000 A51P4 | 2008-04-04 | Other | Other | An isi rep reported that while preparing to start a da vinci s cardiac biventricular lead placement procedure, the camera head and endoscope attached to the endoscopic camera manipulator (ecm) arm would not advance. An isi technical support engineer (tse) assisted via telephone and had the isi rep undock, try to manually move the ecm then restart the system, however, the issue remained. The pt side cart monitor also lost video when the ecm setup joint was moved. The pt was under anesthesia for 30 minutes and the port incisions were already made when the surgeon decided to complete the planned surgical procedure with the site's spare da vinci system. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found system error code #23007 in the system error logs, which is associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm and ecm setup joint (suj). The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #23007 system error code appears when the da vinci s safety system determined that a sensor was out of specified tolerance for agreement on an axis potentiometer, thus causing the system error #23007. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation. During engineering eval, a grinding noise could be heard coming from an axis insertion drive. The top link cap cover was removed and engineering observed that two screws holding the brake were loose, thus causing the grinding noise and system error code #23007 experienced by the customer. The ecm suj was also returned for failure analysis investigation, however, no issues were found during testing. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1036873 | 1036847 | 995218 | 2955842200800232 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-04-25 | 2008-04-28 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-04-04 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | The bipolar forceps instrument was returned with no information provided. No patient harm, adverse outcome or injury was reported. | The instrument was evaluated. Engineering observed one conductor wire broken at the yaw pulley exit. The yaw pulley with broken conductor has a burn mark at the grip base. The wire insulation has small cuts and indentations and is reduced in diameter at the yaw pulley exit. Engineering determined that the reduced diameter may be due to stretching the insulation. The wire strands may not have had proper strain relief at the grip and broke during articulation of the wrist. Electrical continuity failed. Engineering also observed the distal end of the main tube has a few deep scratches with material removed on one side of the tube. Based on the location and randomly oriented appearance, the scratch marks are most likely caused by instrument collisions. No other damage found. The endo-wrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1035052 | 1037164 | 995535 | 2955842200800233 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-05-01 | 2008-04-28 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2008-01-28 | Yes | No | Reuse | 2008
40018005 |  | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, in the middle of coagulating, the surgeon noticed arcing from the monopolar curved scissors instrument (mcs). The surgical staff immediately removed the mcs instrument and found a "little hole" on the tip cover accessory. The planned procedure was completed with a replacement tip cover accessory and without significant delay. No patient harm, adverse outcome or injury was reported. | The accessory and instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for eval, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1035693 | 1037165 | 995536 | 2955842200800220 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-04-29 | 2008-04-28 | No | Yes | Health Professional | 40018005 | No | No | 2008-04-01 | Device Not Retu | Yes | Initial | 2008
40018005 |  | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, the surgeon noticed arcing from the monopolar curved scissors instrument (mcs). The surgical staff immediately removed the mcs instrument and found a hole on the tip cover accessory. The planned procedure was completed with a replacement tip cover accessory and without significant delay. No patient harm, adverse outcome or injury was reported. | The accessory and instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for eval, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1036239 | 1037166 | 995537 | 2955842200800229 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-04-28 | 2008-04-28 | No | Yes | Health Professional | 40018005 | No | No | 2008-04-03 | Device Not Retu | Yes | Initial | 2008
42017908 |  | Malfunction |  | It was reported that during the sterilization process a frayed cable was found on the monopolar curved scissors instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable frayed at the scissor grip. The grip hub has a peak across the cable groove that may have contributed to the fraying. The tube extension is broken at the proximal clevis interface, the clevis is not dislodged. A. 150 by. 250 inch broken piece of the tube extension was found loose in the shipping box. Engineering determined that the tube extension may have been damaged during shipping. Engineering also observed the main tube has a section with material removed all around the tube. The damaged area is parallel to the tube axis and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033339 | 1034834 | 993245 | 2955842200800223 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-03 | Yes | No | Reuse | 2008
40016906 | 2007-04-02 | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure, the harmonic curved shears insert broke and pieces fell inside of the pt. The broken pieces were retrieved and the planned surgical procedure was completed. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering observed that one side of the clap arm attachment hinge feature is bent. Based on the damage to the hinge feature, engineering concluded that overloading likely occurred with the clamp arm in the open position. There may also have been an instrument collision that contributed to the separation of the clamp arm. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1033326 | 1034835 | 993246 | 2955842200800228 | 1 | LFL | Manufacturer | User facility | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-04-15 | Yes | Yes | Initial | 2008
42017908 |  | Malfunction |  | It was reported that there was difficulty opening the tips of the monopolar curved scissors instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly and the tips did not stick when the blades were opened. The instrument passed cuts as intended and electrical continuity passed. Engineering also observed the distal end of main tube has a. 825 inch long,. 230 inch wide section with material removed. The damaged area starts 2. 9 inches above tube extension and is parallel to the tube axis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033323 | 1034836 | 993247 | 2955842200800222 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-04-02 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that tips of the maryland bipolar forceps instrument are not meeting. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was retuned and evaluated. Per the customer reported complaint, engineering confirmed that one grip is bent, causing misalignment of the grips. There is a. 060 inch offset at the tips. The glue joints at the jaw pulleys are intact, however, the most likely cause of the bending is overloading at the tip. Electrical continuity passed. Engineering also observed the distal end of the main tube has one deep scratch perpendicular to the tube axis with material removed. The scratch is. 165 inch long and. 040 inch wide, located. 625 of an inch above the proximal clevis. Based on the location and appearance, the scratches were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033412 | 1034837 | 993248 | 2955842200800221 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-04-02 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the "protected red tip" of the monopolar curved scissor instrument is broken and has no more power at the blade tip when the power is on. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the tube extension is missing a section. 125 inches by. 250 inches at the interface with the proximal clevis. The proximal clevis is slightly dislodged from the tube extension and has one of the thin metal interface walls bent at the damage site. Electrical continuity passed, suggesting the cautery is still function. There are also various char marks on the tube extension, suggesting an arcing event occurred. The proximal clevis also has some charred bio-material adjacent to the proximal clevis charring. The charred material on the proximal clevis indicates that a tip cover accessory was not installed on the instrument when the arcing occurred. It is also possible that high generator settings contributed to the arcing. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033417 | 1034838 | 993249 | 2955842200800218 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-31 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the cable of the monopolar curved scissors instrument is broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the scissor grip and distal idlers, allowing the pulley to spin freely. The grip hub has a peak across the cable groove that may have damaged the cable. Engineering concluded that the cable wear on the grip/pulleys combined with the high torque at the input discs likely contributed to the breakage. Engineering also observed the distal end of main tube has a 2 inch long section with light material removed. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage was most likely caused be a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTOSURGICAL INSTRUMENT | 1033405 | 1034839 | 993250 | 2955842200800215 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-28 | Yes | No | Reuse | 2008
4209307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument has shards coming off of a scratch on the shaft case. No pieces fell into the patient and no further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of main tube has a couple of deep scratches with material removed. Scratches are narrow, and relatively short, about. 135 -. 280 inches. Based on the location and appearance, the tube damage is likely due to instrument collisions. Instrument is fully functional. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1033400 | 1034840 | 993251 | 2955842200800203 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 4209307 | Yes | No | 2008-03-28 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that a pk dissecting forceps instrument did not work properly. There was no impact to the patient when the failure occurred. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests. Instrument recognition and engagement passed. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. The complaint was not confirmed. Engineering also observed the main tube has a 2. 5" long section with light material removed. Damaged are is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. There are also a couple of deep scratches directly above the proximal clevis with material removed. Scratches are not parallel to tube axis and were likely caused by instrument collisions. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033477 | 1034841 | 993252 | 2955842200800224 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-04-02 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that a pk dissecting forceps instrument did not work properly. There was no impact to the patient when the failure occurred. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing. Instrument recognition and engagement passed. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. The complaint was not confirmed. Engineering also observed the distal end of the main tube has material removed from deep scratches and from a scrape mark. Scratch marks are short (under. 200") and not aligned with tube axis. Scrape mark is 1" long and parallel to tube axis. Based on the location, orientation and appearance, damage may have been caused by instrument collisions and a cannula accessory. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions with warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033423 | 1034842 | 993253 | 2955842200800225 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-04-02 | Yes | No | Reuse | 2008
42004905 |  | Malfunction |  | It was reported that a cardiere forceps instrument did not work properly. There was no impact to the patient when the failure occurred. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. The complaint was not confirmed. Engineering also observed the distal end of the main tube has various deep scratches on one side with material removed. Scratches are all under. 150" long and are not parallel to tube axis. Based on the location, orientation and appearance, damage may have been caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDSCOPIC INSTRUMENT | 1033485 | 1034843 | 993254 | 2955842200800226 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42004905 | Yes | No | 2008-04-02 | Yes | No | Reuse | 2008
40018402 |  | Malfunction |  | It was reported that smoke was coming from the wrist of a permanent cautery spatula instrument. The patient was not injured when the instrument began smoking. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the yaw pulley and conductor cap to have localized charring/melting, suggesting an arcing event occurred. Engineering observed a section of the conductor wire missing its insulation, thus creating a path for arcing to occur. There were also signs of distal clevis charring and localized melting of the distal idler/proximal pulleys on the side opposite of the conductor wire. Correlation of this damage to arcing at the jaw pulley/conductor cap can not be determined. High generator settings may also have contributed to arcing. No other damage found. | PERMANENT CAUTERY SPATULA INSTURMENT | ELECTROSURGICAL INSTRUMENT | 1033491 | 1034844 | 993255 | 2955842200800227 | 1 | GEI | Manufacturer | OtherHealth Professional | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 40018402 | Yes | No | 2008-04-08 | Yes | No | Reuse | 2008
42021405 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a cable on a pk dissecting forceps instrument broke. No instrument fragments fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire broken at the jaw pulley exit. One set of grip cables are derailed at the distal idlers. The wire most likely pulled out during articulation of the wrist due to poor strain relief. Cable derailment may have also contributed to wire breakage. Engineering also observed the distal end of the main tube has a couple of deep scratches with material removed. Scratch marks are not parallel to tube axis. Based on the location, orientation and appearance, damage may have been caused by instrument collisions. Electrical continuity failed. No other damage found. The endowrist instruments instruction for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unidentified consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1033493 | 1034845 | 993256 | 2955842200800230 | 1 | GEI | Manufacturer | OtherHealth Professional | UNKNOWN | Initial | 2008-04-25 | 2008-04-25 | No | Yes | Health Professional | 42021405 | Yes | No | 2008-04-04 | Yes | No | Reuse | 2008
| 2008-04-23 | No Answer Provided |  | Vcare perforated uterus during robotic-assisted (using the davinci system) laparoscopic-assisted vaginal hysterectomy with bso (bilateral salpingo oopherectomy). |  | VCARE SMALL | UTERINE MANIPULATORRETRACTOR | 1006665 | 1038948 | 997297 | 1038948 | 1 | LKF | User Facility |  |  | Initial | 2008-04-25 | 2008-04-25 | No | No | Physician |  | Yes | No Answer Provided |  |  |  |  | 2008
42009307 |  | Malfunction |  | It was reported that the fiberglass shaft of the endowrist prograsp forceps instrument is shredding. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube has various small scratches, under. 25 inch long, all around the tube. A couple of the scratches are deep enough to have removed a small amount of material. Based on the location and appearance of the scratches, they were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use ( ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1038101 | 1031313 | 989781 | 2955842200800211 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-03-27 | Yes | No | Reuse | 2008
42017202 |  | Malfunction |  | It was reported that the fiberglass shaft of the fenestrated maryland bipolar instrument is shredding. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube has a 1 inch long by. 200 inch wide section with light material removed. The damaged area starts above the proximal clevis and is parallel to the tube longitudinal axis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1035821 | 1031314 | 989782 | 2955842200800212 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017202 | Yes | No | 2008-03-27 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument was defective. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument moved intuitively with full range of motion and in all directions. The grips open and close properly and recognition and engagement tests passed. Engineering also observed the distal end of main tube has two 2 inch long section, 180 degrees apart, with light material removed. The damaged areas are both parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance of the damage area, engineering determined that the wear pattern is consistent with cannula related tube abrasion. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | EIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037075 | 1031315 | 989783 | 2955842200800217 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-29 | Yes | No | Reuse | 2008
42020504 | 2008-03-06 | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument was not being recognized by the system. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Electric continuity passed. Engineering also observed the distal end of the main tube has various scratch marks with a couple of scratches deep enough to have a small amount of material removed. Based on the location and appearance, the scratches are most likely due to instrument collisions or mishandling during cleaning. No other damage found. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037108 | 1031316 | 989784 | 2955842200800204 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-03-24 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the casing of a monopolar curved scissor instrument was broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Damage may be due to excessive side loading. Engineering also observed the main tube has a 3" long section just below the midpoint with material removed. Damaged area is parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037086 | 1031317 | 989785 | 2955842200800214 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-27 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, plastic around the jaw of a maryland bipolar forceps instrument broke. A broken cautery cable was also reported and that no components fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering inspected the instrument and confirmed that there is material removed from both yaw pulleys due to burning, on the same side as one conductor wire. The conductor wire is detached from the grip as a result. The yaw pulleys exhibit localized melting. Charring/melting is likely due to repeated activation of cautery without tissue between grips. High generator settings may have contributed to the damage. Engineering also observed the distal end of the main tube has a section with light material removal all around the tube. Damaged area is 2. 5 " long, parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037124 | 1031318 | 989786 | 2955842200800209 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-25 | Yes | No | Reuse | 2008
42017206 | 2008-03-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the cautery cable of the maryland bipolar forceps instrument broke, however, no components fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the instrument and confirmed the one conductor wire is broken at the yaw pulley exit. Yaw pulley shows no signs of arcing. Grip cables are not derailed at wrist. Wire insulation does not appear cut or melted. Yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in the yaw. Added range of motion of the grip could create excess tension on the wire, causing it to break. The wire may not have been securely attached to the grip and had poor strain relief. This condition could allow the wire to be pulled out during articulation of the wrist. Electrical continuity failed. Engineering also observed the main tube has a 2. 75" long section with material removed. Damaged area is parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation in underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037137 | 1031319 | 989787 | 2955842200800208 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-25 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that one of the prongs where the electrocautery cord attaches to on the maryland bipolar forceps instrument was broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the top riser pin pushed into back end and the bottom pin is bent. The chassis feature that retains bipolar insert and pins is broken. The instrument was likely mishandled, breaking the chassis feature. Electric continuity passed. Engineering also observed the distal end of the main tube has a couple of deep scratches that have material removed. Based on the location and appearance, the scratches were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use ( ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handling instruments with care. Avoid mechanical shock or stress that can cause damage to the instrument. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1037096 | 1031320 | 989788 | 2955842200800210 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-25 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, plastic around the jaw of a maryland bipolar forceps instrument broke, however, no components fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the yaw pulley was missing a piece of plastic on the opposite side of the conductor wire. The grip has additional range of motion in yaw as a result, however, the conductor wire is intact. Additionally, there is charring on both yaw pulleys at the grip base. The yaw pulleys are fused together from the heat. Charring/melting is likely due to repeated activation of cautery without tissue between grips. Engineering also observed the distal end of the main tube has light material removal. Damaged area is 3" long and parallel to tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1029965 | 1031321 | 989789 | 2955842200800207 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-25 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that during sterilization of a pk dissecting forceps instrument, insulation seemed to be flaking off 6 inches from the instrument tip. Fiberglass can be felt by running your finger over the instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube had a 4" long section directly above the proximal clevis with material removed all around the tube. Damaged area had a rough surface finish parallel to the tube's longitudinal axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. Engineering also observed that one grip was bent, creating a. 045" offset at the tips. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1030023 | 1031322 | 989790 | 2955842200800216 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2008-03-29 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the proximal clevis of a monopolar curved scissors instrument was bent and pieces did not fall into the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the proximal clevis was not bent, however, the tube extension was dislodged from the main tube. The entire tube extension wall was broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Damage may be due to excessive side loading. Engineering also observed the distal end of the main tube had a 2. 5" long section with material removed. Damaged area is parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Electrical continuity passed. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1031859 | 1031323 | 989791 | 2955842200800213 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-18 | 2008-04-18 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-27 | Yes | No | Reuse | 2008
400180 | 2008-03-01 | Other | Other | It was reported that towards the middle of a da vinci s hysterectomy procedure, while performing a right pelvic lymphadenectomy using the monopolar curved scissors instrument (mcs) with the tip cover accessory installed, the patient's right iliac vein and artery were cauterized. The planned procedure was completed with a replacement tip cover accessory and delayed 30 minutes as the surgeon stitched the burnt bowel. No additional pt harm, adverse outcome or injury was reported. | The accessory and instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined, however, the account indicated that the electro surgical generator settings used during the surgical procedure were set at 45 watts which is higher than isi's recommended settings. If the accessory or instrument are returned for eval, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1029213 | 1031777 | 990247 | 2955842200800192 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-04-18 | 2008-04-18 | Yes | Yes | Health Professional | 400180 | No | No | 2008-03-07 | Device Not Retu | Yes | Initial | 2008
IS2000 A51P4 | 2008-03-24 | Other |  | It was reported that during a da vinci s partial nephrectomy surgical procedure, the ecm clutch became stuck and the customer experienced system error code #23020 while docking the pt. The surgeon could not clear the system error code and decided to bring in their spare da vinci s system to complete the planned procedure. This resulted in a 30 minute delay. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code #23020 experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #23020 system error code appears when one of the switches in the manipulator is showing inconsistent signals on it's two switch leads. The ecm was returned to isi for failure analysis investigation. Engineering eval discovered a faulty switch, thus generating the system errors experienced by the customer. As of march 26, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1029207 | 1031778 | 990248 | 2955842200800206 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-04-18 | 2008-04-18 | Yes | Yes | Health Professional | IS2000 A51P4 | Device Returned To Manufacturer | No | 2008-04-02 | Yes | No | Reuse | 2008
IS1200 A43P8 | 2008-03-24 | Other | Other | It was reported that while docking the da vinci surgical sys to the pt for surgery, the customer experienced sys error codes #20008 and #21008 and could not clutch and move an axis on the endoscopic camera manipulator (ecm). An isi technical support engineer (tse) assisted the customer via telephone and requested that the customer restart the sys. The sys restarted and homed successfully without any faults. However, the customer later called back reporting that sys error codes #20008 and #21008 were recurring. The pt had been under anesthesia for approx two hrs, and the port incisions had been created when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. No pt harm, adverse outcome or injury was reported. | Conclusions: the investigation conducted by isi field svc engineering concluded that sys error codes #20008 and #21008 experienced by the customer were associated with a endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The sys was repaired by replacing the affected ecm. The sys alarm (sys generated fault code) functioned as designed and there was no injury to the pt. The #20008 and #21008 sys error codes appeared when the da vinci safety sys determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation. Engineering replaced an axis potentiometer as well as a motor encoder assembly based on a review of the sys error logs and in house testing performed. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1032453 | 1031779 | 990249 | 2955842200800205 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-04-18 | 2008-04-18 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2008-03-24 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-03-19 | Other | Other | It was reported that 2. 5 hours into a da vinci s hysterectomy procedure, the surgeon experienced a lack of light in the surgeon side console (ssc). No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. The customer contacted isi technical support after the procedure was complete. | The investigation conducted by isi field service engineering (fse) discovered the vision system to be within specification. The fse confirmed the surgeon side console (ssc) camera control unit (ccu) settings were correct and verified the light output from the illuminator. The surgeon side console is the control center where the operator sits outside the sterile field and control instruments and a 3d endoscope with his/her hands. The illuminator bulb was found to be near it's end of life which may have contributed to the reported lack of light. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1029204 | 1031785 | 990255 | 2955842200800196 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2008-04-11 | 2008-04-15 | Yes | Yes | Health Professional | IS2000 A51P4 | No | Yes | 2008-03-19 | Yes | No | Reuse | 2008
42009307 | 2008-03-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, pieces of cable from a endowrist prograsp forceps instrument broke off and fell into the pt. As of 2008, no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the instrument and found no evidence of broken or frayed cables. Engineering was unable to confirm the reported problem. Engineering also observed the distal end of the main tube has two distinct sections,. 90" and 1. 05" long x. 095" wide, separated by. 060", with material removed. Damaged areas are roughly parallel to each other and to the tube axis. Wear pattern of both sections have tapered ends, which along with the close spacing is unusual. Material removal may be due to tube pressure against a cannula accessory or another sharp edged item. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1027927 | 1028893 | 987384 | 2955842200800202 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-03-31 | Yes | No | Reuse | 2008
42017908 | 2008-03-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, pieces of cable from a monopolar curved scissors instrument broke off and fell into the pt. As of 2008, no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the instrument and found no evidence of broken or frayed cables. Engineering was unable to confirm the reported problem. Engineering also observed the distal end of the main tube has a 1. 25" long,. 220" wide section directly above the tube extension with material removed. Above this section, 90 degrees away, there is also a 1. 5" section with light material removal. Damaged areas are parallel to tube axis and have a rough surface finished. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | MONOPLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027825 | 1028894 | 987385 | 2955842200800201 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-31 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that a maryland bipolar forceps instrument was not recognized by the da vinci s surgical system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a da vinci s test system to check for recognition. System failed to recognize instrument on multiple installation attempts. Visual inspection revealed a pogo pin being shorter than the others. This pin is likely not making electrical contact with sterile adapter, preventing recognition. Pin is sticking, possibly due to dried chemical residue from cleaning process stuck to it. Other back end components do not show excessive cleaning residue. Engineering also observed the distal end of the main tube has a 4" long section with material removed. Damaged area is parallel to tube's longitudinal axis and may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027832 | 1028895 | 987386 | 2955842200800199 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-31 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that plastic around the bottom of a monopolar curved scissors instrument was broken. The damage was not identified during a procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension broken and missing a. 390" long x. 180" wide piece at the proximal clevis interface. Clevis was dislodged from tube extension. The tube extension may have broken due to excessive side loading or mishandling. Failure likely occurred outside patient, with tip cover removed. Engineering also observed the distal end of main tube has a 4" long section with material removed all around the tube. Damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027818 | 1028896 | 987387 | 2955842200800197 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-20 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure, a cable broke at the tip of the fenstrated bipolar forceps instrument near the shaft. Nothing fell into the pt. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface, resulting in clevis dislodging from the main tube. Clevis is intact, however, the main tube interface wall is collapsed and missing pieces. Breakage likely due to overloading at the tip. One conductor wire is broken at the yaw pulley exit. Engineering also observed the distal end of the main tube has a couple deep scratches directly above proximal clevis. Scratches have material removed and are indicative of instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENSTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027843 | 1028897 | 987388 | 2955842200800195 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-03-19 | Yes | No | Reuse | 2008
42009305 | 2008-02-20 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the endowrist prograsp instrument broke at the tip shaft and had a broken cable. Nothing fell in the pt. No other info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument is broken at the proximal clevis to main tube interface. Proximal clevis is dislodged from main tube as a result. Both the proximal clevis and main tube are missing pieces. The proximal clevis is missing a. 150 inch by. 180 inch piece between the ears. The majority of the main tube interface wall has been cut off. Engineering concluded the breakage is likely due to overloading at the tip. Engineering also found the main tube has a 2 inch long section with material removed. The damaged area is parallel to tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory and it is possible that extra force was required to remove the instrument from the cannula after clevis dislodgment. This extra force may have contributed to tube scraping. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1028050 | 1028898 | 987389 | 2955842200800198 | 1 | GCJ | Manufacturer | Company RepresentativeOther | Other | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42009305 | Yes | No | 2008-03-20 | Yes | No | Reuse | 2008
42017206 | 2007-10-11 | Malfunction |  | It was reported that during a da vinci s radical prostatectomy surgical procedure, the system was not recognizing the maryland bipolar forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the system successfully recognized instrument. Engineering was unable to replicate site's instrument recognition issues. Engineering also observed the distal end of the main tube has a couple of deep scratches with a small amount of material removed. The scratch marks are not aligned with the tube's longitudinal axis, indicating that the damage was caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027729 | 1028899 | 914316 | 2955842200800193 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-17 | Yes | No | Reuse | 2008
40022702 | 2008-03-11 | Malfunction |  | It was reported that during a da vinci hysterectomy surgical procedure, a broken cable was observed on the pk dissecting forceps instrument. No pt harm, adverse outcome or injury was reported. Additional info provided by an isi representative indicated that when the instrument was removed from the pt, the customer noticed a piece of plastic broken off near the tip which was. 25 centimeters by. 25 centimeters long. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. At this time, it also cannot be determined if the missing instrument component fell inside of the pt during the surgical procedure. A follow-up mdr will be submitted if the instrument is rec'd and in-house failure analysis has been completed or if any additional info is provided. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027907 | 1028900 | 987391 | 2955842200800187 | 1 | GEI | Manufacturer | User facilityCompany Representative | Other | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 40022702 | No | No | 2008-03-12 | Device Not Retu | No | Unkown | 2008
40021405 | 2008-03-10 | Malfunction |  | It was reported that during a da vinci sacral colpopexy surgical procedure, a broken cable was observed on the pk dissecting forceps instrument. No pt harm, adverse outcome or injury was reported. Additional info provided by an isi representative indicated that when the instrument was removed from the pt, the customer noticed a piece of plastic broken off near the tip which was. 25 centimeters by. 25 centimeters long. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. At this time, it also cannot be determined if the missing instrument component fell inside of the pt during the surgical procedure. A follow-up mdr will be submitted if the instrument is rec'd and in-house failure analysis has been completed or if any additional info is provided. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1027733 | 1028901 | 987391 | 2955842200800186 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-04-11 | 2008-04-11 | No | Yes | Health Professional | 40021405 | No | No | 2008-03-12 | Device Not Retu | No | Reuse | 2008
IS1200 A4 3P9 | 2008-03-17 | Other | Other | It was reported that prior to starting a da vinci prostatectomy procedure, the customer experienced persistent #20008 system error codes when installing a sterile adapter on a pt side manipulator (psm). An isi tech support engineer (tse) determined that the fault was related to the psm and had the customer shut down the system and check the system cable connectors, all connectors were found to be normal. The tse had the customer exercise the psm axis, power on the system and home, however, when the customer attempts to seat, a sterile adapter on the psm the #20008 system error codes continued to recur. The pt was under anesthesia for approx. Ten minutes when the surgeon decided to abort and reschedule the planned surgical procedure to a later date. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with a pt side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #20008 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". Engineering concluded that a broken clamp screw was allowing an axis to spin freely, thus generating the system errors experienced by the customer. The psm was repaired by replacing the four carriage clamp screws. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1027575 | 1028966 | 987452 | 2955842200800194 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2008-04-11 | 2008-04-11 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2008-03-17 | Yes | No | Reuse | 2008
40018005 | 2008-03-20 | Other | Other | It was reported at the beginning of a da vinci hysterectomy procedure, while taking down the ip on the left side of the ovary, the surgeon noticed the monopolar curved scissors instrument (mcs) sparked. Upon further inspection, the surgeon noticed a burn mark on the pt's uterus where the instrument was being used. The surgeon immediately stopped using the mcs instrument and replaced the mcs tip cover accessory. The planned procedure was completed with the replacement tip cover accessory and without delay. No add'l pt harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has a. 014 inch by. 007 inch oblong hole burned through the silicone. The hole is located. 175 inches above the silicone to ultem sleeve interface, 20 degrees radially from one of the parting lines. There is a mechanical tear at roughly the same distance above the silicone to sleeve interface, located 60 degrees radially from the same parting line. The hole exhibits localized melting and has an adjacent tear that also has localized melting on the outer surface of the silicone, suggesting an arcing event occurred. It was determined that tearing is likely due to instrument collisions. Collisions may have damaged the silicone, locally reducing dielectric strength and creating a path for arcing. High generator setting may have also contributed to arcing. The monopolar curved scissors instrument (mcs) was returned with the tip cover installed. The position of the arced hole appears to line up with a char mark on the distal clevis hub, near the pitch cable. The wrist may have been pitched when arcing occurred, reducing the silicone thickness and dielectric strength. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1027544 | 1028967 | 987453 | 2955842200800200 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-04-11 | 2008-04-11 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-03-21 | Yes | Yes | Initial | 2008
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument proximal clevis was cracked and there were abrasions on its main tube. It was also reported that no instrument components fell inside of the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a. 150" x. 150" piece of the proximal clevis was missing with one of the fracture lines halfway between the proximal clevis ears. The main tube interface wall was intact and the clevis had not dislodged from the tube. Grip may have been overloaded at the tip with the wrist pitched, causing the clevis fracture. Engineering also observed the distal end of the main tube had various deep scratch marks all around the tube with material removed. Scratches are randomly oriented relative to tube axis and were likely caused by instrument collisions or some sort of user mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1024194 | 1025441 | 984051 | 2955842200800188 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42021403 | Yes | No | 2008-03-14 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument does not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have a 1. 5 inch long section with material removed. The damaged area is 4 inches above the proximal clevis, parallel to tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance of the damage, engineering determined that loading on the instrument may have caused high pressure against the cannula, which caused scraping of the tube as it moved back and forth within the cannula. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1024197 | 1025442 | 984052 | 2955842200800182 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-03-10 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument had a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable broken at the scissor grip and distal idler pulley. Grip hub has a small peak across the cable groove that may have abraded cable. Pulley spins freely. Cable broke under tensile loading. Other cables at wrist are not damaged. Cable wear on grip/pulleys combined with high torques at input discs likely contributed to breakage. Engineering also observed the distal end of the main tube has material removed. Damaged area is a 1. 6" long section, parallel to the tube axis. Based on the location and appearance, the damage may have been caused by cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1024200 | 1025443 | 984053 | 2955842200800190 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-03-14 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument does not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performances tests and found that the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed that the distal end of the main tube has various deep scratches on one side of the tube. There is material removed as a result of scratches. The scratches are not parallel to tube axis. Based on the location and appearance, the scratches were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1024206 | 1025444 | 984054 | 2955842200800183 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-03-10 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the tabs popped off the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that both release levers were installed, but flipped 180 degrees. The instrument could not be installed onto a sterile adapter in its current state. Two housing pins on the same side were broken off. Evidence suggests the instrument was mishandled and the housing and levers popped off. Levers were installed incorrectly when housing was put back on. Engineering also observed one side of the main tube has various light scratch marks and also a couple deep scratches at the distal end. The deep scratches have removed material from the tube. Scratches are randomly oriented relative to tube axis, suggesting they were caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1024210 | 1025445 | 984055 | 2955842200800191 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-03-15 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument does not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube has a 2 inch long section with light material removed. The damaged area is 1. 25 inches above the proximal clevis, parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance of the damage, engineering determined that loading on the instrument may have caused high pressure against the cannula, leading to material removal as the instrument moved within the cannula. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1024216 | 1025446 | 984056 | 2955842200800077 | 1 | GCJ | Manufacturer | OtherUser facility | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-03-10 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument made the system go into fault mode. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and confirmed the instrument moved intuitively with full range of motion in all directions. Instrument recognition and engagement passed. A electrocautery cord was attached to the instrument and hooked up to a generator. The cautery feature functioned properly and no faults were created during activation. Engineering also observed the distal end of main tube has two 1. 5 inch long sections, that are 90 degrees apart, with material removed. The sections are 1. 5 inches and 4 inch above the tube extension and parallel to tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1024218 | 1025447 | 984057 | 2955842200800189 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-04-04 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-03-18 | Yes | No | Reuse | 2008
40018005 | 2008-03-10 | Other | Other | It was reported that during a da vinci hysterectomy procedure, during the posterior colpotomy, a pin hole on monopolar curved scissors (mcs) tip cover accessory was observed where electrical energy had shot through. The planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover as having six holes ranging in size from. 010 inches to. 025 inches burned through the silicone. The position of the holes are away from parting lines with distances from silicone to sleeve interface ranging from. 070" to. 615". Four holes are aligned along the same line, with the remaining two holes radially offset by 45 degrees. All holes exhibit localized melting of the silicone, suggesting multiple arcing events occurred. There are also various mechanical tears in the silicone adjacent to a few of the holes. Tearing is likely due to instrument collisions. Collisions may have damaged the silicone, locally reducing dielectric strength and creating a path for arcing. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. The mcs instrument was also returned for evaluation. Engineering found a. 010 inch by. 015 inch char mark on the proximal clevis, which lines up with the hole on the tip cover. Char mark may be the location where energy to escaped. Although tip cover has multiple holes, no further signs of arcing could be found on the instrument. The instrument does not have any burrs or abnormally sharp edges that would damage the silicone. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1029330 | 1025509 | 984119 | 2955842200800185 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2008-04-04 | 2008-04-04 | Yes | Yes | Service Personnel | 40018005 | Yes | No | 2008-03-10 | Yes | Yes | Initial | 2008
IS2000 A51P4 | 2008-03-10 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced system error code 20005. The surgeon had difficulty finding the plain between the prostate, and was concerned since the pt had been under anesthesia for approx four hours. The surgeon decided to convert to an open surgical technique to complete the planned procedure. No pt harm was reported. | The investigation conducted by isi field service engineering determined that the actual system error code experienced by the customer was # 23003. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code # 23003 appears when the da vinci s safety system detects that a system arm was "bumped" or restricted from moving properly. Add'l info provided by the isi rep in attendance at the case indicated that two system arms had collided. System error code # 23003 was cleared and no add'l system issues were encountered. The isi rep in attendance at the case also indicated that the surgeon was uncomfortable as this was their first system procedure performed with no attending proctor and that poor port placement may have contributed to the difficulties experienced by the surgeon. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1027299 | 1025510 | 984120 | 2955842200800184 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-03-28 | 2008-04-04 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No |  | Device Not Retu | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was broken, and no pieces fell inside of the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the top plug riser pin is pushed into the back end of the instrument and the chassis feature that retains the pins is broken. It was determined that the damage was most likely due to mishandling. Engineering also observed the distal end of the main tube has a 3" long section with light material removed. The damaged area is parallel to main tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1022479 | 1022638 | 981309 | 2955842200800161 | 1 | GEI | Manufacturer | OtherUser facility | Unknown | Initial | 2008-03-31 | 2008-03-31 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2008-03-04 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | The site returned a permanent cautery spatula instrument with no information provided. No harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the distal end of the instrument main tube has a 3. 5" long section directly above the proximal clevis with material removed. There is also a narrower, 1. 5" long section with material removed at the midpoint of the tube. Both damaged areas are aligned with each other and parallel to the tube's longitudinal axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. The site also returned a cannula accessory which may have caused the damage observed on the instrument main tube. Additional information regarding the cannula accessory can be found in medwatch mfr report #2955842-2008-00156. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1022477 | 1022639 | 981310 | 2955842200800162 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-03-31 | 2008-03-31 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-03-04 | Yes | No | Reuse | 2008
VER02 8 MM | 2008-03-28 | Malfunction |  | During robotic operative procedure, the plastic hub on the tip of the da vinci hook cautery was noted to be missing. The instrument was removed. A portion of the plastic hub was recovered from the abdomen. Another piece of the hub was not found and presumed to remain in the patient. This re-usable piece of equipment had one 'use' left in its allotted number of uses. |  | DA VINCI ENDOWRIST PERMANENT CAUTERY HOOK | INSTRUMENT SURGICAL ROBOTIC CA | 995121 | 1028543 | 987054 | 1028543 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2008-03-31 | 2008-03-31 | No | Yes | Physician | VER02 8 MM | Yes | No Answer Provided |  |  |  |  | 2008
42017206 | 2007-11-27 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the jaws of the maryland bipolar forceps instrument were misaligned and had a lack of grasping functionality. The planned surgical procedure was completed with a replacement instrument. No add'l information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip is bent, causing side to side misalignment of grips. There is a. 045 inch offset at the tips. The bent grip has a slight separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Electrical continuity passed. Engineering also observed the distal end of the main tube has various deep scratch marks with material removed. The tube has various randomly oriented scratches and one damaged section that is nearly parallel to the tube axis indicating that scratches are likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1029112 | 1022772 | 981443 | 2955842200800178 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-07 | Yes | No | Reuse | 2008
|  | Malfunction |  | It was reported that an instrument was not retracting completely from a system arm. The system passed all verification tests conducted by an isi field service engineer (fse) while onsite. The fse discovered one da vinci s instrument with tube abrasions. Instrument tube abrasions may have been caused by the dented cannula reported under medwatch mfr report #2955842-2008-00163. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for further eval and failure analysis as the account discarded the instrument. Based on the observations of the isi fse, the tube abrasion damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | 8MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1022492 | 1022773 | 981444 | 2955842200800164 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional |  | No | No | 2008-03-04 | Device Not Retu | No | Unkown | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument shaft was splintering. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube (customer referred to as "shaft") had material removed from a combination of deep scratches and scraping. There are various scratch marks all around the tube that are randomly oriented in relation to the tube axis. These scratches were likely caused by instrument collisions. Additionally, there are a few long, narrow sections parallel to the tube axis with material removed. It is possible these may have been caused by a defective cannula. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1026561 | 1022774 | 981445 | 2955842200800172 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-03-06 | Yes | No | Reuse | 2008
420002 |  | Malfunction |  | It was reported that an instrument was not retracting completely from a system arm. The system passed all verification tests conducted by an isi field service engineer (fse) while onsite. The fse discovered one da vinci s instrument with tube abrasions. Instrument tube abrasions may have caused by the dented cannula reported under medwatch mfr report# 2955842-2008-00164. No pt harm, adverse outcome or injury was reported. | The cannula accessory has not been returned for further evaluated and failure analysis as the account discarded the cannula accessory. Per the on site evaluation by an isi fse, the deformation of the tip has the potential to cause main tube scraping in certain loading conditions and may have contributed to the system arm retraction issue experienced by the customer. | 8MM REGULAR CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 1021544 | 1022775 | 981446 | 2955842200800163 | 1 | GCJ | Manufacturer | Company RepresentativeOther | Other | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 420002 | No | No | 2008-03-07 | Device Not Retu | No | Unkown | 2008
42000605 |  | Malfunction |  | It was reported the endo of the tip on the large needle driver instrument is scratched. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the metal components at the tip are not scratched. The instrument moves intuitively in all directions when the input discs are manually rotated. Engineering also observed the distal end of the main tube has a couple of sections with material removed. One section is. 875 inches long by. 180 inches wide and parallel to the tube axis with a rough surface finish. The wear pattern is consistent with cannula related tube abrasions. Above this section there is a broader section that appears to be a combination of scratches and scraping. This section has less material removed. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1021373 | 1022776 | 981447 | 2955842200800171 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-14 | 2008-03-28 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-03-06 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s radical prostatectomy surgical procedure, the monopolar curved scissors instrument stopped working. The account also indicated that this was the third procedural use of this instrument. The planned surgical procedure was completed with another instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering placed the instrument on a da vinci s test sys. Sys recognition passed. The test sys also showed zero uses remaining, indicating that the instrument had expired. Engineering also observed the instrument's main tube had a 6 inch long section, directly above the tube extension, missing material around its perimeter. Wear pattern is similar to cannula related tube abrasions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1021364 | 1022777 | 981448 | 2955842200800175 | 1 | GEI | Manufacturer | User facilityOther | BIOMEDICAL ENGINEER | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-03-06 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s radical prostatectomy surgical procedure, the pins at the end of the maryland bipolar forceps instrument were not "ok" and the bipolar cable could not be used. The planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the bottom plug riser pin is bent so that the top and bottom pins are not parallel. As a result, a bipolar cord cannot be attached to the instrument. It was determined that the damage was most likely due to mishandling where an impact force on pin likely caused bending. Engineering also observed the distal end of main tube has a couple sections parallel to the tube axis with light material removal. Sections are roughly 1. 5" and 1. 8" long, 180 degrees apart. Wear pattern is consistent with cannula related tube abrasions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1021355 | 1022778 | 981449 | 2955842200800174 | 1 | GEI | Manufacturer | User facilityOther | BIOMEDICAL ENGINEER | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-06 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that prior to starting a da vinci s prostatectomy surgical procedure, the monopolar curved scissors instrument did not work properly due to the wheel on the back of the instrument which seems to be stuck. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated by engineering using a da vinci s test system. Instrument recognition and engagement were confirmed. Instrument movement was intuitive, with full range of motion in all directions, indicating that none of the input disks (which the customer referred to as the "wheel") were stuck. Blades opened and closed properly. Latex cut test passed. The alleged failure could not be duplicated. Engineering also observed the distal end of the instrument main tube had a 1" long section with light material removal. Damaged area is 3. 8" above the tube extension, parallel to the tube axis, and has a rough surface finish. Wear pattern on tube is consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1025020 | 1022779 | 981450 | 2955842200800177 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-07 | Yes | No | Reuse | 2008
42017206 | 2008-02-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the maryland bipolar forceps instrument broke. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly and instrument recognition and engagement passed. Engineering also observed the distal end of the main tube has 2 inch long section with material removed. The damaged area is 3. 5 inches above the proximal clevis and is parallel to the main tube axis. The wear pattern is consistent with cannula related tube abrasions. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1025016 | 1022781 | 976474 | 2955842200800166 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-04 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissor instrument did not work. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed distal end of the main tube has a. 175" x. 120" oblong hole burned through it. Tube exhibits localized charring/melting, indicating an arcing event occurred. Instrument was disassembled to find the tube extension seal melted in the same location as the main tube, with one of the hypotubes directly below the tube extension having charred material on it. Main tube has a crack parallel to the tube axis directly below the hole that runs all the way down the tube, creating a pathway for arcing to occur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019522 | 1022782 | 981452 | 2955842200800179 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-03-08 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated by engineering. Instrument was driven on system and moved intuitively with full range of motion in all directions. Grips opened and closed properly. Instrument is functionally acceptable. The alleged failure could not be duplicated. Engineering also observed the distal end of the instrument main tube had a 6" long section directly above the proximal clevis with material removed all around the tube. Damaged area is parallel to tube axis and has a rough surface finish. There is also a 2" long section above the midpoint of the tube with light material removal. Wear pattern in both cases is similar to cannula related tube abrasions. The position of the damaged area above the main tube suggests damage from the internal bowl-tube transition of a cannula. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019521 | 1022783 | 981453 | 2955842200800181 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-08 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument did not work. No additional information was provided. No patient harm adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering verified that one conductor wire is broken at the yaw pulley exit. Yaw pulley shows no signs of arcing. Grip cables are not derailed at wrist. Wire insulation does not appear cut or melted. Electrical continuity failed. Engineering also observed the instruments distal end of the main tube has scratches and scrape marks with material removed. Scrape mark is parallel to tube axis and has a wear pattern consistent with cannula related tube abrasions. Scratch marks were likely caused by instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019580 | 1022784 | 981454 | 2955842200800180 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-03-08 | Yes | No | Reuse | 2008
42017206 | 2008-03-04 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the surgeon noticed shavings from the cannula accessory inside the pt. The instrument and cannula accessory had been in use for 10 mins when the shavings were observed. The shavings were retrieved and the planned procedure was completed without delay using the same instrument and cannula accessory. No pt harm, adverse outcome or injury was reported. | Results and conclusions: the instrument and cannula accessory were returned and evaluated. Per the customer reported complaint, engineering found a 2. 5" long section of material missing from the instrument main tube. Damaged area is located roughly 4" above the proximal clevis, parallel to the tube's longitudinal axis, with a rough surface finish. Wear pattern on tube was consistent with cannula related tube abrasions. The cannula does not show any obvious damage. Engineering determined that the instrument main tube damage may have been caused by high pressure against the cannula accessory during loading. Per the case notes, the instrument "shavings" were successfully retrieved from the pt. There are also a couple of cosmetic, surface scratch marks (mo material removed) below the damaged area that was likely caused by instrument collisions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1020097 | 1022785 | 981455 | 2955842200800165 | 1 | GEI | Manufacturer | OtherCompany Representative | UNKNOWN | Initial | 2008-03-28 | 2008-03-28 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-03-17 | Yes | No | Reuse | 2008
40018005 |  | Other | Other | It was reported that near the end of a vinci total hysterectomy surgical procedure, the surgeon had the wrist of the monopolar curved scissors instrument at a 90 degree angle while completing the posterior cuff of the colpotomy (removal of the uterus from vagina). After a brief period, the surgeon noticed sparks coming out of the wrist of the instrument. Upon further inspection, a hole in the mcs tip cover accessory was discovered. Further, a burn mark was observed on the pt's uterus that was already being removed. The tip cover was immediately removed and replaced with a new tip cover. Using the same instrument, the planned procedure was completed with no delay. No add'l pt harm, adverse outcome or injury was reported. | The msc instrument and tip cover accessory were returned and evaluated. Per the customer reported complaint, engineering observed that the ultem sleeve portion of the tip cover was missing; the returned silicone section had been cut off near the interface between the two materials. The silicone had four distinct holes burned through it, all of which exhibit localized melting consistent with electrical arcing. Holes are oblong and heart shaped, with one located at one parting line and the other three holes adjacent to each other, roughly 90 degrees from the first hole. Severe tears/scratches on the silicone were found near the three holes, indicating possible instrument collisions. Further, several mechanical tears were found at the distal end of the silicone, however, it can not be determined if these tears contributed to the arcing. The instrument contained several char spots on the distal clevis and ear of the proximal clevis, further suggesting arcing was the failure mode. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1022472 | 1023503 | 982194 | 2955842200800170 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-28 | 2008-03-28 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-03-10 | Yes | Yes | Initial | 2008
IS2000 A51P4 | 2008-03-05 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy surgical procedure, the customer was having difficulty positioning the endoscopic camera manipulator (ecm) to the pt due to the ecm clutch button not working. An isi technical support engineer (tse) had the customer restart the system, however, the system would not home and multiple occurrences of system error code # 23020 appeared. The isi tse had the customer exercise and actuate the ecm clutch switch several times as well as restart the system, however, the system error code #23020 recurred. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error code #23020 experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The ecm was returned to isi for failure analysis investigation. The #23020 system error code appears when one of the switches in the manipulator is showing inconsistent signals on it's two switch leads. Engineering eval discovered that the embedded serial board which contains a connection for the port clutch switch had malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1022456 | 1023505 | 982196 | 2955842200800169 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-28 | 2008-03-28 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-03-05 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-03-05 | Other | Other | It was reported that prior to a da vinci s cardiac mitral valve repair procedure, while starting up the system, the customer experienced recurring system error code #25589. An isi technical support engineer had the customer turn off the system and reset the circuit breaker on the pt side cart (psc), however, when the system powered back up, the system error code #25589 reappeared. The planned surgical procedure was completed with the customer's spare da vinci system, with approx. A 30 minute delay. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the remote arm controller (rac) board. The system was repaired by replacing the affected rac. The rac is an electronic module comprised of four printed circuit assemblies and cooling fans. System error code #25589 appears when the da vinci s safety system determines during the power up self test that the remote arm controller board (rac) brakes failed the brake voltage test. Upon determining these conditions, the safety systems put a da vinci s in a "recoverable safe state". The rac was returned to isi for failure analysis investigation. Inspection was performed and faulty chips were discovered and replaced to repair the rac. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1022476 | 1023507 | 982198 | 2955842200800168 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-28 | 2008-03-28 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-03-05 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-03-06 | Other | Other | It was reported that during a da vinci s cardiac surgical procedure, a pt side manipulator (psm) arm would not allow the instrument jaws to fully open and close. The customer tried a different instrument on the psm; however, they experienced the same problem. The procedure was converted to traditional open surgical techniques to finish the planned surgery due to the psm issue experienced by the customer and pt bleeding unrelated to the use of the da vinci s surgical system. No pt harm was reported. | On site investigation and eval was conducted by field service engineering. System verification tests were performed and the customer reported failure mode could not be duplicated. Engineering believes re-seating the sterile adapter accessory would have corrected the issue, however, the right pt side manipulator (psm) was replaced as a precaution. The pt side manipulator is an instrument arm located on the pt side cart which provides the sterile interface for the endowrist instruments. The psm was returned to isi for failure analysis investigation. Mechanical and performance tests were conducted and the customer reported failure mode could not be reproduced. As of 2008, there have been no reported recurrences of the issue at this hospital. Add'l info was provided by the surgeon on march 27, 2008 indicating that the pt recovered successfully. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1022480 | 1023508 | 982199 | 2955842200800176 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-03-28 | 2008-03-28 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | Yes | 2008-03-06 | Yes | No | Reuse | 2008
ISI2000 A51P4 | 2008-03-05 | Other | Other | It was reported that during a da vinci s cardiac surgical procedure, the console surgeon experienced a "pt cart is not detected" message. With the assistance of a technical support engineer, the site attempted to troubleshoot this issue via telephone, however, the issue could not be resolved. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | On site investigation conducted by isi field service engineering (fse) discovered that the cable connecting the pt side cart (psc) to the surgeon console (sc) was defective. This cable passes video, audio, and data during system operation from the psc to the sc. The system was repaired by replacing the cable. As of march 25, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1022502 | 1023510 | 982199 | 2955842200800167 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-28 | 2008-03-28 | Yes | Yes | Health Professional | ISI2000 A51P4 | Yes | No | 2008-03-05 | Yes | No | Reuse | 2008
42009307 |  | No Answer Provided |  | It was reported that the shaft on the endowrist prograsp forceps instrument is "roughed up and leaving fiber slivers on the technician's thumb". No additional information was provided. No patient harm or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of main tube has a 1. 5 inch long area with light material removed. The damaged area is 2. 3 inches above the proximal clevis, parallel to the tube axis, and has a rough surface finish. The wear pattern is consistent with cannula related tube abrasions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Engineering also observed two areas on the tube directly above the proximal clevis ears with material removed due to deep scratches. Engineering determined the scratches are likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1023416 | 1023707 | 982398 | 2955842200800173 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-04-04 | 2008-03-28 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-03-06 | Yes | No | Reuse | 2008
IS2000 A51P2 | 2008-02-20 | Death | Death | On march 5, 2008, intuitive surgical was notified per a customer provided letter: "this gentlemen was diagnosed with severe mitral regurgitation. He was originally evaluated for possible candidacy for transmitral percutaneous repair. He was turned down due to mitral anatomy. The pt was referred here for possible robotic intervention and agreed to proceed with the surgery. Incidently, the pt had a tee which revealed a very calcified valve with severe mitral valve regurgitation and fair overall ejection fraction about 50%. The procedure was performed robotically, initially with failure of the mitral valve repair necessitating further bypass and mitral ring annuloplasty, which also failed-necessitating mitral valve replacement. Multiple transfusions were given intraoperatively (prbc, 16 units, plasma-14 units, cryo-40 units, platelets-90 units). The pt was in the operating room from 8:30am-12 midnight. Ultimately, the pt developed low cardiac output, cardiogenic shock and uncontrollable bleeding diathesis. Demise at 16:35. ". | On march 5, 2008, intuitive surgical spoke to the vice president of quality management at the account regarding the incident. The vice president of quality management provided a letter march 12, 2008, stating that the patient's previous medical history and anatomy contributed to the pt's death. No malfunction of the system was alleged by the hospital. The hospital has continued to use the da vinci s system to perform surgery on patients. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 987618 | 1020383 | 979096 | 2955842200800160 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-21 | 2008-03-26 | Yes | No | Health Professional | IS2000 A51P2 | No | No | 2008-03-05 | No Answer Provi | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument is broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument was returned attached to the cannula accessory and the instrument cannot be removed due to dislodgement of the metal proximal clevis from the main tube. The clevis is intact, however, a section of the main tube interface wall has collapsed and the main tube is missing pieces. Dislodgement is likely due to overloading with a force perpendicular to one ear of the clevis. Engineering also observed the distal end of the main tube has a 1. 5 inch section with light material removal. Damaged area is parallel to main tube and has a wear pattern consistent with cannula related tube abrasions. Based on the location and appearance, however, the cannula does not appear visibly damaged at the tip. High pressure against the inner edge of the cannula likely caused the material removal. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A f/u mdr will be submitted if add'l info is rec'd. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019804 | 1019764 | 978475 | 2955842200800152 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
42020504 | 2008-02-12 | Malfunction |  | It was reported that prior to starting a da vinci s hysterectomy surgical procedure, the fenestrated bipolar forceps instrument was not recognized by the sys. No other info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a test sys to check for recognition, recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. Engineering also observed the distal end of the main tube has various scratch marks, some of which are deep enough to have removed material from the tube. The scratches are randomly oriented, and although some are parallel to main tube axis, the size and relative proximity to each other suggests the scratches were caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to instruments. Do no use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019793 | 1019765 | 978476 | 2955842200800146 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-02-27 | Yes | No | Reuse | 2008
42019407 | 2008-02-14 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure, the mega needle driver instrument would not release from the sterile adapter. The customer pried the housing off in order to remove the instrument from the sterile adapter. The planned surgical procedure was completed with approx a ten min delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument returned with both levers and the housing detached from the chassis. The chassis has a broken piece on one side where the lever installs. Housing has broken snap tabs and housing pins, indicative of mishandling, however, it is unclear if this damage occurred in trying to remove the instrument from the sterile adapter. It is possible the instrument was placed on the sys without levers, since properly installed levers should not cause engagement issues. Engineering also observed the distal end of the main tube has material removed on one side. Damaged area is a combination of scratches and a scrape mark. The scrape mark is parallel to main tube axis and the wear pattern is consistent with cannula related tube abrasions. The scratches are likely due to instrument collisions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. The endowrist instruments instructions for use (ifu) specifically states; general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection from the instrument tip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1019781 | 1019766 | 978477 | 2955842200800144 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42019407 | Yes | No | 2008-02-26 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | It was reported that the permanent cautery spatula instrument has abrasions on its shaft. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one side of the main tube has a 4 inch long section directly above the proximal clevis with material removed. The damaged area is parallel to the tube's longitudinal axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage found. The site has recently returned a defective cannula accessory. The defective cannula accessory was reported under medwatch mfr report #9255842-2008-00156. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018646 | 1019767 | 978478 | 2955842200800157 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-02-29 | Yes | No | Unkown | 2008
42018304 |  | Malfunction |  | It was reported that the permanent cautery hook instrument axises on the back are not spinning. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed and the instrument moved intuitively with full range of motion in all directions. Engineering was not able to replicate the customer complaint of "axises not spinning" engineering also observed that the distal end of the main tube has various deep scratch marks with material removed. The scratches are randomly oriented relative to tube axis, suggesting they were caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such a disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019770 | 1019768 | 978479 | 2955842200800153 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument wire is broken and no pieces fell inside of the pt. No additional info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers and scissor grip. Half of the cable diameter is broken in each location and the pulley spins freely indicating that the cable broke under tensile loading. The other cables at wrist are not damaged. Engineering determined that the cable wear on pulleys and grip combined with high grasping force likely contributed to breakage. Engineering also observed the main tube has sections at the midpoint an proximal end with material removed. Damaged areas are parallel to main tube, but location of wearing is high. The damage area on proximal end is out of range of a cannula tip. Engineering concluded that the material removal may be due to a sharp internal bowl/tube interface, or from the customer's cleaning process. No other damage found. The endowrist instruments instructions for the (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from an other instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018656 | 1019769 | 978480 | 2955842200800150 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-04 | 2008-03-25 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported prior to a da vinci s surgical procedure, the bipolar maryland forceps instrument was not recognized by the system. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a test system to check for recognition and the instrumetn was successfully recognized. The grip and wrist move properly when inputs are manually rotated. Engineering also observed the distal end of the main tube has a 1. 75 inch long section with material removed on one side of the tube. Damaged area is located 2 inches above the proximal clevis, parallel to the tube axis, and has a rough surface finish. The wear pattern is consistent with cannula related tube abrasion. Electrical continuity passed. No other damage found. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018664 | 1019770 | 978481 | 2955842200800142 | 1 | GEI | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-26 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the monopolar curved scissors (mcs) instrument broke. No other info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument does not have any broken components. Instrument recognition and engagement tests passed as well and the instrument moved intuitively with full range of motion in all directions. At the backend, the luer plate is dislodged from the chassis without damage to the chassis feature that retains it. Two housing pins are broken, suggesting the instrument was mishandled. Engineering also observed the distal end of the main tube has a 1. 5 inch long section with material removed. The damaged area is located 3 inches above the tube extension, parallel to main tube axis, and has a wear pattern consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018671 | 1019771 | 978482 | 2955842200800140 | 1 | GEI | Manufacturer | User facilityOther | Physician | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-02-25 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | The site proactively returned a is2000 8mm regular cannula accessory to determine if the cannula may have caused the instrument tube abrasion reported under medwatch mfr report #2955842-2008-00157. No patient harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering observed the cannula is visibly dented at the distal tip and a. 342 gage pin does not fit through the distal tip inner diameter. The deformation of the tip has the potential to cause main tube scraping in certain loading conditions. Damage cannula is likely due to mishandling. No other damage found. | 8MM REGULAR CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 1018676 | 1019772 | 978483 | 2955842200800156 | 1 | GCJ | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42000202 | Yes | No | 2008-02-29 | Yes | No | Unkown | 2008
42017908 | 2008-03-11 | Malfunction | Other | It was reported that after a da vinci s hysterectomy surgical procedure, while the monopolar curved scissors (mcs) instrument was in central processing, one of the blades was found to be broken off and missing. One of the scrub techs stated they saw the blade fall off during the procedure, however, the tech did not provide the info to the surgeon during the case because they thought the fragment was only charred tissue. The account stated that excessive charring occurred during the procedure and that the mcs instrument was used only for cauterization and not for the scissor function. In 2008, the surgeon confirmed via x-ray that the mcs tip is inside of the pt. The surgeon has discussed the event with the pt and at this time no arrangements have been made to remove the mcs tip. | The instrument was retuned and evaluated. Per the customer reported complaint, engineering found one scissor blade is broken off at the cross section of the grip cable crimp. The broken blade was not returned. The cable crimp has slipped out of the crimp pocket due to the breakage. The grip cable is frayed and the crimp is scratched, however, this likely happened after the blade broke and the cable crimp moved across the sharp fractured edges. Site previously reported a broken scissor blade complaint leaving engineering to conclude that based on the similarity, which indicates possible corrosion, the corrosive cracking was likely to have caused the blade failure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1019830 | 1019773 | 978484 | 2955842200800159 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | Yes | Yes | Health Professional | 42017908 | Yes | No | 2008-03-11 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure, the fenestrated bipolar forceps instrument would not grasp. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument grips open and close properly when inputs are manually rotated. The pitch down cable is broken at the distal clevis, likely breaking during tensile loading. The clevis does not have any burrs, therefore, the cable may have worn against the proximal clevis cable hole. Engineering also observed the distal end of the main tube has material removed. The damaged area is parallel to main tube axis and has a wear pattern consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is rec'd. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1018683 | 1019774 | 978485 | 2955842200800143 | 1 | GEI | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-02-26 | Yes | No | Reuse | 2008
42019006 |  | Malfunction |  | It was reported that the cobra grasper instrument has a broken wire. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument does not have any broken cables but has both grip cables on one side of the distal clevis derailed at the idler pulleys. The inner grip close cable is seated on the outermost pulley and the outer grip close cable is exposed on the outside of the pulley. Cables likely stretched from heavy usage, creating excess cable slack. Engineering also observed the distal end of the main tube has a 1. 5 inch long section with material removed. Damaged area is parallel to main tube axis and surface is locally flat due to material removal. The wear pattern is consistent with cannula related tube abrasions. No other damage found. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1018692 | 1019775 | 978486 | 2955842200800154 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42019006 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
42022702 | 2008-02-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the pk dissecting forceps instrument could not be recognized when inserted on the sterile adapter. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a test system to check for recognition and the instrument was successfully recognized. Engineering also observed the distal end of the main tube has various scratch marks on one side, a couple of which are deep enough to have removed material from the tube. Scratches are not aligned with the tube axis and are spaced relatively close together, therefore, scratches were likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018698 | 1019776 | 978487 | 2955842200800141 | 1 | GEI | Manufacturer | DistributorOther | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-02-26 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument wire is broken and no pieces fell inside of the pt. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the proximal pulley and the pulley spins freely. The other cables at wrist are not damaged. Engineering determined that the cable broke under tensile loading and the cable wear on pulleys combined with high blade closing force likely contributed to breakage. Engineering also observed the distal end of main tube has a 3 inch long section with material removed. The damaged area is parallel to the tube axis and has a wear pattern similar to cannula related tube abrasions. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018707 | 1019777 | 978488 | 2955842200800151 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-04-04 | 2008-03-25 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
42014304 | 2008-02-21 | Malfunction |  | It was reported that the 5mm maryland dissector instrument would not grasp. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed the distal end of the main tube has a. 120 inch by. 130 inch section of insulation missing directly above the snake wrist. Engineering concluded that the material may have been removed due to pressure against the distal tip of the cannula accessory, which may have had a sharp inner edge. No other damage found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1018684 | 1019778 | 978489 | 2955842200800147 | 1 | GCJ | Manufacturer | OtherCompany Representative | NOT APPLICABLE | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-02-27 | Yes | No | Reuse | 2008
42020906 |  | Malfunction |  | It was reported that the suture cut needle driver instrument has a broken cable. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers and the pulley spins freely. The other cables at wrist are not damaged. Engineering concluded that the cable broke under tensile loading, and the cable wear on pulleys combined with high grasping force likely contributed to breakage. Engineering also observed that the distal end of main tube has a couple of deep scratched with material removed. Orientation and size of scratches indicated that the damage was caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1018679 | 1019779 | 978490 | 2955842200800139 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42020906 | Yes | No | 2008-02-25 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that when the customer opened the box of the large needle driver instrument, they observed that the shaft sheet was fractured. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument main tube is not fractured, but has a single deep scratch. 5 inches above the proximal clevis with a small amount of material removed the instrument was returned with eight (8) lives remaining indicating that it was used during two (2) surgical procedures. Engineering determined the scratch was likely caused by a collision with a hard object. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1022542 | 1019780 | 978491 | 2955842200800155 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-02-28 | Yes | No | Reuse | 2008
IS1200 A43P8 | 2008-02-26 | Other | Other | It was reported that during the start of a da vinci prostatectomy surgical procedure, the customer experienced a non-recoverable system error code #212. The surgeon decided to abort the planned surgical procedure. The patient incisions were closed and the case was rescheduled. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that system error code #212 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected right mtm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #212 appears when the da vinci safety system determines a voltage tracking fault reported by the digital signal processor (dsp). Upon determining these conditions, the safety system put da vinci in a "non-recoverable safe state". The right mtm was returned to isi for failure analysis investigation. Engineering concluded that an axis motor and encoder assembly had malfunctioned, thus generating the system error experienced by the customer. The rescheduled procedure was successfully completed the same day the affected mtmr was replaced. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1019711 | 1022531 | 981202 | 2955842200800145 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-21 | 2008-03-25 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2008-02-26 | Yes | No | Reuse | 2008
40018005 | 2008-02-27 | Other | Other | It was reported that during a da vinci s proctectomy procedure, the surgeon saw a plume of energy and then smoke in the surgical view due to burnt patient fat. The surgeon immediately stopped the surgical task being performed and removed the monopolar curved scissors instrument (mcs) and inspected the mcs tip cover accessory, at that time a hole in the mcs tip cover was noticed. The mcs tip cover was replaced and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has a 0. 25" diameter hole burned through the silicone. The hole is located along one of the parting lines, roughly. 215 inches above the silicone to ultem sleeve interface. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. Engineering determined the arcing is likely due to insufficient dielectric strength of the silicone caused by a small crack or void. The mcs instrument was also returned for evaluation. Engineering found charred material on one ear of the proximal clevis and on the distal clevis, possibly the location where energy to escaped. The charring on the proximal clevis lines up closely with the hole in the tip cover, however, with the wrist pitched slightly, the hole may also line up with the distal clevis char spots. Distal and proximal clevis do not have any abnormally sharp edges that would damage the silicone. Cables are not frayed and electrical continuity passed. No other damage found. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1019729 | 1022532 | 981203 | 2955842200800149 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-03-21 | 2008-03-25 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-02-27 | Yes | Yes | Initial | 2008
40018005 | 2008-02-29 | Other | Other | It was reported that in the middle of a da vinci proctectomy procedure, while dissecting the prostrate, the surgeon noticed sparks and a burn mark was observed on the prostate in a different location. The surgeon immediately stopped the surgical task being performed and the monopolar curved scissors mcs tip cover accessory was inspected, at that time a hole in the mcs tip cover was noticed. The mcs tip cover was replaced and the planned procedure was completed with no delay. No additional patient harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has three holes burned through the silicone. The holes range in size from. 008 inch in diameter to. 070 inch wide and was oblong in shape. The holes are positioned at. 300 inch,. 540 inch and. 565 inch above the silicone to sleeve interface, with the. 008 inch hole adjacent to the oblong hole. All holes exhibit localized melting of the silicone, indicating multiple arcing events occurred. Engineering determined the arcing is due to insufficient dielectric strength of the silicone, and high cautery settings per the ifu, maximum monopolar setting for valleylab force fx generator in coag fulgurate mode is 38 watts, however, site has been known to use 20 to 50 watts during their surgical procedures. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1019744 | 1022533 | 981203 | 2955842200800158 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-21 | 2008-03-25 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-03-04 | Yes | Yes | Initial | 2008
IS1200 A43P8 | 2008-02-27 | Other | Other | It was reported that during da vinci cardiac atrial septal defect repair procedure, the console surgeon experienced a distorted 2d view in the left eye of the high resolution stereo viewer (hrsv). With the assistance of a technical support engineer, the site attempted to troubleshoot this problem via telephone, however, the issue could not be resolved. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering (fse) discovered that the high magnification camera that captures the image which then gets transferred to the surgeon console's high resolution stereo viewer (hrsv) was defective. The hrsv provides the video image for the surgeon console operator. The surgeon console is the control center where the operator sits outside the sterile field and controls instruments and a 3d endoscope with his/her hands. The system was repaired by replacing the high magnification camera. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1016991 | 1018405 | 977131 | 2955842200800148 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-21 | 2008-03-21 | Yes | Yes | Health Professional | IS1200 A43P8 | No | No | 2008-02-27 | Yes | Yes | Reuse | 2008
|  | Other | Other |  | On january 15, 2008, the site reported experiencing system error code #23 during system start up, prior to starting the procedure. Isi technical support engineer had customer cycle the power on, the pt side cart twice and the system homed and performed normally with no other problems on the second cycle. Approximately 90 minutes later, the customer called back stating that the led went out on one of the pt side manipulator arms (psm). The psm was stowed and the procedure was completed with the remaining two system arms. The investigation conducted by field service engineering (fse) concluded that the system error #23 experienced by the customer was associated with the configured embedded serializer set up joint (cfg, essj). The system was repaired by replacing the affected cfg, essj. System error code #23 is reported by software to denote that the hardware wheel "wdog" has tripped on one of the digital communication links in the system (due to an excessive number of retries on hardware message packets). This means that the system cannot reliably communicate over that digital link and therefore cannot continue normal operation. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. An evaluaiton of the affected psm led was performed by the fse and found no issues. As of march 20, 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 989347 | 1023172 | 969410 | 2955842200800134 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-03-14 | 2008-03-21 | Yes | Yes | Invalid Data |  | No | No Answer Provided | 2008-02-22 | Device Not Retu | No | Reuse | 2008
PS2000 | 2008-03-04 | Malfunction |  | Physician was using intuitive da vinci robot when camera arm stopped working during a case and listed non-recoverable error 23020 fault. Camera arm was replaced next day with a re-manufactured arm and tested by intuitive engineer. Same arm failed again just prior to the start of next scheduled case. Engineer replaced a second camera arm. As a precaution, two other robotic arms where replaced to match software version of the third. ======================manufacturer response for da vinci robot endoscopic instrument control system: manufacturer was sympathetic and regretted circumstances surrounding robot arm failing twice in a row. Manufacturer offered to add hospital to their "customer care plan," which is designed to increase customer confidence by having service engineer download machine values after each use. In theory, downloaded values will identify issues prior to them happening. |  | DA VINCI ROBOT ENDOSCOPIC INSTRUMENT CONTROL SYSTEM | ARM ROBOTIC DEVICE | 991749 | 1025515 | 984125 | 1025515 | 1 | NAY | User Facility |  |  | Initial | 2008-03-18 | 2008-03-21 | No | Yes | Physician | PS2000 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
IS1200 A43P8 | 2008-01-25 | Other | Other | It was reported that during a da vinci sigmoidectomy surgical procedure, while releasing the colon, a sys pt side manipulator (psm) was working intermittently. No pt harm, adverse outcome, or injury was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering determined that the system, and the affected pt side manipulator (psm). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The psms are working to spec which leads to isi field engineering to conclude that the customer reported failure mode may be user preference related. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1014956 | 1016399 | 975190 | 2955842200800120 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-14 | 2008-03-14 | Yes | Yes | Health Professional | IS1200 A43P8 | No | No | 2008-01-25 | Device Not Retu | No | Reuse | 2008
40018005 | 2008-02-22 | Other | Other | It was reported that towards the end of a da vinci s myomectomy procedure, while the surgeon was attempting to cauterize tissue in the patient's abdomen, he noticed sparking coming from an area around the uterus. Upon further inspection, the surgeon located a burn mark on the patient's uterus. The surgeon immediately stopped using and the monopolar curved scissors instrument (mcs) and the mcs tip cover accessory was inspected, at that time, a hole in the mcs tip cover was noticed. The mcs tip cover was replaced and the planned procedure was completed with approximately a 15 minute delay. No additional patient harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed four distinct holes in the silicone. Three holes are spaced relatively close together on one side of the silicone, and the fourth hole is roughly 90 degrees from the other holes. The holes range in size from. 010" to. 030" and are. 150" to. 315" above the silicone to sleeve interface. Localized melting of silicone suggests failure is due to multiple arcing events through the sleeve. Also observed were a couple of scratches on the inside surface of the silicone near the three holes. Arcing may be due to prolonged cautery in combination with damage in silicone. Engineering believes it is possible instrument collisions may have damaged the silicone on the inner or outer surface, weakening the electrical insulating capacity and creating a path for energy to escape. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1016156 | 1016400 | 976464 | 2955842200800131 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-02-22 | Yes | Yes | Initial | 2008
IS2000 A51P2 | 2008-02-18 | Other | Other | It was reported that prior to a da vinci s prostatectomy surgical procedure, while docking the system to the patient, the customer experienced recurring system error code #25589. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the remote arm controller board (rac1). The system was repaired by replacing the affected remote arm controller board (rac). The remote arm controller (rac) is an electronic module comprised of four printed circuit assemblies (pcas) and cooling fans. In 2008, the site experienced system error code #25589. An isi field service engineer replaced the endoscopic camera manipulator (ecm) arm, the rac, and two patient side power distribution boards (ppd). The system functioned as designed and there was no injury to the patient. System error code #25589 appears when the da vinci s safety system determines during the power up self test that the remote arm controller board (rac) brakes failed the brake voltage test. Upon determining these conditions, the safety system put da vinci s in a "recoverable safe state. " the rac was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. The problem could have resulted from a fault in the power subsystem that supplies voltage to the brakes and the ppd that was replaced would have addressed this potential cause. The brake itself could have had a problem and the ecm that was replaced would have addressed this last potential cause. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1016170 | 1016401 | 975192 | 2955842200800121 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-14 | 2008-03-14 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2008-02-18 | Yes | No | Reuse | 2008
IS1200 A43P8 | 2008-02-21 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, the customer experienced a persistent #20009 system error codes. An isi technical support engineer determined that the fault was related to a system arm and had the customer undock and exercise the affected arm multiple times and reboot the system, however, the 20009 system error code continued to recur as soon as the arm was moved. The pt was under anesthesia and port incisions were made when the surgeon decided to abort and reschedule the planned surgical procedure to a later date. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with a patient side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The psm was returned to isi for failure analysis investigation. Engineering discovered both 20009 and 21009 in the error logs. The 20009 abd 21009 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety system put da vinci in a "recoverable safe state". Engineering concluded that the motor for the psm axis two had malfunctioned, thus generating the system errors experienced by the customer. The psm was repaired by replacing the encoder and motor assembly for axis two. As of march 12, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1016317 | 1016404 | 975195 | 2955842200800129 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-03-14 | 2008-03-14 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2008-02-21 | Yes | No | Reuse | 2008
40017905 |  | Malfunction |  | It was reported that the orange cover is cracked on the monopolar curved scissors instrument. No additional info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension broken and missing a. 180 inch by. 130 inch piece at the proximal clevis interface. The clevis is dislodged from the tube extension. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a pt during a surgical procedure. Engineering also observed the main tube has a deep scratch just below the midpoint with material removed. The scratch is. 400 inch long and is diagonal to the tube's longitudinal axis. The main tube damage was likely caused by inadvertent contact with another object, based on the location. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT ACC | 1018355 | 1017732 | 976462 | 2955842200800126 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 40017905 | Yes | No | 2008-02-21 | Yes | No | Reuse | 2008
42017908 | 2008-01-11 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, while picking up tissue, one side of the tip of the monopolar curved scissors instrument "cracked" and fell into the pt. The fragment was immediately removed. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported that "one side of the tip is broken", however, failure investigation confirmed that the left blade is broken off at the cross section of the grip cable crimp. Customer was likely referring to the scissor blade when stating that the "one side of the tip is broken". The broken blade was not returned, per the case notes, the broken instrument component was successfully retrieved from the pt. The break plane of the blade is not totally flat as there are a couple of changes in elevation along with a crack at the crimp pocket. Intact blade shows signs of light pitting on the non-cutting edge near the tip, indicating possible corrosion which may have caused corrosive cracking, causing the blade failure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016184 | 1017733 | 976463 | 2955842200800130 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-22 | Yes | No | Reuse | 2008
40018005 | 2008-02-19 | Malfunction |  | It was reported that during a da vinci s oophorectomy surgical procedure a burn hole in the monopolar curved scissors (mcs) tip cover accessory was observed. The mcs tip cover was replaced and the planned procedure was completed without a delay. No pt harm, adverse outcome or injury was reported. | The msc tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found no holes in the silicone and no signs of melting anywhere on the tip cover that would indicate an arcing event occurred. The tip cover does have a mechanical tear in the silicone at the distal tip, however, there is no material removed. The tear is. 150 inch long and parallel to the longitudinal axis, offset by. 030 inch from one of the parting lines. It was determined the tear is likely due to repeated opening and closing of the monopolar curved scissors instrument. The silicone may have had a small crack or flaw at the distal tip that propagated proximally during articulation of the wrist. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1016201 | 1017734 | 976464 | 2955842200800135 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-02-19 | Yes | Yes | Initial | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument wire is broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip close cable is broken at the distal idlers and the pulley spins freely. Engineering believes the cable broke under tensile loading. Other cables at wrist are not damaged. The cable wear on pulleys combined with high blade closing force may have contributed to the breakage. Engineering also observed the distal end of the main tube has a section with material removed all around the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. The damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016345 | 1017735 | 976465 | 2955842200800137 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-23 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that pk dissecting forceps instrument was not recognized by the davinci s sys. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed instrument on sys and recognition was confirmed. Instrument moved intuitively with full range of motion in all directions. Engineering also observed the distal end of the main tube has various scratch marks, a couple of which have removed material from the main tube. Most scratches are randomly oriented relative to the tubes longitudinal axis, suggesting they were caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016770 | 1017736 | 976466 | 2955842200800125 | 1 | GEI | Manufacturer | Other | Other | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-02-20 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the housing from the maryland bipolar forceps instrument came off. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the luer plate is dislodged from the chassis and pushed into the back end of the instrument. The chassis feature that retains the luer plate is not damaged, suggesting the dislodgement occurred due to the housing coming off. Engineering believes the housing came off due to mishandling. Engineering also observed the main tube has two distinct sections with material removal. One section is directly above the proximal clevis, and the other is roughly 4. 5 inches above the proximal clevis. They are offset by 90 degrees from each section. Both damaged areas are parallel to the tube axis, and has a rougher surface finish than the rest of the tube. The wear pattern is typical of damage caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPTS INSTRUMEENT | ELECTROSURGICAL INSTRUMENT | 1016249 | 1017737 | 976467 | 2955842200800136 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017206 | Yes | Unknown | 2008-02-23 | Yes | No | Reuse | 2008
4017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument does not work. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed scissors do not cut cleanly through. 006 inches of latex. Latex gets snagged at scissor tips. The blades are not damaged, but blade edges are slightly rounded at the tips, negatively affecting cut performance. The blade crossover at the tips appears less than usual. At the back end, three of the four housing snaps are broken, engineering believes it is due to mishandling. Electrical continuity passed. Engineering also observed the distal end of the main tube has a three inch long section directly above the extension with material removed. The damaged area is parallel to the tube axis and has a wear pattern similar to cannula related tube abrasions. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016355 | 1017738 | 976468 | 2955842200800128 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 4017908 | Yes | No | 2008-02-21 | Yes | No | Reuse | 2008
40022702 |  | Malfunction |  | It was reported that during a da vinci surgical procedure, the pk dissecting forceps instrument did not work. The procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire is broken at the yaw pulley exit. Yaw pulley shows no signs of arcing. Grip cables are not derailed at wrist. Wire insulation does not appear cut or melted. The wire likely was not securely attached to the grip and had poor strain relief, and pulled out during articulation of the wrist. Electrical continuity failed. Engineering also observed the distal end of main tube has various deep scratches with material removed. Scratches are not aligned with tube axis. Size and orientation of scratch marks suggest they were caused by multiple instrument collisions. No other damage found. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1016782 | 1017739 | 976469 | 2955842200800122 | 1 | GEI | Manufacturer | Other | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 40022702 | Yes | No | 2008-02-19 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tips of the monopolar curved scissors instrument will not close. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the scissor blades do not close due to damage at the base of one blade. One blade edge has two sections with material removed at the base of the blade and the blade catches on one of the chipped edges of the damaged blade when closing. Engineering also observed a damaged section on one side of the tube with a 2 inch long scraped section and various scratch marks surrounding it. The scratches are likely due to instrument collisions and based on the location and appearance of the scrape marks, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is rec'd. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016808 | 1017740 | 976470 | 2955842200800123 | 1 | GEI | Manufacturer | Other | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-19 | Yes | No | Reuse | 2008
42017206 | 2007-11-28 | Malfunction |  | It was reported that during a da vinci s prostectomy procedure the coating from the maryland bipolar forceps instrument was failing. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was retuned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of the main tube has material removed due to combination of scratches and scraping on one side of the tube. There are various deep scratches concentrated in a 1. 8 inch long section, which is. 5 inches above the proximal clevis. The scratches are not aligned with the tube axis, which engineering believes were caused by instrument collisions. Above the scratch marks, the tube exhibits a 2. 25 inch long wide wear pattern consistent with cannula related tube abrasion. The damaged section is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. The damage may have been caused by a defective cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended. Engineering also observed the yaw pulleys have charring at the grip base with one yaw pulley missing a section to expose the cable crimp and the bare wire at the yaw pulley exit. The cable crimp is not completely retained within yaw pulley and there fore grasping function is not precise. It is unclear how yaw pulley section broke. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1018629 | 1017741 | 976471 | 2955842200800124 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-20 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that cables of the endowrist prograsp forceps instrument are loose. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip open cable is frayed at the distal idlers, but the grips can still open and close. Engineering believes the cable may have derailed onto the pulley edge during articulation of wrist and became frayed. Engineering also observed the distal end of the main tube has two sections located roughly 90 degrees apart with light material removal. Both sections are parallel to the tube axis, and have a rougher surface finish than the rest of the tube. The tube damage may be due to a defective cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional info is received. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1016263 | 1017742 | 976472 | 2955842200800132 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-02-22 | Yes | No | Reuse | 2008
IS1200 A43P8 | 2008-01-22 | Malfunction | Other | It was reported that after a da vinci sacrocolpopexy surgical procedure, the surgical staff found the pt's skin surrounding the right port incision where the cannula accessory was inserted, to be "burned". The surgeon used a scalpel to remove the damaged skin, roughly a 1 mm area, and closed the port site. The case was not significantly lengthened due to this event. The isi clinical sales rep present for the case spoke to the surgeon a week after the procedure, at which time the surgeon indicated that the pt was doing fine with no other external trauma issues. No add'l pt harm, adverse outcome or injury was reported. | The instruments and accessories used during the reported case have not been returned or evaluated at this time. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the pt side manipulator arms to ground, through the cannula accessory to the pt, if the pt is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. The correct electrical surgical unit (esu) settings have been confirmed at this site, during the procedure. The reported failure could not be confirmed nor denied. A f/u mdr will be submitted if add'l info is rec'd. As of march 12, 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1016759 | 1017743 | 962495 | 2955842200800119 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-03-14 | 2008-03-14 | Yes | Yes | Health Professional | IS1200 A43P8 | No | No | 2008-01-22 | Device Not Retu | No | Unkown | 2008
42017206 | 2008-02-14 | Malfunction |  | It was reported that the tips of the maryland bipolar forceps instrument are bent and do not meet. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one grip is bent, causing side to side misalignment of the grips. The grips are offset at the tips by. 075 inches. The bent grip has the glue joint intact, however, engineering believes the failure is likely due to overloading at the tp. Electric continuity passed. Engineering also observed the distal end of the main tube has a 2 inch long section with light material removed on one side of the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Engineering believes the damage may be caused by a defective cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | IMARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016208 | 1017744 | 976474 | 2955842200800133 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-22 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy surgical procedure, the pk dissecting forceps instrument was found to have a broken cable. Nothing fell into the pt. The procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip cable sticking out at the wrist, but not visibly damaged, suggesting cable failure at the back end. Housing was removed to find one grip cable broken at the proximal end of the hypotube. Evidence suggests failure due to chemical attack. Site has admitted to using peroxide in cleaning process, but has now stopped. Peroxide is known to damage cables. Instrument was likely exposed to peroxide, reducing cable tensile strength due to corrosion. Engineering also observed the distal end of main tube has material removed from what appears to be a combination of deep scratches and a scrape mark. Tube has various scratch marks that are randomly oriented, suggesting they were caused by instrument collisions. There is also one scrape mark directly above the proximal clevis with wear pattern similar to a cannula related tube abrasion. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Do not expose instrument to hydrogen (h2o), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1018622 | 1017745 | 976475 | 2955842200800134 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-02-23 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the bipolar connector is bent on the fenestrated bipolar forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the top plug riser pin is pushed into back end and the bottom pin is bent. The chassis feature that retains pins is broken. Instrument was likely mishandled, breaking the chassis feature. Electrical continuity passed. Engineering also observed the distal end of the main tube has material removed from a combination of scratches and scrape marks. Tube damage is on one side of the tube and has a rough surface finish. Scratches are likely due to instrument collisions. Scrape marks have a wear pattern parallel to tube axis, suggesting they could have been caused by a defective cannula. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instrument instructions or use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1016215 | 1017746 | 976476 | 2955842200800127 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-14 | 2008-03-14 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-02-21 | Yes | No | Reuse | 2008
420179 | 2008-02-15 | Injury | Disability | It was reported that during the isolation of the left ureter during a da vinci s left ureteral stricture procedure, a hole had developed in the tip cover of the monopolar curved scissors instrument, allowing energy to escape, burning a hole in the ureter. Upon surgical analyses, it was confirmed that the ureter was non viable. Based on the condition of the ureter and the pt's age, the site's surgical staff determined that the pt's best option would be to undergo a nephrectomy. Surgical staff did not believed that the pt's best option would be to undergo a secondary surgery to attempt to repair the ureter. It was reported that the pt recovered well and was discharged 2 days after the surgery without complications. | The site has not returned the instrument for eval. If the instrument is returned, failure analyses will be performed and a f/u mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 979520 | 1009715 | 968623 | 2955842200800111 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | 420179 | No | Yes | 2008-02-15 | Device Not Retu | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument would not cut. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the scissors do not fully close due to damage on one blade. One blade has a burr on the shearing surface just below the midpoint that acts as a mechanical stop for the other blade. As a result, scissors cannot cut latex through entire blade length. Burr appears to have been caused by shearing action of blades. Electrical continuity passed. Engineering also observed the main tube to have an 1. 2 inch long section with light material removal on one side of the tube. The damaged area is 3. 5 inches above tube extension, parallel to tube axis and has a rougher surface finish than the remainder of the tube. Tube damage may be caused by a defective cannula. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1007247 | 1009794 | 968701 | 2955842200800116 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument would not cut. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the scissors do not cut cleanly through. 006 inch of latex. The latex gets snagged at scissor tips. The blades are not damaged. The blade edges are slightly rounded at the tips and are negatively affecting cut performance. The decreased cable tension and or belleville spring force may also contribute to poor cutting. The electrical continuity passed. Engineering also observed the distal end of the main tube to have a section with material removed all around the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. Based on the location and appearance, the damage may have been caused by a defective cannula. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1007243 | 1009795 | 968702 | 2955842200800117 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument would not open. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions. The blades open and close properly. The latex cut test passed. Engineering also observed the distal end of the main tube to have a 2 inch long section with material removed, located 1. 35 inches above the tube extension. The damaged area covers one side of the tube, is parallel to tube axis, and has a rough surface finish. Engineering believes the damage may be due to a defective cannula. Add'l investigation is underway regarding the cannula accessories. Engineering also observed a couple of deep scratches on the tube, roughly 180 degrees from the scraped area. The scratches also have material removed and are not aligned with tube axis, suggesting they were caused by instrument collisions. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1007236 | 1009796 | 968703 | 2955842200800118 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-06 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument is splintered. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have a 1. 2 inch long section with material removed on one side of the tube. The damaged area is 4 inches above the tube extension, parallel to the tube's longitudinal axis, and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1006288 | 1009797 | 968704 | 2955842200800109 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-15 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the cable of the large needle driver instrument was broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip close cable is broken at the distal idlers and the pulley spins freely. Engineering believes the cable broke under tensile loading. Other cables at wrist are not damaged. The cable adjacent to the broken cable is currently derailed and seated on the outermost distal idler pulley. The cable derailment may have contributed to the breakage. Engineering also observed the distal end of the main tube to have a 1. 5 inch long, by. 250 inch wide section with material removed on one side of the tube. The damaged area is parallel to the tube axis, and has a rougher surface finish than the rest of the tube. The damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1007282 | 1009798 | 968705 | 2955842200800107 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-02-14 | Yes | No | Reuse | 2008
40017908 | 2008-01-17 | Malfunction |  | It was reported that the cable of the monopolar curved scissors instrument tip broke during the cutting of tissue. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is frayed in two locations, at the scissor grip and the distal idlers the blades still open and close. The scissor grip has a peak on the cable groove near the cable fray site. The wear grip/pulleys combined with high grasping/cutting force has likely contributed to fraying. Engineering also observed the main tube to have material removed all around the tube. One of the damaged areas starts above the midpoint of the tube, has a rough surface finish, and is parallel to the tube axis. The wear pattern is consistent with cannula related tube abrasion, but the position on the tube is too high for damage to be cause by the distal tip of cannula. The damage may be caused by a cannula defect at the bowl/tube interface. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1007261 | 1009811 | 960830 | 2955842200800103 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-02-12 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the wrist of the pk dissecting forceps instrument tips are misaligned. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips to have minimal side to side misalignment, however, the teeth do not mesh properly. One grip is roughly. 020 inches shorter than the other, resulting in peaks of teeth making contact. Minimal bending of one grip does not account for height difference. Per the customer reported event description, the pt was not impacted by this malfunction. Engineering also observed the distal end of the main tube to have a scratch/gouge mark with material removed from the tube, located 1. 5 inches above the proximal clevis. Orientation of the scratch is not parallel to the tube axis. There are also a couple of light scratch marks below this scratch. Scratches were likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) does contain a handling precaution, as identified in the following test: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1015676 | 1009813 | 968720 | 2955842200800108 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-02-14 | Yes | No | Reuse | 2008
42020906 |  | Malfunction |  | It was reported that the suturecut needle driver instrument would not open properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips to open and close properly when inputs are manually rotated. The instrument could not be placed on the sys because on lever is missing and the other is not installed correctly, creating interference with the sterile adapter. The housing is intact, but likely came off in order for lever to fall out. Engineering believes it is likely the instrument was mishandled, popping the housing off and causing both levers to fall out. The housing was put back on with one lever incorrectly installed, rotated 180 degrees. Engineering also observed the distal end of the main tube to have 2 inch long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Damage may be caused by a defective cannula. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submit. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1013719 | 1009814 | 968721 | 2955842200800115 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42020906 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017206 | 2008-01-18 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the cable of the monopolar curved scissors instrument broke. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one conductor wire is broken at the yaw pulley exit, however, the grip cables are not derailed at the wrist. The yaw pulley shows no signs of arcing and the wire insulation does not appear cut or melted. The wire likely was not securely attached to the grip and had poor strain relief and pulled out during articulation of the wrist. Electrical continuity failed. Engineering also observed the distal end of the main tube has a 1. 75 inch long section with material removed on one side of the tube. Damaged area is located 3. 5 inches above the proximal clevis, parallel to the tube axis, and has a rough surface finish. The wear pattern is consistent with cannula related tube abrasion. No other damage found. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submit. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1013712 | 1009815 | 968722 | 2955842200800104 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-13 | Yes | No | Reuse | 2008
40016906 | 2007-11-28 | Malfunction |  | It was reported that during a da vinci prostatectomy surgical procedure the harmonic curved shears insert accessory broke and a component fell into the pt. The broken component was retrieved from the pt and the planned surgical procedure was completed. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint engineering found the clamp arm is missing from the insert. One side of the attachment feature, which actuates up and down, is bent and the actuating piece is stuck in the "clamp arm fully open" position. Based on the appearance of the damage, engineering concluded that the instrument clamp arm was hyperextended while open, which bent the attachment and caused the clamp arm to detach. Per the case notes, the missing accessory component that fell inside of a pt was successfully retrieved. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1013779 | 1009816 | 968723 | 2955842200800102 | 1 | LFL | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-02-11 | Yes | Yes | Initial | 2008
42009307 |  | Malfunction |  | It was reported that the shaft of the endowrist prograsp forceps instrument is splintered. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube has a few scratch and gouge marks with a small amount of material removed. There are at least four to five distinct marks within a 1 inch long area with varying degrees of tube damage. Although some of the defects are parallel to the tube axis, the size and spacing between marks is not typical of cannula related tube abrasion. The tube damage was likely caused by instrument collision or some kind of mishandling during cleaning. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1013725 | 1009817 | 968724 | 2955842200800110 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2008-02-15 | Yes | No | Reuse | 2008
42018304 | 2008-02-13 | Other | Other | It was reported that during a da vinci s colon resection procedure the brown, tan wires on one side of the cautery hook broke off and fell into the pt while the surgeon was attempting a bowel dissect. The piece was not retrieved. The instrument was removed and the planned surgical procedure was completed with a replacement instrument of the same type. This event lengthened the procedure by 30 mins. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal clevis to have one entire ear broken off. The conductor cap is also missing. The missing ear piece is roughly. 215" x. 280". The yaw pulley is no longer retained in the clevis as a result of the breakage. The clevis fractured at the base of the ear. The location of fracture is very near the inner corner, which is a stress concentration area when excessively loaded. It is likely that the tip was overloaded, stressing the inner corner of the clevis and leading to fracture. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1006922 | 1009826 | 968731 | 2955842200800105 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | 42018304 | Yes | Yes | 2008-02-13 | No Answer Provi | No | Reuse | 2008
IS1200 A43P8 | 2008-02-11 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, a black line appeared in the middle of the left side screen. While troubleshooting with the assistance of an isi technical support engineer, the issue was resolved by shutting down the left camera controller unit and reseating the camera cables. The site called back to advise that the issue re-occurred with the addition of double vision in the viewer. The site was unable to resolve the issue after add'l troubleshooting. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field svc engineering concluded that the problem occurred due to incorrect cable connections completed by the site's biomed after replacement of a broken coax cable while replacing an assistants monitor. The sys was repaired by reconnecting the sys cables to its normal state. | DA VINICI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1006938 | 1009827 | 968732 | 2955842200800101 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | Yes | 2008-02-11 | Yes | No | Reuse | 2008
40017208 | 2007-12-21 | Other |  | It was reported that during a 6 week post-operation visit, after a da vinci prostectomy procedure, the dr noticed a burn mark on the pt at the trocar site. The burn mark was located on the pt's abdomen. The site advised that they did not deviate from the recommended setup instructions for the esu (electrosurgical unit). | An isi rep visited the site and performed a visual inspection on all three of the site's bipolar maryland forceps instruments. No defects were observed on any of the instruments. The root cause of the customer reported failure mode cannot be determined. The site indicated that they will be returning the instrument to isi for add'l eval. A f/u mdr will be submitted once the instrument is rec'd and in-house failure analysis has been completed. The following medwatch reports listed below were also sent to the fda for the two other bipolar maryland forceps instruments evaluated by the isi rep: mfr report # 2955842-2008-00112; mfr report # 2955842-2008-00114. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1006942 | 1009828 | 968733 | 2955842200800113 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | 40017208 | Yes | No | 2008-02-07 | Yes | No | Unkown | 2008
40017208 | 2007-12-21 | Other |  | It was reported that during a 6 week post-operation visit, after a da vinci prostectomy procedure, the dr noticed a burn mark on the pt at the trocar site. The burn mark was located on the pt's abdomen. The site advised that they did not deviate from the recommended setup instructions for the esu (electrosurgical unit). | An isi rep visited the site and performed a visual inspection on all three of the site's bipolar maryland forceps instruments. No defects were observed on any of the instruments. The root cause of the customer reported failure mode cannot be determined. The site indicated that they will be returning the instrument to isi for add'l eval. A f/u mdr will be submitted once the instrument is rec'd and in-house failure analysis has been completed. The following medwatch reports listed below were also sent to the fda for the two other bipolar maryland forceps instruments evaluated by the isi rep: mfr report # 2955842-2008-00112; mfr report# 2955842-2008-00113. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1006962 | 1009829 | 968733 | 2955842200800114 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | 40017208 | Yes | No | 2008-02-07 | Yes | No | Unkown | 2008
40017208 | 2007-12-21 | Other |  | It was reported that during a 6 week post-operation visit, after a da vinci prostectomy procedure, the doctor noticed a burn mark on the patient at the trocar site. The burn mark was located on the patient's abdomen. The site advised that they did not deviate from the recommended setup instructions for the esu (electrosurgical unit). | An isi representative visited the site and performed a visual inspection on all three of the site's bipolar maryland forceps instruments. No defects were observed on any of the instruments. The root cause of the customer reported failure mode cannot be determined. The site indicated that they will be returning the instrument to isi for additional evaluation. A follow-up mdr will be submitted once the instrument is received and in-house failure analysis has been completed. The following medwatch reports listed below were also sent to the fda for the two other bipolar maryland forceps instruments evaluated by the isi representative. Mfr report #'s 2955842-2008-00113, -00114. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1007616 | 1009830 | 968733 | 2955842200800112 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | 40017208 | Yes | No | 2008-02-07 | Yes | No | Unkown | 2008
IS1200 A43P8 | 2008-02-13 | Other | Other | It was reported that prior to starting a da vinci mitral valve repair procedure, the surgical staff reported the image was blurry with wavy lines on all monitors. Intuitive surgical's technical support engineer (tse) had the customer check all cable connections and camera control unit (ccu) settings. The vision cart power was cycled and the red indicator light for the black balance was solid on both the right and the left ccu. The customer then switched their high magnification camera to a wide-angle camera and was able to complete the planned surgical procedure with no further issues. The patient was in the room and under anesthesia for approximately one hour while the tse assisted the customer in resolving the vision issue. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by isi field service engineering concluded that the wavy image experienced by the customer was associated with camera control units (ccu). The ccus control the quality of the video images. Two ccus are included with the da vinci system, one each for the left and right optical paths. The field service engineer determined that the site's ccus were incompatible with the high magnification camera and repaired the system by replacing the ccus with an upgraded compatible version. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1007613 | 1009831 | 968736 | 2955842200800106 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-03-07 | 2008-03-07 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | Yes | 2008-02-13 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the teneculum forceps instrument does not hold tissue well. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions. The cable tension is acceptable and fangs are not bent. No issues with grasping tissues were observed. Engineering also observed the distal end of the main tube to have 3-4 deep scratches on one side of the tube. Material has been removed as a result of the scratch marks. Scratches are not parallel to the tube axis, suggesting they were caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENECULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1001664 | 1006157 | 965088 | 2955842200800092 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42004905 |  | Malfunction |  | It was reported that the housing of the cadiere forceps instrument was separated. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the housing was not separated from the chassis, however, there is a diagonal gouge mark at one of the rear housing corners. The luer plate is dislodged from the chassis and pushed into the back-end. The chassis feature that retains the luer plate is broken on one side of the inner wall. The flush port may have been loaded non-axially during cleaning, causing the chassis to break. There is also a. 500 inch by. 200 inch piece missing from the chassis above the bipolar insert hole. The bottom plane of the chassis has a deep scratch mark diagonally across the rear corner. Angle of the housing and chassis damage are similar, suggesting the instrument may have been wedged between something and was forced out. Evidence suggests instrument was mishandled. One of the rear grip input dogs is not subflush to the bottom plane of the chassis, making instrument installation and removal difficult. Engineering also observed the main tube has material removed at the distal end. Damaged area is parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | CADIERE FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1001654 | 1006158 | 965089 | 2955842200800090 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-02-22 | 2008-03-03 | No | Yes | Health Professional | 42004905 | Yes | No | 2008-02-07 | Yes | No | Initial | 2008
40011007 |  | Malfunction |  | It was reported that the wrist of the precise bipolar pyramid tip instrument was broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to be broken at the proximal clevis to main tube interface. The clevis is not dislodged from the main tube, but is missing a. 250" x. 160" piece between the ears. One of the ears is cracked and the piece is only retained by the proximal pin. The main tube interface wall is intact. The damage to the clevis may be from overloading or from a high impact collision. Per the customer reported event description, the patient was not impacted by this event. Engineering also observed the distal end of main tube to have various scratch marks roughly 180 degrees from the missing clevis piece. A couple of scratches are deep enough to have removed material from the tube. Orientation and size of scratches is consistent with instrument collisions. At the back end, one of the bipolar pins is bent, suggesting mishandling. Electrical continuity passed, no other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1001674 | 1006159 | 965090 | 2955842200800095 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 40011007 | Yes | No | 2008-02-08 | Yes | No | Reuse | 2008
42017908 | 2008-01-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the monopolar curved scissors instrument was dull. The planned surgical procedure was completed successfully. No pt harm, adverse outcome or injury was reported. No add'l info was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to be within specification. Cut tests and electrical continuity testing were completed successfully. Engineering also observed the distal end of the main tube has a 1 inch long by. 160 wide section with material removed on one side of the tube. Damaged area is 2 inches above the tube extension, parallel to the tube axis, and has a rough surface finish than the rest of the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002522 | 1006162 | 960830 | 2955842200800088 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument does not hold tissue well. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tips to open and close properly when inputs are manually rotated. The tips are aligned and cable tension is acceptable. Engineering also observed the distal end of the main tube to have a couple of sections with material removed. The sections are roughly 180 degrees apart and range in size from. 100" w x 1. 8" l to. 325" w x 2" l. Both scraped areas are parallel to the tube axis and have a rough surface finish. There are also a couple of deep scratches on the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collision. Add'l investigation is underway regarding the cannula accessories. No other damge was found. A follow-up mdr will be submitted if add'l info is received. No other damage found. Scratch marks are likely due to instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002527 | 1006163 | 965094 | 2955842200800091 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42011704 |  | Malfunction |  | It was reported that the housing of the 5mm needle driver instrument is separating. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one of the rear housing snaps along with two housing pins on the same side are broken off. Damaged side of housing can be lifted, exposing inner components. Damage is likely due to mishandling. Engineering also observed the main tube is missing a. 75 inch diameter piece of insulation, located roughly 5 inches above the snake wrist. There are also two very small. 15 inch by. 30 inch pieces of insulation missing. 100 inches above the main damaged area. Larger size breakage may be due to pressure against a sharp edge of a 5mm cannula accessory. No other damage found. The damage on the instrument was found to be created by a defective 5mm cannula that has since been replaced as part of a voluntary recall. All defective 5mm cannula have been recalled and replaced at this site. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1003095 | 1006164 | 965095 | 2955842200800089 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-02-29 | 2008-03-03 | No | Yes | Health Professional | 42011704 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar maryland instrument is splintered. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a 1. 7 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Additionally, engineering observed, that one grip is severely bent causing side to side misalignment of grips, resulting in a. 150 inch offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Electrical continuity passed. No other damage was found. | FENESTRATED BIPOLAR MARYLAND INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1003101 | 1006165 | 965096 | 2955842200800083 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-02-05 | Yes | No | Unkown | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument does not open. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to moved intuitively with a full range of motion in all directions. The jaws opened and closed properly. The customer complaint was unable to be replicate. Engineering also observed the distal end of the main tube to have two. 125" long scratch marks located. 250" above the proximal clevis. The scratches are deep enough to have removed a small amount of material from the tube. The scratches are not parallel to the tube axis, suggesting they were caused by instrument collisions. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instruments intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1003071 | 1006167 | 965098 | 2955842200800086 | 1 | GCJ | Manufacturer | OtherUser facility | Other | Initial | 2008-03-03 | 2008-03-03 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-02-06 | Yes | No | Reuse | 2008
42000605 | 2007-11-15 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the large needle driver instrument is not recognized by the system when inserted on the arm. The planned surgical procedure was completed. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a system to check for recognition and found the instrument successfully recognized. Results: engineering also observed the distal end of main tube has a section with a heavy concentration of deep scratches. The scratches are only on one side of the tube and have removed material from the tube. Based on the quantity, size, and orientation, the scratch marks are most likely caused by instrument collisions. No other damage was found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instruments intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1003078 | 1006168 | 924417 | 2955842200800093 | 1 | GCJ | Manufacturer | User facilityDistributor | UNKNOWN | Initial | 2008-02-22 | 2008-03-03 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-02-07 | Yes | No | Reuse | 2008
42017908 | 2008-02-08 | Malfunction |  | On 2/26/08, medwatch uf/importer report was received: "this was a pt with a known history of prostate cancer, undergoing a robotic-assisted laparoscopic radical prostatectomy and left pelvic lymph node dissection when the physicians assistant commented that he smelled garlic. Intra operatively, it was immediately noticed that the monopolar scissors flamed, sparked and smoked at the joint when activated by the surgeon. Instrument was immediately removed from the pt. Per the surgeon there was no harm to the pt in any way. There was no thermal effect from the incident. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no". | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INTRUMENT | ELECTROSURGICAL INSTRUMENT | 976021 | 1055003 | 966950 | 2955842200800096 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-29 | 2008-03-03 | No | Yes | Health Professional | 42017908 | No | No | 2008-02-08 | Device Not Retu | No | Unkown | 2008
IS2000 A5 1P2 |  | Death |  |  | No conclusion can be drawn to the product, as it has not been returned for evaluation. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 951649 | 1008770 | 936935 | 2955842200800099 | 1 | NAY | Manufacturer | Health ProfessionalOther | Nurse | Initial |  | 2008-02-29 | Yes | No | Health Professional | IS2000 A5 1P2 |  | No Answer Provided | 2008-02-01 | Device Not Retu | No | Reuse | 2008
IS2000 A51P2 |  | Death |  |  |  | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 948563 | 1008772 | 938085 | 2955842200800098 | 1 | NAY | Manufacturer | Health ProfessionalOther | Nurse | Initial |  | 2008-02-29 | No |  |  | IS2000 A51P2 |  | Yes | 2008-02-01 | Device Not Retu | No | Reuse | 2008
IS2000 A51P4 | 2008-01-24 | Death | Death | In 2008, after a single vessel small thoracotomy da vinci s procedure, the patient expired while weaning off of bypass. Prior to the procedure, the patient appeared to have a heparin reaction and experienced a hypotensive event with a systolic pressure of 39mm. The site initiated a cardiopulmonary bypass run and completed the procedure without event. Wean from bypass was unsuccessful as patient exhibited stone heart characteristics. No malfunction of the davinci s system was alleged by the hospital, and the information provided indicates that this event is related to the inherent risks associated with bypass and the patient's medical history. | In 2008, intuitive surgical has requested additional information related to this event from this site. A follow-up mdr will be submitted if additional information becomes available. The hospital has continued to use the da vinci s system to perform surgery on patients. As of 2008, no adverse events or product problems have been experienced with the site's da vinci s system. As of 2008, no similar instance of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 971608 | 1002352 | 961423 | 2955842200800073 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-02-22 | 2008-02-22 | Yes | No | Health Professional | IS2000 A51P4 | No | No | 2008-01-24 | No | No | Reuse | 2008
40016906 | 2008-01-21 | Malfunction |  | It was reported that while sealing gastric vessels and dissection of the stomach during a da vinci rectocele repair surgical procedure, the jaw of the harmonic curved shears (hcs) insert accessory used in conjunction with the harmonic curved shears instrument, became unhinged and fell into the pt after approx 2 hours of use. The broken piece was retrieved and the surgical procedure continued with a replacement hcs insert. After approx 2 hours of use, the second hcs insert broken off and fell into the pt. The second broken piece was retrieved and the surgical procedure continued with a third replacement hcs insert. A third hcs insert broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement hcs insert. Procedure was successfully completed and was not significantly lenghtened due to incident. No pt harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering found the clamp arm missing from the curved shears insert. Both sides of the attachment features on the shaft are bent and deformed. Based on the features of the damage, engineering concluded that the instrument clamp arm was hyperextended while open or that a twisting action was attempted with the arm closed. The hinge joint could not withstand the loading and failed, causing clamp arm to detach. Per the case notes, the missing accessory component that fell inside of a pt during a surgical procedure, was retrieved. The following medwatch reports listed below were also sent to the fda as they relate to this event. Mfr report #2955842-2008-00066 and mfr report #2955842-2008-00068. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC ACCESSORY | 1004413 | 1006029 | 964962 | 2955842200800067 | 1 | LFL | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-15 | 2008-02-22 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-01-21 | Yes | Yes | Initial | 2008
40016906 | 2008-01-21 | Malfunction |  | It was reported that while sealing gastric vessels and dissection of the stomach during a da vinci rectocele repair surgical procedure, the jaw of the harmonic curved shears (hcs) insert accessory used in conjunction with the harmonic curved shears instrument, became unhinged and fell into the pt after approx 2 hours of use. The broken piece was retrieved and the surgical procedure continued with a replacement hcs insert. After approx 2 hours of use, the second hcs insert broke off and fell into the pt. The second broken piece was retrieved and the surgical procedure continued with a third replacement hcs insert. A third hcs insert broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement hcs insert. Procedure was successfully completed and was not significantly lengthened due to incident. No pt harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering found the clamp arm missing from the curved shears insert. Both sides of the attachment features on the shaft have minimal damage, however, the curved blade has three distinct scratch marks on the clamp arm side. Based on the scratches and features of the damage, engineering concluded that the instrument collisions may have hyperextended the clamp arm, causing it to detach. Per the case notes, the missing accessory component that fell inside of a pt during a surgical procedure, was retrieved. The following medwatch reports listed below were also sent to the fda as they relate to this event. Mfr report #2955842-2008-00066 and mfr report #2955842-2008-00067. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC ACCESSORY | 1004417 | 1006031 | 964962 | 2955842200800068 | 1 | LFL | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-15 | 2008-02-22 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-01-21 | Yes | Yes | Initial | 2008
42020504 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar forceps instrument was scratched. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have various deep scratch marks that have removed material from the tube. The majority of scratches are concentrated on one side of the tube, however, there is deep scratch/gouge oriented 90 degrees from the main damaged area. Scratch marks are not aligned with longitudinal tube axis, suggesting they were caused by multiple instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002189 | 1006032 | 964963 | 2955842200800079 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-25 | 2008-02-22 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No | 2008-01-31 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the tips of the maryland bipolar forceps instrument do not meet. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips are properly aligned at the tips, in accordance with specification. The teeth mesh and the offset between grips is within specification. Performance testing found the instrument grips open and close properly when placed on a system. Engineering also observed the distal end of the main tube to have various deep scratch marks with material removed from the tube. Scratches are not parallel to the tube axis, suggesting they were caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002194 | 1006033 | 964964 | 2955842200800074 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-22 | 2008-02-22 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-01-29 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the tips of the bipolar maryland forceps instrument did not align. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip to be bent, causing side to side misalignment of grips. There is a. 020" offset at the tips. The bent grip does not have separation of the glue joint, however, likely cause of bending is still overloading at the tip. Engineering also found the distal end of the main tube to have a 2. 5" section directly above the proximal clevis with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. There are various scratch marks 180 degrees from the scraped section. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collision. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1002199 | 1006035 | 964966 | 2955842200800076 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-25 | 2008-02-22 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-01-29 | Yes | No | Reuse | 2008
40016906 | 2008-01-21 | Malfunction |  | It was reported that while sealing gastric vessels and dissection of the stomach during a da vinci rectocele repair surgical procedure, the jaw of the harmonic curved shears (hcs) insert accessory used in conjunction with the harmonic curved shears instrument, became unhinged and fell into the patient after approximately 2 hours of use. The broken piece was retrieved and the surgical procedure continued with a replacement hcs insert. After approximately 2 hours of use, the second hcs insert broke off and fell into the patient. The second broken piece was retrieved and the surgical procedure continued with a third replacement hcs insert. The third hcs insert broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement hcs insert. Procedure was successfully completed and was not significantly lengthened due to incident. No patient harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering found the clamp arm missing from the curved shears insert. Both sides of the attachment features on the shaft are bent and deformed. Based on the features of the damage, engineering concluded that the instrument clamp arm was hyperextended while open or that a twisting action was attempted with the arm closed. The hinge joint could not withstand the loading and failed, causing clamp arm to detach. Per the case notes, the missing accessory component that fell inside of a patient during a surgical procedure, was retrieved. The following medwatch reports listed below were also sent to the fda as they relate to this event. Mfr report #2955842-2008-00067; mfr report #2955842-2008-00068. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC ACCESSORY | 1002211 | 1006037 | 964962 | 2955842200800066 | 1 | LFL | Manufacturer | User facility | Other | Initial | 2008-02-15 | 2008-02-22 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-01-24 | Yes | Yes | Initial | 2008
42017206 |  | Malfunction |  | It was reported that the tips of the bipolar maryland forceps instrument did not align. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tips to not be misaligned. The teeth mesh properly and there is no side to side offset on grips. The top plug riser pin is pushed into the back end. The chassis feature that retains pins is broken. Instrument was likely mishandled, breaking the chassis feature. Engineering also found the main tube to have a couple different sections below the midpoint with material removed on one side of the tube. The damaged areas ranges from 1. 5-2" in length and are parallel to tube axis. The surface finish is rough due to glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. Electrical continuity passed. No other damage found. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1002173 | 1006039 | 964970 | 2955842200800075 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-25 | 2008-02-22 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-01-29 | Yes | No | Reuse | 2008
42017908 | 2008-01-24 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the monopolar curved scissors instrument was not recognized. The planned surgical procedure was completed successfully. No patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a system to check for recognition and found the instrument successfully recognized. Performance tests were conducted and the instrument to move intuitively with full range of motion in all directions. Cut tests and electrical continuity testing were completed successfully. Engineering also observed the main tube has a 1. 5 inch long section with material removed on one side of the tube. Damaged area is 4. 2 inches above the tube extension, parallel to the tube axis, and has a slightly rougher surface finish than the rest of the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002177 | 1006040 | 964971 | 2955842200800085 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-22 | 2008-02-22 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-02-04 | Yes | No | Reuse | 2008
40000620 | 2007-05-31 | Malfunction |  | Per case report: intraoperative breakage of needle driver jaw during robotic-assisted laparoscopic radical prostatectomy published january, 2008: during a da vinci prostatectomy procedure, during urethrovesical anastomosis, the jaw of the large needle driver instrument was missing from the operative field. The surgical staff paused the procedure and carefully inspected the operative field. The broken jaw was found on top of the bladder and was immediately removed. No intraoperative or perioperative complications were experienced and postoperative recovery was uneventful. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint that the instrument did not work, engineering found the grip (referred to as "jaw" by customer) is completely broken off at the grip base exposing the cable crimp. Grip was likely cracked along the thin wall where the cable crimp inserts. A likely failure mode is the high drive torque combined with the high gripping force of the needle driver likely caused the crack to propagate. Gripping and twisting action would likely create a fracture at the grip base. No other damage found. Per the referenced article, the missing instrument component that fell inside of the patient during a surgical procedure, was removed. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1002267 | 1006041 | 964972 | 2955842200800078 | 1 | GCJ | Manufacturer | User facilityDistributor | Other | Initial | 2008-02-22 | 2008-02-22 | No | Yes | Health Professional | 40000620 | Yes | Yes | 2008-01-31 | Yes | No | Unkown | 2008
42020702 | 2008-01-29 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure and while the physician was grasping the uterus, the tip of the tenaculum forceps instrument broke, resulting in pieces falling into the patient. The pieces were retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the wrist of the instrument to be cut off at the proximal clevis. Both the proximal clevis and the main tube interface wall were returned with missing pieces. The proximal clevis is missing roughly 3/4 of the section that surrounds the main tube interface wall. The likely failure mode is a break at the proximal clevis to main tube interface. The clevis became dislodged from the tube as a result of breakage and the wrist was likely cut off to aid in removal of instrument from cannula, but without any additional information from the site this cannot be concluded any further. No other damage found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1002162 | 1006042 | 964973 | 2955842200800084 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-02-25 | 2008-02-22 | No | Yes | Health Professional | 42020702 | Yes | No | 2008-02-01 | Yes | No | Unkown | 2008
IS2000 A51P4 | 2008-01-25 | Other | Other | It was reported that 2 hours into a da vinci s prostatectomy surgical procedure, the customer experienced recurring system error codes #201 and #264. The site was advised by an isi regional support specialist to exercise all axis on the right master tool manipulator (mtmr) and to restart the system. The system error code #201 continued to occur and could not be overridden. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master manipulators at the surgeon's console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr and multiple servo driver pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error codes #201 and #264 appear when the davinci s safety system determines a voltage tracking fault reported by the digital signal processor (dsp). Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the reported system errors, the encoder assembly and a motor assembly were replaced as a precautionary measure. The servo driver pca board was also returned to isi, however, failure analysis has not been completed at this time. A follow-up mdr will be submitted once evaluation is completed. As of 2008, there have been reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1003880 | 1006071 | 965002 | 2955842200800070 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-02-15 | 2008-02-22 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-01-25 | Yes | No | Reuse | 2008
49917208 | 2008-01-21 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, while dissecting the seminal vesicles, the maryland bipolar forceps instrument would not articulate. The customer stated that the instrument functioned properly at the beginning of the procedure. The surgeon decided to abort the procedure and convert to traditional open surgical techniques. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable derailed and frayed at the distal idlers, however, the grips are still able to open and close. The conductor wire adjacent to the frayed cable is broken at the yaw pulley exit, the yaw pulley shows no signs of arcing. Additionally, the wire insulation does not appear cut or melted. Engineering believes the likely failure was due to the wire not securely attached to the grip and therefore caused poor stain relief and pulled out during articulation of the wrist. Electrical continuity failed. No other damage found. The following medwatch report listed below was also sent to the fda as it relates to this event. Mfr. Report #2955842-2008-00081. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002541 | 1006073 | 965005 | 2955842200800080 | 1 | GEI | Manufacturer | User facilityDistributor | Other | Initial | 2008-02-22 | 2008-02-22 | Yes | Yes | Health Professional | 49917208 | No | No | 2008-02-01 | Device Not Retu | No | Reuse | 2008
40017908 | 2008-01-21 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, during the transect of the bladder neck, the maryland bipolar forceps instrument jaws could not be opened and cauterization did not function. The customer stated that the instrument functioned properly at the beginning of the procedure. The surgeon decided to abort the procedure and convert to traditional open surgical techniques. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is broken at the scissor grip and a couple of the cable strands extend past the grip and break at the distal idlers. The forceps blades will not open due to the broken grip cable. Engineering observed the grip hub has a peak across the cable groove at the cable break site. The cable wear against the grip likely contributed to breakage. Electrical continuity passed. Engineering also observed the main tube has a couple of sections with light material removed parallel to the tube axis. There is a section below the midpoint of the tube. Wear patter is similar to cannula related tube abrasions, however, part of the section above the midpoint it out of range of the distal tip of the cannula. Based on the location and appearance, the damage may be due to the customer's cleaning process. No other damage was found. A follow-up mdr will be submitted if additional info is rec'd. The following medwatch report listed below was also sent to the fda as it relates to the event mfr. Report #2955842-2008-00080. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1002542 | 1006074 | 965005 | 2955842200800081 | 1 | GEI | Manufacturer | User facilityDistributor | Other | Initial | 2008-02-22 | 2008-02-22 | Yes | Yes | Health Professional | 40017908 | No | No | 2008-02-01 | Device Not Retu | No | Reuse | 2008
IS2100 A51P4 | 2008-01-10 | Other | Other | It was reported that while preparing to da vinci s system for surgery, the endoscopic camera manipulator (ecm) arm insertion axis was stuck. Isi technical support engineer instructed the customer to reboot the system and found error code #23007 appeared, also, the self test did not start. The system was restarted and the user interface panel led lights began to continuously blink. The pt had been under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found when system the homed error code #23013 appeared. This error is associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. System error code #23013 appears when the da vinci s safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer, were out of specified tolerance for agreement. Upon determining this condition, the safety system put the da vinci s system in a "recoverable safe state". The system error code #23007 experienced by the customer appears when a sensor is out of specified tolerance for agreement on an axis potentiometer, thus causing the system error. The ecm was returned to isi for failure analysis investigation. Based on the system error logs, a potentiometer assembly and a motor assembly were replaced. The s5vp pca board was replaced to address the power up issue experienced by the customer. The pca board has been returned but failure analysis has not been completed. A follow-up mdr will be submitted once failure analysis is complete. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1002547 | 1006076 | 965007 | 2955842200800082 | 1 | NAY | Manufacturer | User facilityDistributor | UNKNOWN | Initial | 2008-02-21 | 2008-02-22 | Yes | Yes | Health Professional | IS2100 A51P4 | Yes | No | 2008-01-10 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the precise bipolar pyramid tip instrument does not work correctly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to move intuitively with full range of motion in all directions. Engineering also observed the main tube has a 2 inch long section with material removed on one side of the tube. Damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A f/u mdr will be submitted if add'l info is received. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999089 | 1001753 | 960824 | 2955842200800061 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-15 | 2008-02-19 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-01-22 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the tip of the maryland bipolar forceps instrument was bent. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tip to not be bent. The grips are properly aligned. There is charring along the outer edge of one yaw pulley and a burned spot that is divided between both yaw pulleys. One of the grip cables has a burn spot directly above the yaw pulley charring. Another energy instrument may have caused the burn spots. It is unclear whether the charring on the outer edge was caused by this instrument or a different instrument. Engineering also observed that the distal end of the main tube also has a 2. 25" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A f/u mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999059 | 1001754 | 960825 | 2955842200800060 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-19 | 2008-02-19 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-01-22 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument does not grasp. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to move intuitively with full range of motion in all directions. The instrument jaws open and close properly and grasp without problems. Engineering also observed the distal end of the main tube has a 1. 3 inch long by. 0170 inch wide section with material removed on one side of the tube. Damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1000000 | 1001755 | 960826 | 2955842200800064 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-02-15 | 2008-02-19 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-01-23 | Yes | No | Reuse | 2008
42000202 |  | Malfunction |  | The cannula accessory was returned by request from isi as the site reported that their 8mm cannula accessories are breaking apart. No patient harm, adverse outcome or injury was reported. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the weld cracked around the circumference of the tube, however, the base does not separate at the weld. Engineering also observed the cannula to be visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument main tube scraping in certain loading conditions. No other damage was found. The cannula accessory instructions for use (ifu) specifically states: general precautions and warnings. Do not use a cannula with imperfections on its distal end. Examples of defects include a rough edge, dents, or an out-of-round shape. Use of a defective cannula may result in instrument damage, including scraping of the instrument shaft and particulate falling into the patient. Inspect the cannula prior to use for any defects. - cannula defects can be caused by dropping the cannula or handling it roughly. To inspect the cannula, hold the cannula up close and visually inspect for defects. | 8MM CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 999900 | 1001756 | 960827 | 2955842200800062 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-02-15 | 2008-02-19 | No | Yes | Health Professional | 42000202 | Yes | Yes | 2008-01-23 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had a broken housing. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one lever to be detached. The housing is intact, but likely popped off in order for the lever to fall out. The instrument was likely mishandled and the housing was then put back on without the lever. Engineering also observed the distal end of the main tube to have a 3. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999910 | 1001757 | 960828 | 2955842200800065 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-02-19 | 2008-02-19 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-01-24 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not recognized by the da vinci s surgical system. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering performed system recognition testing with the instrument and instrument recognition passed. Engineering also observed the main tube has various deep scratches with material removed on one side at the distal end of the tube. Based on the appearance and randomly oriented location of the scratches, the scratch marks are most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999964 | 1001758 | 960829 | 2955842200800059 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-02-15 | 2008-02-19 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-01-21 | Yes | No | Reuse | 2008
42017908 | 2008-01-15 | Malfunction |  | It was reported that the jaws of the monopolar curved scissors instrument do not open. It was reported that the patient was not impacted. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering performed functional testing and found the instrument to move intuitively with a full range of motion in all directions. The jaws opened and closed fine. The instrument was placed on the system for a latex cut test. The scissors do not cut cleanly through. 006" latex. The latex gets snagged at scissor tips. One blade exhibits pitting. Engineering also observed the distal end of the main tube to have a couple of deep scratch marks with material removed. The scratches are located on one side of the tube and are not parallel to the tube axis, suggesting they were caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999940 | 1001759 | 960830 | 2955842200800063 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-02-19 | 2008-02-19 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-01-23 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tips of the monopolar curved scissors instrument will not open and close. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the jaws opened and closed without issue. Engineering also observed the tube extension is slightly dislodged from the main tube and the entire tube extension wall is broken at the keys. The distal end of the main tube and the tube extension have burned sections. The discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure and the tube extension breakage may have created a path for arcing to occur. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 999953 | 1001760 | 960831 | 2955842200800072 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-02-15 | 2008-02-19 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-01-21 | Yes | No | Reuse | 2008
40009307 |  | Malfunction |  | It was reported that the wrist of the prograsp forceps instrument does not move freely. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions. No increased or abnormally high friction was noticed. Engineering also observed the distal end of the main tube to have various deep scratch marks with material removed on one side of the tube. Scratches are randomly oriented to tube, suggesting they are caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 999923 | 1001761 | 960832 | 2955842200800069 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-02-01 | 2008-02-19 | No | Yes | Health Professional | 40009307 | Yes | Yes | 2008-01-25 | Yes | No | Reuse | 2008
IS1200 A4 3P8 | 2008-01-25 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the site experienced difficulties removing the instrument from a pt side manipulator. The instrument was inserted about two inches past the cannula tip but was not responding to the surgeon's commands. The message on the screen indicated the instrument was not past the cannula tip. The surgical staff was unable to physically remove the instrument as the instrument tip was fully open and could not pass through the cannula. Using a laparoscopic tool via the accessory port, the site was able to close the instrument tip enough to allow it to pass through the cannula. An isi regional support specialist attempted to troubleshoot the issue via telephone and requested that the site replace the sterile adapter. The site advised that they would convert to open techniques to finish the planned surgery. The pt had been under anesthesia for 3. 5 hours. No pt harm was reported. | The investigation conducted by isi field service engineering found the system to perform to isi standards. Performance testing was conducted and the reported complaint could not be replicated. As of february 19, 2008, there have been no reported recurrences of the issue at the hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1001578 | 1003519 | 962495 | 2955842200800071 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-02-19 | 2008-02-19 | Yes | Yes | Health Professional | IS1200 A4 3P8 | Yes | No | 2008-01-25 | Yes | No | Reuse | 2008
IS2000 A51P2 | 2008-01-11 | Death | Death | On january 17, 2008, intuitive surgical was notified that following a da vinci s hysterectomy surgical procedure, the pt suffered from sepsis and expired four days post da vinci assisted surgery. | On january 18, 08, intuitive surgical spoke to the physician who performed the procedure. The physician stated the procedure was complex due to that pt's previous medical history and anatomy, however, the case was completed successfully with the da vinci s system and no malfunction of the system was alleged by the hosp. The physician indicated the pt died from sepsis due to a defect in the small bowel, however, it is unclear if the perforation in the bowel occurred during or after the surgical procedure. In 2008, intuitive surgical spoke to the attending nurse present during the procedure. The nurse stated that when the surgical staff removed the trocar at the end of the procedure, the pt's bowel appeared normal. Additionally, the surgical staff observed no issues during the surgery which could have caused or contributed to the pt's adverse outcome. The hosp has continued to use the da vinci s system to perform surgery on patients. As of 2008, no adverse events or product problems have been experienced with the site's da vinci s system. As of 2008, no similar instances of this event have been reported to intuitive surgical. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 967285 | 997372 | 956455 | 2955842200800056 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-02-08 | 2008-02-12 | Yes | No | Health Professional | IS2000 A51P2 | No | No | 2008-01-17 | No | No | Reuse | 2008
IS1200 A43P8 | 2008-01-14 | Other | Other | It was reported that during a da vinci prostatectomy procedure the console surgeon experienced sys vision issues that created difficulties in grasping tissue. With the assistance of a regional support specialist, the site attempted to troubleshoot this problem via telephone, however, the issue could not be resolved. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Conclusions: the investigation conducted by isi field svc engineering (fse) discovered that the vision cables used to connect the video signal distribution (vsd) and the high resolution stereo viewer (hrsv) were crossed. The sys was repaired by placing the cables in the proper configuration. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1001661 | 1002859 | 961906 | 2955842200800030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-02-11 | 2008-02-12 | Yes | No | Health Professional | IS1200 A43P8 | Yes | No | 2008-01-14 | Yes | No | Reuse | 2008
IS2000 A5 1P4 | 2008-01-16 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced a non-recoverable sys error code #25589. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Conclusions: the investigation conducted by field svc engineering concluded that sys error code #25589 experienced by the customer was associated with the remote arm controller (rac). The rac consists of 5 pcas that provide control of the sys instrument and camera arms. The sys was repaired by replacing the affected rac. The sys alarm (sys generated fault code) functioned as designed and there was no injury to the pt. Sys error code #25589 appears when the da vinci s safety sys determines that the brakes failed the brake voltage test during the rac power up self test. Upon determining this condition, the safety systems put the da vinci s in a "recoverable safe state. " the rac was returned to isi for failure analysis investigation. Inspection was performed and faulty chips on the drive modules were replaced to repair the rac. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1001260 | 1002861 | 961908 | 2955842200800053 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-02-11 | 2008-02-12 | Yes | Yes | Service Personnel | IS2000 A5 1P4 | Yes | No | 2008-01-16 | Yes | No | Reuse | 2008
420179 | 2008-02-08 | Malfunction |  | This was a patient with a known history of prostate cancer, undergoing a robotic-assisted laparoscopic radical prostatectomy and left pelvic lymph node dissection when the physicians assistant commented that he smelled garlic. Intra-operatively, it was immediately noticed that the monopolar scissors flamed, sparked and smoked at the joint when activated by the surgeon. Instrument was immediately removed from the patient. Per the surgeon there was no harm to the patient in any way. There was no thermal effect from the incident. |  | DAVINCI | INSTRUMENT ROBOTIC SURGICAL MO | 976021 | 1008021 | 966950 | 1008021 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2008-02-12 | 2008-02-12 | No | Yes | Physician | 420179 | Yes | No Answer Provided |  |  |  |  | 2008
IS1200 A4 3P8 | 2008-01-18 | Other | Other | It was reported that while preparing the da vinci surgical system for surgery, system error codes # 20008 and # 21008 appeared several times and was unable to be overridden. The pt had been under anesthesia for approx 45 mins and the pt port incisions had been created when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that system error codes # 20008 and # 21008 experienced by the customer were associated with a pt side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering evaluation confirmed that axis one, stage two cables were broken. It was determined that the cables most likely broke during shipment. The system alarm ( system generated fault code) functioned as designed and there was no injury to the pt. The 20008 and 21008 system error codes appeared when the da vinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor ( potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a " recoverable safe state". As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1001170 | 1002914 | 961961 | 2955842200800057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-02-08 | 2008-02-11 | Yes | Yes | Health Professional | IS1200 A4 3P8 | Yes | Yes | 2008-01-18 | Yes | No | Reuse | 2008
IS12000 A43P8 | 2008-01-17 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, the right system arm was not allowing the instrument jaws to open and close properly, and the surgeon found it difficult to articulate the jaws during wrist movement. The instrument sterile adapter was replaced, however, these issues continued to recur. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | On site investigation and evaluation was conducted by field service engineering. System verification tests were performed and the customer reported failure mode could not be duplicated. The right pt side manipulator (psm) was replaced as a precaution. The pt side manipulator is an instrument arm located on the pt side cart which provides the sterile interface for the endowrist instruments. The psm was returned to isi for failure analysis investigation. Mechanical and electrical performance tests were conducted and the customer reported failure mode could not be reproduced. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1000525 | 1002915 | 961962 | 2955842200800055 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-02-08 | 2008-02-11 | Yes | Yes | Health Professional | IS12000 A43P8 | Yes | No | 2008-01-17 | Yes | No | Reuse | 2008
42004905 |  | Malfunction |  | It was reported that the shaft of the cadiere forceps instrument had scratches on it, however, the patient was not impacted. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube to have a 4" long section directly above the proximal clevis with material removed. There are also some light scratch marks in the same damaged area. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collision. Additional investigation is underway regarding the cannula accessories. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992063 | 990876 | 950030 | 2955842200800044 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-02-04 | 2008-02-01 | No | Yes | Health Professional | 42004905 | Yes | No | 2008-01-09 | Yes | No | Initial | 2008
42009307 |  | Malfunction |  | It was reported that while completing a da vinci s prostatectomy surgical procedure the surgical staff noticed "black carbon fiber" from the large needle driver instrument and the endowrist prograsp forceps instrument had fallen into the patient. The area was irrigated, but not all of the "black carbon fibers" were removed. Incident did not significantly lengthen the procedure. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992074 | 990877 | 950031 | 2955842200800040 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 42009307 | No | No | 2008-01-07 | Device Not Retu | No | Unkown | 2008
42022702 |  | Malfunction |  | It was reported that the shaft of the pk dissecting forceps instrument had scratches on it, however, the patient was not impacted. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a couple of deep scratch marks on one side of the tube. Material has been removed due to scratches. Scratches are likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 992079 | 990878 | 950032 | 2955842200800043 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2008-02-01 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2007-01-09 | Yes | No | Initial | 2008
42000605 |  | Malfunction |  | It was reported that while completing a da vinci s prostatectomy surgical procedure, the surgical staff noticed "black carbon fiber" from the large needle driver instrument and the endowrist prograsp forceps instrument had fallen into the patient. The area was irrigated, but not all of the "black carbon fibers" were removed. Incident did not significantly lengthen the procedure. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992216 | 990879 | 950033 | 2955842200800041 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 42000605 | No | No | 2008-01-07 | Device Not Retu | No | Unkown | 2008
42017206 |  | Malfunction |  | It was reported that a wire came out of the fenestrated maryland bipolar instrument. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found both bipolar pins and insert are missing causing the conductor wires to stick out of the back end of the instrument. Engineering also observed the instrument main tube has a 2 inch long section with material removed. Damaged area is 2. 5 inches above the proximal clevis and is parallel to the tube axis. Based on the location cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 992219 | 990880 | 950034 | 2955842200800046 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-01-10 | 2008-02-01 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2008-01-10 | Yes | No | Reuse | 2008
40018402 | 2008-01-10 | Malfunction |  | It was reported that during the beginning of a da vinci myomectomy procedure, while dissecting fibroids from the uterus, the permanent cautery spatula instrument stopped working. The surgical staff then noticed the instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the instrument from the cannula accessory. The instrument was replaced with another instrument of the same type. The surgical staff then noticed after a few minutes that the replacement permanent cautery spatula instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the replacement instrument from the cannula accessory. | The instrument and detached wrist were returned and evaluated. Per the customer reported complaint, engineering observed the proximal clevis is missing a. 200 inch by. 115 inch piece between the ears and pieces of the main tube interface wall are missing. Based on the engineering evaluation, a likely failure mode was overloading of the instrument at the tip, causing the breakage. No other damage found. Per the case notes, no broken instrument components fell into the patient. The following medwatch reports listed below were also sent to the fda are related to this case. 2955842-2008-00048, 2955842-2008-00050. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 992254 | 990881 | 950040 | 2955842200800049 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 40018402 | Yes | Yes | 2008-01-10 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument did not work. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have a section just below the midpoint with material removed on one side of the tube. The damaged area is 1. 6" long and. 150" wide and parallel to tube axis. The surface finish is rough due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992143 | 990882 | 950036 | 2955842200800047 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-02-04 | 2008-02-01 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-01-10 | Yes | No | Reuse | 2008
40020707 | 2008-01-10 | Malfunction |  | It was reported that during the beginning of a da vinci myomectomy procedure, while dissecting fibroids from the uterus, the permanent cautery spatula instrument stopped working. The surgical staff then noticed the instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the instrument from the cannula accessory. The instrument was replaced with another instrument of the same type. The surgical staff then noticed after a few minutes that the replacement permanent cautery spatula instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the replacement instrument from the cannula accessory. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tip is not bent. Engineering also observed one grip cable is broken near the proximal clevis cable hole and the inside of the clevis hole has a metal strand stuck to it. Based on the engineering evaluation, a likely failure mode was tensile loading of the instrument combined with high grasping force, causing cable wear on the proximal clevis hole. No other damage found. The following medwatch reports listed below were also sent to the fda and are related to this case. 2955842-2008-00048, 2955842-2008-00049. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992262 | 990883 | 950037 | 2955842200800050 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 40020707 | Yes | Yes | 2008-01-10 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument has a broken cable. No pieces of the instrument fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation discovered one conductor wire to be broken at the yaw pulley exit. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. The added range of motion of the grip likely created excess tension on the wire, causing it to break. Engineering also located at the back end, both levers are detached and one snap tab and adjacent pin are broken off, suggesting the instrument was likely mishandled. Electrical continuity failed. The distal end of the main tube to have a 2" long section with light material removal all around the tube. Surface finish is slightly rougher than the rest of the tube in the affected area. Below the scraped section, there is a. 800" long scratch mark parallel to the tube axis with material removed. Based on the location and appearance, the damge may have been caused by a cannula accessory or instrument collision. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 992258 | 990884 | 950038 | 2955842200800039 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-02-04 | 2008-02-01 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-01-05 | Yes | No | Reuse | 2008
42019004 |  | Malfunction |  | It was reported that one of the silver levers fell out of the cobra grasper instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed both release levers and the housing detached. The four housing pins and one snap tab are broken off, indicating that the instrument was mishandled. It was determined that the instrument housing most likely popped off during mishandling causing the levers to fall out. Engineering also observed the distal end of the main tube has a deep scratch above the proximal clevis with material removed. Based on the location and appearance, the scratch mark is most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 992266 | 990885 | 950039 | 2955842200800051 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 42019004 | Yes | Yes | 2008-01-11 | Yes | No | Reuse | 2008
40018402 | 2008-01-10 | Malfunction |  | It was reported that during the beginning of a da vinci myomectomy procedure, while dissecting fibroids from the uterus, the permanent cautery spatula instrument stopped working. The surgical staff then noticed the instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the instrument from the cannula accessory. The instrument was replaced with another instrument of the same type. The surgical staff then noticed after a few minutes that the replacement permanent cautery spatula instrument was broken at the elbow. As a result, the surgical staff undocked the system arm and cut the instrument tip off in order to remove the replacement instrument from the cannula accessory. | The instrument and detached wrist were returned and evaluated. Per the customer reported complaint, the proximal clevis is intact and pieces of the main tube interface wall are missing. Based on the engineering evaluation, a likely failure mode was overloading of the instrument at the tip, causing the breakage. No other damage found. Per the case notes, no broken instrument components fell into the patient. The following medwatch reports listed below were also sent to the fda and are related to this case. 2955842-2008-00049. 2955842-2008-00050. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 993420 | 990886 | 950040 | 2955842200800048 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Health Professional | 40018402 | Device Returned To Manufacturer | Yes | 2008-01-10 | Yes | No | Reuse | 2008
IS2000 A51P2 | 2008-01-08 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced a system error 22003. With the support of an isi regional service specialist, the site attempted to troubleshoot the system error without success. The site converted to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that system error code #22003 experienced by the customer was associated with the right gimbal master tool manipulator (mtm). The master tool manipulator refers to the master manipulators at the surgeon's console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The right gimbal is a component within the mtmr, which contains 3 of 7 axis. The system was repaired by replacing the affected right gimbal mtm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #22003 appears when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety system put davinci s in a "recoverable safe state". The right gimbal mtm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer assembly had malfunctioned, thus generating the system error experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 992768 | 991135 | 950289 | 2955842200800042 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-04 | 2008-02-01 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-01-08 | Yes | No | Reuse | 2008
PS2000 | 2008-01-15 | Malfunction |  | During a robotic assisted mitral valve repair surgery, the robotic instrument arm which holds the left atrial retractor had an error code and was no longer usable. The patient was on the heart lung machine at this time and the heart had already been opened. The robotic instrument was removed, the arm moved out of the way, and a manual atrial retractor system was used instead. The patient spent additional time on the heart lung machine due to this issue. At this time, the patient is doing fine. What is significant is that this arm had malfunctioned prior to this surgery and the intuitive technician had calibrated and repaired the robot and returned it to service. This was the first surgery since the repair and the arm malfunctioned. |  | DA VINCI S MODEL ROBOT | ROBOTIC SURGICAL SYSTEM | 978029 | 1010484 | 969410 | 1010484 | 1 | NAY | User Facility |  | Nurse | Initial | 2008-02-01 | 2008-02-01 | No | Yes | Physician | PS2000 | Yes | Yes |  |  |  |  | 2008
| 2007-12-21 | Injury | Required Intervention | Robot assisted prostatectomy had frequent error messages from robot. Davinci rep called but unable to fix problem; error codes began more frequently and resulted in inability to override. Robotic assist procedure aborted and surgeon proceeded with an open prostatectomy instead. Follow-up: manufacturer repaired robot in 2007. |  | DAVINCI S ROBOT | SURGICAL ROBOT | 960945 | 991148 | 950302 | MW5005223 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2008-01-14 | 2008-01-30 | No | Yes | Health Professional |  | No | No |  |  |  |  | 2008
ISI2000 A43P8 | 2008-01-03 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure the customer experienced recurring system error code # 212. The site was advised by isi regional service support to undock the patient and reboot the system. The system error code # 212 continued to occur and could not be overridden. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the three master manipulators at the surgeon's console. The mtm is assigned to the surgeon's left hand (mtml), one to his right (mtmr), and one to control the camera video position (mtmv). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #212 appears when the davinci safety system determined a voltage tracking fault reported by the digital signal processor (dsp). Upon determining these conditions, the safety systems put davinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Based on the system error logs, an encoder assembly and a motor assembly were replaced. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 993384 | 988965 | 948139 | 2955842200800036 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-02-04 | 2008-01-28 | Yes | Yes | Health Professional | ISI2000 A43P8 | Device Returned To Manufacturer | No | 2008-01-03 | Yes | No | Reuse | 2008
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had a cable loose. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire is broken at the yaw pulley exit. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion. Engineering also observed the distal end of the main tube has a few deep scratch marks with material removed. Based on the random location and appearance, the scratch marks were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments of other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 988934 | 985929 | 945137 | 2955842200800035 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-25 | 2008-01-25 | No | Yes | Health Professional | 42021403 | Yes | No | 2008-01-03 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | The large needle driver instrument was returned to isi with problem stated as unk. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a tooth on the internal roll pulley tube hitting a grip cable, as the roll axis rotates, causing a clicking sound at the back end of the instrument. Roll axis motion is not affected by this. Engineering also observed the distal end of the main tube has an arced shaped deep scratch located. 650 inches above the proximal clevis, with material removed. The orientation of the scratch is perpendicular to the tube axis, suggesting the scratch may be due to an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 988919 | 985930 | 945138 | 2955842200800037 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-25 | 2008-01-25 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-01-04 | Yes | No | Reuse | 2008
IS2000 A51P2 | 2007-07-16 | Death | DeathRequired Intervention | Medwatch uf/importer report was rec'd at isi on september 28, 2007. Intuitive surgical was previously informed of this event on july 25, 2007. It was also reported that the pt was later brought back to the or in 2007, to drain clots which had formed, however, the da vinci sys was not used for this procedure. The pt died four days post da vinci assisted surgery. | On august 17, 2007, intuitive surgical spoke to the physician who performed the procedure. The physician indicated that the pt laceration (the tear at the left atrium) could have occurred by a number of different causes. The physician also indicated that the pt died of severe heart failure and pulmonary edema, however, the physician does not believe that the pt death was related to the laceration or repair. No malfunction of the sys was alleged by the hosp. On january 22, 2008, a hosp rep indicated that no conclusion can be drawn regarding any association with the pt laceration and the add'l procedure performed in 2007, as decreased function and left ventricular failure is a possibility with any cardiac surgery, regardless of a laceration. Per customer request one month later, an intuitive surgical field engineer tested the sys and verified the sys was operating within normal specs. No issues with the sys were found and the sys was verified for normal operation. The hosp has continued to use the da vinci sys to perform surgery on pts. As of the date of this report, no adverse events or prod problems have been experienced with the site's da vinci sys which may have caused or contributed to a death or serious injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 959958 | 991415 | 894316 | 2955842200800038 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-01-25 | 2008-01-25 | Yes | No | Health Professional | IS2000 A51P2 | Yes | Yes | 2007-07-25 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that shards were found on the shaft of the pk dissecting forceps instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a 1. 5" long,. 200" wide section with material removed on one side of the tube. The damaged area is parallel to the tube axis and located. 250" above the proximal clevis. The surface finish of the tube is rough in the damaged area due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 956224 | 985899 | 945108 | 2955842200800034 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-01-24 | 2008-01-24 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2008
42009305 |  | Malfunction |  | It was reported that at the end of a da vinci s prostatectomy procedure, when trying to remove the endowrist prograsp instrument from the trocar, the instrument became stuck. As a result, the surgical staff forcefully removed the instrument and upon doing so, the tip broke and pieces fell into the pt. The broken instrument components were found and removed from the pt and the planned surgical procedure was completed with another instrument of the same type. During sterilization, after the procedure, it was reported that the instrument tip became completely severed from the shaft during the sterilization process. No pt harm, adverse outcome or injury was reported. | The instrument and detached wrist were returned and evaluated. Per the customer reported complaint, engineering observed the cables have been cut at the distal end. The main tube interface wall is missing a section and the proximal clevis is missing the section that mates with the main tube interface wall. Based on the engineering evaluation, a likely failure mode was overloading of the instrument at the tip, causing the breakage. Per the case notes, the broken instrument components were successfully retrieved from the pt. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 989936 | 989244 | 948417 | 2955842200800032 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-18 | 2008-01-18 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-12-28 | Yes | No | Reuse | 2008
| 2007-12-10 | Malfunction |  | A piece broke off the cadiere forceps for the da vinci robot, but did not break in the patient. It was a not a clean, straight break, but left a very jagged end. The break was not in the metal tip of the forceps, but rather in the plastic, about 0. 5 inch up from where the metal tip is attached to the long plastic handle. |  | DA VINCI CADIERE FORCEPS | ROBOTIC ASSISTED SURGICAL INST | 968102 | 999480 | 958551 | 999480 | 1 | NAY | User Facility |  |  | Initial | 2007-12-20 | 2008-01-18 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2008
IS1200 A43P8 | 2007-12-19 | Other | Other | It was reported that during a da vinci surgical procedure the customer experienced multiple occurrences of system error code #20013. The procedure was aborted after port incisions had been created and after the pt had been under anesthesia for one hour and ten mins. The procedure was rescheduled. No pt harm was reported. | The investigation conducted by field service engineering found that the system error code # 20013 was experienced by the customer as well as system error code # 21013. The field service engineer concluded that these system error codes were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The # 20013 and # 21013 system error code appeared when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation. Based on the system error logs, a potentiometer assembly and a motor assembly were replaced. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 988741 | 981087 | 940447 | 2955842200800027 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-01-15 | 2008-01-15 | Yes | Yes | Service Personnel | IS1200 A43P8 | Yes | Yes | 2007-12-19 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2007-12-27 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon experienced a difference in focus between the left and the right eyes in the surgeons console. The site changed da vinci surgical systems to successfully complete the procedure which created a 30 min delay. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi engineering concluded that the focus issue experienced by the customer was due to a defective camera head. The system was repaired by replacing the camera head. A camera head converts the reflected photons passing through the endoscope into analog video signals and performs focus adjustment. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 985490 | 981090 | 940450 | 2955842200800029 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-01-15 | 2008-01-15 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | Yes | 2007-12-27 | No | No | Reuse | 2008
420179 | 2007-12-19 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument broke and fragments fell into the pt. The planned surgical procedure was completed with another instrument of the same type and was lengthened by 20 minutes. No pt harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | MONOPULAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 982615 | 979308 | 938939 | 2955842200800028 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-11 | 2008-01-11 | No | Yes | Health Professional | 420179 | No | No | 2007-12-19 | Device Not Retu | No | Unkown | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument is broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed all components functions properly. Instrument moved intuitively with full range of motion in all directions. Engineering also observed the distal end of the main tube has various scratch marks all around the tube. Some of the scratches have removed material from the tube. Based on the location and appearance, the scratch marks are most likely due to instrument collisions. No other damage found. The endowrists instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 982625 | 979309 | 938940 | 2955842200800024 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-11 | 2008-01-11 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-12-18 | Yes | No | Reuse | 2008
400180 | 2007-12-18 | Malfunction |  | It was reported that during a da vince s prostatectomy procedure the surgeon was using the monopolar curved scissors instrument to cut tissue intra-abdominally, when a spark was noticed on the screen. The instrument was replaced with another of the same type and the planned procedure was completed with approx a 10 minute delay. The site indicated that the 8mm monopolar curved scissors tip cover accessory used in conjunction with the instrument during the procedure was thrown away. No pt harm, adverse outcome or injury was reported. | As indicated in the case notes, the tip cover accessory was discarded by the site, therefore, the reported failure could not be confirmed nor denied. The monopolar curved scissors instrument used in conjunction with the tip cover accessory during the reported case has not been returned or evaluated at this time, see mdr 2955842-2008-00025. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 982638 | 979310 | 938941 | 2955842200800033 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-11 | 2008-01-11 | No | Yes | Health Professional | 400180 | No | No | 2007-12-18 | Device Not Retu | Yes | Initial | 2008
42017908 | 2007-12-18 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon was using the monopolar curved scissors instrument to cut tissue intra-abdominally, when a spark was noticed on the screen. The instrument was replaced with another of the same type and the planned procedure was completed with approx a 10 minute delay. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation. The reported failure could not be confirmed nor denied. The tip accessory used in conjunction with this instrument during the reported case has not been returned or evaluated at this time, see mdr 2955842-2008-00033. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 981536 | 979311 | 938942 | 2955842200800025 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-11 | 2008-01-11 | No | Yes | Health Professional | 42017908 | No | No | 2007-12-18 | Device Not Retu | No | Unkown | 2008
400183 |  | Malfunction |  |  | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 951219 | 983785 | 926335 | 2955842200800020 | 1 | GEI | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial |  | 2008-01-11 | No |  | Invalid Data | 400183 |  | No | 2007-11-19 | Device Not Retu | No | Initial | 2008
42009305 |  | Malfunction |  | It was reported that after sterilization the endowrist prograsp instrument was found with "fibers coming out from the shaft". No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube has two sections, 180 degrees apart, with material removed. Sections are parallel to tube axis and roughly 1. 3 inches long and. 125 inches wide. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Engineering also observed the proximal clevis is missing a. 160 inch by. 200 inch piece, but the clevis is not dislodged from the main tube. It appears that the clevis may have broken due to mishandling of the instrument. No other damage was found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 996311 | 997864 | 956938 | 2955842200800052 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-08 | 2008-01-11 | No | Yes | Health Professional | 42009305 | Yes | No | 2008-01-16 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the shaft of the pk dissecting forceps instrument has scratches on it. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube has various randomly oriented deep scratches with material removed from the main tube. Based on the location and appearance, the scratch marks are most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instruments intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument. Engineering also observed the instrument is missing one release lever. The housing is intact but likely came off in order for the lever to fall out, indicating that the instrument ws mishandled. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 996305 | 997865 | 956939 | 2955842200800045 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-08 | 2008-01-11 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-01-16 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that after sterilization the large needle driver instrument was found with "fibers coming out from the shaft". No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube has a 2. 5 inch long by. 250 inch wide section with material removed on one side. Damaged area is parallel to the tube axis and has a rough surface finish due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Also observed were faint scratch marks on the opposite side of the tube. These scratches are likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instruments intraoperatively. This may result in damage to the instruments or other unintended consequences, such as the disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 996295 | 997866 | 956940 | 2955842200800031 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-08 | 2008-01-11 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-01-16 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument would not open and close properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to move intuitively with full range of motion in all directions. The instrument jaws open and close properly and no issues were found with the friction and cable tensioning passed. Engineering also observed the distal end of the main tube has a. 750 inch long by. 220 inch wide section with material removed on one side of the tube. Damaged area is parallel to the tube axis and has a rough surface finish due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 996288 | 997869 | 956943 | 2955842200800054 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-02-08 | 2008-01-11 | No | Yes | Health Professional | 42022702 | Yes | No | 2007-01-17 | Yes | No | Reuse | 2008
| 2007-08-20 | Death | DeathHospitalization | Pt underwent robotic tecab (totally endoscopic coronary arteries bypass) with 2 internal mammary arteries, using the da vinic s robot and coalescent "u" clips to facilitate distal anastomoses, on a beating heart. On coronary angiography 5 days later, there was 80% stenosis of one distal anastomosis, and 40% stenosis of another. |  | DA VINCI S | ROBOT | 951660 | 976910 | 936934 | MW5004896 | 1 | NAY | Voluntary |  | Physician | InitialFollowup | 2007-12-11 | 2008-01-09 | Yes | Yes | Other |  | No | Yes |  |  |  |  | 2008
| 2007-08-16 | Death | Death | Attempted tecab (totally endoscopic coronary artery bypass) surgery using the devinic s robot and faculated by the coalesent u clip on a beating heart. Only 2 out of 4 arteries were bypassed. Pt arrested at the kind of the proceduse. Cpr was initiated and the chest opened. Pt was connected to cadiopulomonary bypass and additional bypass grafts performed. Pt died on the operating room table. |  | DA VINCI S ROBOT |  | 951649 | 976911 | 936935 | MW5004897 | 1 | NAY | Voluntary |  | Physician | InitialFollowup | 2007-12-12 | 2008-01-09 | Yes | Yes | Other |  | No | Yes |  |  |  |  | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument had a broken cable. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | Conclusions & results - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable to be broken in two locations, at the distal idlers and the scissor grip. Pulley spins freely. The cable broke under tensile loading. Other cables at wrist are not damaged. Engineering also observed the main tube to have material removed all around the tube from the midpoint to the tube extension interface. The surface finish is very rough due to glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980586 | 976693 | 936747 | 2955842200800004 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-07 | 2008-01-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-12-08 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was dull, and a cable is loose or broken at the wrist. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the maryland bipolar forceps instrument was returned and evaluated. Engineering observed one grip close cable to be frayed at the distal idlers. A single strand of cable sticks out just before the idler pulleys. Grips can still open and close. The cable wear on pulleys may have contributed to fraying. Engineering also observed the distal end of the main tube to have a section with various deep scratches concentrated on one side of the tube. Material is removed from tube as a result of scratches. Scratches are not aligned with tube axis, suggesting they are due to instrument collisions. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980594 | 976694 | 936748 | 2955842200800023 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-07 | 2008-01-07 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-12-18 | Yes | No | Reuse | 2008
420179 | 2007-12-14 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon was using the monopolar curved scissors instrument when the cautery burned thru the wrist of the monopolar curved scissors tip cover accessory. The tip cover was replaced to successfully complete the procedure. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2008-00019. | Conclusions - the instrument was not returned for evaluation. The reported failure could not be confirmed nor denied. The tip accessory used in conjunction with this instrument during the reported case was returned and evaluated. See mdr 2955842-2008-000xx for failure analysis results. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980614 | 976695 | 936749 | 2955842200800018 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2008-01-07 | 2008-01-07 | No | Yes | Health Professional | 420179 | No | No | 2007-12-14 | Device Not Retu | No | Unkown | 2008
40018005 |  | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon was using the monopolar curved scissors instrument when the cautery burned thru the wrist of the monopolar curved scissors tip cover accessory. The tip cover was replaced to successfully complete the procedure. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2008-00018. | Conclusions - the tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover was returned with four holes burned through the silicone. The holes range in size from. 020 to. 050" and are located from. 185" to. 350" above the silicone to the sleeve interface. All holes exhibit localized melting and are scattered multiple around of the perimeter of the silicone, suggesting multiple arcing events occurred. The silicone also has various mechanical tears in the general vicinity of the holes. The tears are likely due to instrument collisions. The tip cover may have had tears in the current hole locations which created a path for arcing. Evidence of tearing may have been destroyed by the creation of the hole after an arcing event. The plastic sleeve also has various scratch event. The plastic sleeve also has various scratch marks concentrated on one side of the sleeve. These scratches may also be signs of instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 980835 | 976696 | 936750 | 2955842200800019 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-01-04 | 2008-01-07 | No | Yes | Health Professional | 40018005 | Yes | No | 2007-12-14 | Yes | Yes | Initial | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument did not work. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions all components function properly. Engineering also observed the distal end of the main tube to have various scratch marks concentrated on one side of the tube. A couple of scratches are deep enough to have removed material from the tube. Scratches are not aligned with the tube axis, suggesting they are due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 980601 | 976697 | 936751 | 2955842200800014 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-01-07 | 2008-01-07 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-12-14 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during sterilization of the maryland bipolar forceps instrument the flush port fell out. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the instrument was returned and evaluated. Per the customer reported complaint engineering observed the luer plate to be dislodged from the chassis and pushed into the back end. The chassis feature that retains the luer plate is broken at the inner wall. Excessive axial force at the flush ports likely cause breakage. Engineering concluded that the likely cause of failure is mishandling/misuse in central processing. Engineering also observed distal end of the main tube to have deep scratch marks, 180 degrees apart, with material removed. Scratches are not aligned with the tube axis, suggesting they are due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FOCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980611 | 976699 | 936753 | 2955842200800022 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-12-07 | 2008-01-07 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-12-18 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2007-12-14 | Other | Other | It was reported that during a da vinci procedure, the customer experienced multiple occurrences of system error code #20008. Initially, the site was able to override the error and continue with the case but the error recurred and could not be overridden. No pt harm, adverse outcome, or injury was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering concluded that system error codes #20009 and #21009 were experienced by the customer and not system error code #20008. This system error code was associated with a pt side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error codes #20009 and #21009 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. Upon determining this condition, the safety system put da vinci in a "recoverable safe state". Based on the system error logs two motor encoders were replaced. As of january 07, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 983202 | 976851 | 936901 | 2955842200800017 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-07 | 2008-01-07 | Yes | Yes | Service Personnel | IS1200 A43P9 | Device Returned To Manufacturer | No | 2007-12-14 | Yes | No | Reuse | 2008
42004904 | 2007-10-31 | Malfunction |  | The robotic forceps was placed inside the pt during the procedure, and the end of the forcep broke off into the pt. The one piece was retrieved and the forcep removed. The trocar remained on the forcep and was unable to remove the trocar from the forcep. While attempting to remove the trocar, that was needed for the remainder of the case, two other pieces broke off. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). ". | The instrument was not returned for eval. The root cause to the customer reported failure mode cannot be determined. If additional info becomes available a follow-up mdr will be submitted. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 949979 | 1009545 | 941716 | 2955842200800011 | 1 | GCJ | Manufacturer | User facility | RISK MANAGER | Initial |  | 2008-01-07 | No | Yes | Health Professional | 42004904 | No | Yes | 2007-12-12 | Device Not Retu | No | Unkown | 2008
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a broken wire. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no broken conductor wires or damaged cables. Engineering also observed the instrument is broken at the proximal clevis to main tube interface with the clevis dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces. One of the clevis pieces is fractured at the midpoint between the ears, indicating overloading at the tip occurred with the instrument wrist pitched. The distal end of the main tube has a 1. 5 inch long,. 250 inch wide section with material removed on one side of the tube. Damaged area is located 3. 25 inches above the clevis and has a rough surface finish due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977844 | 975226 | 935506 | 2955842200800013 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-12-13 | Yes | No | Reuse | 2008
42017908 | 2007-11-09 | Malfunction | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was dissecting nodes, a burn on the patient's bowel was noticed. Upon further inspection, the surgeon found a crack on the tip cover accessory of the monopolar curved scissors instrument. The instrument was removed and the tip cover was replaced. The bowel was repaired by over sewing the burn and the procedure was completed with an approximate delay of one hour. No additional patient harm, adverse outcome or injury was reported. | The site returned 2 monopolar curved scissors instruments, making it unclear as to which instrument was used with the tip cover accessory. The instrument was returned and evaluated. Per the customer reported complaint, engineering found no char marks or any other signs of arcing on the metal components at the wrist of the instrument. The metal components do not have any abnormally sharp edges. Latex cut test and electrical continuity passed. All components functioned properly. Medwatch mfr report #2955842-2007-00494 was sent to the fda for the other monopolar curved scissors instrument returned and medwatch mfr report #2955842-2007-00495 was sent to the fda for the monopolar curved scissors tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977854 | 975227 | 924415 | 2955842200800008 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2008-01-04 | 2008-01-04 | Yes | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-12-11 | Yes | No | Reuse | 2008
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a broken wire. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire is broken at the yaw pulley exit. The yaw pulley shows no signs of arcing. Electrical continuity failed. Engineering also observed the instrument main tube has a. 750 inch long,. 180 inch wide section with material removed parallel to the tube axis. Damaged area is located 4. 5 inches above the clevis and has a rough surface finish, due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977863 | 975228 | 935508 | 2955842200800012 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-12-13 | Yes | No | Reuse | 2008
40020502 | 2007-12-04 | Malfunction |  | It was reported that during a da vinci hysterectomy surgical procedure, while dissecting tissue, the fenestrated bipolar forceps instrument tip broke and pieces fell into the patient. The surgical staff immediately removed the instrument and the trocar from the patient. The broken instrument components were found and removed from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface and the clevis is dislodged from the main tube as a result. Both the proximal clevis and the main tube are missing pieces. Based on the engineering evaluation, a likely failure mode was overloading of the instrument at the tip, causing the breakage. Per the case notes, the broken instrument components were successfully retrieved from the patient. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977692 | 975229 | 935509 | 2955842200800005 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-01-04 | 2008-01-04 | No | Yes | Health Professional | 40020502 | Yes | No | 2007-12-10 | Yes | No | Reuse | 2008
40000620 | 2007-12-10 | Malfunction |  | It was reported that near the end of a da vinci prostatectomy surgical procedure, during suturing, a piece of the tip of the large needle driver instrument broke off and fell into the patient. The broken piece was removed causing approximately a two minute delay. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the carbide insert has detached from one of the grips, the detached insert was returned with the instrument. The brazing surface of the grip does not have full area coverage of the filler metal. The insert brazing surface has almost no traces of filler metal and is discolored at the center, indicating the absence of filler metal. It was concluded that the grip was inadequately brazed during manufacturing leading to shearing failure during suturing. No other damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 977685 | 975230 | 935510 | 2955842200800010 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 40000620 | Yes | No | 2007-12-12 | Yes | No | Reuse | 2008
40018005 |  | Malfunction | Other | It was reported that smoke was coming out of the wrist of the monopolar curved scissors instrument (mcs), which had a tip cover accessory attached to it. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2008-00001. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover to have a. 020" x. 035" oblong hole in the silicone, located. 260" above the sleeve to silicone interface. The hole exhibits localized melting at the outer surface of the silicone, indicating electrical energy arced through the silicone. The tip cover may have had a cut from an instrument collision at the hole site, which would weaken dielectric strength of silicone and create a path for arcing. Evidence of the cut may have been removed by the arcing. The monopolar curved scissors instrument returned with tip cover did not have any obvious anomalies that would cause hole. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 977738 | 975231 | 935511 | 2955842200800002 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-01-04 | 2008-01-04 | No | Yes | Health Professional | 40018005 | Yes | No | 2007-12-06 | Yes | Yes | Initial | 2008
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire is broken at the yaw pulley exit. Engineering also observed the distal end of the main tube has various scratch marks concentrated on one side of the tube. Some of the scratches have removed material from the tube. Based on the location and appearance, the scratch marks are most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977677 | 975232 | 935512 | 2955842200800015 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-12-14 | Yes | No | Reuse | 2008
42017206 | 2007-11-15 | Malfunction |  | It was reported that the maryland bipolar forceps instrument has a broken cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed no broken cables or conductor wires. Engineering also observed the distal end of the main tube has a combination of deep scratches and scrape marks on the main tube. Material is removed as a result of the damage. The wear pattern of tube scraping is consistent with damage from a defective cannula. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Scratch marks are likely due to instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from. Electrical continuity passed. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977716 | 975233 | 935513 | 2955842200800007 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-12-11 | Yes | No | Reuse | 2008
42020503 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument has "a lot of wear at the tip". No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube has a combination of deep scratches and scrape marks on one side of the main tube. Material is removed as a result of the damage. The wear pattern of tube scraping is consistent with damage from a defective cannula. Additional investigation is underway regarding the cannula accessories and a follow-up mdr will be submitted if additional information is received. Scratch marks are likely due to instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from. Engineering also observed one conductor wire is broken at the yaw pulley exit. Wire may not have been securely attached to grip and broke during articulation of the wrist. Electrical continuity failed. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977694 | 975234 | 935514 | 2955842200800006 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-01-04 | 2008-01-04 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-12-11 | Yes | No | Reuse | 2008
40017908 |  | Malfunction | Other | It was reported that smoke was coming out of the wrist of the monopolar curved scissors instrument (mcs), which had a tip cover accessory attached to it. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2008-00002. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed no charring on any of the components. A tip cover was returned with a hole burned through the silicone, which is likely why the instrument was smoking, but the mcs has no abnormally sharp edges or burrs that would cut the tip cover. The instrument placed on the system and driven. Performance testing was performed and the instrument was found to move intuitively with a full range of motion in all directions. All components function properly. A latex cut test and electrical continuity passed. No physical damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 977688 | 975235 | 935515 | 2955842200800001 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-12-06 | 2008-01-04 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2007-12-06 | Yes | No | Reuse | 2008
IS2000 A5 1P2 | 2007-12-06 | Malfunction | Other | It was reported that after a da vinci s hysterectomy procedure, the surgical staff found the pt's skin surrounding the port incisions where the cannula accessories were, was "charred". The surgeon indicated that the pt required no additional medical treatment and felt fine the following day. No additional pt harm, adverse outcome or injury was reported. | Investigation was performed by field service engineering. Since the site's last service event, the field service engineer found that the vision cart had been re-configured and rewired. The customer mounted a surge protector to the vision cart which was powering the isolation transformer, illuminator, assistant monitor, insufflator and cautery system. The field service engineer re-configured and rewired the vision cart to isi specifications and recommended to the hosp staff that the cautery system, and the assistant monitor be powered separately from the vision cart, due to the power demands of these components. A full electrical safety test was conducted as well as a full preventative maintenance service and no issues were encountered. The da vinci s surgical system and the electrical surgical unit (esu) system passed the ground bond and leakage current electrical safety test; however, cautery generators are known to product high frequency leakage currents. It is possible for these currents to pass through the pt side manipulator arms to ground, through the cannula accessory to the pt. Re-configuration of the vision cart could have aggravated this high frequency leakage. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 979779 | 975624 | 932816 | 2955842200800003 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-01-04 | 2008-01-04 | Yes | No | Health Professional | IS2000 A5 1P2 | Yes | No | 2007-12-07 | Yes | No | Reuse | 2008
IS2000 A51P2 | 2007-12-14 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the customer experienced multiple occurrences of system error code #23008. The pt was under anesthesia for 3 hours and 20 minutes and the port incisions were already made when the surgeon decided to abort the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the customer was associated with an endoscopic camera manipulator (ecm). The endoscopic camera manipulator is the camera arm located on the pt side cart, that provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The da vinci s safety systems determined that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state". Testing performed by engineering discovered that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrences of the issues at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSRUMENT CONTROL S | 979046 | 975778 | 940451 | 2955842200800016 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2008-01-04 | 2008-01-04 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | Yes | 2007-12-14 | Yes | No | Reuse | 2008
|  | No Answer Provided | [5 empty] | 953677 | 986136 | 921279 | 2955842200700470 | 1 | NAY | Manufacturer |  | Invalid Data | Initial | 2008-01-04 | 2008-01-04 | No |  | Invalid Data |  |  | No Answer Provided | 2007-12-18 | No Answer Provi | No An |  | 2008
420172 | 2007-10-05 | Other | Other | It was reported that after a da vinci prostatectomy surgical procedure, the physician noticed exterior burn markings above and medial to the working port on the left of the pt's abdomen. The area was reported as a superficial second degree burn, erythema in color. The surgical staff could not remember which instrument was used in this port and did not return instruments for eval despite isi's request. No additional info was reported. The site reported a second occurrence which occurred the same day. See mdr 2955842-2007-00549. | The instrument was never returned for eval. The customer reported, complaint can not be confirmed or denied at this time. If the instrument is returned a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 973669 | 972033 | 914202 | 2955842200700550 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-28 | 2007-12-28 | Yes | Yes | Health Professional | 420172 | No | No | 2007-10-09 | Device Not Retu | No | Unkown | 2007
420172 | 2007-10-05 | Other | Other | It was reported that after a da vinci prostatectomy surgical procedure, the physician noticed exterior burn markings above and medial to the working port on the left of the pt's abdomen. The area was reported as a superficial second degree burn, erythema in color. The surgical staff could not remember which instrument was used in this port and did not return instruments for eval despite isi's request. No additional info was reported. The site reported a second occurrence which occurred the same day. See mdr 2955842-2007-00550. | The instrument was never returned for eval. The customer reported complaint can not be confirmed or denied at this time. If the instrument is returned a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 973687 | 972035 | 914202 | 2955842200700549 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-28 | 2007-12-28 | Yes | Yes | Health Professional | 420172 | No | No | 2007-10-09 | Device Not Retu | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. , as this site had previously returned instruments which showed indications of cannula related tube abrasion. These were reported via the following accessory reports: mfr. Report # 2955842-2007-00239, -00355, -00356, -00357. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr. Report #'s 2955842-2007-00531, -00532. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 977410 | 971088 | 931820 | 2955842200700533 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-27 | 2007-12-27 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-12-04 | Yes | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. , as this site had previously returned instruments which showed indications of cannula related tube abrasion. These were reported via the following accessory reports: mfr. Report # 2955842-2007-00239, -00355, -00356, -00357. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr. Report #'s 2955842-2007-00531, -00533. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 977404 | 971089 | 931821 | 2955842200700532 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-27 | 2007-12-27 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-12-04 | Yes | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. , as this site had previously returned instruments which showed indications of cannula related tube abrasion. These were reported via the following accessory reports: mfr. Report # 2955842-2007-00239, -00355, -00356, -00357. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr. Report #'s 2955842-2007-00532, -00533. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 977399 | 971090 | 931822 | 2955842200700531 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-27 | 2007-12-27 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-12-04 | Yes | No | Unkown | 2007
42018304 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the permanent cautery hook instrument was "slipping out of the sleeve and would not articulate". The surgical staff also noticed a small piece of the instrument had fallen inside the pt, which was removed and discarded. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the ceramic sleeve is intact and the hook is not twisted. Engineering also observed that the instrument is broken at the proximal clevis to main tube interface, with the clevis dislodged from the main tube as a result. Both the proximal clevis and the main tube are missing pieces. Per the case notes, the broken instrument fragment was successfully retrieved from the pt. The clevis has various scratch marks all around the outer diameter. The location and appearance of the scratches suggests the damage may be due to instrument collisions with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980041 | 971688 | 932366 | 2955842200700535 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2007
42020503 | 2007-11-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy surgical procedure, the surgical staff noticed the fenestrated bipolar forceps instrument had a loose cable. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed no loose or damaged cables. One conductor wire is broken at the yaw pulley exit, the yaw pulley shows no signs of arcing. The proximal clevis is missing a. 250 inch by. 200 inch piece at the main tube interface. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a pt during a surgical procedure. Engineering also observed the distal end of the main tube has sections with material removed. Damaged areas vary in length and width, all parallel to the tube axis, with the wear pattern consistent with damage from a defective cannula. No other damage was found. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 973643 | 971690 | 932368 | 2955842200700534 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the endowrist prograsp instrument tip broke. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tip is not broken. One grip cable is frayed and a strand is sticking out at the instrument wrist. A nick on the pulley edge was observed, likely caused by contact with another instrument, causing damage to the grip cable. Engineering also observed the distal end of the main tube has a 1 inch long. 110 inch wide section with light material removed on one side of the tube. Damage occurred 1. 35 inches above the proximal clevis and the wear pattern is consistent with damage from a defective cannula. No other damage was found. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 973637 | 971691 | 932369 | 2955842200700530 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2007-12-03 | Yes | No | Reuse | 2007
42020503 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument is not working with the light cord. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire is broken at the yaw pulley exit. Both yaw pulleys have charring at the grip base, with material removed due to burning. Engineering also observed the distal end of the main tube has various scratch marks with material removed from the tube. The location and appearance of the scratches suggests the damage may be due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 973738 | 971692 | 932370 | 2955842200700529 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-11-27 | Yes | No | Reuse | 2007
42009305 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the surgical staff noticed the instrument tip broke off. The broken component was found and removed from the pt without incident. The planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument cables are cut at the distal end and the wrist is detached from the housing. The proximal clevis is broken and missing a piece below one ear. Most of the main tube interface wall has been cut off. Based on the engineering eval, a likely failure mode was a break at the proximal clevis to main tube interface, and the wrist was cut off to remove the instrument from the cannula accessory. Per the case notes, the broken instrument component was successfully retrieved from the pt. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 973666 | 971693 | 932371 | 2955842200700523 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42009305 | Device Returned To Manufacturer | No | 2007-11-28 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument tips are not aligned. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed grips are aligned and that the teeth mesh properly and the grips are not bent. Engineering also observed the distal end of the main tube has various scratch marks on one side with a material removed from the tube. The location and appearance of the scratches suggests the damage may be due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 980049 | 971695 | 932373 | 2955842200700521 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2007-11-27 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument does not hold needles properly. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument moved intuitively with full range of motion in all directions and there were no issues observed with grasping. Engineering also observed the distal end of the maintube has a 2. 5 inch long section with material removed all around the main tube. Damage occurred 3 inches above the proximal clevis and the wear pattern is consistent with damage from a defective cannula. No other damage was found. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 973650 | 971696 | 932374 | 2955842200700520 | 1 | GCJ | Manufacturer | Health Professional | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | Yes | 2007-11-26 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that a frayed cable was found on the pk dissecting forceps instrument. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip open cable is frayed at the distal idlers. One cable strand sticks out at the yaw pulley. The cable likely frayed because it derails from the proximal pulley as the wrist articulates, causing abrasion on pulley edges. Engineering also observed the main tube to have various deep scratches at the distal end. Some of the scratches have removed material from the tube. Scratch marks are likely due to instrument collisions. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 980056 | 971697 | 932375 | 2955842200700540 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-12-05 | Yes | No | Reuse | 2007
42004904 |  | Malfunction |  | It was reported that shards were found on the shaft of the cadiere forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a 2. 5" long,. 275" wide section with material removed on one side of the tube. The damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. Engineering also observed the main tube to have a couple of scratch marks likely from instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 973740 | 971698 | 932376 | 2955842200700536 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42004904 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2007
42022702 |  | Malfunction |  | It was reported that shards were found on the shaft of the pk dissecting forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip to be bent, creating a. 030" offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed the distal end of main tube to have a 4" long section with material removed all around the tube. The damaged area is gray in color and has a rough surface finish due to the glass fibers being exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 973747 | 971699 | 932377 | 2955842200700537 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2007
42004905 |  | Malfunction |  | It was reported that shards were found on the shaft of the cadiere forceps instrument. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube to have a couple of narrow sections concentrated on one side of the tube with material removed. The damaged sections range from. 350-. 950" in length and. 060-. 075" in width. The material removal is not severe, but there is a change in surface roughness compared to the remainder of the tube. The damaged sections are parallel to the tube axis, however, there are gaps between the sections, suggesting multiple instances of scraping occurred. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 974003 | 971700 | 932378 | 2955842200700539 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42004905 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Initial | 2007
42004905 |  | Malfunction |  | It was reported that shards were found on the shaft of the cadiere forceps instrument. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed. The distal end of the main tube to have a deep scratch with material removed 2" above the proximal clevis. The scratch is. 130" long, not parallel to tube axis, and was likely caused by an instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 974008 | 971701 | 932379 | 2955842200700538 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42004905 | Device Returned To Manufacturer | No | 2007-12-04 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the jaws on the large needle driver instrument would not close. No other info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing. The instrument was found to move intuitively with full range of motion in all directions. The jaws opened and closed properly with no issues. The distal end of the main tube has a 2" long,. 220" wide section with material removed on one side of the tube. The damaged area is located 2" above the proximal clevis and has a rough surface finish. The damaged area is not fully parallel to the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional info is received. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 974017 | 971702 | 932380 | 2955842200700522 | 1 | GCJ | Manufacturer | Health Professional | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-11-27 | Yes | No | Reuse | 2007
42017907 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument broke where the shaft meets the wrist. The instrument was intact without loss of pieces. The procedure was successfully completed using another instrument of the same type. No further info was available. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension to be dislodged from the main tube. The keys on the tube extension have broken off. As a result, the tube extension can be rotated 360 degrees. One side of the tube extension is slightly melted at the break point, indicating that arcing had occurred. The main tube does not have any signs of arcing. Engineering concluded that it was likely that the break occurred before the tube extension was burned, as the break would create a path for electrical energy to escape. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 974041 | 971703 | 932381 | 2955842200700546 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42017907 | Device Returned To Manufacturer | No | 2007-07-23 | Yes | No | Reuse | 2007
42017908 | 2007-10-31 | Malfunction |  | It was reported that a hole was found in the carbon fiber of the monopolar curved scissors instrument. The hole was reported as being located 1 cm above the metal ring. No further info was available. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have a. 095" hole burned through the main tube. The hole is located roughly. 5" above the tube extension. The instrument was disassembled to find the distal end of the tube extension burned in the same place as the hole. The hypotubes and cap coupling tube have charring, with the cap coupling tube exhibiting localized melting. Looking on the inside of the main tube, a micro-crack was found that runs from the distal end of the tube to the bottom of the hole. This crack may have created a path for electrical energy to escape. The crack is usually found in all mcs instruments with this failure mode. The site also returned two tip covers, neither of which show any arcing damage. Failure at main tube is likely unrelated to tip covers. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 974051 | 971704 | 914303 | 2955842200700547 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-21 | 2007-12-21 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-11-01 | Yes | No | Initial | 2007
IS1200 A3P8 | 2007-12-05 | Other | Other | It was reported that during a da vinci prostatectomy procedure the site experienced several unknown system error codes. It was also reported that the patient was experience undisclosed issues. The procedure was converted to traditional open techniques to successfully complete the planned surgical procedure. No patient harm was reported. No additional information was provided. | The investigation conducted by field service engineering concluded that the system error code #21013 experienced by the customer was associated with the patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The psm was replaced and returned to isi for additional evaluation. The psm was returned to isi for failure analysis investigation. System error code #21013 is generated when an error is detected as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). Testing revealed no trouble could be found. The psm performed within specification. As of 2007 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 983342 | 972169 | 932812 | 2955842200700541 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-12-21 | 2007-12-21 | Yes | Yes | Health Professional | IS1200 A3P8 | Yes | Yes | 2007-12-05 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-12-05 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy surgical procedure the customer experienced 20013 system error codes. The patient was under anesthesia for 50 minutes and port incisions were made when the surgeon decided to abort the procedure. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with a patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The psm was returned to isi for failure analysis investigation. Engineering discovered both 20013 and 21013 in the error logs. The 20013 and 21013 system error codes appeared when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety system put davinci in a "recoverable safe state". Engineering concluded that the encoder and potentiometer assemblies had malfunctioned, thus generating the system errors experienced by the customer. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 982728 | 972171 | 932812 | 2955842200700542 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-21 | 2007-12-21 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-12-05 | Yes | No | Reuse | 2007
IS2000 A51P3 | 2007-12-06 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy surgical procedure, the customer found the master tool manipulator (mtm) left grip spring and lever came loose. The pt was under anesthesia for approximately 40 minutes when the decision was made to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering concluded that the pin that holds the lever and spring for the left master tool manipulator (mtml) had fallen out. The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Engineering discovered one of the mtml levers missing, therefore, the mtml grip was rebuilt. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 974918 | 972173 | 932816 | 2955842200700543 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2007-12-21 | 2007-12-21 | Yes | Yes | Health Professional | IS2000 A51P3 | Device Returned To Manufacturer | Yes | 2007-12-06 | Yes | No | Reuse | 2007
42017206 |  | Malfunction |  | An isi representative visiting the site discovered multiple da vinci s 8mm instruments with scratches on the main tube and advised that the account return them to isi for evaluation. | The maryland bipolar forceps instrument was returned and evaluated. Engineering found the distal end of the main tube has a section directly above the proximal clevis with various scratch marks, creating a rough surface and a small amount of material to be removed from the main tube. The scratches are concentrated on side of the tube and are likely due to inadvertent contact with other instruments. Engineering also observed a 3 inch section above the proximal clevis with light wearing. All cannulae at this site have been evaluated and were found to be within specification. Engineering believes a possible source of this damage is due to instrument shafts rubbing together. No other damage found. The endowrist instruments instructions for use (ifu) specifically states; general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972288 | 966987 | 928201 | 2955842200700514 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-18 | 2007-12-17 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-11-19 | Yes | No | Unkown | 2007
420179 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was damaged due by a damaged cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation. The customer reported complaint could not be confirmed nor denied. The site did return 6 cannula for investigation. Engineering found no damage on the cannulae that would cause or contribute to the damage of instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972289 | 966988 | 928202 | 2955842200700525 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2007-12-14 | 2007-12-17 | No | Yes | Health Professional | 420179 | No | No | 2007-11-16 | No | No | Unkown | 2007
42009307 |  | Malfunction |  | An isi representative visiting the site discovered multiple da vinci s 8mm instruments with scratches on the main tube and advised that the account return them to isi for evaluation. | The endowrist prograsp instrument was returned and evaluated. Engineering found the distal end of the main tube to have various scratch marks. A couple of the scratches have material removed from the main tube. The scratches are not aligned with the tube axis and are likely due to inadvertent contact with other instruments and not a defective cannula. No other damage found. The endowrist instruments instructions for use (ifu) specifically states; general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 972296 | 966989 | 928203 | 2955842200700515 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-12-18 | 2007-12-17 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-11-19 | Yes | No | Unkown | 2007
42000604 |  | Malfunction |  | An isi representative visiting the site discovered multiple da vinci s 8mm instruments with scratches on the main tube and advised that the account return them to isi for evaluation. | The large needle driver instrument was returned and evaluated. Engineering found the distal end of the main tube has a 2. 3 inch long by. 250 inch wide section with material removed. The scratches are concentrated on the side of the tube and are likely due to inadvertent contact with other instruments. Damaged area is parallel to the tube axis. All cannulae at this site have been evaluated and were found to be within specification. Engineering believes a possible source of this damage is due to instrument shafts rubbing together. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 972261 | 966990 | 928204 | 2955842200700516 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-12-18 | 2007-12-17 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-11-19 | Yes | No | Unkown | 2007
420172 |  | Malfunction |  | It was reported that the fenestrated maryland bipolar instrument was damaged due by a damaged cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation. The customer reported complaint could not be confirmed nor denied. The site did return 6 cannulae for investigation. Engineering found no damage on the cannulae that would cause or contribute to the damage of instruments. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972246 | 966991 | 928205 | 2955842200700526 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2007-12-14 | 2007-12-17 | No | Yes | Health Professional | 420172 | No | No | 2007-11-16 | No | No | Unkown | 2007
42014204 | 2007-11-19 | Malfunction |  | It was reported that during a da vinci assisted transoral tumor resectioning procedure the insulation on the main tube of the 5mm monopolar cautery instrument began to splinter. It was reported by an isi clinical sales representative who was present during the procedure, that the instrument had been used more aggressively than usual as the physician experienced difficulties accessing the operative area. The splintered insulation was cut off by a nurse and the procedure was completed without any further issues. No material fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluated. The customer reported complaint can not be confirmed nor denied. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972265 | 966992 | 928206 | 2955842200700518 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-17 | 2007-12-17 | No | Yes | Health Professional | 42014204 | No | No | 2007-11-19 | No | No | Reuse | 2007
42017908 | 2007-10-31 | Malfunction |  | It was reported that a hole was found in the carbon fiber of the monopolar curved scissors instrument. The hole was reported as being located 1 cm above the metal ring. No further information was available. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube to have a. 125" diameter hole burned through the main tube. The hole is located. 560" above the interface with the tube extension. The tube extension seal is burned in the same location as the hole. A crack was found on the main tube that runs from the distal end of the main tube to the bottom of the hole. The crack likely created a path for arcing. The crack may have been inherent in the main tube. A cracked main tube is a common feature with this failure mode. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972275 | 966993 | 914303 | 2955842200700524 | 1 | GEI | Manufacturer | Company Representative | UNKNOWN | Initial | 2007-12-17 | 2007-12-17 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-11-01 | Yes | No | Initial | 2007
42020503 |  | Malfunction |  | It was reported that during a da vinci surgical procedure, the "jaw" of the fenestrated bipolar forceps instrument broke off. The fragment was immediately retrieved. No patient harm, adverse outcome or injury was reported. | The fenestrated bipolar forceps instrument was returned and evaluated. The customer reported that the "jaw broke off," however, the failure investigation confirmed that the jaw and pulleys that the jaw is pinned to were not broken. The investigation did reveal that the instrument was broken at the proximal clevis to the main tube interface and the clevis was dislodged from the main tube as a result. Both the proximal clevis and main tube were missing pieces. Per the case notes, the broken instrument component was successfully retrieved from the patient. Customer was likely referring to the proximal clevis and main tube interface when stating that the "jaw broke off. " the distal end of the main tube has various scratch marks on one side of the tube with a small amount of material removed. The location and appearance of the scratches suggests the damage may be due to instrument collisions with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972234 | 966994 | 928208 | 2955842200700517 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-18 | 2007-12-17 | No | Yes | Health Professional | 42020503 | Yes | No | 2007-11-19 | Yes | No | Reuse | 2007
420006 |  | Malfunction |  | It was reported that the large needle driver instrument was damaged due by a damaged cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation. The customer reported complaint could not be confirmed nor denied. The site did return 6 cannulae for investigation. Engineering found no damage on the cannulae that would cause or contribute to the damage of instruments. | LARGE NEEDLE DRIVER INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 972238 | 966995 | 928209 | 2955842200700527 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2007-12-14 | 2007-12-17 | No | Yes | Health Professional | 420006 | No | No | 2007-11-16 | No | No | Unkown | 2007
42017205 |  | Malfunction |  |  | The instrument has not been returned for evaluated. The customer reported complaint can not be confirmed nor denied. If the instrument is returned for eval a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 946318 | 978467 | 926859 | 2955842200700519 | 1 | GEI | Manufacturer | Other | Invalid Data | Initial |  | 2007-12-17 | No |  | Invalid Data | 42017205 |  | No | 2007-11-20 | Yes | No | Unkown | 2007
IS1200 A43P8 | 2007-06-28 | Other | Other | It was reported that prior to beginning the da vinci prostatectomy procedure, the customer experienced system error code #201 when powering up the system. The system was restarted several times, however, system error code # 201 still recurred. The patient was under anesthesia when the surgeon decided to complete the procedure using the site's spare da vinci surgical system, causing an approximate delay of one hour. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by isi field service engineering concluded that system error code # 201 experienced by the customer may have been associated with a remote interface adapter (ria) pca and/or multiple servo driver (msd) pca. The system was repaired by replacing the affected ria and msd pcas. System error code #201 implies that the system diagnostics indicates a motor was not responding properly to the commanded voltage. The system alarm (system generated fault code) functioned as designed, and there was no injury to the patient. Upon determining this condition, the safety system put da vinci in a "non-recoverable safe state". The ria pca and msd pca were returned to isi for evaluation. Functional and system tests were performed and engineering was unable to reproduce the customer reported failure mode, therefore, the root cause of the customer reported failure mode cannot be definitively determined. As of december 14, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 974289 | 965658 | 848474 | 2955842200700379 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-04 | 2007-12-14 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-06-28 | Yes | No | Reuse | 2007
ISI200 A43P8 | 2007-11-14 | Other | Other | It was reported that 30 minutes into a da vinci prostatectomy procedure and after some dissection had been completed, one patient side manipulator (psm) became sluggish. The site rebooted the system, however, they could not complete homing and the error messages indicated "arms forced or not free to move". The procedure was aborted and the planned surgery procedure was rescheduled. No patient harm was reported. | The investigation conducted by isi field device engineering consisted of reviewing the system error logs and conducting performance tests. There were no error codes that appeared to have caused the system to fault as the site initially reported. The site was not clear on the exact problem and could not identify a specific arm that was causing the problems. There were numerous 20003 and 21003 error codes for both psm1 and psm2 found on the error logs. Reported symptoms could not be replicated. Psm1 was replaced based on the sequence of the error logs. A following case was performed with no issues. The psm was returned for evaluation. Engineering underwent additional performance testing and could not find any problem that would cause the arm to feel sluggish. It was determined that the system error 20003/21003 were "cascading" errors and not related to the event. The root cause of the customer reported failure mode could not be determined. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 974297 | 965659 | 927017 | 2955842200700505 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-12-14 | 2007-12-14 | Yes | Yes | Health Professional | ISI200 A43P8 | Device Returned To Manufacturer | No | 2007-11-14 | Yes | No | Reuse | 2007
IS2000 A5 1P3 | 2007-11-14 | Other | Other | It was reported that prior to surgery and while docking the da vinci s surgical system for a mitra valve repair procedure, the site experienced issues with a patient side manipulator (psm). The procedure was converted to mini thoracotomy techniques to finish the planned surgery prior to the site contacting isi to troubleshoot the issues. No patient harm was reported. | The investigation conducted by isi field service engineering consisted of reviewing the system error logs and conducting performance testing. In reviewing the system error codes, it was found that an instrument engagement issues was experienced as system fault codes were experienced on the patient side manipulator (psm) 2 and 3. The error indicates that no cannula was present on the device. The 31009 shows that only two of the three tool sensors were sensed on another device after all of the tool sensors were previously detected. A system error code was also experienced. This code indicates that the system lost contact with the dallas chip on the second device. Later they got a code on the second device that also indicated that no cannula was present. The isi field service engineer also conducted performance testing which resulted in no trouble being found. Field service engineering concluded that the event experienced by the site may have been due to user error. As of late 2007, the site has continued to use the da vinci s surgical system without a recurrence of this issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 974353 | 965660 | 927018 | 2955842200700501 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-12-14 | 2007-12-14 | Yes | Yes | Health Professional | IS2000 A5 1P3 | No | No | 2007-11-14 | Yes | No | Reuse | 2007
400179 | 2007-11-14 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, arcing was observed through the tip cover accessory of the monopolar curved scissors instrument. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below were also sent to the fda as they relate to this event. #2955842-2007-00503. | The instrument has not been returned for eval. A follow-up mdr will be submitted if the monopolar curved scissors instrument is returned and failure analysis has been completed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 967445 | 962816 | 924399 | 2955842200700502 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-13 | 2007-12-10 | No | Yes | Health Professional | 400179 | No | No | 2007-11-14 | Device Not Retu | No | Unkown | 2007
400180 |  | Malfunction | Other | It was reported that a spark at the tip of the monopolar curved scissors instrument occurred, and a hole in the tip cover accessory was noticed. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2007-00513. | The tip cover accessory has not been returned for eval. A follow-up mdr will be submitted if the monopolar curved scissors instrument is returned and failure analysis has been completed. The monopolar curved scissors instrument used in conjunction with the tip cover accessory, was returned and evaluated. Per the customer reported complaint, engineering observed the instrument has a. 060 inch by. 120 inch black char mark on the proximal clevis. There are two small char spots roughly. 060 inches above the large char mark. Engineering concluded this evidence suggests an arcing event occurred. Electrical continuity passed. No other damage was found. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 968804 | 962817 | 924400 | 2955842200700512 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-15 | 2007-12-10 | No | Yes | Health Professional | 400180 | No | No | 2007-11-16 | Device Not Retu | Yes | Initial | 2007
40018005 | 2007-11-09 | Malfunction | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was dissecting nodes, a burn on the pt's bowel was noticed. Upon further inspection, the surgeon found a crack on the tip cover accessory of the monopolar curved scissors instrument. The instrument was removed and the tip cover was replaced. The bowel was repaired by over sewing the burn and the procedure was completed with an approx delay of one hr. No add'l pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. #2955842-2007-00494. | The instrument tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has a. 010 inch hole burned through the silicone with localized melting. Examination of the hole under magnification shows the hole is located roughly. 240 inches above the plastic sleeve to silicone interface. The melting of the silicone suggests an arching event occurred. The hole appears to have originally been cut, as there is a cut bisecting one side of the hole. There is a mechanical tear. 060 inches above the hole on the same side and a. 140 inch long tear 180 degrees from the hole, roughly. 300 inches above the silicone to sleeve interface. Damage to silicone is likely due to instrument collisions. Engineering concluded that a cut in the silicone may have created a path for arcing. The plastic sleeve is cracked along the longitudinal axis, however, this may have occurred during return shipping. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT | 968172 | 962818 | 924401 | 2955842200700495 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-09 | 2007-12-10 | Yes | Yes | Health Professional | 40018005 | Device Returned To Manufacturer | No | 2007-11-12 | Yes | Yes | Initial | 2007
40018005 | 2007-11-14 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure arcing was observed through the tip cover accessory of the monopolar curved scissors instrument. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below were also sent to the fda as they relate to this event. #2955842-2007-00502. | The instrument tip cover accessory has not been returned for eval. A follow-up mdr will be submitted if the tip cover accessory is returned and failure analysis has been completed. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 968145 | 962819 | 924402 | 2955842200700503 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-13 | 2007-12-10 | No | Yes | Health Professional | 40018005 | No | No | 2007-11-14 | Device Not Retu | Yes | Initial | 2007
42017204 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument shaft is dented. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the main tube does not have anything characteristic to a "dent", however, engineering also observed a section 3 inches above the proximal clevis with material removed all around the tube. Damaged area is 3 inches long. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968271 | 962820 | 924403 | 2955842200700500 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-14 | 2007-12-10 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-11-14 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument has broken wires. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed both grip cables attached to one grip are broken. The grip close cable is broken at the distal idlers and the grip open cable is broken before the distal idlers on the cable groove. It appears that the cables broken under tensile loading. Engineering also observed the distal end of the main tube had deep scratches on one side with a minimal amount of material removed. The location and appearance of the scratches suggests the damage may be due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 968263 | 962821 | 924404 | 2955842200700506 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-15 | 2007-12-10 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-11-15 | Yes | No | Initial | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is not alligned. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension dislodged from the main tube. The entire extension wall is broken at the base and as a result, the tube extension can be rotated 360 degrees. Damage may be due to excessive side loading. Engineering also observed the main tube has a 2 inch long section just below the midpoint with material removed all around the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. The dislodged tube extension may also have contributed to the main tube damage, as instrument would have been more difficult to remove from the cannula due to breakage. No other damage was found. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968258 | 962822 | 924405 | 2955842200700493 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial |  | 2007-12-10 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-11-12 | Yes | No | Unkown | 2007
40017208 |  | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, the maryland bipolar forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip is bent, causing misalignment of grips. Bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. There is also an offset at the tips. Engineering also observed the distal end of main tube has 2. 5 inch long section located 1 inch above the proximal clevis with material removed on one side of the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 967651 | 962830 | 924413 | 2955842200700496 | 1 | GEI | Manufacturer | ForeignUser facility | Other | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 40017208 | Yes | No | 2007-11-13 | Yes | No | Reuse | 2007
42017908 |  | Malfunction | Other | It was reported that a spark at the tip of the monopolar curved scissors instrument occurred, and a hole in the tip cover accessory was noticed. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. # 2955842-2007-00512. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument has a. 060 inch by. 120 inch black char mark on the proximal clevis. There are two small char spots roughly. 060 inches above the large char mark. Engineering concluded this evidence suggests an arcing event occurred. The tip cover accessory was not returned for evaluation, however, assuming the tip cover was installed, the position of the large char mark lines up with the silicone to sleeve interface of the tip cover accessory. The tip cover may have had a cut at the silicone to sleeve interface that created a path for arcing. Electrical continuity passed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 967637 | 962831 | 924414 | 2955842200700513 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-15 | 2007-12-10 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-11-16 | Yes | No | Reuse | 2007
420179 | 2007-11-09 | Malfunction | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was dissecting nodes, a burn on the patient's bowel was noticed. Upon further inspection, the surgeon found a crack on the tip cover accessory of the monopolar curved scissors instrument. The instrument was removed and the tip cover was replaced. The bowel was repaired by over sewing the burn and the procedure was completed with an approximate delay of one hour. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it relates to this event. # 2955842-2007-00495. | The instrument has not been returned for evaluation. A follow-up mdr will be submitted if the monopolar curved scissors instrument is returned and failure analysis has been completed. The monopolar curved scissors tip cover accessory used in conjunction with the instrument was returned and evaluated. Engineering evaluation of the tip cover accessory found evidence that arcing had occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 967664 | 962832 | 924415 | 2955842200700494 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-09 | 2007-12-10 | Yes | Yes | Health Professional | 420179 | No | No | 2007-11-12 | Device Not Retu | No | Unkown | 2007
42000604 | 2007-11-14 | Malfunction |  | It was reported that after cleaning, it was noticed that shavings came off the shaft of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube has a 1 inch long,. 125 inch wide section with light material removed on one side of the tube. Damage occurred 5 inches above the proximal clevis and the wear pattern is consistent with damage from a defective cannula. No other damage was found. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 967668 | 962834 | 924417 | 2955842200700508 | 1 | GCJ | Manufacturer | Health Professional | Nurse | Initial | 2007-12-13 | 2007-12-10 | No | Yes | Health Professional | 42000604 | Yes | Yes | 2007-11-15 | Yes | No | Initial | 2007
42017906C | 2007-11-14 | Malfunction |  | It was reported that during a da vinci bladder augmentation procedure, the surgical staff noticed a "gear" had fallen off of the monopolar curved scissors instrument and fell into the patient's bladder. The fragment was immediately retrieved. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. If the instrument is received, failure analysis will be completed and a follow-up mdr will be submitted. Per the case notes, the broken instrument component was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 967669 | 962835 | 924418 | 2955842200700504 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-14 | 2007-12-10 | No | Yes | Health Professional | 42017906C | No | No | 2007-11-14 | Device Not Retu | No | Unkown | 2007
42021403 |  | Malfunction |  | It was reported that the tip of the pk dissecting forceps instrument became bent. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tip is not bent. There is a small offset at the tips of the grips, however, this is due to a broken section of yaw pulley. That allows movement of the grip within the yaw pulley. One conductor wire is broken at the yaw pulley exit. The yaw pulley shows no signs of arcing. The grip cables are not derailed at the wrist. The wire insulation does not appear cut or melted. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that added range of motion of the grip likely created excess tension on the wire, causing it to break. The proximal clevis is missing a. 160" x. 250" at the main tube interface, between the ears, but the clevis is not dislodged. Electrical continuity failed. Engineering also observed the main tube to have a section with material removed on one side of the tube. The affected area is 4" long, parallel to the tube axis, with a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968275 | 962836 | 924419 | 2955842200700509 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 42021403 | Yes | No | 2007-11-16 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | The pk dissecting forceps instrument was returned without information. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed one conductor wire to be broken at the yaw pulley exit. The yaw pulley shows no signs of arcing. The grip cables are not derailed at the wrist. The wire insulation does not appear cut or melted. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. The added range of motion of the grip likely created excess tension on the wire, causing it to break. Electrical continuity failed. Engineering also observed the distal end of main tube has a 2. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish, with glass fibers exposed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | PK DISSECTING FORCEP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968278 | 962837 | 924420 | 2955842200700497 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-11-13 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that the cautery connector of the pk dissecting forceps instrument was broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire to be broken at the yaw pulley exit. The yaw pulley has a burn mark directly above the break site. The other yaw pulley has a burn mark on the same side as the conductor break. The wire insulation has localized charring/melting at break site. The broken wire may have caused burn marks on the yaw pulleys. The wire may have broken due to catching on components during the wrist articulation. Electrical continuity failed. No other damage found. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 966862 | 962838 | 924421 | 2955842200700499 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-11-14 | Yes | No | Reuse | 2007
42018304 |  | Malfunction |  | An isi representative evaluated the site's da vinci s 8mm instruments to determine if the 8mm cannula accessories may have caused instrument tube abrasions similar to medwatch mfr report # 2955842-2007-00438. | The permanent cautery hook instrument was returned and evaluated. Engineering found the distal end of the main tube to have various scratch marks. A couple of the scratches have removed material from the main tube. All scratches are found on one side of the tube, but in different locations. The scratches are not aligned with the tube axis and are likely due to inadvertent contact with other instruments and not a defective cannula. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 966901 | 962839 | 924422 | 2955842200700510 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 42018304 | Yes | No | 2007-11-16 | Yes | No | Unkown | 2007
42018404 | 2007-11-16 | Malfunction | Other | It was reported that while transecting the bladder during a da vinci prostatectomy procedure, the wrist of the permanent cautery spatula instrument began smoking. The instrument was replaced with an instrument of the same type to complete the procedure. No add'l pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed that the instrument has localized melting/charring of the plastic idler pulley. The cable is derailed at the same idler pulley. Additionally, the conductor wire has damaged insulation with bare wire exposed indicating arcing occurred. There is also a char mark on the yaw pulley above the damaged wire. The wire may have been damaged by another instrument. Idler pulley has chipped/bent edges away from the melted section that may have come from another instrument hitting the pulley. Cable derailment likely occurred after initial conductor damage, and wire arced to the cable and surrounding areas of the idler pulley and yaw pulley when cautery was activated. The ceramic sleeve has hairline cracks; however, this is likely unrelated to the arcing. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968182 | 963342 | 924865 | 2955842200700511 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-12-10 | 2007-12-10 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2007-11-16 | Yes | No | Initial | 2007
42000604 | 2007-11-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the large needle driver instrument "tip cracked all the way around". As a result the surgical staff had difficulty removing the instrument. The planned surgical procedure was completed with an instrument of the same type with approx a 30 minute delay. The site indicated that no instrument fragments of the broken tip fell into the pt. No pt harm, adverse outcome, or injury was reported. | The instrument has not been returned for eval. If the instrument is received, failure analysis will be completed and a follow-up mdr will be submitted. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 970908 | 964595 | 924417 | 2955842200700507 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-14 | 2007-12-10 | Yes | Yes | Health Professional | 42000604 | No | No | 2007-11-15 | Device Not Retu | No | Unkown | 2007
IS2000 A51P3 | 2007-11-13 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, approx 20 minutes into the procedure, the pt side cart (psc) switched to battery power. The procedure was completed with the site's spare da vinci s surgical system, causing approx a 40 minute delay. No pt harm, adverse outcome, or injury was reported. | The investigation conducted by isi field service engineering concluded that the pt side cart (psc) power supply was defective. The system was repaired by replacing the psc power supply. Field service verified system operations and completed field verification tests before returning the system to the customer. As of late 2007, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 970903 | 964596 | 927018 | 2955842200700498 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-12 | 2007-12-10 | Yes | Yes | Health Professional | IS2000 A51P3 | Yes | No | 2007-11-13 | Device Not Retu | No | Reuse | 2007
IS2000 A51P4 | 2007-10-19 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the customer experienced system error code #23 which could not be overridden. The surgical staff restarted the system, however, system error code #23 recurred. No patient harm, adverse outcome, or injury was reported. The procedure was completed with the site's spare da vinci surgical system, causing approximately a one hour delay. | The investigation conducted by isi field service engineering concluded that system error code #23 experienced by the customer was associated with the cable connecting the patient side cart (psc) and surgeon side console (ssc). The system was repaired by cleaning the fiber optic connectors on the cable. System powered on homed without any errors or faults. System fault #23 appears when the davinci s safety system defects an internal system communication error. In this case the davinci s safety system determined that the psc and ssc were not communicating properly. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put davinci s in a "non-recoverable safe state". As of 2007, there have been no reported recurrences of the issue at this hospital. The da vinci s surgical system instructions for use (ifu) specifically states: remove the cables' protective caps, inspect the cable connectors and system connectors for debris or bent pins. | D VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 966794 | 959775 | 921529 | 2955842200700492 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2007-12-04 | 2007-12-04 | Yes | Yes | Health Professional | IS2000 A51P4 | No | Yes | 2007-10-19 | Device Not Retu | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering completed performance testing and observed that the scissors do not cut cleanly through. 006" latex. The latex gets pinched at the scissor tips. One blade edge has a deformation very near the tip that prevents the blades from fully closing. The damage to the blade may be due to mishandling/misuse. Electrical continuity passed. Engineering also observed the distal end of the main tube to have a 2" long section located 3" above the tube extension with material removed all around the tube. The glass fibers are exposed, making the surface finish rough in the damaged area. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965406 | 957477 | 919298 | 2955842200700483 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-03 | 2007-12-03 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2007-11-07 | Yes | No | Reuse | 2007
42019004 |  | Malfunction |  | The cobra grasper instrument was returned with no info provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with a full range of motion in all directions. The distal end of main tube has 4 long, narrow sections with material removed. Three of the sections are concentrated on one side of the tube. The width of scrapes and spacing suggest separate occurrences of scraping. There are also various scratch marks in the same general area of the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collisions. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 966885 | 957478 | 919299 | 2955842200700488 | 1 | GCJ | Manufacturer | Unknown | Physician | Initial | 2007-12-03 | 2007-12-03 | No | Yes | Health Professional | 42019004 | Yes | Yes |  | Yes | No | Reuse | 2007
42018303 |  | Malfunction |  | It was reported that an axis of the permanent cautery hook instrument axis is not moving. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch input rotates but movement is not translated to the instrument wrist due to a broken pitch cable at the distal clevis. Engineering also observed the main tube is damaged on the same side as cable breakage where the distal end has scratches perpendicular to the tube axis with a small amount of material removed. Engineering concluded that the scratches are likely due to inadvertent contact with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965414 | 957479 | 919300 | 2955842200700480 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2007-12-02 | 2007-12-03 | No | Yes | Health Professional | 42018303 | Device Returned To Manufacturer | No | 2007-11-02 | Yes | No | Reuse | 2007
42018303 |  | Malfunction |  | It was reported that the permanent cautery hook instrument has a broken cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed half the pitch down cable is broken at the proximal clevis cable hole and the remaining cable breaks within the proximal clevis. The conductor wire insulation adjacent to the broken cable is cut, exposing bare wire. The proximal clevis has localized melting and charring near the conductor wire damage, indicating that the instrument may have arced during a surgical procedure, the insulation damage most likely created a path for arcing. Engineering concluded that the insulation damage may have been an instrument collision. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The proximal clevis is missing a. 270 inch by. 150 inch piece on the side opposite to the melting. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965421 | 957480 | 919301 | 2955842200700475 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2007-12-02 | 2007-12-03 | No | Yes | Health Professional | 42018303 | Device Returned To Manufacturer | No | 2007-11-02 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument moved intuitively with full range of motion and all components functioned per specification. Engineering also observed the distal end of the main tube has a section with material removed all around the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 965431 | 957481 | 919302 | 2955842200700484 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-12-07 | 2007-12-03 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2007-11-07 | Yes | No | Reuse | 2007
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument jaws do not approximate. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the teeth on the grips do not mesh properly. One grip is approximately. 005 inches longer than the other, resulting in an offset tooth pitch. Engineering also observed the distal end of the main tube, directly above the proximal clevis, has material removed around 3/4 of the outer diameter. The location and appearance of the scratches suggests the damage may be due to instrument collisions and cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965424 | 957482 | 919303 | 2955842200700486 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-07 | 2007-12-03 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2007-11-07 | Yes | No | Reuse | 2007
40000619 | 2007-11-07 | Malfunction |  | It was reported that during a davinci prostatectomy procedure, the surgical staff noticed the large needle driver instrument jaw was broken and that a fragment had fallen into the patient. The fragment was retrieved. No patient harm was reported. Additionally no post operative complications have been reported by this account. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one entire jaw of the instrument had broken off, the broken piece was not returned. Jaw is broken off at the thin wall which may have had a void thus reducing load bearing capacity. The backside of the grip has a gouge, likely caused from instrument collisions, which may have contributed to the grip failure. Per the case notes, the broken instrument component was successfully retrieved from the patient. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 956821 | 957483 | 919304 | 2955842200700490 | 1 | GCJ | Manufacturer | ForeignDistributor | Other | Initial | 2007-12-06 | 2007-12-03 | No | Yes | Health Professional | 40000619 | Device Returned To Manufacturer | No | 2007-11-07 | Yes | No | Unkown | 2007
42017205 | 2007-11-06 | Malfunction |  | It was reported that during the middle of a da vinci prostatectomy surgical procedure the surgical staff noticed the instrument proximal clevis had fallen into the patient. The broken component was found and removed. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface with pieces missing. The proximal clevis is dislodged from the main tube as a result of the breakage. Both the proximal clevis and the main tube are missing pieces. A. 200 inch by. 220 inch piece of the clevis was returned as well as a larger piece of the proximal clevis, together these pieces cover the entire missing section. The fractured plane of a broken proximal clevis piece indicates that breakage is likely due to overloading at the tip with a pitched wrist. Per the case notes, the broken instrument component was successfully retrieved from the patient. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 956836 | 957484 | 919305 | 2955842200700482 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-12-05 | 2007-12-03 | No | Yes | Health Professional | 42017205 | Device Returned To Manufacturer | No | 2007-11-06 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that during an inspection prior to a procedure the endowrist prograsp instrument was found to have burrs on it. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube to have a 2" long,. 200" wide section directly above the proximal clevis with material removed. The damaged area is parallel to tube's longitudinal axis. Additionally, within the damaged area there is a single deep scratch. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collision. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. The endowrist instruments for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 966961 | 957485 | 919306 | 2955842200700485 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-12-03 | 2007-12-03 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-11-06 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-11-08 | Other | Other | It was reported that during the da vinci procedure the system endoscope became foggy. The procedure was converted to traditional open surgical techniques to finish the planned surgery prior to contacting isi for troubleshooting. No patient harm was reported. | The investigation conducted by isi field service engineering via telephone call with the site after the procedure discovered that condensation on the camera head lens was the cause of the fogging experienced by the customer. The system was repaired by cleaning the camera head lens. The customer converted the procedure to a traditional open technique before calling technical support. Additional investigation performed by an isi product manager concluded that the site used the endoscope, while it was still wet after sterilization. Once the wet endoscope was connected to the camera, moisture condensed on the camera lens causing the endoscope to appear foggy. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 966837 | 959791 | 921545 | 2955842200700489 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-12-03 | 2007-12-03 | Yes | Yes | Health Professional | IS1200 A43P8 | No | No | 2007-11-08 | Yes | No | Reuse | 2007
40000120 | 2007-11-09 | Other | Other | It was reported that during a da vinci ureteral reimplantation procedure, the potts scissors instrument became stuck in the sterile adapter accessory, and the surgeon was unable to remove the instrument. As a result the surgical staff removed the instrument and trocar concurrently from the patient. The planned surgical procedure was completed with approximately a 35 minute delay. No patient harm, adverse outcome, or injury was reported. | The instrument and sterile adapter accessory were returned and evaluated. Per the customer reported complaint, engineering installed and removed the instrument on the sterile adapter accessory multiple times on a system. No difficulties were experienced when removing the instrument with the wrist straightened and the blades open before removal. The instrument moved intuitively with full range of motion. Engineering concluded that the instrument wrist was most likely articulated when the blades were closed, causing the instrument to be difficult to remove due to high torque on the input discs and high cable tension. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings always straighten the instrument wrist under endoscopic visualization prior to removal of the instrument through the cannula. | POTTS SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 966985 | 959795 | 921549 | 2955842200700491 | 1 | LRW | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-12-08 | 2007-12-03 | Yes | Yes | Health Professional | 40000120 | Yes | No | 2007-11-09 | Yes | No | Reuse | 2007
42011003 | 2007-11-07 | Other |  | It was reported that during a da vinci hysterectomy procedure, during the vaginal cuff process, the precise bipolar forceps instrument proximal clevis broke and pieces fell into the patient. The surgical staff attempted to remove the instrument, but had difficulty removing the instrument from the cannula. As a result, the surgical staff had to remove the instrument, cannula and sterile adapter concurrently. The fragmented components were found and removed from the patient's anatomy. The procedure was completed with another instrument of the same type and without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface, and as a result, the clevis is dislodged from the main tube. Both proximal clevis and main tube are missing pieces. A. 175" x. 240" piece of the clevis was returned with the instrument, however, the clevis is still missing another section of a slightly smaller size. One of the fracture surfaces of the broken clevis piece is located halfway between the ears, suggesting the instrument was overloaded with the wrist pitched. One of the grip cables had derailed at the distal idlers and one conductor wire is broken at the yaw pulley exit, likely after breakage of proximal clevis. The main tube has various cosmetic scratches (no material removed) perpendicular to the tube axis that cover the entire tube length. No other damage found. Per the case notes, the broken instrument component was successfully retrieved from the patient. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 966996 | 959797 | 921551 | 2955842200700487 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-12-07 | 2007-12-03 | No | Yes | Health Professional | 42011003 | Yes | No | 2007-11-07 | Yes | No | Reuse | 2007
| 2007-11-19 | Malfunction |  | A small piece of plastic was found inside the patient during a robotic laparoscopic procedure. The physician went back in to determine if any additional pieces could be found. No additional pieces were found. It was determined that the piece of plastic that was initially found was from the permanent cautery hook. It appears to be similiar to a washer that surrounded a pin at the end of the cautery. Manufacturer response for permanent cautery hook 8 mm, da vinci: please send it back to the factory so we can evaluate it. |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 934579 | 964894 | 926335 | 964894 | 1 | GEI | User Facility |  | BIOMEDICAL ENGINEER | Initial | 2007-11-29 | 2007-11-29 | No | Yes | Physician |  | Device Returned To Manufacturer | No |  |  |  |  | 2007
40017104 | 2007-10-29 | Other | Other | It was reported that during a da vinci coronary artery bypass graft surgical procedure, while performing the lima take down, a nurse noticed the "gear" of the instrument in the patient's abdomen. The surgeon removed the instrument component and completed the procedure with another instrument of the same type. The planned surgical procedure was delayed approx 30 minutes. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is missing both distal idler pulleys (which the customer referred to as the "gear") on one side of the distal clevis, only the outer idler pulley was returned. Further evaluation determined that the distal clevis idler pin was improperly swaged and as a result, the pulley may have had room to slip out of the pin. Engineering also observed a broken conductor wire. No other damage was found. Per the case notes, the broken instrument component was successfully retrieved from the pt. | MICROBIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 953915 | 955424 | 917287 | 2955842200700467 | 1 | GEI | Manufacturer | Distributor | UNKNOWN | Initial | 2007-11-28 | 2007-11-28 | Yes | Yes | Health Professional | 40017104 | Device Returned To Manufacturer | No | 2007-10-29 | Yes | No | Reuse | 2007
31146505 | 2007-09-10 | Injury | Other | It was reported that during a da vinci nissen fundoplication surgical procedure, "the pt was burned from the heat emitted from the endoscope. The tip of the 12mm 30 degree endoscope heated up enough to cause thermal injury to this pt's liver. " the light source (illuminator) setting was at 100%. The site indicates that they, "have since run the light source at 50%, and have not had any issues with inadequate illumination, fogging of the endoscope, or thermal injury. " was received at isi on 10/10/07, regarding this incident. | The endoscope was not returned for eval, however, isi representatives conducted an investigation at the hosp on 10/18/07. Currently, the da vinci s surgical system user manual specifically states: note: to minimize fogging, maintain heating of the endoscope tip by setting the illuminator brightness to 100% and using the level or iris knobs on the ccus to adjust the brightness of the surgical image. Caution: if the endoscope is working close to tissue, reduce light source output by turning the aperture knob on the illuminator in a counter-clockwise direction pressing the down arrow on the illuminator. This will prevent tissue damage due to excessive heat. Isi contacted the hosp via letter and clarified that the light at the tip of the endoscope can be intense enough to cause tissue damage and that a safe working distance needs to be maintained between the endoscope tip and tissue. Isi also clarified that when working close to tissue with the endoscope tip, the illuminator brightness may need to be set to a lower level to reduce tip heating; however this can result in tip fogging. As of 11/27/07, there have been no reported recurrences of the issue at this hosp. | 30 DEGREE SCHOELLY STEREO ENDOSCOPE | ENDOSCOPE | 895395 | 961602 | 885342 | 2955842200700474 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-11-27 | 2007-11-28 | Yes | Yes | Health Professional | 31146505 | No | Yes | 2007-09-26 | No | No | Unkown | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument has a bent tip. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tip was not bent. All components were intact and functioned properly. Engineering also observed the distal end of the main tube to have a 1 inch section directly above the proximal clevis, with various deep scratches concentrated on one side of the tube. A minimal amount of material was removed. The location and appearance of the scratches suggests the damage may be due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 951474 | 952912 | 914883 | 2955842200700469 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-29 | 2007-11-21 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-10-30 | Yes | No | Reuse | 2007
400093 | 2007-10-24 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure while the surgeon was cauterizing, an arc of energy from the monopolar curved scissors instrument came through its tip cover accessory striking the endowrist prograsp instrument. Additionally the surgical staff noticed the endowrist prograsp instrument was cracked at the proximal clevis, making it difficult to remove from the cannula. While attempting to remove the prograsp instrument a fragment fell into the pt and was immediately removed, causing approximately a 15 minute delay. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below were also sent to the fda as they related to this event. #2955842-2007-00477 and #2955842-2007-00478. | The instrument tip cover accessory has not been returned for eval. A follow-up mdr will be submitted if the tip cover accessory is returned and failure analysis has been completed. Per the case notes, the broken instrument component was successfully retrieved from the pt. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 965201 | 952913 | 914884 | 2955842200700479 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-11-21 | 2007-11-21 | No | Yes | Health Professional | 400093 | No | No | 2007-10-24 | Device Not Retu | No | Unkown | 2007
400179 | 2007-10-24 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure while the surgeon was cauterizing, an arc of energy from the monopolar curved scissors instrument came through its tip cover accessory striking the endowrist prograsp instrument. Additionally the surgical staff noticed the endowrist prograsp instrument was cracked at the proximal clevis, making it difficult to remove from the cannula. While attempting to remove the prograsp instrument a fragment fell into the pt and was immediately removed, causing approximately a 15 minute delay. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below were also sent to the fda as they related to this event. #2955842-2007-00477 and #2955842-2007-00479. | The instrument tip cover accessory has not been returned for eval. A follow-up mdr will be submitted if the tip cover accessory is returned and failure analysis has been completed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965225 | 952915 | 914886 | 2955842200700478 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-11-21 | 2007-11-21 | No | Yes | Health Professional | 400179 | No | No | 2007-10-24 | Device Not Retu | No | Unkown | 2007
420179 | 2007-10-29 | Malfunction | Other | It was reported that during nucleation of the myoma during a da vinci myomectomy procedure, small tears were noticed in the monopolar curved scissors instrument tip cover accessory. The tip cover had been in use for approximately 2 hrs when the tears were observed. The tip cover was removed and replaced the planned surgical procedure was successfully completed without significant delay. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00466. | The instrument was not returned for evaluation. Engineering observed multiple signs of arcing and mechanical tearing on the tip cover accessory used in conjunction with the instrument, which may have caused a path for arcing to occur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 965185 | 952916 | 914303 | 2955842200700465 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-28 | 2007-11-21 | No | No | Health Professional | 420179 | No | No | 2007-10-29 | Device Not Retu | No | Unkown | 2007
42017908 |  | Malfunction |  | It was reported the monopolar curved scissors instrument was dull. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the instrument was able to cut per specification. No damage to the blades was observed and electrical continuity passed. The customer reported failure mode cannot be confirmed. Engineering also observed the main tube has 2 inch long section of wearing on one side of the tube. The section is located 3. 35 inches above the tube extension and a minimal amount of material has been removed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional info is rec'd. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 951442 | 952917 | 914888 | 2955842200700463 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-11-29 | 2007-11-21 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2007-10-29 | Yes | No | Reuse | 2007
400180 | 2007-10-24 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure while the surgeon was cauterizing, an arc of energy from the monopolar curved scissors instrument came through its tip cover accessory striking the endowrist prograsp instrument. Additionally, the surgical staff noticed the endowrist prograsp instrument was cracked at the proximal clevis, making it difficult to remove from the cannula. While attempting to remove the prograsp instrument a fragment fell into the pt and was immediately removed, causing approximately a 15 minute delay. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below were also sent to the fda as they related to this event. #2955842-2007-00478 and #2955842-2007-00479. | The instrument tip cover accessory has not been returned for eval. A follow-up mdr will be submitted if the tip cover accessory is returned and failure analysis has been completed. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 965218 | 952918 | 914889 | 2955842200700477 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-11-21 | 2007-11-21 | No | Yes | Health Professional | 400180 | No | No | 2007-10-24 | Device Not Retu | Yes | Initial | 2007
40018005 | 2007-10-29 | Malfunction | Other | It was reported that during nucleation of the myoma during a da vinci myomectomy procedure, small tears were noticed in the monopolar curved scissors instrument tip cover accessory. The tip cover and had been in use for approximately 2 hrs when the tears were observed. The tip cover was removed and replaced and the planned surgical procedure was successfully completed without significant delay. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00465. | The instrument tip cover accessory was returned and evaluated. Per the reported complaint, engineering observed multiple signs of arcing and mechanical tearing. The plastic sleeve has charring and a. 020 inch hole at the silicone sleeve interface. It appears that the hole is a result of arcing, as localized melting of silicone is exhibited. In addition, there is a. 040 inch long mechanical tear adjacent to this hole. There is a slight separation of silicone adjacent to the hole, making the area more susceptible to arcing. Also observed, is a section of silicone with 3-4 holes ranging in size from. 005 inch to. 015 inch that are all connected by a common t-shaped mechanical tear. There is one single. 020 inch hole located. 215 inch above the sleeve with clear evidence of silicone melting and no adjacent tearing. The various holes and tears would suggest multiple arcing events occurred, with tearing likely preceding the arcing. Tearing of silicone, possibly from contact with other instruments or with the tip of the cannula may have created a path for arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 965157 | 952920 | 914304 | 2955842200700466 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-28 | 2007-11-21 | No | No | Health Professional | 40018005 | Yes | No | 2007-10-29 | Yes | Yes | Initial | 2007
42004904 |  | Malfunction |  | It was reported that the tip of the cadiere forceps instrument did not operate properly. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to be broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result. The clevis is intact, however, the main tube interface wall has collapsed inwards, likely due to overloading at the tip. Engineering also observed the distal end of the main tube to have material removed. There are scratch marks not aligned with the tube axis located directly above the proximal clevis. These marks are likely due to instrument collisions. Directly above the scratched area there is a 2. 75" long section with material scraped off, exposing the glass fibers. A 1. 75" long,. 070" wide streak with glass fibers exposed is located 180 degrees from the other scraped section. The tube damage may have occurred when attempting to remove the instrument from the cannula. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 952195 | 952922 | 914893 | 2955842200700476 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-21 | 2007-11-21 | No | Yes | Health Professional | 42004904 | Yes | No | 2007-10-22 | Yes | No | Reuse | 2007
| 2007-10-31 | Malfunction |  | The robotic forceps was placed inside the patient during the procedure, and the end of the forceps broke off into the patient. The one piece was retrieved and the forceps removed. The trocar remained on the forceps and was unable to remove the trocar from the forceps. While attempting to remove the trocar, that was needed for the remainder of the case, two other pieces broke off. |  | DA VINCI CADIERE FORCEPS | FORCEPS ROBOTIC | 949979 | 982442 | 941716 | 982442 | 1 | GCJ | User Facility |  |  | Initial | 2007-11-19 | 2007-11-21 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2007
40017206 | 2007-10-05 | No Answer Provided | Other | It was reported that during a davinci prostatectomy procedure, a "wire broke" on the maryland bipolar forceps instrument. The procedure was successfully completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire to be broken at the yaw pulley exit. The yaw pulley shows no signs of arcing. The grip cables are not derailed at the wrist. The wire insulation does not appear cut or melted. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Added range of motion of the grip likely created excess tension on the wire, causing it to break. Electrical continuity failed. Engineering also observed the distal end of the main tube to have various light scratch marks on one side of the tube, likely due to inadvertent contact with other instruments. A minimal amount of material has been removed from a couple of the scratches which have gone deep enough to expose glass fibers. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. As of november 19, 2007, there have been no reported recurrences of the issue at this hospital. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 966091 | 952198 | 914202 | 2955842200700458 | 1 | GEI | Manufacturer | ForeignUser facility | Other | Initial | 2007-11-19 | 2007-11-19 | No | No | Health Professional | 40017206 | Device Returned To Manufacturer | No | 2007-10-25 | Device Not Retu | No | Unkown | 2007
40009307 | 2007-10-16 | Malfunction |  | It was reported that during the da vinci prostatectomy procedure, the proximal clevis of the endowrist prograsp instrument broke into pieces which fell into the patient. The pieces were retrieved and the procedure was completed using a replacement instrument of the same type. It was also reported that the cannula was stuck on the instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. If the instrument is received, failure analysis will be completed and a follow-up mdr will be submitted. Per the case notes, the broken instrument component was successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 947020 | 952301 | 914302 | 2955842200700464 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 40009307 | No | Yes | 2007-10-26 | Device Not Retu | No | Unkown | 2007
420179 | 2007-10-30 | Malfunction |  | It was reported that a representative returned a monopolar curved scissors instrument tip cover accessory used during a davinci s surgical procedure at the site for evaluation. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00473. | The instrument was not returned for evaluation. Engineering observed a sign of arcing and a mechanical tear on the tip cover accessory used in conjunction with the instrument, which may have caused a path for arcing to occur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947013 | 952302 | 914303 | 2955842200700472 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 420179 | No | No | 2007-10-29 | Device Not Retu | No | Invalid Data | 2007
40018005 | 2007-10-30 | Malfunction |  | It was reported that a representative returned a monopolar curved scissors instrument tip cover accessory used during a davinci s surgical procedure at the site for evaluation. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00472. | During evaluation, engineering observed the tip cover to have a. 025" diameter hole exhibiting localized melting of the silicone. The hole is located 30 degrees radially from one silicone parting line and. 420" above the silicone to sleeve interface. There is a. 050" long mechanical tear at the distal tip of the silicone. The silicone may have had a small nick on either the inner or outer surfaces of the silicone at the hole site. Evidence of the nick would be destroyed by the localized melting of the silicone and the creation of the hole. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 947002 | 952303 | 914304 | 2955842200700473 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2007-10-30 | Yes | Yes | Initial | 2007
42017206 |  | Malfunction |  | It was reported that the end of the shaft of the maryland bipolar forceps instrument had cracked. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the end of the shaft not to be cracked. This instrument has the metal proximal clevis and thicker main tube interface wall. Engineering also observed the distal end of the main tube to have various scratch marks concentrated in a 2" long section on one side of the tube. The scratches are not aligned with the tube axis and may be caused by inadvertent contact with other instruments. A few of the scratches are deep enough to remove a minimal amount of material from the main tube. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947182 | 952305 | 914306 | 2955842200700461 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-10-27 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument was scratched. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed the distal end of the main tube to have a section 1. 5" above the proximal clevis with glass fibres exposed and material removed on one side of the tube. The damaged area is roughly 2" long x. 200" wide, aligned with the tube axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 946984 | 952306 | 914307 | 2955842200700456 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-10-22 | Yes | No | Reuse | 2007
42011004 |  | Malfunction |  | It was reported that the precise bipolar pyramid tip forceps instrument has a broken wire. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip cable frayed at the yaw pulley. Engineering also observed the distal end of the main tube has various scratch marks. There is also a 3 inch long section 2. 5 inches above the prox clevis with a small amount of material removed. The appearance of the damage suggests that the scratches may have been caused by inadvertent contact with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947149 | 952307 | 914308 | 2955842200700452 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42011004 | Yes | No | 2007-10-20 | Yes | No | Reuse | 2007
42017907 |  | Malfunction |  | It was reported that the jaws of the monopolar curved scissors instrument did not open or close. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the jaws to open and close fine when inputs are manually rotated. The tube extension is dislodged from the main tube. The keys on the tube extension have broken off. As a result, the tube extension can be rotated 360 degrees. Engineering concluded that the tube extension may have dislodged due to excessive side loading. Engineering also observed the distal end of the main tube to have two distinct sections with material removed. One section is. 900" long, directly above the tube extension, and has scratch marks with scraping superimposed on each other. Roughly 2" above the tube extension and 90 degrees from the first damaged section is a 1. 75" long section with light wearing on one side of the tube. A minimal amount of material is removed from this section. Based on the location and appearance, the damage may have been caused by a cannula accessory and instrument collisions. Add'l investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if add'l info is received. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result. ". | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 949706 | 952308 | 914309 | 2955842200700462 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42017907 | Yes | No | 2007-10-29 | Yes | No | Reuse | 2007
42018404 | 2007-10-26 | Malfunction |  | It was reported that during a da vinci s procedure a part of the jaw of the permanent cautery spatula instrument broke off and fell into the patient. The procedure was successfully completed. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the spatula portion of the tip to have broken off. The missing piece is roughly. 125" x. 090". The location of breakage appears to be where the formed spatula transitions into a flattened spatula shape. This is likely a stress concentration area. The break plane has a slight curvature towards one side, suggesting overloading occurred on the side of the spatula opposite the curvature. The ceramic sleeve has a scratch on the same side where the overloading may have occurred. Engineering concluded that it is possible the spatula was overloaded as a result of a severe instrument collision. Electrical continuity passed. No other damage found. The site did not indicate whether the fragmented component was retrieved from the patient. Attempts have been made to obtain add'l info without success. A follow-up mdr will be submitted if add'l info is received. | PERMANENT CAUTERY SPATUAL INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 949710 | 952309 | 914310 | 2955842200700460 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42018404 | Yes | No | 2007-10-26 | Yes | No | Reuse | 2007
40021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a missing plug. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the cautery connector plug was not missing as reported. Engineering also observed that the main tube has various light scratch marks and light scraping, with a minimal amount of material removed. The appearance of the damage suggests that the scratches and scrapes may have been caused by inadvertent contact with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947138 | 952310 | 914311 | 2955842200700459 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-11-21 | 2007-11-19 | No | Yes | Health Professional | 40021403 | Yes | No | 2007-10-25 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument was broken at the wrist. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed all components at the wrist are intact. All components functioned properly. Engineering also observed the distal end of the main tube has a 1. 25 inch long section, 0. 5 inches above the proximal clevis, with various deep scratches concentrated on one side of the tube with minimal material removed. Scratches are not aligned with tube axis, suggesting they may have been caused by inadvertent contact with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 947170 | 952311 | 914312 | 2955842200700451 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-10-20 | Yes | No | Reuse | 2007
42017206 | 2007-10-18 | Malfunction |  | On october 18, 2007, isi representatives evaluated the site's da vinci s 8mm instruments to determine if the 8mm cannula accessories may have caused instrument tube abrasions similar to medwatch mfr report #2955842-2007-00438. | Considerable visible scratches were found on the maryland bipolar forceps instrument main tube at the distal end near the metal wrist section. The main tube scratches are consistent with collisions between two instruments inside the body and may have been created by instrument collisions over multiple cases. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947156 | 952312 | 914313 | 2955842200700470 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-09 | 2007-11-19 | No | Yes | Health Professional | 42017206 | No | No | 2007-10-18 | Yes | No | Unkown | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the instrument moved intuitively with full range of motion, however, the blade edges are slightly rounded at the tips which can negatively affect cutting performance by pinching material during use. Engineering also observed that one side of the main tube has a 1 inch section with material removed directly above the tube extension. Also, there is a 2. 5 inch long section with light scraping 3 inches above the tube extension. Based on the location and appearance, the damage may have been caused by a cannula accessory. The site has previously returned defective cannula accessories which were reported via mdr mfr report #2955842-2007-00381 and 2955842-2007-00382. The instrument may have been used with these cannula during a surgical procedure. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947174 | 952313 | 914314 | 2955842200700455 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-11-21 | 2007-11-19 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-10-22 | Yes | No | Reuse | 2007
42009307 | 2007-10-18 | Malfunction |  | In 2007, isi representative evaluated the site's da vinci s 8mm instruments to determine if the 8mm cannula accessories may have caused instrument tube abrasions similar to medwatch mfr report #2955842-2007-00438. | Visible scratches were found on the instrument main tube at the distal end near the metal wrist section. The main tube scratches are consistent with collisions between two instruments inside the body and may have been created by instrument collisions over multiple cases. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 947057 | 952314 | 914315 | 2955842200700471 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-09 | 2007-11-19 | No | Yes | Health Professional | 42009307 | No | No | 2007-10-18 | Yes | No | Unkown | 2007
40017205 | 2007-10-09 | Malfunction |  | It was reported that the maryland bipolar forceps instrument showed 0 uses left and then changed to having 4 uses left. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a tool reader. The reader confirmed that 4 used remain. The instrument was then placed on the system. The system did not show 0 used left. The same result occurred when the instrument was placed on a different patient side manipulator (psm). The instrument was driven and moved intuitively with full range of motion in all directions. All components function properly. Engineering also observed the distal end of the main tube to have light scratch marks directly above the proximal clevis, likely due to instrument collisions. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings : handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947029 | 952315 | 914316 | 2955842200700454 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-19 | 2007-11-19 | No | Yes | Health Professional | 40017205 | Yes | No | 2007-10-20 | Yes | No | Reuse | 2007
42020503 |  | Malfunction |  | It was reported that the top of the fenestrated bipolar forceps instrument was bent and broken. The site reported that this damage did not occur during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a break at the proximal clevis to main tube interface. The clevis is dislodged from main tube as a result. Both the proximal clevis and the main tube are missing pieces. Per the complaint notes, the instrument damage did not occur during a surgical procedure. Engineering also observed the distal end of the main tube to have a 2 inch section of localized scratch marks directly above the proximal clevis. Scratches may be due to instrument collisions causing a minimal amount of material to be removed from the main tube. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings : handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947043 | 952316 | 914317 | 2955842200700453 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-29 | 2007-11-19 | No | Yes | Health Professional | 42020503 | Yes | No | 2007-10-20 | Yes | No | Reuse | 2007
42009305 | 2007-10-22 | Other | Other | It was reported that during a da vinci surgical procedure while performing an instrument exchange, the surgical staff had difficulty removing the endowrist prograsp instrument. As a result, the surgical staff had to remove the instrument and cannula together. Upon removal of the instrument, it was also noticed a piece of the proximal clevis was missing. After further inspection of the pt's anatomy, the fragmented component was found and removed. The planned surgical procedure was completed without any delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface with pieces missing. The proximal clevis is dislodged from the main tube as a result of the breakage. Both the proximal clevis and the main tube are missing pieces. A broken piece of the proximal clevis was returned in a separate container and covers the entire missing section. Per the case notes, the broken instrument component was successfully retrieved from the pt. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 949515 | 952458 | 914884 | 2955842200700457 | 1 | GCJ | Manufacturer | User facility | Nurse | Initial | 2007-11-21 | 2007-11-19 | Yes | Yes | Health Professional | 42009305 | Yes | Yes | 2007-10-22 | Yes | No | Reuse | 2007
42000604 | 2007-09-21 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure the surgical staff noticed that a portion of the tip of the large needle driver instrument was missing. The site in unclear if the missing piece came off before or during the procedure, however, no pieces were observed inside of the patient. The planned surgical procedure was successfully completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering noticed that the distal half of the carbide insert was missing from one grip. Engineering evaluation observed that the brazing surface shows little presence of filler metal and the fracture plane of the insert is relatively straight, indicating possible shearing failure. Engineering also observed the outer surfaces of the grips (backside) to show various scratch marks, suggesting possible aggressive usage. It was concluded that the inadequate braze and aggressive usage most likely contributed to the insert falling off. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 952228 | 951637 | 913651 | 2955842200700442 | 1 | GCJ | Manufacturer | ForeignUser facility | Other | Initial | 2007-11-14 | 2007-11-14 | No | Yes | Health Professional | 42000604 | Yes | Yes | 2007-10-15 | Yes | No | Reuse | 2007
42020702 |  | Malfunction |  | It was reported that the tenaculum forceps instrument has a cracked proximal clevis, however, pieces did not fall inside of pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument broken at the proximal clevis to main tube interface. Both the proximal clevis and main tube were missing pieces. Per the case notes, no instrument fragments fell inside of the pt. Engineering also observed the distal end of the main tube has various deep scratches with material removed. The appearance of the damage suggests that the scratches may have been caused by inadvertent contact with other instruments. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 948046 | 950928 | 912961 | 2955842200700450 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2007-11-13 | 2007-11-13 | No | Yes | Health Professional | 42020702 | Yes | No | 2007-09-01 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that one of the wires on the monopolar curved scissors instrument had broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the grip close cable to be broken in two locations, at the scissor grip and at the distal idlers. Roughly half of the cable diameter is broken at each location. The cable broke under tensile loading. Other cables at the wrist are not damaged. Engineering also observed the distal end of the main tube to have a 2 inch long section, 3 inches above the tube extension, with material removed all around the tube. Damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 943580 | 947389 | 909460 | 2955842200700445 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-12 | 2007-11-12 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-10-19 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-10-18 | Other | Other | It was reported that during a da vinci radical hysterectomy surgical procedure, the customer experienced recoverable 23008 system error codes after experiencing difficulty handling the second arm of the system. After the dissection of the nodes, ureters, and salpingoophorectomias a nonrecoverable 23008 system error code was experienced. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The pcm was returned to isi for failure analysis investigation. Engineering discovered 23008 in the error logs. The 23008 system error code appeared when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". Engineering concluded that the potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 944708 | 947724 | 909796 | 2955842200700446 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2007-11-12 | 2007-11-12 | Yes | Yes | Health Professional | IS2000 A51P2 | Device Returned To Manufacturer | No | 2007-10-19 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-10-16 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, the customer experienced a 23013 system error code that they were unable to clear. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The ecm was returned to isi for failure analysis investigation. Engineering discovered both 230013 and 23008 in the error logs. The 23013 and 23008 system error codes appeared when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". Engineering concluded that the motor/ encoder and potentiometer assemblies had malfunctioned, thus generating the system errors experienced by the customer. As of november 12, 2007, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 944379 | 947725 | 909796 | 2955842200700444 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-12 | 2007-11-12 | Yes | Yes | Health Professional | IS2000 A51P2 | Device Returned To Manufacturer | No | 2007-10-16 | Yes | No | Reuse | 2007
42001102 |  | Other | Other | In 2007, isi rep evaluated the site's is2000 5mm regular cannula accessories to determine if the cannula may have caused the instrument tube abrasion reported under medwatch mfr report #2955842-2007-00438. | Two of the site's four cannulae were found to have a sharp inside edge, which have the potential to cause instrument main tube scraping. Medwatch mfr report #2955842-2007-00449 was submitted for the other damaged cannula accessory. As of 2007, there have been no reported recurrences of the issue at this hospital. | CANNULA 5MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 942055 | 947726 | 909797 | 2955842200700448 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-09 | 2007-11-12 | No | Yes | Health Professional | 42001102 | No | No | 2007-10-18 | Yes | No | Unkown | 2007
42001102 |  | Other | Other | On oct 18, 2007, isi representatives evaluated the site's is2000 5mm regular cannula accessories to determine if the cannulae may have caused the instrument tube abrasion reported under medwatch mfr report #2955842-2007-00438. | Two of the site's four cannulae were found to have a sharp inside edge which have the potential to cause instrument main tube scraping. Medwatch mfr report #2955842-2007-00448 was submitted for the other damaged cannula accessory. As of 2007, there have been no reported recurrences of the issue at this hospital. | CANNULA 5MM REGULAR ISI2000 | ENDOSCOPIC ACCESSORY | 942058 | 947727 | 909798 | 2955842200700449 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-09 | 2007-11-12 | No | Yes | Health Professional | 42001102 | No | No | 2007-10-18 | Yes | No | Unkown | 2007
IS1200 A4 3P8 | 2007-10-19 | Other | Other | It was reported that prior to beginning the da vinci cardiac procedure, the system surgeon console right eye view was foggy and the left eye had a pink tint. The system endoscope was removed and cleaned which resolved the fogging issue in the right eye, however, the pink tint in the left eye remained. The pt was under anesthesia, however, no surgical tasks had been performed with the system when the surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | The investigation conducted by isi field service engineering discovered that the settings of the camera control units (ccus) were incorrect and that the black and white balance had not been performed. The isi field service engineer also found that the system illuminator light output was low. The system was repaired by replacing the affected illuminator. As of 2007, there have been no reported recurrence of the issue at this hospital. The insite vision system user manual specifically states: setting the black and white balance. Automatic black-balance control (abc) should be set prior to using the da vinci system. It should be set prior to setting white-balance. Automatic white balance control (awc) should be set prior to using the da vinci system each day, or if an endoscope change has caused a color distortion in the image. The user manual also provides the default setting for use for the ccus. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 943003 | 947728 | 909799 | 2955842200700447 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial |  | 2007-11-12 | Yes | Yes | Health Professional | IS1200 A4 3P8 | Device Returned To Manufacturer | No | 2007-10-19 | Yes | No | Reuse | 2007
IS2000 A5 1P3 | 2007-10-16 | Other | Other | It was reported that during a da vinci surgical procedure, the customer experienced recurring system error code # 23013. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer were associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The mtm was returned to isi for failure analysis investigation. System error code #23013 appears when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". Engineering evaluation found that an axis motor encoder and potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. As of november 15, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 965867 | 947729 | 909796 | 2955842200700443 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-11-15 | 2007-11-12 | Yes | Yes | Service Personnel | IS2000 A5 1P3 | Device Returned To Manufacturer | No | 2007-10-16 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-10-15 | Other | Other | It was reported that while preparing the da vinci surgical system for surgery, a drape became stuck in one of the system arms. When the surgical staff attempted to remove the drape, they experienced system error code #20008. Closer inspection of the stuck drape revealed a broken cable on the system arm. The pt had been administered anesthesia when the surgeon decided to abort the planned surgical procedure and rescheduled it for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that system error code #20008 experienced by the customer was associated with a pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering eval confirmed that an axis cable had derailed, causing the cables to fall off the guide pulleys. Based on the complaint notes, it was determined that the cable most likely became derailed when the customer was attempting to remove the stuck drape. System error code #20008 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. Upon determining this condition, the safety system put davinci in a "recoverable safe state". As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 944533 | 947730 | 909801 | 2955842200700441 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-11-14 | 2007-11-12 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | Yes | 2007-10-15 | Yes | No | Reuse | 2007
420143 | 2007-09-10 | Other | Other | It was reported that during a da vinci nissen fundoplication surgical procedure small black pieces of "debris" were seen on the pt's anatomy. The pieces were removed, the abdomen was irrigated and suctioned to remove all small pieces. The surgical staff removed and inspected robotic instruments and found that three of the instruments appeared to be scraped, but intact. Further investigation by the surgical staff, under a microscope, found that the maryland forceps instrument had scuffed and scratched areas that may have material missing. | The instrument was not returned to isi for eval; however, isi representatives conducted an investigation at the hosp in 2007, and found the following: the marks on the main tube of the instrument were consistent with marks caused by the instrument main tube scraping against the inside edge of a cannula accessory. The small black pieces of "debris" have been identified as pieces of the instrument main tube that were skived off by the sharp edge of a cannula accessory. On site eval of the hospital's 5mm cannula accessories was performed. During eval, two defective cannulae were discovered. As of 11/9/07, there have been no reported recurrences of the issue at this hosp. The following medwatch reports listed below were also sent to the fda regarding these defective cannulae. 2955842-2004-00048, 2955842-2007-00449. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 894950 | 952442 | 884811 | 2955842200700438 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-11-09 | 2007-11-12 | Yes | Yes | Health Professional | 420143 | No | Yes | 2007-09-26 | Yes | No | Unkown | 2007
420172 | 2007-09-27 | Malfunction |  | The patient was having a robot assisted laparoscopic radial prostatectomy. During the procedure, a piece of the maryland bipolar forceps fell off of the instrument into the operating field. The surgeon was able to retrieve and remove the metal piece, which looked like a washer from the end of the instrument. The instrument was removed from the surgical field and tagged for return to intuitive surgical, inc. For investigation and replacement. There was no injury to the patient as a result of this incident. |  | 8MM MARYLAND BIPOLAR FORCEPS | ATTACHMENT SURGICAL ROBOT | 935227 | 965488 | 926859 | 965488 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2007-09-28 | 2007-11-06 | No | Yes | Physician | 420172 | No | No Answer Provided |  |  |  |  | 2007
420172 | 2007-09-27 | Malfunction |  | The patient was having a robot assisted laparoscopic radial prostatectomy. During the procedure, a piece of the maryland bipolar forceps fell off of the instrument into the operating field. The surgeon was able to retrieve and remove the metal piece, which looked like a washer from the end of the instrument. The instrument was removed from the surgical field and tagged for return to intuitive surgical, inc. For investigation and replacement. There was no injury to the patient as a result of this incident. |  | 8MM MARYLAND BIPOLAR FORCEPS | ATTACHMENT SURGICAL ROBOT | 935228 | 965489 | 926858 | 965489 | 1 |  | User Facility |  | RISK MANAGER | Initial | 2007-09-28 | 2007-11-06 | No | Yes | Physician | 420172 | No | No Answer Provided |  |  |  |  | 2007
420172 | 2007-09-27 | Malfunction |  | The patient was having a robot assisted laparoscopic radial prostatectomy. During the procedure, a piece of the maryland bipolar forceps fell off of the instrument into the operating field. The surgeon was able to retrieve and remove the metal piece, which looked like a washer from the end of the instrument. The instrument was removed from the surgical field and tagged for return to intuitive surgical, inc. For investigation and replacement. There was no injury to the patient as a result of this incident. |  | 8MM MARYLAND BIPOLAR FORCEPS | ATTACHMENT SURGICAL ROBOT | 935229 | 965490 | 926859 | 965490 | 1 |  | User Facility |  | RISK MANAGER | Initial | 2007-09-28 | 2007-11-06 | No | Yes | Physician | 420172 | No | No Answer Provided |  |  |  |  | 2007
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument has a broken cable towards the forcep. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw, possibly creating excess tension on the wire causing it to break. Engineering also observed the distal end of the main tube has a 1. 5" long section with material removed on one side of the tube. Based on the location and appearance, the damage may have been caused by a cannula accessory. Various small scratches were also observed, not aligned with the tube axis, below the damaged distal end area. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 940712 | 938344 | 900417 | 2955842200700433 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-11-05 | 2007-11-05 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-10-06 | Yes | No | Reuse | 2007
40009308 |  | Malfunction |  | It was reported that the endowrist prograsp instrument has cables sticking out. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch up cable was frayed at the distal clevis. Engineering noticed a scratch mark on the clevis adjacent to where the fray is located. Cable may have been nicked when clevis was scratched. Engineering also observed that the distal end of main tube has two distinct sections with material removed. Both sections have wear patterns parallel to the tube's longitudinal axis. Based on the location and appearance, the damage may have been caused by a cannula accessory. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 940638 | 938345 | 900418 | 2955842200700432 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-11-05 | 2007-11-05 | No | Yes | Health Professional | 40009308 | Yes | No | 2007-10-06 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the instrument was able to cut cleanly through. 006" latex. No damage to the blades was observed and electrical continuity passed. Engineering also observed that one side of the main tube has a section 4. 75" above the tube extension with material removed. The damaged area is. 8" long x. 2" wide, gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 940633 | 938346 | 900419 | 2955842200700398 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-11-05 | 2007-11-05 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-10-06 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull and not functioning properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the instrument was not able to cut cleanly through. 006" latex. The latex gets snagged at the scissor tips. The blades are not damaged, however, the blade edges are slightly rounded at the tips, negatively affecting cut performance. Electrical continuity passed. Engineering also observed that one side of the distal end of the main tube has a long section with material removed on one side of the tube. The damaged area goes all the way up to the midpoint of the tube. The damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 940647 | 938348 | 900421 | 2955842200700434 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-11-05 | 2007-11-05 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2007-10-08 | Yes | No | Reuse | 2007
31146402 | 2007-10-02 | Other |  | It was reported that during port placement for a davinci prostatectomy procedure, the customer experienced poor vision. The 0 degree schoelly stereo endoscope was removed and it was noticed that the lens was missing. A search of the pt's anatomy was performed, however, the missing lens could not be located. The endoscope was removed and replaced and the planned surgical procedure was completed with approx a 20 min delay. No pt harm, adverse outcome or injury was reported. Additionally, no post operative complications have been reported by this account. | The endoscope has been returned to the original equipment manufacturer (oem) for evaluation and failure analysis has not been completed yet. A follow-up mdr report will be submitted upon the oem's conclusion of the investigation. | 0 DEGREE SCHOELLY STEREO ENDOSCOPE | ENDOSCOPE | 935345 | 938425 | 900498 | 2955842200700436 | 1 | NAY | Manufacturer | User facilityCompany Representative | Nurse | Initial | 2007-11-07 | 2007-11-05 | No | Yes | Health Professional | 31146402 | Yes | Yes | 2007-10-10 | No | No | Unkown | 2007
IS2000 A51P2 | 2007-10-08 | Other | Other | It was reported that prior to starting a da vinci surgical procedure the customer experienced a recurring system error code #23013. The surgical staff shut down and restarted the system several times, however, this action was unsuccessful in resolving the system error code. The pt had been under anesthesia for 30 mins and the port incisions were not yet made when the surgeon decided to abort the planned surgical procedure and reschedule the surgery for a later date. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system error code #23013 experienced by the customer was associated with a pt side manipulator. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error code #23013 appears when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state. " based on the system error logs and axis potentiometer and motor encoders were replaced. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 935356 | 938427 | 900500 | 2955842200700435 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2007-11-07 | 2007-11-05 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-10-08 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-10-09 | Other | Other | It was reported that during a da vinci prostatectomy procedure the customer experienced system error codes #20008 adn #20009 on one of the system arms. The system was shut down and rebooted, however, once an instrument was inserted into psm2, the same system error codes occurred. Due to the recurring system error codes the procedure was aborted and rescheduled for a later date. No pt harm was reported. | The investigation conducted by isi field service engineering concluded that the system error code #20008 and #20009 experienced by the customer was associated with a pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error codes #20008 and #20009 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. Upon determining this condition, the safety system put davinci in a "recoverable safe state. " based on the system error logs an axis potentiometer and motor encoders were replaced. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 935368 | 938428 | 900501 | 2955842200700437 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2007-11-07 | 2007-11-05 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-10-09 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-10-11 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure the customer experienced multiple occurrences of system error code #20013. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error code #20013 experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The ecm was returned to isi for failure analysis investigation. Engineering discovered both #20013 and #21013 system error codes in the error logs. The #20013 system error code appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state. " engineering evaluation discovered that a potentiometer assembly did not meet performance specifications, thus generating the system errors experienced by the customer. There have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 935614 | 938429 | 900501 | 2955842200700440 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2007-11-10 | 2007-11-05 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | Yes | 2007-10-11 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-10-11 | Other | Other | It was reported that prior to beginning a cardiac procedure the right camera controller unit (ccu) lost power and was unable to be powered back on. The site made the decision to abort the procedure after the pt had been administered anesthesia. No pt harm was reported. | The investigation conducted by field service engineering concluded that the loss of video experienced by the customer was associated with the right camera controller unit (ccu). The system was repaired by replacing the affected ccu. There have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 935476 | 938430 | 900503 | 2955842200700439 | 1 | KOG | Manufacturer | User facility | Unknown | Initial | 2007-11-05 | 2007-11-05 | No | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-10-11 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-10-04 | Other | Other | It was reported that during a da vinci prostatectomy procedure the customer experienced system error code #20008 which could not be overridden. The system was shut down and rebooted, however, the same system error code recurred. The site shut down the system for a second time and the system would not power up when turned on. No patient harm, adverse outcome, or injury was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering concluded that system error code #20008 experienced by the customer was associated with a patient side manipulator (psm) and that the power on failure was associated with the primary power supply (pps). The system was repaired by replacing the affected psm and pps. The psm was returned to isi for failure analysis investigation. Under normal circumstances, every 2000x fault also has a 2100x fault with additional information. Together the two fault codes describe completely what the machine records. For the type of faults that result in a 20008/21008 fault pair, the system safety software determined that the robotic joint angular measure as recorded by the primary sensor (encoder) and secondary sensor (potentiometer) differed by more than the specified tolerance. During testing and evaluation, engineering found that a motor encoder axis pinion gear had shifted approx 90 degrees from the shaft due to a failing loctite bond, thus generating the system error codes. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put davinci in a "recoverable safe state". The pps was also returned to isi. Isi forwarded the pps to the vendor for failure analysis investigation. If a malfunction cannot be confirmed, a follow-up mdr will be submitted. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 934445 | 936578 | 898652 | 2955842200700428 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-11-02 | 2007-11-01 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-10-04 | Yes | No | Reuse | 2007
42018904 | 2007-10-01 | Other | Other | It was reported that during a da vinci prostatectomy procedure, one of the system arms stopped working. The surgical staff inspected the fenestrated grasper instrument and found wires hanging out of the tip. While attempting to remove the instrument, the surgical staff noticed that "shards" of material had fallen into the pt. The surgical staff had some difficulty removing all of the "shards" due to their small size. Upon removal of the instrument, it was also noticed that pieces of the proximal clevis were missing. After further inspection of the pt's anatomy, two pieces of the proximal clevis were found and removed. The planned surgical procedure was completed with an approximate delay of 30 minutes. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument is broken at the proximal clevis to main tube interface with pieces missing. The proximal clevis is dislodged from the main tub as a result of the breakage. Both the proximal clevis and the main tube are missing pieces. It was determined that the small "shards" noticed by the customer are most likely pieces of the main tube interface wall. The materials used in the intuitive surgical double fenestrated grasper instruments have been tested for biocompatibility and the materials used for the proximal clevis and main shaft were found to be non-cytotoxic, non-sensitizing, non-toxic, non-hemolytic, non-pyrogenic and appropriately biocompatible for their intended use. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 924599 | 935968 | 898042 | 2955842200700421 | 1 | GCJ | Manufacturer | Health ProfessionalCompany Representativ | Other | Initial | 2007-10-31 | 2007-10-31 | Yes | Yes | Health Professional | 42018904 | No | No | 2007-10-01 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing on the instrument and found the scissors cut cleanly through. 006" latex. The blades are not damaged. Engineering also observed the distal end of main tube to have a 2. 5" long section located 2. 5" above the tube extension with material removed on one side of the tube. The damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932561 | 935389 | 897463 | 2955842200700430 | 1 | GEI | Manufacturer | Health Professional | Unknown | Initial | 2007-10-29 | 2007-10-29 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-10-05 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument did not work. No additonal information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire to have a cut near the yaw pulley exit, exposing the bare wire. Engineering also observed the distal end of the main tube to have a 4. 5" long section directly above the tube. The damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932571 | 935390 | 897464 | 2955842200700429 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-10-29 | 2007-10-29 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-10-05 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument jaws were not working properly. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed all instrument components were functioning properly and the instrument moved intuitively with full range of motion. The customer reported failure mode could not be confirmed. During evaluation, engineering observed a 3" section on the instrument main tube which had deep scratches and material removed. The damaged area was located on the distal end of the main tube, directly above the proximal clevis. Based on the location and appearance of the damage, the scratches may have been caused by inadvertent contact with other instruments throughout the life span of the instrument. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932581 | 935391 | 897465 | 2955842200700425 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-11-01 | 2007-10-29 | No | Yes | Health Professional | 42021405 | Yes | Yes | 2007-10-03 | Yes | No | Reuse | 2007
42017908 | 2007-09-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the proximal clevis of the monopolar curved scissors instrument broke. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the proximal clevis, which is made of metal for the instrument, is not broken. Instead, the tube extension is dislodged from the main tube. The keys on the tube extension have broken off, resulting in tube extension rotating 360 degrees. Breakage may be due to overloading. Engineering also observed the distal end of the main tube to have a section 2 inches above the tube extension with material removed all around the tube. The damaged area is gray in color and has a rough surface finish. The damage may be caused by a defective cannula. There are also various deep scratches below the section with material removed. Electrical continuity passed. The following medwatch reports were created fro the defective cannula previously returned from this site: mfr report # 2955842-2006-00300, -00301, -00302, -00303, -00304, and -00305. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932651 | 935392 | 897466 | 2955842200700424 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-29 | 2007-10-29 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-10-02 | Yes | No | Initial | 2007
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument would not cauterize. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results found that electrical continuity passed and the conductor wires show no signs of damage. The customer reported failure mode could not be confirmed. Also during evaluation, engineering observed a 3" section on the instrument main tube which had material removed. The damaged area was located on the distal end of the main tube, directly above the proximal clevis. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. Engineering also observed that one grip was bent, causing misalignment of the grips. No other damage was found. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932547 | 935393 | 897467 | 2955842200700423 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-10-31 | 2007-10-29 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-10-02 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument had a bent cautery connector. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the bottom bipolar pin to be bent up towards the top pin so that a bipolar cord cannot be installed. Damage to pin is likely due to mishandling. Engineering also observed the distal end of the main tube to have a 4. 5" long section directly above the proximal clevis with material removed on one side of the tube. The damaged area is gray in color and has a rough surface finish. Based on the location and appearance, the damage may have been caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if additional information is received. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932533 | 935394 | 897468 | 2955842200700427 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-29 | 2007-10-29 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-10-04 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument flush port fell out of the housing. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the flush tube missing from the instrument, the luer plate dislodged from the chassis and loose in the back end. It appears that the damage was most likely due to a use issue during instrument cleaning. Also during evaluation, engineering observed the distal end of the main tube had various deep scratch marks with material removed. Based on the appearance of the damage, the scratches may have been caused by inadvertent contact with other instruments. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 932683 | 935395 | 897469 | 2955842200700422 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-10-31 | 2007-10-29 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-10-02 | Yes | No | Reuse | 2007
42020501 |  | Malfunction |  | It was reported that the tip of the fenestrated bipolar forceps instrument broke during sterilization. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed all components at the wrist to be intact, however, the grips are slightly misaligned. The luer plate is dislodged from the chassis. The chassis feature that retains the luer plate is broken. Instrument was likely due to a use issue during the cleaning process. Engineering also observed the distal end of the main tube to have a section 3 inches above the proximal clevis with deep scratches. The scratches have removed material from the tube. Scratches are concentrated in a 1. 5" long area on one side of the tube only. The scratch marks are not parallel to the tube axis, suggesting they may have been caused by inadvertent contact with other instruments. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 932678 | 935396 | 897470 | 2955842200700426 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-10-29 | 2007-10-29 | No | Yes | Health Professional | 42020501 | Yes | No | 2007-10-04 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-09-26 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced multiple occurrences of system error code # 20008, which they were initially able to override. Surgical staff was eventually unable to override the error code despite an isi representative's attempt to assist the site in troubleshooting the issue. The patient had been under anesthesia for 45 mins and the patient port incisions had already been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code #20008 experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Testing performed by engineering discovered a broken motor gear clamp screw on an axis of the mtml, thus generating the system error code #20008 experienced by the customer. System alarm (system generated fault code) functioned as designed, and there was no injury to the patient. The system error code #20008 appears when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining this condition, the safety system da vinci in a "recoverable safe state". As of oct 23, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 934290 | 934479 | 894183 | 2955842200700419 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-10-26 | 2007-10-26 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | Yes | 2007-09-26 | Yes | No | Reuse | 2007
IS1200 A49P8 | 2007-09-25 | Other | Other | It was reported that during preparation for a da vinci prostatectomy procedure, one of the systems patient side manipulators was not recognizing instruments. The surgical staff shut down and restarted the system, however, this action was unsuccessful in resolving the instrument recognition issue. The patient had been under anesthesia for 1. 5 hrs and the patient port incisions had already been made when the surgeon decided to abort the planned surgical procedure and reschedule the surgery for a later date. No patient harm, adverse outcome or injury was reported. | The evaluation conducted by isi field service engineering could not reproduce the customer reported failure mode, however, the system error logs were reviewed and it was observed that system error code #31009 had occurred multiple times, indicating that an instrument was not being detected by the system. Investigation of the instrument sterile adapters used on the site's system revealed that many of the sterile adapters were being used past their designated life. The system tracks the usage of each sterile adapter, and displays the number of uses left at the end of each procedure. It was determined that using a sterile adapter past its designated usage most likely contributed to the customer reported failure mode. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 934280 | 934473 | 894183 | 2955842200700420 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-10-25 | 2007-10-25 | Yes | Yes | Health Professional | IS1200 A49P8 | Yes | No | 2007-09-25 | Yes | No | Reuse | 2007
42018403 |  | Malfunction | Other | It was reported that during a da vinci thymectomy procedure, a burn hole was noticed through protective thermal cover of the cuttery spatula instrument, while dissecting the thymus gland. The surgeon noticed an arcing mark on the thymus gland after removal. The instrument was replaced with an instrument of the same type to complete the procedure. No additional pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the ceramic sleeve missing a. 060" x. 060" piece at the distal end. No evidence of a hole burned through the ceramic sleeve was found. Failure mode is believed to be due to mechanical fracture, not electrical burning. The larger diameter section of the ceramic sleeve has various scratch marks, suggesting multiple instances of inadvertent contact with other instruments. Breakage of ceramic sleeve may be due to the impact from an instrument collision. Electrical continuity passed. No other damage found. There are no indications from the site that the missing instrument components fell inside of the pt during the surgical procedure. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA | ELECTROSURGICAL INSTRUMENT | 928904 | 933211 | 895338 | 2955842200700415 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-26 | 2007-10-24 | No | Yes | Health Professional | 42018403 | Yes | No | 2007-09-26 | Yes | No | Initial | 2007
42017206 |  | Malfunction |  | It was reported that the wrist of the bipolar maryland forceps instrument does not rotate properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering eval results found that all instrument components were functioning properly and the instrument moved intuitively with full range of motion in all directions. The customer reported failure mode could not be confirmed. During eval, engineering also observed a 1. 5" section on one side of the distal end of the instrument main tube which had some material removed. Based on the location and appearance, the damage may have been caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 929387 | 932033 | 894175 | 2955842200700416 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-26 | 2007-10-22 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2007-09-27 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the maryland bipolar instrument was not working properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, that the instrument is not working, engineering observed that the instrument had sustained cosmetic damage, but was unrelated to the reported failure. During eval, engineering observed the distal end of the main tube had material removed on one side. Based on the location and appearance, the damage may have caused by a cannula accessory. Add'l investigation is underway regarding the cannula accessories. Also observed were deep scratches in the same localized area, likely caused by contact from other instruments. No other damage was found. A follow-up mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 930242 | 932034 | 894176 | 2955842200700409 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-24 | 2007-10-22 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-09-24 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had a broken/loose cable. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed one conductor wire to have a small cut on the insulation. The cut is located near the wire hole in the proximal clevis. Engineering also the distal end of the main tube to have a large section with scraping all around the tube. The damaged area ranges from directly above the proximal clevis to nearly the midpoint of the tube. The scraped section is gray in color and has a rough surface finish. Engineering has concluded that the damage may be caused by a defective cannula. Electrical continuity passed. No other damage found. The following medwatch reports listed below will also be sent to the fda. These reports only pertain to the cannulae returned from this specific site: 2955842-2007-00389, 2955842-2007-00390. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 927261 | 932037 | 894179 | 2955842200700410 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-22 | 2007-10-22 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-09-24 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-09-26 | Other | Other | It was reported that during the da vinci prostatectomy surgical procedure, the customer experienced system error codes #21002 and #21008. No patient harm was reported. The procedure was converted to traditional open surgical technique to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. System error codes #21002 and #21008 are sympathetic errors to two sensors not being in agreement and appear when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (hall effects sensor) and the secondary sensor (hall effects sensor), were out of specified tolerance for agreement. Testing revealed that the hall effect sensor circuit malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. As of october 22, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 934371 | 932041 | 894183 | 2955842200700413 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2007-10-22 | 2007-10-22 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-09-26 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-09-27 | Other | Other | It was reported that during the da vinci cardiac surgical procedure, the customer experienced recurring faults. The site's biomedical engineer reviewed the system's error logs from the procedure and noticed multiple occurrences of system error code # 21013. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error #21013 experienced by the customer was associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error code #21013 is generated when an error is detected as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). Testing revealed an axis potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed, and there was no injury to the pt. As of oct. 18, 2007, there have been no reported recurrences of the issue as this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 929435 | 932042 | 894183 | 2955842200700414 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-10-26 | 2007-10-22 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-09-27 | Yes | No | Reuse | 2007
400169 | 2007-08-30 | Malfunction |  | It was reported that during the surgical procedure one of the tips of the harmonic curved shears insert broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was successfully completed with a replacement insert. No pt harm, adverse outcome or injury was reported. | The instrument accessory has not yet been returned for eval. Once the instrument accessory is returned and evaluated, a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC ACCESSORY | 946915 | 978923 | 889814 | 2955842200700418 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-10-18 | 2007-10-22 | No | Yes | Health Professional | 400169 | No | Yes |  | No | Yes | Initial | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00306, -00307, -00308, -00309, -00310, -00311, 2955842-2007-00401, -00402. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a 2" section of scraping on one side of the distal end of the instrument main tube. There were also a couple of deep scratches in the same area as the scrape marks. It was determined that the damage to the main tube was most likely caused by a defective cannula accessory. The following medwatch reports were created for the defective cannula previously returned from this site: mfr. Report #2955842-2006-00300, -00301, -00302, -00303, -00304, -00305. Engineering also conducted performance testing and observed that material became snagged at the instrument tips when attempting to cut latex. Electrical continuity passed and no other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 935450 | 930991 | 893123 | 2955842200700403 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-19 | 2007-10-18 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-09-20 | Yes | No | Initial | 2007
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00306, -00307, -00308, -00309, -00310, -00311, 2955842-2007-00402, -00403. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various scratches on the distal end of the instrument main tube with material removed. The scratches are located 2" above the proximal clevis. The scratches are not parallel to the tube's longitudinal axis, indicating that they were most likely caused by inadvertent contact with another instrument. No other damage was found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 935455 | 930992 | 893124 | 2955842200700401 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-10-19 | 2007-10-18 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-09-20 | Yes | No | Initial | 2007
42017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00306,mfr. Report #2955842-2006-00307, mfr. Report #2955842-2006-00308, mfr. Report #2955842-2006-00309, mfr. Report #2955842-2006-003010, mfr. Report #2955842-2006-003011, mfr. Report #2955842-2006-00401, mfr. Report #2955842-2006-00403. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the distal end of the instrument main tube with light scraping, and a small amount of material removed directly above the proximal clevis. It was determined that the damage to the main tube was most likely caused by a defective cannula accessory. No other damage was found. The following medwatch reports were created for the defective cannulae previously returned from this site: mfr. Report #2955842-2006-00300, mfr. Report #2955842-2006-00301, mfr. Report #2955842-2006-00302, mfr. Report #2955842-2006-00303, mfr. Report #2955842-2006-00304,mfr. Report #2955842-2006-00305. | BIPOLAR MARYLAND FORCEPS INSTRUMENTS | ELECTROSURGICAL INSTRUMENT | 926323 | 930993 | 893125 | 2955842200700402 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-19 | 2007-10-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-09-20 | Yes | No | Reuse | 2007
42020904 |  | Malfunction |  | It was reported that the suturecut needle driver instrument had a broken casing. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument housing (which the customer referred to as the "casing") and chassis were intact, with no evident signs of damage. During evaluation, engineering observed a 2" section on the instrument main tube which had material removed. The damaged area was located on the distal end of the main tube, directly above the proximal clevis and covered half of the outer diameter. It was determined that the damage to the main tube was most likely caused by a defective cannula accessory. No other damage was found. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 926813 | 930994 | 893126 | 2955842200700407 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2007-10-19 | 2007-10-18 | No | Yes | Health Professional | 42020904 | Device Returned To Manufacturer | No | 2007-09-21 | Yes | No | Reuse | 2007
31195402 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As part of an investigation regarding cannula related tube abrasions. | The cannula accessory was evaluated. Engineering observed that the cannula accessory is out of round at the distal tip. It appears that the long side of the distal tip has been slightly dented. A. 342 gage pin does not fit through the distal tip inner diameter. Engineering concluding that the damage to the cannula is likely due to mishandling. The deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | 38 OD CANNULA | ENDOSCOPIC ACCESSORY | 926805 | 930995 | 893127 | 2955842200700412 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-10-18 | 2007-10-18 | No | Yes | Health Professional | 31195402 | Yes | No | 2007-08-30 | Yes | No | Unkown | 2007
31195402 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As part of an investigation regarding cannula related tube abrasions. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a flat spot at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Engineering has concluded that the damage to the cannula is likely due to mishandling. Deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | 38 OD CANNULA | ENDOSCOPIC ACCESSORY | 926796 | 930997 | 893129 | 2955842200700411 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-10-18 | 2007-10-18 | No | Yes | Health Professional | 31195402 | Yes | No | 2007-08-30 | Yes | No | Unkown | 2007
IS1200 A43P8 | 2007-09-18 | Other | Other | It was reported that during the da vinci prostatectomy surgical procedure, the customer experienced recurring system error code #20013. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. A review of the system error logs indicate that system error code #21013 was also generated. System error codes are generated when an error is detected as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). Testing revealed an axis potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. As of october 11, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 905002 | 929974 | 892104 | 2955842200700394 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-10-17 | 2007-10-15 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-09-18 | Yes | No | Reuse | 2007
42017908 | 2007-09-19 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, smoke was noticed around the tip of the monopolar curved scissors instrument while performing tissue dissection. The instrument was removed and replaced with an instrument of the same type. The tip cover accessory was discarded at the site. The procedure was successfully completed without significant delay. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00399. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a. 060" x. 080" black mark on the proximal clevis, centered at one of the corners of the slots that expose the silicone overmolding. The black mark appears to be charred bio material. Examination of the silicone under magnification shows localized melting of the silicone, indicating an arcing event occurred. The location of the mark lines up approx with the sleeve to silicone interface on the tip cover. It is possible that the tip cover may have had a cut in the silicone that created a path for arcing. Per the case notes, the site discarded the tip cover and did not notice any damage to it upon visual examination. However, the melting of the silicone combined with the location of the mark suggest the tip cover would have had some kind of damage near the silicone to sleeve interface. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 930983 | 930605 | 892743 | 2955842200700400 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-09-19 | Device Not Retu | No | Unkown | 2007
40021405 | 2007-09-14 | Malfunction |  | It was reported that during a da vinci total hysterectomy procedure, the proximal clevis of a pk dissecting forceps instrument broke while trying to remove the lymph nodes. Instrument fragments fell into the patient and were retrieved. The procedure was completed without delay using a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed no broken pieces or visible damage on the proximal clevis of the instrument. One of the yaw pulleys is missing a 1/4" x 1/8" piece, which was returned taped to the housing. As a result of the breakage, the cable crimp is no longer retained in the yaw pulley and the conductor wire has pulled out of the grip. It was concluded that the yaw pulley may have had a crack which propagated with heavy usage. Electrical continuity failed. No other damage found. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 929243 | 930606 | 892738 | 2955842200700396 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional | 40021405 | Device Returned To Manufacturer | No | 2007-09-18 | Yes | No | Reuse | 2007
| 2007-09-19 | Malfunction |  | It was reported that during a prostatectomy procedure, the smoke was noticed around the tip of the monopolar curved scissors instrument while performing a dissection. The instrument was removed and replaced with an instrument of the same type. The tip cover accessory was discarded at the site. The procedure was successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00400. | The accessory was not returned for evaluation, therefore, the root cause of the customer reported failure mode could not be determined. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 929248 | 930608 | 892740 | 2955842200700399 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional |  | No | No | 2007-09-19 | Device Not Retu | Yes | Initial | 2007
420179 |  | Malfunction | Other | It was reported that during a da vinci prostatectomy procedure, a burn hole was noticed on the tip cover accessory of the monopolar curved scissors instrument while attempting to cauterize tissue. Upon removing the instrument,the site noticed the patient's iliac vein had been damaged. The iliac vein was repaired using a suture and the tip cover accessory was replaced with a tip cover of the same type to complete the procedure. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00405. | The instrument was not returned for evaluation. A tear was found during failure analysis of the tip cover accessory used in conjunction with the instrument, which may have caused a path for arcing to occur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 929259 | 930609 | 892741 | 2955842200700404 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-15 | 2007-10-15 | Yes | Yes | Health Professional | 420179 | No | No | 2007-09-20 | Device Not Retu | No | Unkown | 2007
40018005 | 2007-09-17 | Malfunction |  | It was reported that during a da vinci radical nephrectomy procedure, a burn hole was noticed on the tip cover accessory of the monopolar curved scissors instrument while the instrument was inside the patient. The instrument had been in use for 92 minutes when the event occurred. The tip cover was replaced with a tip cover of the same type and the procedure was successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00395. | The instrument accessory was returned and evaluated. Per the reported complaint, engineering observed small pinholes burned through the silicone. In addition, there are various tears in the silicone. A hole, roughly. 020" in diameter, was located. 150" above the gray plastic sleeve. Material is removed from the outer surface of silicone at the hole, suggesting heat/electrical energy caused the hole. There is also a tear cutting through the middle of the hole which may have created a path for arcing. There are two mechanical tears in the silicone above the hole, and two larger tears at the tip of the silicone. Tears are likely due to contact with another instrument. Video of procedure shows various instances of the tip cover accessory being nicked by other instruments and monopolar curved scissors instrument articulated in ful pitch and yaw, stretching the silicone. At 1. 5 hours into procedure, video shows two successive instance of arcing though the tip cover. Observation of the tip cover position during arching appears to indicate electrical energy went through the mechanical tears, not the hole. Video does not show any other clear evidence of arching. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 929881 | 930610 | 892740 | 2955842200700408 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional | 40018005 | Device Returned To Manufacturer | No | 2007-09-17 | Yes | Yes | Initial | 2007
| 2007-09-19 | Malfunction |  | It was reported that during a da vinci radical nephrectomy procedure, a burn hole was noticed on the tip cover of the monopolar curved scissors instrument while the instrument was inside the patient. The instrument had been in use for 92 minutes when the event occurred. The tip cover was replaced with a tip cover of the same type and the procedure was successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00408. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode could not be determined. Please see mdr report #2955842-2007-00408. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 929266 | 930611 | 892743 | 2955842200700395 | 1 | GEI | Manufacturer | Company Representative | UNKNOWN | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional |  | No | No | 2007-09-17 | Device Not Retu | No | Unkown | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument would not register on the sterile adapter. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer complaint, engineering placed the instrument on a system and checked for recognition. The system successfully recognized the instrument, showing 5 uses left. A magnet is installed and pogo pins do not stick when actuated. Engineering also observed a single deep scratch at the distal end of the main tube, 1/4" above the proximal clevis. Material has been removed due to the scratch. The scratch is not parallel to the longitudinal axis of the tube, suggesting it was likely caused by an instrument collision. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 926245 | 930612 | 892744 | 2955842200700397 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-09-19 | Yes | No | Reuse | 2007
4200935 |  | Malfunction |  | It was reported that the jaws of the endowrist prograsp instrument do not move properly. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the instrument to move intuitively with full range of motion in all directions. All components function properly. Engineering also observed the distal end of the main tube to have various scratches. Two scratches are deep enough to remove a small amount of material. Scratches are not parallel to tube's longitudinal axis indicating that the scratches are likely due to inadvertent contact ("swordfighting") with other instruments. The instrument has 1 use remaining and has more than average wear and tear on tube. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 926892 | 930613 | 892745 | 2955842200700406 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Health Professional | 4200935 | Yes | No | 2007-09-21 | Yes | No | Reuse | 2007
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument would not open or close. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the instrument was not recognized upon initial installation, however, subsequent installations passed. The pogo pins do not stick and heights are fine. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed that the main tube has a section 4 inches above the proximal clevis with light scraping all around the outer diameter. The damaged area is lighter in color than the remainder of the tube and has a slightly rough surface finish. Enginerring concluded that the damge may have been caused by a defective cannula. Electrical continuity passed. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 925164 | 930614 | 892746 | 2955842200700392 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-10-15 | 2007-10-15 | No | Yes | Other | 42017206 | Yes | No | 2007-09-17 | Yes | No | Reuse | 2007
40018005 |  | Malfunction | Other | It was reported that during a da vinci prostatectomy procedure, a burn hole was noticed on the tip cover accessory of the monopolar curved scissors instrument while attempting to cauterize tissue. Upon removing the instrument, the site noticed the patient's iliac vein had been damaged. The iliac vein was repaired using a suture and the tip cover was replaced with a tip cover accessory of the same type to complete the procedure. No additional patient harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00404. | The instrument tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover has roughly a. 025" diameter hole burned through the silicone. The hole is located roughly. 200" above the plastic sleeve to silicone interface. Examination of the hole under magnification shows localized melting of the silicone, suggesting heat/electrical energy caused the hole. There is a tear adjacent to the center of the hole that may have been caused by inadvertent contact from another instrument. Engineering concluded that this tear may have created a path for arcing if it became filled with body fluids and cautery was activated. There is also a small tear at the distal tip of the silicone, very near one of the parting lines. No other damage found. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 925177 | 930615 | 892747 | 2955842200700405 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-15 | 2007-10-15 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2007-09-20 | Yes | Yes | Initial | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found that the instrument was able to cut per specification. No damage to the blades was observed and electrical continuity passed. Engineering also observed that the distal end of the instrument main tube had a section all around the tube with material removed, located immediately above the tube extension. The tube damage may have been caused by a defective cannula. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 926238 | 930604 | 892736 | 2955842200700393 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-10-17 | 2007-10-11 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-09-17 | Yes | No | Reuse | 2007
420172 | 2007-10-09 | Malfunction |  | This is the second report of small black pieces of an unidentified foreign substance found during a robotic procedure. It was noted that there were some collisions between the instruments, which occurred prior to the discovery of the pieces. Post procedure, there was some visible damage to the instruments. One of the possibilities was that the black debris could have originated from the instruments. Collisions of the instruments are avoided in the or by correctly placing the port locations and positioning of robotic surgical arms. Surgeons must be aware of where instrument shafts are for both robotic and non-robotic instruments, and the patient side assistant and/or other personnel in the room should be making the console side surgeon aware of when collisions occur, and what devices are colliding. |  | MARYLAND BIPOLAR | ATTACHMENT SURGICAL ROBOT | 929655 | 959494 | 921279 | 959494 | 1 | NAY | User Facility |  |  | Initial | 2007-10-11 | 2007-10-11 | No | Yes | Physician | 420172 | Yes | No Answer Provided |  |  |  |  | 2007
420172 | 2007-10-09 | Malfunction |  | This is the second report of small black pieces of an unidentified foreign substance found during a robotic procedure. It was noted that there were some collisions between the instruments, which occurred prior to the discovery of the pieces. Post procedure, there was some visible damage to the instruments. One of the possibilities was that the black debris could have originated from the instruments. Collisions of the instruments are avoided in the or by correctly placing the port locations and positioning of robotic surgical arms. Surgeons must be aware of where instrument shafts are for both robotic and non-robotic instruments, and the patient side assistant and/or other personnel in the room should be making the console side surgeon aware of when collisions occur, and what devices are colliding. |  | MARYLAND BIPOLAR | ATTACHMENT SURGICAL ROBOT | 929655 | 959494 | 921279 | 959494 | 1 | NAY | User Facility |  |  | Initial | 2007-10-11 | 2007-10-11 | No | Yes | Physician | 420172 | Yes | No Answer Provided |  |  |  |  | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned per the request of isi as the site had previously returned instruments with damaged main tubes which may have been caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda. Regarding these instruments: mfr. Report#: 2955842-2007-00272, 2955842-2007-00194, 2955842-2007-00195, 2955842-2007-00131. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was visibly dented at the distal tip and it appears that the damage was most likely caused user mishandling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. The site also returned another damaged cannula accessory and mfr. Report# 2955842-2007-390 was submitted for this accessory. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 911059 | 926336 | 888510 | 2955842200700389 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-10-13 | 2007-10-10 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-09-13 | Yes | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned per the request of isi as the site had previously returned instruments with damaged main tubes which may have been caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda. Regarding these instruments: mfr. Report#: 2955842-2007-00272, 2955842-2007-00194, 2955842-2007-00195, 2955842-2007-00131. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a subtle flat spot at the distal tip, causing out of roundness and it appears that the damage was most likely caused by user mishandling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. The site also returned another damaged cannula accessory and mfr. Report# 2955842-2007-389 was submitted for this accessory. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 911122 | 926337 | 888511 | 2955842200700390 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-10-10 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-09-13 | Yes | No | Unkown | 2007
420093 | 2007-09-12 | Malfunction |  | It was reported that during the da vinci supracervical hysterectomy procedure, an error message was noticed and the system arms would not move. The surgical staff examined each system arm to remove and reset the instruments installed and had difficulty removing the endowrist prograsp instrument from one of the arms. The surgical resident attempted to grab the endowrist prograsp instrument with another instrument to remove it from the pt, but was unsuccessful. The instrument was removed with the trocar still attached and the instrument appeared to be shredded upon removal. Upon redocking, a metal instrument component was observed in the pt and was retrieved. An instrument of the same type was installed on the system arm and the planned surgical procedure was completed after a 45 minute delay. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. Once the instrument has been received and failure analysis has been completed, a follow-up mdr will be submitted. Per the case notes, the broken instrument component was successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 912716 | 926339 | 888513 | 2955842200700387 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2007-10-12 | 2007-10-10 | No | Yes | Health Professional | 420093 | No | Yes | 2007-09-12 | No | No | Unkown | 2007
42020702 | 2007-09-10 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure the shaft of the tenaculum forceps instrument broke when the surgical staff attempted to force it out of the system. Instrument fragments fell inside of the patient and were retrieved. The planned surgical procedure was completed after a 15-30 minute delay. No patient harm, adverse outcome or injury was reported. | The tenaculum forceps instrument has not been returned for evaluation. Once the instrument has been received and failure analysis has been completed, a follow-up mdr will be submitted. Per the case notes, the instrument fragments were successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENTS | ENDOSCOPIC INSTRUMENT | 912131 | 926341 | 888515 | 2955842200700383 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-10-10 | 2007-10-10 | No | Yes | Health Professional | 42020702 | No | No | 2007-09-10 | No | No | Unkown | 2007
IS1200 A43P8 | 2007-09-10 | Other | Other | It was reported that during a da vinci prostatectomy procedure the customer experienced system error code # 20008, which could not be overridden. The system power was recycled, however, this action did not resolve the issue. System homing could not be completed and the customer again experienced system error code # 20008. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error code #20008 experienced by the customer was indicative of a mismatch between the primary and backup sensor for a particular axis associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Testing performed by engineering discovered that a motor encoder assembly on an axis of the mtmr had malfunctioned, thus generating the system error code #20008 experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error code #20008 appeared when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder), and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system put davinci in a "recoverable safe state". As of october 5, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 911905 | 926812 | 888986 | 2955842200700380 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-10-10 | 2007-10-10 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-09-10 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-09-12 | Other | Other | It was reported that during the surgical procedure, the customer experienced recurring system error code #20008 which could not be overridden. No patient harm, adverse outcome, or injury was reported. The procedure was completed with the sites other da vinci surgical system, causing approximately a one hour delay. | The investigation conducted by field service engineering concluded that the system error code #20008 experienced by the customer was indicative of a mismatch between the primary and backup sensor for a particular axis associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Testing performed by engineering discovered that a motor encoder assembly on an axis of the mtml had malfunctioned, thus generating the system error code #20008 experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error code #20008 appeared when the davinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system put davinci in a "recoverable safe state". As indicated in the complaint notes, the planned surgical procedure was completed with the site's other da vinci surgical system. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 940695 | 926814 | 888986 | 2955842200700388 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-12 | 2007-10-10 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-09-12 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-09-13 | Other | Other | It was reported that during a da vinci hysterectomy surgical procedure, the customer experienced a recoverable system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The psm was returned to isi for failure analysis investigation. Engineering discovered both in the error logs. The both system error codes appeared when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put davinci s in a "recoverable safe state". Engineering concluded that the potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. As of october 5, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 940723 | 926815 | 888989 | 2955842200700391 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-13 | 2007-10-10 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-09-13 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the fiber glass on the shaft of the endowrist prograsp instrument was splintering. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed that the distal end of the main tube has three adjacent scrape marks with material removed. Large slivers of material have been removed so that the tube's outer diameter is no longer round at the damaged areas. Scrape marks range from. 75" to 1. 5" in length and are all roughly. 100" wide at their maximum width. All scrape marks decrease in width at the ends. Positioning and shape of scrapes suggests scrapes did not occur all at once, but rather one at a time. Damage does not appear to be consistent with cannula related tube abrasion. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 915071 | 925175 | 887345 | 2955842200700384 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-10 | 2007-10-08 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2007-09-11 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | It was reported that there is an indentation inside of the cannula accessory. No patient harm, adverse event or injury was reported. | The instrument was returned and evaluated. Engineering observed the cannula to be visibly dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. The damage to the cannula is likely due to mishandling. The deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 914992 | 925176 | 887346 | 2955842200700382 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-10 | 2007-10-08 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-09-10 | Yes | No | Unkown | 2007
42009305 |  | Malfunction |  | It was reported that a cable on the endowrist prograsp instrument was frayed. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed one grip close cable is derailed and frayed near the distal idlers. The cable derailment may have contributed to the fraying as the cable is rubbing on the outer edges of the pulley as grips move. Engineering also observed that at the distal end of the main tube, material had been removed in a couple of different locations, creating a localized rough surface finish. Damage may be due to a defective cannula. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 915062 | 925177 | 887347 | 2955842200700378 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-10 | 2007-10-08 | No | Yes | Health Professional | 42009305 | Device Returned To Manufacturer | No | 2007-09-10 | Yes | No | Reuse | 2007
42018304 | 2007-09-10 | Malfunction |  | It was reported that during the da vinci prostatectomy procedure the yellow piece on the distal tip of the permanent cautery hook instrument broke off while the surgeon was attempting to dissect tissue. The broken pieces fell inside of the patient and were retrieved. The instrument was removed and the planned surgical procedure was completed with a replacement instrument of the same type and without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned. A follow-up mdr will be submitted upon receipt and evaluation of the instrument. As indicated in the case notes, all broken instrument components were successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 915052 | 925179 | 887349 | 2955842200700386 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-10 | 2007-10-08 | No | Yes | Health Professional | 42018304 | No | No | 2007-09-11 | No | No | Unkown | 2007
42000202 |  | Malfunction |  | It was reported that there is an indentation inside of the cannula accessory. No pt harm, adverse event or injury was reported. | The instrument was returned and evaluated. Engineering observed the cannula to be visibly dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. The damage to the cannula is likely due to mishandling. The deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 914973 | 925180 | 887350 | 2955842200700381 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-10 | 2007-10-08 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-09-10 | Yes | No | Unkown | 2007
42000604 |  | Malfunction |  | It was reported that the fiber glass on the shaft of the large needle driver instrument was splintering. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed the main tube to have a section located 4 inches above the wrist with material removed around half of the outer diameter. The damaged area is roughly 2 inches in length, gray in color, and has a rough surface finish. Engineering concluded that the damage may be caused by a defective cannula. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 915041 | 925181 | 887351 | 2955842200700385 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-10-11 | 2007-10-08 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-09-11 | Yes | No | Reuse | 2007
| 2007-08-24 | Injury | Required InterventionHospitalization | I had a hysterectomy in 2007 using the davinci instrument. A screw fell out of the instrument into my abdomen which resulted in the surgery taking 3 additional hours whereas the surgeon made the decision to leave the screw in my abdomen. Since the surgery, the screw remains -6 weeks post op-, i have had 2 bouts of gastroparesis which required a re-hospitalization- gastroparesis continues at 6 weeks post op-, vaginal bleeding -continues-, left leg numbness and delayed wound healing-continues 6 weeks post op-. Dates of use: 2007. Diagnosis or reason for use: hysterectomy. |  | DA VINCI SURGICAL INSTRUMENT | NONE | 901330 | 926161 | 888335 | MW5004023 | 1 | KOG | Voluntary |  | Patient | Initial | 2007-10-04 | 2007-10-04 | Yes | Yes | Health Professional |  |  | Yes |  |  |  |  | 2007
42017908 |  | Malfunction |  | It was reported that toward the end of an appendectomy procedure, the monopolar curved scissors instrument tip cover accessory slid off the monopolar curved scissors instrument and fell into the pt. The tip cover was retrieved and placed back on the instrument to successfully complete the planned procedure without significant delay. Upon completion of the procedure, the tip cover accessory was discarded by this site. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00377. | The instrument was not returned for evaluation, therefore the root cause of the customer reported failure mode could not be determined. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 910667 | 922764 | 884958 | 2955842200700376 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-02 | 2007-10-02 | No | Yes | Health Professional | 42017908 | No | No | 2007-09-07 | No | No | Unkown | 2007
40018005 |  | Malfunction |  | It was reported that while attempting to remove the monopolar curved scissors instrument during a radical cystectomy procedure, staff noticed a hole burned through the side of the monopolar curved scissor instrument tip cover accessory. The tip cover accessory was removed and replaced to successfully complete the procedure without delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. Once the accessory has been received and failure analysis has been completed, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 910674 | 922766 | 884960 | 2955842200700375 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-10-02 | 2007-10-02 | No | Yes | Health Professional | 40018005 | No | No | 2007-09-06 | No | Yes | Initial | 2007
420179 |  | Malfunction |  | It was reported that while attempting to remove the monopolar curved scissors instrument during a radical cystectomy procedure, staff noticed a hole burned through the side of the monopolar curved scissors instrument tip cover accessory. The tip cover accessory was removed and replaced to successfully complete the procedure without delay. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore the root cause of the customer reported failure mode could not be determined. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 910669 | 922767 | 884961 | 2955842200700374 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-02 | 2007-10-02 | No | Yes | Health Professional | 420179 | No | No | 2007-09-06 | Device Not Retu | No | Initial | 2007
42018404 |  | Malfunction |  | It was reported that the flush port of the permanent cautery spatula instrument fell inside of the housing. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the luer plate was dislodged from the chassis and pushed into the back end. The chassis feature that retains the luer plate was broken at the inner wall. A horizontal crack was discovered on the back face of the chassis, next to where the luer plate seats, indicating mishandling during the cleaning process. Engineering also observed that the instrument had material removed at the distal end of the main tube, creating a rough surface finish on the tube outer diameter. This damage is all around the tube and may have been caused by a defective cannula accessory. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 895792 | 922371 | 884563 | 2955842200700371 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-10-01 | 2007-10-01 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2007-09-01 | Yes | No | Reuse | 2007
42018404 |  | Malfunction |  | It was reported that the plastic flush tube is hanging out the back of the permanent cautery spatula instrument. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that 2" of the instrument flush tube was sticking out of the luer plate. The flange on the flush tube was smaller than usual and had been flattented. The back face of the housing and a cut/indentation and a section next to the cut bluges outward, indicating that the instrument was most likely mishandled. Engineering also observed scraping all around the distal end of the main tube, creating a rough surface finish. It was determined that this damage may have been caused by a defective cannula accessory. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 895788 | 922372 | 884564 | 2955842200700372 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-10-01 | 2007-10-01 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2007-09-01 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As this site had previously returned an instrument which showed indications of cannula related tube abrasion. See medwatch mfr. Report # 2955842-2007-00255. | The accessory was returned and evaluated. Engineering observed that the cannula inner diameter had a consistent cylindrical pattern of scratching/machining marks at the tip. It was determined that the marks have high spots which have the potential to damage instrument shafts/tubes in certain loading conditions. No other damage was found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 897471 | 922373 | 884565 | 2955842200700370 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-10-01 | 2007-10-01 | No | Yes | Health Professional | 42000202 | Yes | Yes | 2007-09-01 | Yes | No | Unkown | 2007
IS2000 A51P2 | 2007-08-24 | Injury | Required Intervention | Isi received a report that during a da vinci total hysterectomy procedure, the surgeon removed his hands from the master tool manipulator (mtm); at the same time, all evidence indicates the surgeon did not remove his head from the surgeon's console, with the result that the pt side manipulator ("psm") remained engaged. Upon removal of the surgeon's hands from the mtm, the psm dropped sharply, causing the scissors instrument, which was installed on the psm to inadvertently pierce the pt's uterer. The uterer was, thereafter, stented and successfully sutured by a urologist. | Investigation was conducted by an isi field service engineer (fse). The isi fse performed system verification tests, resulting in the system passing in accordance to isi standards. No trouble was found with the system, and the isi fse was unable to duplicate the customer reported complaint. Additional investigation performed by an isi systems analyst concluded that it is highly probable that the surgeon inadvertently bumped or moved the master after he removed his hands from the control loops, or let go of the master handle with some residual velocity. Because, his head was still detected by the viewer head sensors, the system responded to the master motion appropriately by moving the slave in kind. As of september 27, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 895154 | 921167 | 883361 | 2955842200700363 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-09-28 | 2007-09-28 | Yes | No | Health Professional | IS2000 A51P2 | No | No | 2007-08-28 | Yes | No | Reuse | 2007
42021405 |  | Malfunction |  | It was reported that one of the fix handles is missing from the pk dissecting forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was evaluated. Engineering observed that the instrument has both levers installed, and no signs of damage to the housing were seen. Performance testing was conducted and the instrument moved intuitively with full range of motion in all directions. All components function properly, and the customer reported failure mode could not be confirmed. Engineering also observed that the main tube has a section located 3 inches from the wrist with material removed all around the tube. The damaged area is gray in color and has a rough surface finish. The damage may be caused by a defective cannula. Per the original reported complaint, the account did not report material being removed from the instrument or material falling into the patient. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 905715 | 921519 | 883713 | 2955842200700369 | 1 | GEI | Manufacturer | ForeignUser facility | Other | Initial | 2007-09-27 | 2007-09-28 | No | Yes | Health Professional | 42021405 | Device Returned To Manufacturer | No | 2007-08-30 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As this site had previously returned an instrument, which showed indications of cannula related tube abrasion. See medwatch mfr. Report #2955842-2007-00220. | The accessory was returned and evaluated. Engineering observed that the cannula was had a flat spot at the distal tip and the inner diameter was slightly out of round. The damage appears to have been most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr. Report #2955842-2007-00364, mfr. Report #2955842-2007-00365. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 905704 | 921520 | 883714 | 2955842200700366 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-09-28 | 2007-09-28 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-29 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As this site had previously returned an instrument, which showed indications of cannula related tube abrasion. See medwatch mfr. Report # 2955842-2007-00220. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip, and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr report # 2955842-2007-00364 and 2955842-2007-00366. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 905694 | 921521 | 883715 | 2955842200700365 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-09-28 | 2007-09-28 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-29 | Yes | No | Reuse | 2007
42017206 | 2007-08-29 | Malfunction |  | It was reported that during the middle of the da vinci prostatectomy procedure, the washer of the maryland bipolar forceps instrument fell inside of the patient's pelvic cavity while the surgeon was attempting to dissect the nerve bundle. The washer was immediately retrieved using a large needle driver instrument. The maryland bipolar forceps instrument was removed and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | He instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one of the distal idler pulleys (which the customer referred to as the "washer") slipped off of the distal clevis idler pin. The pulley was returned with the instrument. Further engineering evaluation discovered that the idler pin appears to have a non-uniform swage, with one section not mushroomed out, reducing the swage diameter of the pin. As a result, the pulley may have had room to slip out of the pin. It was also observed that the other idler pin had a similar section that was not mushroomed out, although it appears that more material was displaced in this pin. It was determined that assembly error most likely contributed to the customer reported failure mode. Per the case notes, the detached instrument component was successfully retrieved from the patient. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 947559 | 921522 | 883716 | 2955842200700368 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-09-28 | 2007-09-28 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | Yes | 2007-08-30 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned by request from intuitive surgical, inc. As this site had previously returned an instrument, which showed indications of cannula related tube abrasion. See medwatch mfr. Report #2955842-2007-00220. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip, and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the cannula accessories returned by request from intuitive surgical, inc. From this specific site: mfr. Report #2955842-2007-00365, mfr. Report #2955842-2007-00366. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 905685 | 921523 | 883717 | 2955842200700364 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-09-28 | 2007-09-28 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-29 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument would not be accepted by the system's patient side manipulator (psm). No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned with a sterile adapter that has been cut from the drape. Per the customer reported complaint, engineering conducted performance testing on both the instrument and the sterile adapter. Engagement of the instrument was successful, without difficulties installing or removing it from the sterile adapter. The engagement check was repeated on three different psm's with the same results. The instrument moved intuitively with full range of motion in all directions. All components function properly. Latex cut test and electrical continuity passed. Engineering also observed that the instrument main tube has a 2 inch long section with light scraping, located just below the midpoint of the tube. The surface finish is rougher than the remainder of the tube. The tube damage may be caused by a defective cannula. No ther damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 906616 | 921489 | 883683 | 2955842200700367 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-27 | 2007-09-27 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-08-29 | Yes | No | Reuse | 2007
Y1865 | 2007-09-10 | Malfunction |  | During the nissen fundoplication procedure, it was felt that the patient was burned from the heat emitted from the endoscope. The light source was run at 100% and it is perhaps safer to run the device at 50%. This device is specifically designed for use with the davinci s system, and is manufactured with two regular endoscopes bonded together. This was the first case performed using these instruments and therefore, running the source at 100% was not out of the ordinary. Our facility was not instructed to start running the device at a lower percentage. The user manual instructs users to set up and white balance the s system at 100%, and does not state what the illuminator should be set to during the procedure. The user manual does state however that "to minimize fogging, maintain heating of the endoscope tip by setting the illuminator brightness to 100% and using the level or iris knobs on the ccu's to adjust the brightness of the surgical image". There was no indication from the manufacturing representative, who was in the room at the time of the incident, that the system may be faulty. They later encouraged us to use the illuminator at 100% during future procedures. |  | DAVINCI S | ENDOSCOPE | 895395 | 923148 | 885342 | 923148 | 1 | KOG | User Facility |  |  | Initial | 2007-09-27 | 2007-09-27 | No | Yes | Physician | Y1865 | Yes | No Answer Provided |  |  |  |  | 2007
IS1200 A4 3P8 | 2007-08-25 | Other | Other | It was reported that during the da vinci prostatectomy procedure the customer experienced issues with a pt side manipulator (psm) where the sys recognized the instrument installed on the psm, but the psm would not enter into following mode. Multiple instruments and instrument sterile adapters were tried on the psm, however, the psm would still not enter into following mode. No sys error codes were generated for this issue. An isi field svc engineer (fse) verified with the customer via telephone that psm was assigned to the correct master tool manipulator (mtm) and requested that the customer press the emergency power off button and then the fault override button, which was unsuccessful in resolving the issue. The procedure was converted to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. | An isi fse conducted add'l investigation and evaluated the sys. Sys testing was performed and the sys passed all required tests. The functionality between the mtms and the psms worked as intended and the icon selection change feature functioned normally. The isi fse was unable to reproduce the customer reported failure mode and the root cause of the customer reported failure mode is inconclusive. As of sept. 24, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 898380 | 921732 | 883919 | 2955842200700362 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-24 | 2007-09-24 | Yes | Yes | Health Professional | IS1200 A4 3P8 | No | Yes | 2007-08-25 | No | No | Reuse | 2007
42000202 |  | Malfunction |  | While not in use, the cannula accessory was found to be damaged by an isi clinical sales representative. No pt harm, adverse event or injury was reported. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was visibly dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. The outer diameter/lip has material chipped off. Engineering concluded that the damage to cannula is likely due to mishandling. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 901716 | 918511 | 880745 | 2955842200700360 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-08-29 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00340, #2955842-2007-00341, #2955842-2007-00342. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was not bent at the distal tip. A. 342 gage pin fits through the inner diameter. There is a small step at the bowl/tube interface. The step is located on half of the inner diameter and has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr #2955842-2007-00160. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 901728 | 918512 | 880746 | 2955842200700343 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00340, #2955842-2007-00341, #2955842-2007-00343. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was visibly dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Engineering concluded that the damage to the cannula is likely due to mishandling. The deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00160. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900828 | 918513 | 880747 | 2955842200700342 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00340, #2955842-2007-00342, #2955842-2007-00343. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was not bent at the distal tip. A. 342 gage pin fits through the inner diameter. There is a small step at the bowl/tube interface. The step is located on half of the inner diameter. The step has the potential to cause scraping in certain loading condition. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00160. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900840 | 918514 | 880748 | 2955842200700341 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00344, #2955842-2007-00345. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a small step at the bowl/tube interface. The step is located on half of the inner diameter and has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900846 | 918515 | 880749 | 2955842200700346 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00341, #2955842-2007-00342, #2955842-2007-00343. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was visibly dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Engineering concluded that the damage of the cannula is likely due to mishandling. Deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00160. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900853 | 918516 | 880750 | 2955842200700340 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00347, #2955842-2007-00348, #2955842-2007-00349, #2955842-2007-00350. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was not bent at distal tip. A. 342 gage pin does not slide smoothly through the distal tip inner diameter. Inspection under magnification reveals a small deformation at the lip of the distal tip, creating a tiny raised edge on the inner diameter. The deformation has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00143. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900862 | 918517 | 880751 | 2955842200700351 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00347, 2955842-2007-00349, 2955842-2007-00350, 2955842-2007-00351. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was visibly dented at distal tip. A. 342 gage pin does not slide through the distal tip inner diameter. Inspection under magnification reveals a small deformation at the lip of the distal tip, creating a tiny raised edge on the inner diameter. The deformation has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr # 2955842-2007-00143. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900873 | 918518 | 880752 | 2955842200700348 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00347, #2955842-2007-00348, #2955842-2007-00350, #2955842-2007-00351. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was not bent at distal tip. A. 342 gage pin fits through the inner diameter. There is a small step at the bowl / tube interface. Step is located on half of the inner diameter. The step has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00143. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900883 | 918519 | 880753 | 2955842200700349 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00347, #2955842-2007-00348, #2955842-2007-00349, #2955842-2007-00351. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was not bent at distal tip. A. 342 gage pin fits through the inner diameter. There is a step located on half of the inner diameter. The step has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr# 2955842-2007-00143. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900888 | 918521 | 880755 | 2955842200700350 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00348, #2955842-2007-00349, #2955842-2007-00350, #2955842-2007-00351. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a dent at the distal tip. A. 342 gage pin does fit through the distal tip inner diameter. There is a small step at the bowl/tube interface. The step is located on half of the inner diameter and has the potential to cause scraping in certain loading conditions. No other damage found. This site has returned an instrument with a damaged main tube. See mdr # 2955842-2007-00143. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900905 | 918522 | 880756 | 2955842200700347 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00344 and 2955842-2007-00346. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was dented at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Damage to the cannula is likely due to mishandling. Deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900896 | 918523 | 880757 | 2955842200700345 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00345, #2955842-2007-00346. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a dent at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Damage to the cannula is likely due to mishandling. Deformation of the tip has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902972 | 918524 | 880758 | 2955842200700344 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-22 | Yes | No | Unkown | 2007
40015403 | 2007-08-22 | Malfunction |  | It was reported that during the da vinci myomectomy procedure, the permanent cautery spatula instrument broke while the surgeon was attempting to dissect fibroids. Three broken pieces fell inside of the pt and were immediately retrieved. The instrument was removed and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, no conclusion can be drawn regarding the root cause of the customer reported failure mode. A follow-up mdr will be submitted upon receipt and evaluation of the instrument. Per the case notes, the broken instrument components were successfully retrieved from the pt. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 901594 | 919522 | 881734 | 2955842200700352 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 40015403 | No | No | 2007-08-22 | No | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. Medwatch mfr. Report #2955842-2007-00354 was also sent to the fda. This report only pertains to the field removal action from this specific site. | The accessory was returned and evaluated. Engineering observed that the cannula is slightly out of round at the distal tip and there is a slight flat spot on the short end of the distal tip. It appears that the damage was most likely caused by user handling. There is also a shiny spot on the inner diameter below the short end of the distal tip which may possibly be a small burr and has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned an instrument with scraping. The following medwatch report was sent to the fda regarding this instrument: mfr report #2955842-2007-00142. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 901591 | 919523 | 881735 | 2955842200700353 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-23 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. Medwatch mfr. Report # 2955842-2007-00353 was also sent to the fda. This report only pertains to the field removal action from this specific site. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface which is located around all around the inner diameter. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned an instrument with scraping. The following medwatch report was sent to the fda regarding this instrument: mfr report #2955842-2007-00142. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 901677 | 919524 | 881736 | 2955842200700354 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000401 | Device Returned To Manufacturer | Yes | 2007-08-23 | Yes | No | Unkown | 2007
42009304 |  | Malfunction |  | It was reported that the endowrist prograsp instrument was broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed one grip close cable derailed at the distal idlers and that the cable was exposed on the outer pulley surface. The instrument grips still opened and closed, but movement may have been affected. One rear housing snap was also broken. Engineering also observed a 2" section on the distal end of the instrument main tube with material removed, located directly above the proximal clevis. It was determined that this damage may have been caused by defective cannula accessory. No other damage was found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 901618 | 919525 | 881737 | 2955842200700357 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | Yes | 2007-08-23 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was broken. No additional information was provided. No patient harm, adverse outcom or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed that the monopolar cautery plug was missing from the back end and the nut that retains the plug was loose inside of the chassis. Engineering also observed a section on the distal end of the instrument main tube with material removed, located directly above the tube extension. It was determined that this damage may have been caused by a defective cannula accessory. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 901679 | 919526 | 881738 | 2955842200700356 | 1 | GEI | Manufacturer | User facilityHealth Professional | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | Yes | 2007-08-23 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00334, mfr report #2955842-2007-00335. Mfr report #2955842-2007-00336, mfr report # 2955842-2007-00337, mfr report # 2955842-2007-00338. | The accessory was returned and evaluated. Engineering observed that the cannula was severely dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site was previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902402 | 919527 | 881739 | 2955842200700339 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00334, mfr report #2955842-2007-00335. Mfr report #2955842-2007-00336, mfr report # 2955842-2007-00337, mfr report # 2955842-2007-00339. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. Engineering also observed two additional deformed sections on the lip of the cannula tip, however, the damage is restricted to the outer diameter and would not contribute to instrument scraping. No other damage was found. The site was previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902500 | 919528 | 881740 | 2955842200700338 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00334, mfr report #2955842-2007-00335. Mfr report #2955842-2007-00336, mfr. Report # 2955842-2007-00338, mfr. Report # 2955842-2007-00339. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902391 | 919529 | 881741 | 2955842200700337 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00334, mfr report #2955842-2007-00335. Mfr report #2955842-2007-00337, mfr. Report # 2955842-2007-00338, mfr. Report # 2955842-2007-00339. | The accessory was returned and evaluated. Engineering observed that the cannula distal tip appears to be slightly out of round, indicating possible mishandling. Engineering also discovered a small deformation at the lip of the distal tip, creating a tiny raised edge on the inner diameter. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902475 | 919531 | 881743 | 2955842200700336 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00334, mfr report #2955842-2007-00336. Mfr report #2955842-2007-00337, mfr. Report # 2955842-2007-00338, mfr report # 2955842-2007-00339. | The accessory was returned and evaluated. Engineering discovered a small deformation at the lip of the distal tip, creating a tiny raised edge on the inner diameter. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902510 | 919532 | 881744 | 2955842200700335 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42017904 |  | Malfunction |  | It was reported that the tip of the monopolar curved scissors instrument was broken and the break did not occur inside of the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed that the instrument tube extension was dislodged from the main tube. The keys on the tube extension were broken off, and a result, the tube extension can be rotated 360 degrees. Engineering also observed a 6" section with material removed on the outer diameter of the instrument main tube. It was determined that this damage may have been caused by a defective cannula accessory. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 901636 | 919533 | 881745 | 2955842200700355 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | Yes | 2007-08-23 | Yes | No | Reuse | 2007
40018005 | 2007-06-22 | Malfunction |  | It was reported that during the da vinci total hysterectomy procedure, it was noticed after the monopolar curved scissors instrument was taken out of the patient that a hole had burnt through the cleear protective cover of the instrument. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. Medwatch mfr. Report #2955842-2007-00330, which is related to this event, was also sent to the fda. | The instrument was not returned for evaluation, however, the tip cover accessory used in conjunction with the instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed a. 030" diameter hole -. 25" above the plastic sleeve. The hole had material removed on the outer surface of the silicone, indicating that the damage was heat/energy related. Just below the hole there are some partial tears/scratches on the inner diameter of the silicone which may have come from the instrument blades. Engineering also discovered a tear at the distal tip of the silicone which is not heat related. The silicone may have had prior damage or inherent defects that allowed energy to escape at the hole sites. Additionally, the plastic sleeve is cracked along its entire length and in two additional locations, however, it was determined that this damage most likely occured during shipping and /or handling. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 901675 | 919534 | 881746 | 2955842200700331 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 40018005 | Device Returned To Manufacturer | No | 2007-06-22 | Yes | Yes | Initial | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr report # 2955842-2007-00335, mfr report #2955842-2007-00336. Mft report #2955842-2007-00337, mft. Report # 2955842-2007-00338, mfr. Report # 2955842-2007-00339. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly dented at the distal tip and it appears that the damage was most likely caused by user handling. It was determined that the tip deformation has the potential to cause instrument scraping in certain loading conditions. No other damage was found. The site has previously returned scraped instruments where it was determined the scraping was most likely caused by a potentially defective cannula accessory. The following medwatch reports were sent to the fda regarding these instruments. Mfr report #2955842-2007-00040. Mfr report # 2955842-2007-00220. Mfr report# 2955842-2007-00183. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 902374 | 919536 | 881748 | 2955842200700334 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-08-22 | Yes | No | Unkown | 2007
42009304 | 2007-06-11 | Malfunction |  | It was reporeted that the tip of the endowrist prograsp instrument broke and pieces had fallen into the patient. All broken pieces were removed and the planned surgical procedure was completed. No furhter details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported complaint cannot be determined. A follow-up mdr will be submitted upon return and evaluation of the instrument. Per the complaint notes, all broken pieces were successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 901641 | 919537 | 881749 | 2955842200700358 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 42009304 | No | No | 2007-08-23 | No | No | Unkown | 2007
400179 | 2007-06-22 | Malfunction |  | It was reported that during the da vinci total hysterectomy procedure, it was noticed after the monopolar curved scissors instrument was taken out of the patient that a hole had burnt through the clear protective cover of the instrument. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. Medwatch mfr. Report #2955842-2007-00331, which is related to this event, was also sent to the fda. | The instrument was not returned for evaluation, however, the tip cover accessory used in conjunction with the instrument was returned and evaluated. See medwatch mfr. Report #2955842-2007-00331 for the engineering evaluation results of the 8mm mcs tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 901628 | 919538 | 846391 | 2955842200700330 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-21 | 2007-09-21 | No | Yes | Health Professional | 400179 | No | No | 2007-06-22 | No | No | Unkown | 2007
42017907 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed that the monopolar cautery plug was missing from the back end, and the nut that retains the plug was loose inside of the chassis causing the plug to become unscrewed. Engineering also observed a section on the distal end of the instrument main tube with material removed all around the tube. The damaged area is gray in color and has a rough surface finish. Damage may be caused by a defective cannula. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 901393 | 919685 | 881899 | 2955842200700361 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | UNKNOWN | Initial | 2007-09-24 | 2007-09-21 | No | Yes | Health Professional | 42017907 | Device Returned To Manufacturer | No | 2007-08-25 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-08-24 | Other | Required Intervention | It was reported that during a procedure, the left camera controller lost power. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system power loss failure experienced by the customer was associated with the camera, left camera controller unit, and the multiple servo driver, pca. The system was repaired by replacing the affected components listed above. As of september 20, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 909782 | 921722 | 883909 | 2955842200700359 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-09-20 | 2007-09-21 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-08-24 | Yes | No | Unkown | 2007
42017204 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation discovered one broken conductor wire at the yaw pulley exit. The yaw pulley was missing a piece opposite of the conductor break, which allowed the grip to move within the pulley and have a larger range of yaw motion. The added range of motion of the grip most likely created excess tension on the wire, thus causing it to break. Per the complaint record, there was no indication that the instrument damage occurred during a surgical procedure. Engineering also observed a section with material removed near the distal end of the instrument main tube. The tube had a localized rough surface finish covering the entire outer diameter. It was determined that this damage may have been caused by the defective cannula accessory. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968823 | 919607 | 881821 | 2955842200700333 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-20 | 2007-09-20 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-08-21 | Yes | No | Reuse | 2007
42017908 | 2007-08-09 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, half of the monopolar curved scissors instrument tip broke off and fell inside of the patient while the surgeon was attempting to cut tissue. The instrument had been in use for approx 30 minutes when the issue occurred. The broken tip was retrieved and the planned surgical procedure was completed without significant delay. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with one of the blades broken off at the cross-section of the grip cable crimp. The broken blade was also returned. Engineering evaluation observed that the break plane was nearly flat and the fractured blade had pitting on the non-cutting edge, indicating possible corrosion. Per the case notes, the broken instrument component was successfully retrieved from the patient. Engineering also observed two sections on the distal end of the instrument main tube with material removed. The distal end of the tube extension also had a crack at one of the keys with a small piece missing next to the crack. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 968647 | 919608 | 881822 | 2955842200700299 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-20 | 2007-09-20 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-08-20 | Other | Other | It was reported that during the da vinci prostatectomy procedure, the customer experienced a system error code #23. The system was restarted, however, the same system error code recurred. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error code #23 experienced by the customer was a hardware communication fault associated with a vision input output (vio) pca. The system was repaired by replacing the affected vio pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of september 18, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 899313 | 918048 | 880278 | 2955842200700332 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-19 | 2007-09-19 | Yes | Yes | Health Professional | IS2000 A51P2 | Device Returned To Manufacturer | No | 2007-08-20 | Yes | No | Reuse | 2007
400180 |  | Malfunction |  | It was reported that toward the end of an appendectomy procedure, the monopolar curved scissors instrument tip cover accessory slid off the monopolar curved scissors instrument and fell into the pt. The tip cover was retrieved and placed back on the instrument to successfully complete the planned procedure without significant delay. Upon completion of the procedure, the tip cover accessory was discarded by this site. No pt harm, adverse outcome or injury was reported. The following medwatch report listed below was also sent to the fda as it related to this event. #2955842-2007-00376. | The accessory was not returned for evaluation, therefore the root cause of the customer reported failure mode could not be determined. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTURMENT ACC | 910647 | 921908 | 884090 | 2955842200700377 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-10-02 | 2007-09-19 | No | Yes | Health Professional | 400180 | No | No | 2007-09-07 | No | Yes | Initial | 2007
40018005 | 2007-06-07 | Malfunction |  | It was reported that during the da vinci total hysterectomy procedure, it was noticed during suturing and as the monopolar curved scissors instrument was being cleaned that the monopolar curved scissors instrument tip cover accessory was charred on the gray area and had three small holes on the plastic portion. No burn or char marks on the patient were observed. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed without significant delay. The site indicated that they would not be returning the monopolar curved scissors instrument for evaluation as the instrument was not set aside and they could not determine which instrument was used during the surgical procedure. The site also indicated and that they felt that the instrument did not contribute to the tip cover accessory damage. Medwatch mfr. Report #2955842-2007-00327, which is related to this event, was also sent to the fda. | The case notes indicated that the instrument would not be returned; however, the tip cover accessory used in conjunction with the instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation discovered a. 040" hole at the silicone to sleeve interface, located between the silicon parting lines. Charring was observed on the plastic sleeve below the hole. Engineering also found three similar sized holes and a couple of tears in the silicone. 25" above the interface and near one of the parting lines. All holes have material removed on the outer surface of the silicone, indicating that the damage is heat/energy related. The silicone may have had prior damage or inherent defects that allowed energy to escape at the hole sites. Additionally, the plastic sleeve is cracked along its entire length, however, it was determined that this damage most likely occurred during shipping and/or handling. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 898624 | 915948 | 878182 | 2955842200700328 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-18 | 2007-09-18 | No | Yes | Health Professional | 40018005 | Yes | No | 2007-06-08 | Yes | Yes | Initial | 2007
420179 | 2007-06-07 | Malfunction |  | It was reported that during the da vinci total hysterectomy procedure, it was noticed during suturing and as the monopolar curved scissors instrument was being cleaned that the monopolar curved scissors instrument tip cover accessory was charred on the gray area and had three small holes on the plastic portion. No burn or char marks on the patient were observed. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed without significant delay. The site indicated that they would not be returning the monopolar curved scissors instrument for evaluation as the instrument was not set aside and they could not determine which instrument was used during the surgical procedure. The site also indicated and that they felt that the instrument did not contribute to the tip cover accessory damage. Medwatch mfr. Report #2955842-2007-00328, which is related to this event, was also sent to the fda. | The case notes indicated that the instrument would not be returned, however, the tip cover accessory used in conjunction with the instrument was returned and evaluated. See medwatch mfr. Report #2955842-2007-00328 for the engineering evaluation results of the 8mm mcs tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 898633 | 915949 | 838161 | 2955842200700327 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-18 | 2007-09-18 | No | Yes | Health Professional | 420179 | No | No | 2007-06-08 | No | No | Unkown | 2007
IS2000 A51P2 | 2007-08-20 | Other | Other | It was reported that during the surgical procedure, the customer experienced a 23013 system error code which could not be overridden. The procedure was aborted before the da vinci s surgical system could be used but after the patient was under anesthesia for 30 minutes and after the port incisions had been created. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designated and there was no injury to the patient. The 23013 system error code appeared when the davinci s safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci s in a "recoverable safe state". The ecm returned to isi for failure analysis investigation. Engineering discovered that the potentiometer assembly had malfunctioned, thus generating the system errors experienced by the customer. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 927652 | 915993 | 880278 | 2955842200700329 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-18 | 2007-09-18 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-08-20 | Yes | No | Reuse | 2007
40019002 |  | Malfunction |  | It was reported that the edges of the tip of the cobra grasper instrument were rough. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not observe any rough edges on the instrument wrist components (which includes the tip), therefore, the customer reported complaint cannot be confirmed. Engineering also observed a single deep scratch on the instrument main tube, located 1 inch above the wrist. The scratch had removed some material and based on the location of the scratch, engineering determined that the scratch was most likely caused by a collision with another instrument. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890721 | 912971 | 875204 | 2955842200700326 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-17 | 2007-09-17 | No | Yes | Health Professional | 40019002 | Yes | No | 2007-08-19 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the tip of the monopolar curved scissors instrument wouldn't open fully. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00323 and #2955842-2006-00324. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and confirmed that material becomes pinched at the instrument tips when trying to cut. Visual inspection revealed that one blade edge is chipped at the base of the blade and the blade edges were slightly rounded at the tips, which negatively affected cutting performance. Engineering also observed that the instrument main tube had various light scratches at the distal end and veery light abrasion was also noticed at the midpoint of the tube. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 898083 | 916344 | 878578 | 2955842200700322 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-09-17 | 2007-09-17 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-08-18 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the tip of the monopolar curved scissors instrument wouldn't open fully. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr report # 2955842-2006-00322, mfr report# 2955842-2006-00323. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and observed that the instrument moved intuitively with full range of motion in all directions and was able to cut per specification. The customer reported failure mode could not be confirmed. Engineering also observed a section on the distal end of the instrument main tube with material removed. The damage is located 3. 5" above the tube extension. It was determined that this failure mode may have been caused by the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 902832 | 916346 | 878580 | 2955842200700324 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-17 | 2007-09-17 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2007-08-18 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the tip of the monopolar curved scissors instrument wouldn't open fully. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related compliants used at the same facility. Mfr. Report #2955842-2006-00322 & #2955842-2006-00324. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and observed that the instrument moved intuitively with full range of motion in all directions and was able to cut per specification. The customer reported failure mode could not be confirmed. Engineering also observed a 4" section on the distal end of the instrument main tube with material removed all around the out diameter. The damage is located 2" above the tube extension. It was determined that this failure mode may have been caused by the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 935422 | 916347 | 878581 | 2955842200700323 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-09-17 | 2007-09-17 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-08-18 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the cautery plug missing from the back end. The nut that retains the plug is loose in the chassis resulting in the plug becoming unscrewed. Electrical continuity passed. Engineering also observed that material had been removed from the outer diameter of the main tube, located three inches above the tube extension. The tube extension also has a hairline crack between the keys at the distal end. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 902839 | 916348 | 878582 | 2955842200700325 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-17 | 2007-09-17 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-08-18 | Yes | No | Reuse | 2007
COMPANY REFERENCE 420143 | 2007-09-10 | No Answer Provided |  | Unidentified black pieces seen laying in patient's open abdominal wound during robotic surgical procedure. Unclear as to source of debris. All other surgical instruments inspected after surgery with no source identified. |  | DAVINCI S MARYLAND DISSECTOR | INSTRUMENT SURGICAL ROBOTIC | 894950 | 922619 | 884811 | 922619 | 1 | NAY | User Facility |  |  | Initial | 2007-09-13 | 2007-09-14 | No | No | Physician | COMPANY REFERENCE 420143 | No | No Answer Provided |  |  |  |  | 2007
42009305 | 2007-08-14 | Malfunction |  | It was reported that during a prostatectomy procedure, the wrist of the endowrist prograsp instrument broke resulting in a piece falling inside the patient. The piece was successfully retrieved using standard laparoscopic instruments and thrown away at the site. The planned surgical procedure was completed without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned for eval. Per the customer reported complaint, engineering observed that the instrument was broke at the proximal clevis to main tube interface, resulting in the clevis becoming dislodged from main tube. Both the proximal clevis and the main tube are missing pieces. The proximal clevis is missing a piece below the ear and also at the bottom section between the ears. The entire main tube interface wall appears to have been cut off, likely to aid in removal of instrument from cannula. Engineering concluded that the instrument may have been overloaded at the tip. No other damage found. Per the case notes, the broke piece was successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 893958 | 914324 | 876558 | 2955842200700320 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2007-09-13 | 2007-09-13 | No | Yes | Health Professional | 42009305 | Device Returned To Manufacturer | Yes | 2007-08-14 | Yes | No | Reuse | 2007
40016906 | 2007-08-06 | Malfunction | Other | It was reported that during a hysterectomy procedure, the harmonic curved shears insert broke and pieces fell inside of the pt. The broken pieces were retrieved and the procedure was completed with a replacement insert. The procedure was lengthened by 20 to 30 minutes and it is unclear if all pieces were retrieved. No pt harm, adverse outcome or injury was reported. | The instrument accessory has not yet been returned for evaluation. Once the instrument accessory is returned and evaluated, a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERT 5 PACK | ENDOSCOPIC ACCESSORY | 893944 | 914330 | 876564 | 2955842200700318 | 1 | LFL | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-13 | 2007-09-13 | Yes | Yes | Health Professional | 40016906 | No | No | 2007-08-16 | No | Yes | Initial | 2007
40017908 | 2007-08-16 | Malfunction |  | It was reported that during a prostatectomy procedure the tip of the monopolar curved scissors instrument broke off and fell into the patient. The component was retrieved and the instrument was exchanged with an instrument of the same type to successfully complete the procedure without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one of the instrument blades to be broken off at the cross-section of the grip cable crimp. The broken blade was returned and it was observed that the break plane of the blade is not totally flat, as there are a couple of changes in elevation along with a crack at the crimp pocket. Both blades show signs of pitting on the non-cutting edge, indicating possible corrosion. Per the case notes, the broken piece was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 893935 | 914332 | 876566 | 2955842200700319 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-09-04 | 2007-09-13 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | Yes | 2007-08-16 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation on a instrument returned with a damaged ceramic sleeve from this site. Reference mdr# 2955842-2007-00228. | The accessory was returned and evaluated. Engineering observed chips/burrs on the inner diameter near the inside lip. Damage to inside lip may create sharp edges which have the potential to cause scraping in certain loading conditions. Engineering concluded that the damage to the inner lip appears to be due to mishandling. Engineering also concluded that it is not likely that the burrs in the cannula contributed to the damaged ceramic sleeve in mdr# 2955842-2007-00228, as the cannula was dimensionally in spec. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 893924 | 914335 | 876569 | 2955842200700314 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-13 | 2007-09-13 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-15 | Yes | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation on two instruments returned from this site with tube damage. Reference the following mdr's. 2955842-2007-00206, 2955842-2007-00243. | The cannula accessory was evaluated. Engineering observed a dent on the long side of the distal tip of the cannula. Engineering concluded that the dent was likely caused by user mishandling. Engineering also observed a burr on the inner diameter of the distal tip which has the potential to cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 893913 | 914337 | 876571 | 2955842200700315 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-13 | 2007-09-13 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-08-15 | Yes | No | Unkown | 2007
IS2000 A51P2 | 2007-08-16 | Other | Other | It was reported that during the surgical procedure, the customer experienced a non-recoverable #23 system error code which could not be overridden. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error code #23 experienced by the customer was a hardware communication fault associated with a setup joint pca or setup joint cable harness inside of the patient side manipulator. The system was repaired by replacing the affected setup joint pca and setup joint cable harness. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of september 13, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 894499 | 914418 | 876652 | 2955842200700317 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-13 | 2007-09-13 | Yes | Yes | Health Professional | IS2000 A51P2 | No | No | 2007-08-16 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-08-14 | Other | Other | During the docking of the da vinci surgical system, the site experienced a 23008 system error code. With the support of an isi regional support specialist the site troubleshot the system and cleared the fault. They continued with the planned surgical procedure and the error re-appeared resulting in the surgeon converting the procedure to open surgical techniques for completion. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The system alarm (the 23008 system error code) functioned as designed and there was no injury to the patient. The 23008 system error code appears when the da vinci s safety systems determined that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining, this condition, the safety systems put da vinci s in a "recoverable safe state". This check may result from a failed encoder, the failed potentiometer, a broken link or momentary impact on the arm by a customer (e. G. , bumping the arm). The mtml was returned to isi for failure analysis investigation. Engineering discovered a defective potentiometer assembly within the mtml, thus generating the system error experienced by the customer. As of september 13, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 905870 | 914419 | 876652 | 2955842200700321 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-13 | 2007-09-13 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-08-14 | Yes | No | Reuse | 2007
400169 | 2007-08-30 | Malfunction |  | A harmonic curved shears insert disposable accessory was being used for the da vinci surgical system robotic assisted procedure. During the procedure, one of the tips broke off during use. Both the large intact piece and the single broken piece were found and removed. The patient was not harmed. A new harmonic curved shears insert was obtained and used for the remainder of the procedure. The company representative was contacted and product will be returned to company for analysis. |  | DA VINCI ENDOWRIST | ROBOTIC DEVICE CURVED SHEARS | 899705 | 927660 | 889814 | 927660 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-09-06 | 2007-09-06 | No | Yes | Invalid Data | 400169 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2007
PS2000 | 2007-07-16 | Injury | Required Intervention | Patient was undergoing a robotic mitral valve repair. Ports were placed in the 2nd, 4th, and 6th intercostal space, a working port in the 4th intercostal space. Additional ports for the clamp and pericardial stay sutures were placed. After completing the repair of the mitral valve, the left atrium was then closed using suture. However, upon attempting to close the superior margin of the wound, it became obvious that the left atrial wall had been disrupted. Because of requirements for better visualization of this problem, a clamp was then introduced and the aorta was cross clamped. Cardioplegia was then given with good cardioplegic arrest. Examination of left atrium showed a tear which had occurred in the hood of the atrium medial to the left atrial appendage extending down towards the mitral annulus approximately 3 cm in length. This was a through-and-through tear. Therefore, a closure was made using a suture in a running fashion, double layer, and then tied with the robot. An excellent closure could be seen from the inside and the left atriotomy appeared, because of this closure, not to be adequate for a simple running closure. A patch was then brought into place, trimmed, and this was then sewn using a suture and tied. The aortic cross clamp was then removed using the aortic vent to remove any co2 and good closure of the patch was noted. However, there was sufficient bleeding that good visualization from the limited access was not adequate. Therefore the robot was undocked and removed. A sternotomy incision was then made for further repair. It is possible that the port placement effected the event. The robotic arms were never seen to collide but this could have occurred which resulted in pressure on the retractor. |  | DA VINCI S | ROBOTIC DEVICE ROBOT | 904384 | 932174 | 894316 | 932174 | 1 | NAY | User Facility |  | Nurse | Initial | 2007-09-06 | 2007-09-06 | Yes | No | Physician | PS2000 | Yes | Yes |  |  |  |  | 2007
PS2000 | 2007-07-16 | Injury | Required Intervention | Patient was undergoing a robotic mitral valve repair. Ports were placed in the 2nd, 4th, and 6th intercostal space, a working port in the 4th intercostal space. Additional ports for the clamp and pericardial stay sutures were placed. After completing the repair of the mitral valve, the left atrium was then closed using suture. However, upon attempting to close the superior margin of the wound, it became obvious that the left atrial wall had been disrupted. Because of requirements for better visualization of this problem, a clamp was then introduced and the aorta was cross clamped. Cardioplegia was then given with good cardioplegic arrest. Examination of left atrium showed a tear which had occurred in the hood of the atrium medial to the left atrial appendage extending down towards the mitral annulus approximately 3 cm in length. This was a through-and-through tear. Therefore, a closure was made using a suture in a running fashion, double layer, and then tied with the robot. An excellent closure could be seen from the inside and the left atriotomy appeared, because of this closure, not to be adequate for a simple running closure. A patch was then brought into place, trimmed, and this was then sewn using a suture and tied. The aortic cross clamp was then removed using the aortic vent to remove any co2 and good closure of the patch was noted. However, there was sufficient bleeding that good visualization from the limited access was not adequate. Therefore the robot was undocked and removed. A sternotomy incision was then made for further repair. It is possible that the port placement effected the event. The robotic arms were never seen to collide but this could have occurred which resulted in pressure on the retractor. |  | DA VINCI S | ROBOTIC DEVICE ROBOT | 904384 | 932174 | 894316 | 932174 | 1 | NAY | User Facility |  | Nurse | Initial | 2007-09-06 | 2007-09-06 | Yes | No | Physician | PS2000 | Yes | Yes |  |  |  |  | 2007
42011004 |  | Malfunction |  | It was reported that after sterilization of the precise bipolar pyramid tip forceps instrument it was observed that the tip of the instrument was bent. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument broken at the proximal clevis to main tube interface resulting in the clevis to become dislodged from main tube. The clevis is intact, however, the main tube interface wall has collapsed, possibly due to overloading at the tip. One grip is bent slightly and one conductor wire is broken at the yaw pulley exit. Electrical continuity failed. Engineering also observed the main tube to have a couple of sections with deep scratches and scraping. Damage occurs from the midpoint to the distal end of the tube. Material has been removed from the tube. Engineering concluded that damage is possibly caused by a defective cannula. No other damage found. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INST | ELECTROSURGICAL INSTRUMENT | 900253 | 908767 | 871001 | 2955842200700297 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42011004 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation related to an instrument returned form this site with cannula related tube abrasions. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the instruments and cannulae returned from this specific site: 2955842-2007-00300, 00301, 00303, 00304, 00305, 00306, 00307, 00308, 00309, 00310, 00311. | The cannula accessory was evaluated. Engineering inspection of the cannula found the inner diameter to have a consistent cylindrical pattern of scratching/machining marks at the tip. The surface finish appears rough as a result and marks may have high spots that have the potential to damage instrument shafts in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 899513 | 908768 | 871002 | 2955842200700302 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-09 | Yes | No | Unkown | 2007
42020702 |  | Malfunction |  | It was reported that after sterilization of the tenaculum forceps instrument it was observed that the tip of the instrument had broken off. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument ws returned and evaluated. Per the customer reported complaint, engineering found the wrist detached and the main tube interface wall cut off. The proximal clevis is missing pieces all around the mating section. Engineering concluded that the likely failure mode is a break at the tube to proximal clevis interface resulting in the site cutting the wrist to remove the instrument from the cannula. Engineering also found the distal end of the main tube to have scraping, with material removed. Damage may be due to a defective cannula. No other damage found. This site has previously returned defective cannula. Referenced below are the previously reported cannula related mdr's for this site. 2955842-2007-00084, 00085, 00086, 00087, 00088, 00089, 00090, 00091. | TENECULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 900262 | 908769 | 871003 | 2955842200700298 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42020702 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation related to instruments returned from this site with cannula related tube abrasions. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the instruments and cannulae returned from this specific site: 2955842-2007-00301, 00302, 00303, 00304, 00305, 00306, 00307, 00308, 00309, 00310, 00311. | The cannula accessory was evaluated. Engineering inspection of the cannula under magnification found a deformation on the long side of the distal tip. Only the outer diameter and the top of the lip are deformed, however, the impact from the damage may have caused an out of round condition that may cause scraping in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 899520 | 908771 | 871005 | 2955842200700300 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000202 | Device Returned To Manufacturer | No | 2007-08-09 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: mfr report # 2955842-2007-00289, 2955842-2007-00290. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly deformed at the distal tip and a burr was discovered on the inner diameter of the tip. Engineering also observed a small ledge on the inner diameter of the cannula at the bowl/tube interface. It was determined that the tip deformation, burr and ledge have the potential to cause instrument scraping in certain loading conditions. Site had previously returned a scraped instrument, see mfr report # 2955842-2007-00250. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 897171 | 908773 | 871007 | 2955842200700291 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-08-06 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: mfr. Report# 2955842-2007-00290, mfr. Report# 2955842-2007-00291. | The accessory was returned and evaluated. Engineering observed a small ledge on the inner diameter of the cannula at the bowl/tube interface. It was determined that the ledge has the potential to cause instrument scraping in certain loading conditions. Site had previously returned a scraped instrument, see mfr. Report #2955842-2007-00250. | ICANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 893632 | 908774 | 871008 | 2955842200700289 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-08-06 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: mfr. Report# 2955842-2007-00289, mfr. Report# 2955842-2007-00291. | The accessory was returned and evaluated. Engineering observed that the cannula was visibly deformed at the distal tip. Engineering also observed a small ledge on the inner diameter of the cannula at the bowl/tube interface. It was determined that the tip deformation and ledge have the potential to cause instrument scraping in certain loading conditions. Site had previously returned a scraped instrument, see mfr. Report #2955842-2007-00250. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 893622 | 908775 | 871009 | 2955842200700290 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-08-06 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument did not work. No further information was available. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the main tube at the distal end, above the main tube interface indicating that arcing had occurred during a previous surgical procedure. The hypotubes have charred debris at approx the same location as the hole. The main tube has a crack running from the hole to the distal end of the tube. The crack is offset from one of the slots. There is another crack adjacent to the hole/crack combo that lines up with the edge of one of the slots. Engineering concluded that the crack may have been inherent in the tube and possibly propagated with overloading and likely created a path for arcing and allowing electrical energy to escape through the main tube. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 893613 | 908776 | 871010 | 2955842200700292 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2007-08-06 | Yes | No | Reuse | 2007
42018304 |  | Malfunction |  | It was reported that the permanent cautery hook instrument did not work. No further information was available. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring as the distal clevis on the outer face of one ear and at the hub indicating that arcing had occurred during a previous surgical procedure. The conductor wire has cut insulation next to the hub, which coincides with the most severe localized charring. Engineering concluded that the damage to the conductor wire likely created a path for arcing. Engineering also observed the grip cable on the same side as the conductor wire had derailed at the distal idler pulley. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 893604 | 908777 | 871011 | 2955842200700293 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42018304 | Yes | No | 2007-08-06 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the large needle driver instrument had abrasions on its main tube, which were caused by a defective cannula. No add'l info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report# 2955842-2006-00306, -00307, -00308, -00309, and -0310. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the distal end of the instrument main tube with material removed, approx 1" above the proximal clevis. About 90 degrees from this section, engineering observed another section where material had been removed. It was determined that the damage to the main tube was most likely caused by a defective cannula accessory. No other damage was found. The following medwatch reports were created for the defective cannula previously returned from this site: mfr. Report# 2955842-2006-00300, mfr. Report# 2955842-2006-00301, mfr. Report# 2955842-2006-00302, mfr. Report# 2955842-2006-00303, mfr. Report# 2955842-2006-00304, mfr. Report# 2955842-2006-00305. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 891972 | 909215 | 871449 | 2955842200700311 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-08-09 | Yes | No | Initial | 2007
42017908 |  | Malfunction |  | It was reported that the large needle driver instrument had abrasions on its main tube, which were caused by a defective cannula. No add'l info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report# 2955842-2006-00306, -00307, -00308, -00309, and -0311. | The instrument was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed nor denied. The following medwatch reports were created for the defective cannulae previously returned from this site: mfr. Report# 2955842-2006-00300, -00301, -00302, -00303, -00304, and -00305. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 891965 | 909216 | 871450 | 2955842200700310 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42017908 | No | No | 2007-08-09 | Device Not Retu | No | Unkown | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument had abrasions on its main tube, which were caused by a defective cannula. No add'l info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report# 2955842-2006-00306, -00307, -00308, -00310, and -0311. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a deep scratch on the distal end of the instrument main tube with a small amount of material removed, directly above the proximal clevis. A couple of light scratches were also found at the distal end of the main tube, however, no material had been removed. Based on the orientation of the scratches, it was determined that the damage to the main tube was most likely caused by inadvertent contact with another instrument and not by a defective cannula. No other damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 891956 | 909217 | 871451 | 2955842200700309 | 1 | KOG | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00306, mfr. Report #2955842-2006-00307, mfr. Report #2955842-2006-00309, mfr. Report #2955842-2006-00310, mfr. Report #2955842-2006-00311. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on one side of the distal end of the instrument main tube with material removed, approximately. 5" above the tube extension. It was determined that the damage to the main tube was most likely caused by a defective cannula accessory. The following medwatch reports were created for the defective cannulae previously returned from this site: mfr report #2955842-2006-00300, mfr report #2955842-2006-00301, mfr report #2955842-2006-00302, mfr report #2955842-2006-00303, mfr report #2955842-2006-00304, mfr report #2955842-2006-00305. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 900259 | 909218 | 871452 | 2955842200700308 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the bipolar marylabnd forceps instrument had a broken cautery connector. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the top plug riser pin was pushed into the back end and the chassis feature that retains the pins was broken. During eval, engineering also observed deep scratches on the distal end of the instrument main tube, above the proximal clevis. Some material had been removed as a result of the scratches. No other damage was found. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 889886 | 909219 | 871453 | 2955842200700296 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-08-08 | Yes | No | Reuse | 2007
40018402 | 2007-07-03 | Malfunction |  | It was reported that during the surgical procedure, the plastic on the top of the permanent cautery spatula instrument broke. A broken piece fell inside of the patient's thorax and was retrieved. The instrument was removed and the planned surgical procedure was completed with another instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument ceramic sleeve was missing a piece at the distal end where the outer diameter is smaller. The instrument had heavy scratch marks on both sides, indicating aggressive cleaning techniques which may have contributed to the instrument failure mode. No other damage was found. Per the case notes, the broken instrument component was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 897163 | 909220 | 871454 | 2955842200700294 | 1 | GEI | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 40018402 | Yes | No | 2007-08-06 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Reports: #2955842-2006-00306, 2955842-2006-00308, 2955842-2006-00309, 2955842-2006-00310, 2955842-2006-00311. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the distal end of the instrument main tube with material removed, approximately 2" above the proximal clevis. It was determined that this damage to the main tube was most likely caused by a defective cannula accessory. Engineering also observed a deep scratch approx. 5" above the proximal clevis. It was determined that the scratch was most likely caused by a collision with another instrument. No other damage was found. The following medwatch reports were created for the defective cannulae previously returned from this site: 2955842-2006-000300, 2955842-2006-000301, 2955842-2006-000302, 2955842-2006-3003,2955842-2006-000304, 2955842-2006-000305. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 897155 | 909221 | 871455 | 2955842200700307 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation related to an instrument returned from this site with cannula related tube abrasions. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the instruments and cannulae returned from this specific site: 2955842-2007-000300, 2955842-2007-000301, 2955842-2007-000302, 2955842-2007-000303, 2955842-2007-000304, 2955842-2007-000305, 2955842-2007-000306, 2955842-2007-000307, 2955842-2007-000308, 2955842-2007-000309, 2955842-2007-000310, 2955842-2007-000311. | The cannula accessory was evaluated. Engineering inspection of the cannula found the inner diameter to have a consistent cylindrical pattern of scratching/machining marks at the tip. The surface finish appears rough as a result and marks may have high spots that have the potential to damage instrument shafts in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 897134 | 909223 | 871457 | 2955842200700305 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-08-09 | Yes | No | Unkown | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation related to an instrument returned from this site with cannula related tube abrasions. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the instruments and cannula returned from this specific site: 2955842-2007-000300, 2955842-2007-000301, 2955842-2007-000302, 2955842-2007-3003, 2955842-2007-305, 2955842-2007-306, 2955842-2007-307, 2955842-2007-308, 2955842-2007-309, 2955842-2007-310, 2955842-2007-000311. | The cannula accessory was evaluated. Engineering inspection of the cannula found the inner diameter to have a consistent cylindrical pattern of scratching/machining marks at the tip. The surface finish appears rough as a result and marks may have high spots that have the potential to damage instrument shafts in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 900194 | 909224 | 871458 | 2955842200700304 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-08-09 | Yes | No | Unkown | 2007
42017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument had abrasions on its main tube, which were caused by a defective cannula. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00307, mfr. Report #2955842-2006-00308, mfr. Report #2955842-2006-00309,mfr. Report #2955842-2006-00310,mfr. Report #2955842-2006-00311. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed two sections on the distal end of the instrument main tube with material removed, directly above the proximal clevis. Engineering also observed another section approximately three inches above the proximal clevis where material had also been removed. It was determined that the abrasions to the main tube were most likely caused by a defective cannula accessory. No other damage was found. The site has previously returned defective cannulae and the following medwatch reports were created for the defective cannulae: mfr. Report #2955842-2006-00300, mfr. Report #2955842-2006-00301,mfr. Report #2955842-2006-00302,mfr. Report #2955842-2006-00303,mfr. Report #2955842-2006-00304,mfr. Report #2955842-2006-00305. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 899746 | 909225 | 871459 | 2955842200700306 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-08-09 | Yes | No | Reuse | 2007
42000202 |  | Malfunction |  | The cannula accessory was returned as part of an investigation related to an instrument returned from this site with cannula related tube abrasions. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the instruments and cannulae returned from this specific site: 2955842-2007-000300, 2955842-2007-000301, 2955842-2007-000302, 2955842-2007-0003004, 2955842-2007-000305, 2955842-2007-000306, 2955842-2007-000307, 2955842-2007-000308, 2955842-2007-300009, 2955842-2007-000310. | The cannula accessory was evaluated. Engineering inspection of the cannula found the inner diameter to have a consistent cylindrical pattern of scratching/machining marks at the tip. The surface finish appears rough as a result and marks may have high spots that have the potential to damage instrument shafts in certain loading conditions. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 897143 | 909226 | 871460 | 2955842200700303 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-09-04 | 2007-09-04 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-08-09 | Yes | No | Unkown | 2007
40017905 |  | Other | Other | It was reported that during the beginning of the da vinci prostatectomy procedure, it was noticed that the monopolar curved scissors instrument was dull. The instrument was replaced with another monopolar curved scissors instrument which was also dull. The planned surgical procedure was delayed 30-40 mins. No patient harm, adverse outcome or injury was reported. Mfr. Report# 2955842-2007-00313 was created for the second monopolar curved scissors instrument used during the surgical procedure. | The instrument was not returned for evaluation, therefore, no conclusion can be drawn regarding the customer reported complaint. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 892083 | 909887 | 872121 | 2955842200700312 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-04 | 2007-09-04 | Yes | Yes | Health Professional | 40017905 | No | No | 2007-08-10 | No | No | Unkown | 2007
40017908 |  | Other | Other | It was reported that during the beginning of the da vinci prostatectomy procedure, it was noticed that the monopolar curved scissors instrument was dull. The instrument was replaced with another monopolar curved scissors instrument which was also dull. The planned surgical procedure was delayed 30-40 mins. No patient harm, adverse outcome or injury was reported. Mfr. Report# 2955842-2007-00312 was created for the initial monopolar curved scissors instrument used during the surgical procedure. | The instrument was not returned for evaluation, therefore, no conclusion can be drawn regarding the customer reported complaint. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 893804 | 909888 | 872122 | 2955842200700313 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-09-04 | 2007-09-04 | Yes | Yes | Health Professional | 40017908 | No | No | 2007-08-10 | No | No | Unkown | 2007
42017206 | 2007-07-20 | Malfunction |  | It was reported that the tip of the maryland bipolar forceps instrument was bent. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip to be bent, causing misalignment of the grips. There is an offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed scraping around the main tube located 5 inches from the wrist. The damaged area is gray in color can has a rough surface finish. Engineering concluded that the damage may be caused by a defective cannula. Electrical continuity passed. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1418561 | 907117 | 869364 | 2955842200700286 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-29 | 2007-08-31 | No | Yes | Health Professional | 42017206 | Yes | No | 2007-08-01 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-07-31 | Other | OtherRequired Intervention | It was reported that during the surgical procedure, the customer was experiencing a nonrecoverable 20013 fault error code after several arm collisions. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered that two potentiometer assemblies had malfunctioned, thus generating the system error experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of august 29, 2007, there have been no reported recurrences of the issues at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 889322 | 908306 | 870542 | 2955842200700285 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-30 | 2007-08-31 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | No | 2007-07-31 | Yes | No | Reuse | 2007
42000604 | 2007-07-25 | Malfunction |  | It was reported that during a prostatectomy procedure, while sewing the dorsal vein, the surgeon noticed that the tip of the large needle driver instrument had broken off and fallen inside of the patient. The piece was retrieved and the procedure was successfully completed with a replacement instrument of the same type. The tip of the instrument was discarded at the site. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with the carbide insert missing from one grip. Engineering evaluation observed that the brazing surface shows very little presence of filler metal on the outer edges. This discovery leads engineering to conclude that the carbide insert may have not been properly brazed onto the grip, thus causing the carbide insert to loosen and fall off. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 899987 | 906848 | 869097 | 2955842200700295 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2007-08-30 | 2007-08-30 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-08-02 | Yes | No | Reuse | 2007
40017907 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull and corroded. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and the instrument failed to cut as intended. Engineering observed that one scissor blade was deformed at the tip and appeared to have partially melted, preventing the scissors from fully closing. Engineering also observed pitting on the surface of the damaged blade, which is not present on the other blade. The damaged blade was darker at the tip, indicating that the instrument may have arced during a previous surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 899992 | 906850 | 869099 | 2955842200700287 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-30 | 2007-08-30 | No | Yes | Health Professional | 40017907 | Yes | No | 2007-08-02 | Yes | No | Reuse | 2007
40018402 | 2007-07-31 | Malfunction |  | It was reported that the patient had strong adhesion and during the surgical procedure, the permanent cautery spatula instrument broke "during preparation. " a broken piece fell inside of the patient's thorax and was retrieved. The instrument was removed and the planned surgical procedure was completed with another instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the celvis was dislodged from the main tube both the proximal clevis and main tube were missing pieces. The instrument cables were also derailed at the distal idler. No other damage was found. Per the case notes, the broken instrument component was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 900000 | 906851 | 869100 | 2955842200700284 | 1 | GEI | Manufacturer | ForeignHealth Professional | Physician | Initial | 2007-08-30 | 2007-08-30 | No | Yes | Health Professional | 40018402 | Yes | Yes | 2007-07-31 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-07-24 | Other | Other | It was reported that during a da vinci prostatectomy procedure the customer experienced 20013 faults and resistance in vertical wrist movement. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for further failure analysis investigation. During failure analysis of the returned psm, engineering concluded that the root cause for the system fault was a defective potentiometer assembly. A replacement potentiometer assembly was installed to resolve this problem. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As a of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 879251 | 902896 | 865190 | 2955842200700269 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2007-08-22 | 2007-08-24 | Yes | Yes | Health Professional | IS1200 A43P8 | Device Returned To Manufacturer | Yes | 2007-07-24 | Yes | No | Reuse | 2007
40020503 | 2007-07-27 | Malfunction | Other | It was reported that after a total hysterectomy surgical procedure lasting 1. 5 hours, the physician noticed a burn on the pt's skin around the area where the instrument trocar was placed. Add'l info has been requested from the site without success. No permanent pt injury was reported. | The fenestrated bipolar forceps instrument has not yet been returned for evaluation. A follow up mdr will be submitted once the instrument has been received and evaluated. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 900490 | 903181 | 865474 | 2955842200700279 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-24 | 2007-08-24 | Yes | Yes | Health Professional | 40020503 | Yes | No | 2007-07-27 | No | No | Unkown | 2007
400179 |  | Malfunction |  | It was reported that two of the monopolar curved scissors instrument tip cover accessories had a hole in them. No add'l info was provided other than the site stating that this had occurred 5 or 6 times prior to this reported complaint. It is unk if more than one monopolar curved scissors instrument had been used during these events. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda as they related to this event. #2955842-2007-00268, #2955842-2007-00282. | The instrument was not returned; however, two tip cover accessories used with the instrument were returned and evaluated. Engineering's evaluation found both tip covers to have a hole burned through the silicone. The holes have a crater-like appearance, with material removed on the outer surface of the silicone. The holes appear to be energy/heat related. The silicone may have had prior damage or inherent defects that caused energy to escape at the hole sites. Engineering also observed the tip covers to have one or more tears at or near the tip of the silicone, likely due to inadvertent contact with a sharp object. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 900467 | 903183 | 865476 | 2955842200700268 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-08-22 | 2007-08-24 | No | Yes | Health Professional | 400179 | No | No | 2007-07-23 | Device Not Retu | No | Unkown | 2007
40018005 |  | Malfunction |  | It was reported that two of the monopolar curved scissors instrument tip cover accessories had a hole in them. No add'l info was provided other than the site stating that this had occurred 5 or 6 times prior to this reported complaint. It is unk if more than one monopolar curved scissors instrument had been used during these events. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda as they related to this event. #2955842-2007-00268, #2955842-2007-00282. | The instrument was not returned; however, two tip cover accessories used with the instrument were returned and evaluated. Engineering's evaluation found both tip covers to have a hole burned through the silicone. The holes have a crater-like appearance, with material removed on the outer surface of the silicone. The holes appear to be energy/heat related. The silicone may have had prior damage or inherent defects that caused energy to escape at the hole sites. Engineering also observed the tip covers to have one or more tears at or near the tip of the silicone, likely due to inadvertent contact with a sharp object. | HOT SHEARS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 900483 | 903184 | 865477 | 2955842200700283 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-08-22 | 2007-08-24 | No | Yes | Health Professional | 40018005 | Yes | No | 2007-07-23 | Device Not Retu | Yes | Unkown | 2007
40018005 |  | Malfunction |  | It was reported that two of the monopolar curved scissors instrument's tip cover accessories had a hole in them. No add'l info was provided other than the site stating that this had occurred 5 or 6 times prior to this reported complaint. It is unk if more than one monopolar curved scissors instrument had been used during these events. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda as they related to this event. #2955842-2007-00268, #2955842-2007-00283. | The instrument was not returned, however, two tip cover accessories used with the instrument were returned and evaluated. Engineering's evaluation found both tip covers to have a hole burned though the silicone. The holes have a crater-like appearance, with material removed on the outer surface of the silicone. The holes appears to be energy/heat related. The silicone may have had prior damage or inherent defects that caused energy to escape at the hole sites. Engineering also observed the tip covers to have one or more tears at or near the tip of the silicone, likely due to inadvertent contact with a sharp object. | HOT SHEARS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 882021 | 903185 | 865478 | 2955842200700282 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-08-22 | 2007-08-24 | No | Yes | Health Professional | 40018005 | Device Returned To Manufacturer | No | 2007-07-23 | Device Not Retu | Yes | Unkown | 2007
40020705 | 2007-05-18 | Malfunction |  | It was reported that 3 hours into the total hysterectomy procedure the black tip of the tenaculum forceps instrument disconnected from the shaft while the surgeon was manipulating the patient's uterus. The tip was being held by a couple of wires and pk dissecting forceps instrument was used to help align the tip for removal from the trocar. During the alignment, a cable of the pk dissecting forceps instrument broke and a wire from the tenaculum forceps instrument broke and fell inside of the patient. The broken wire was retrieved and the planned surgical procedure was completed with another tenaculum forceps instrument. It was also reported that a black piece of the tenaculum forceps instrument was missing and the surgical staff was unsure if the missing instrument component fell inside of the patient during the surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the wrist detached. Per the customer reported complaint, engineering evaluation observed that the instrument proximal clevis was missing half of the bottom section which mates with the main tube. The missing piece is approximately. 25" x. 125". The main tube is also missing a smaller piece at one of the t-tab slots. Engineering determined that the failure mode was most likely a break in the main tube to proximal clevis interface. | TENACULUM FORCEPS INSTRUMENTS | ENDOSCOPIC INSTRUMENT | 890463 | 903215 | 865508 | 2955842200700267 | 1 | GDW | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 40020705 | Yes | No | 2007-05-18 | Yes | No | Reuse | 2007
42014204 |  | Malfunction |  | It was reported that a wire in the 5mm monopolar cautery instrument tip was frayed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not observe any frayed wires at the instrument tip and none of the wires at the wrist showed any damage. Engineering also observed small sections of instrument tube insulation removed in three locations. There were also various scratches and indentations in the tube insulation. It was determined that this failure mode was most likely caused by a potentially defective cannula accessory, which may remove material from an instrument during use. The site has previously returned instruments with damaged insulation. The following medwatch reports were submitted for these instruments: mfr. Report #2955842-2007-00223; mfr report #2955842-2007-00224; mfr report #2955842-2007-00225; mfr report #2955842-2007-00252. The site has also returned defective cannula accessories. The following medwatch reports were submitted for the defective cannulae: mfr report #2955842-2007-00226; mfr report #2955842-2007-00227. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 904993 | 903216 | 865509 | 2955842200700270 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42014204 | Yes | No | 2007-07-25 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the wire connecting the cautery to the tip of the maryland bipolar forceps instrument had broken. The site experienced a loss of heat on the cautery. No additionally information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire broken at the yaw pulley exit. The insulation had not been cut, however, the yaw pulley is missing a piece opposite the conductor break which may have allowed extra tension on the wire, thereby causing it to break. Engineering also observed various deep scratches concentrated in a small area on the distal end of the main tube with small amounts of material removed. It was determined that the damage to the tube is likely caused by another instrument. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890453 | 903217 | 865510 | 2955842200700274 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-07-27 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the sheath of the monopolar curved scissors instrument was slivering. The site evaluated their cannula accessories and no defects were found. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section with scraping on the distal end of the instrument main tube. The scraping was approximately 3" above the tube extension and located all around the outer diameter. It was determined that this failure mode may have been caused by a potentially defective cannula accessory. The site has previously returned defective cannula accessories. As no damage was found to the cannulae currently being used by the site, it was determined that this instrument may have been used with the previously returned defective cannula accessories based on the instrument manufacture date and number of uses remaining. The following medwatch reports were submitted for the defective cannulae previously returned from the site: mfr. Report #2955842-2007-00021, mfr. Report #2955842-2007-00022, and mfr. Report #2955842-2007-00023. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890426 | 903218 | 865511 | 2955842200700272 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-07-26 | Yes | No | Reuse | 2007
42017908 |  | Malfunction |  | It was reported that during the prostatectomy procedure the wrist of the monopolar curved scissors instrument cracked and pieces fell inside of the patient's anatomy. The broken pieces were retrieved from the patient. The surgical staff was unable to remove the instrument from the cannula accessory, therefore, the cannula and instrument were removed from the patient concurrently. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tube extension was broken and missing a piece at the proximal clevis was dislodged from the main tube. Engineering also observed various small indentations on both blade edges, making closing of the blades difficult. No other damage was found. Per the case notes, all broken pieces were successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890444 | 903219 | 865512 | 2955842200700278 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42017908 | Yes | No | 2007-07-27 | Yes | No | Reuse | 2007
42018404 | 2007-07-24 | Malfunction |  | It was reported that after 1. 5 hours into a total hysterectomy procedure, the ceramic tip of the permanent cautery spatula instrument cracked and fell off. The instrument was replaced with an instrument of the same type to assist in the close of the vaginal cuff. Upon removal of the second permanent cautery spatula instrument, it was observed that the ceramic portion of the tip was also missing. It is believed that the pieces fell inside of the patient, however, the broken pieces could not be located. The procedure was completed without significant delay. As of 08/23/2007, no patient harm, adverse outcome of injury have been reported. The following medwatch report listed below has also been sent to the fda as it pertains to this event. 2955842-2006-00276. | Per the customer reported complaint, the instrument was returned and evaluated. Engineering evaluation observed a missing piece that broke off of the distal part of the ceramic sleeve, exposing part of the spatula shaft. The spatula has scratches on both sides, suggesting aggressive cleaning. The ceramic sleeve is cracked at the base, however, this crack appears to have propagated from a crack at the distal end of the sleeve. This could be a result of possible overloading at the tip. Electrical continuity passed. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890388 | 903220 | 865513 | 2955842200700277 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-23 | 2007-08-23 | No | No | Health Professional | 42018404 | Yes | No | 2007-07-24 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information as provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering performed functional testing and found the scissors to cut cleanly with no damage found on the blades. Electrical continuity passed. Engineering also observed a section on the distal end of the main tube to have a high concentration of deep scratches. The scratches appear on the side of the tube and material has been removed. Engineering concluded that damage may be due to contact from other instruments. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890420 | 903221 | 865514 | 2955842200700275 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-07-27 | Yes | No | Reuse | 2007
40018302 |  | Malfunction |  | It was reported that after sterilization, charring was observed on the hook end of the permanent cautery hook instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed heavy charring on the inner section of one ear of the distal clevis and also at the yaw pulley/conductor cap. The conductor wire is broken with melted insulation as it curls into the yaw pulley. Most of the conductor cap has melted. Engineering has concluded that the wire may have had prior damage, causing electrical energy to escape through the damaged section. The hook has been twisted and is no longer parallel with the ears of the distal clevis. Electrical continuity failed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890406 | 903222 | 865515 | 2955842200700271 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 40018302 | Yes | No | 2007-07-26 | Yes | No | Reuse | 2007
40018005 | 2007-07-27 | Malfunction |  | It was reported that during a prostatectomy procedure the tip cover accessory being used with the monopolar curved scissors instrument cracked down the middle and fell into the patient. This occurred after 1. 5 hours of use and while the physician was attempting to transect the bladder neck. The procedure was successfully completed without significant delay by replacing the instrument and tip cover with an instrument and tip cover of the same type. The removed tip cover was discarded at the site. No additional information has been reported. No patient harm, adverse outcome of injury was reported. | The hot shears tip cover accessory was discarded by the site and not returned for evaluation. The failure mode related to the customer reported complaint can not be verified. | TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 890380 | 903223 | 865516 | 2955842200700281 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 40018005 | No | No | 2007-07-27 | No | Yes | Initial | 2007
42018404 | 2007-07-24 | Malfunction |  | It was reported that after 1. 5 hours into a total hysterectomy procedure the ceramic tip of the permanent cautery spatula instrument cracked and fell off. The instrument was replaced with an instrument of the same type to assist in the close of the vaginal cuff. Upon removal of the second permanent cautery spatula instrument, it was observed that the ceramic portion of the tip was also missing. It is believed that the pieces fell inside of the patient, however, the broken pieces could not be located. The procedure was completed without significant delay. As of 2007, no patient harm, adverse outcome of injury have been reported. The following medwatch report listed below has also been sent to the fda as it pertains to this event. 2955842-2006-00277. | Per the customer reported complaint, the instrument was returned and evaluated. Engineering evaluation observed a missing piece that broke off of the distal part of the ceramic sleeve, exposing part of the spatula shaft. The missing piece is about 1/8" x 1/16". The spatula has scratches on both sides, suggesting aggressive cleaning. The ceramic sleeve also has various scratches, indicated by shiny streaks on the sleeve. Engineering also observed deep scratches on the main tube one inch above the wrist. One pitch cable is broken at the distal clevis. Electrical continuity passed. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890367 | 903224 | 865513 | 2955842200700276 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-23 | 2007-08-23 | No | Yes | Health Professional | 42018404 | Yes | No | 2007-07-24 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-07-27 | Other | Other | It was reported that during the surgical procedure the customer experienced a recurring 20008 system error code which could not be overridden. Due to the recurring system error code, poor port placement and patient anatomy, the procedure was converted to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering observed that the cables for an axis of the psm were off of the pulleys, thus generating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. Per the case notes, in addition to the system errors experienced by the customer, poor port placement and pt factors also contributed to the surgeon's decision to convert to open surgical techniques to complete the planned procedure. The da vinci surgical system user manual specifically states, "the initial port location should be selected giving consideration to the procedure, specific anatomy and the camera/endoscope to be used. " as of aug 22, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 890276 | 903277 | 865190 | 2955842200700280 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-08-23 | 2007-08-23 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-07-27 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-07-26 | Other | Other | It was reported that during the da vinci prostatectomy procedure the customer experienced a 21013 system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Based on the system logs, a motor encoder and potentiometer were replaced. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of 2007, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 890275 | 903279 | 865190 | 2955842200700273 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-08-23 | 2007-08-23 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | Yes | 2007-07-26 | Yes | No | Reuse | 2007
42004804 | 2007-07-19 | Malfunction |  | It was reported that after two hours in use and while dissecting tissue during a thymectomy procedure, the tip of the long tip forceps instrument broke resulting in pieces falling into the pt's anatomy. It is believed that all pieces have been retrieved. The instrument was removed and replaced with another long tip forceps instrument to successfully complete the procedure without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface resulting in the clevis becoming dislodged from main tube. Pieces are missing between the ears as well as below one of the ears. The instrument was likely overloaded at the tip. Engineering also observed that the main tube has various small scratches at distal end. No other damage found. Per the case notes, the site indicated that they believe all broken pieces were successfully retrieved from the pt. | LONG TIP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890277 | 899280 | 861581 | 2955842200700266 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-20 | 2007-08-17 | No | Yes | Health Professional | 42004804 | Device Returned To Manufacturer | No | 2007-07-20 | Yes | No | Reuse | 2007
42009305 |  | Malfunction |  | It was reported that the cable of the endowrist prograsp instrument derailed from the pulley. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable had become derailed at the distal idler pulley which may have occurred while in max pitch and yaw position. The grip can still open and close; however, the cable is exposed and in danger of becoming frayed. Engineering also observed that the distal end of the instrument's main tube had material removed on one side, which may have been due to a defective cannula accessory. No other damage was found. The site will be contacted to request the return of all their cannula accessories. If one or more cannula are found to be defective in a way that could contribute to the removal of material from an instrument, additional reports will be submitted. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890290 | 899281 | 861582 | 2955842200700264 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-17 | 2007-08-17 | No | Yes | Health Professional | 42009305 | Device Returned To Manufacturer | Yes | 2007-07-20 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument was not working properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned with the flush tube missing and the luer plate detached from the chassis. The outer wall of the chassis feature that retains the luer plate was broken. Engineering eval also observed the material had been removed from the distal end of the instrument main tube. It was determined that this failure mode may have been caused by a potentially defective cannula accessory. The site will be contacted to request the return of their cannula accessories for eval. If one or more cannulae are found to be defective in a way that could contribute to the removal of material from an instrument, add'l reports will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890288 | 899282 | 861583 | 2955842200700265 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-17 | 2007-08-17 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-07-20 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that there was scraping on the insulation of the jaw tip of the new endowrist prograsp instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a small scratch on the instrument main tube, approximately. 5" from the wrist. Engineering also observed that a small amount of material had been removed. The scratch does not appear to have been caused by a defective cannula. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INTRUMENT | 878135 | 899254 | 861555 | 2955842200700262 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-08-16 | 2007-08-16 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2007-07-17 | Yes | No | Initial | 2007
IS2000 A51P2 | 2007-07-18 | Other | Required Intervention | It was reported that during the surgical procedure, the customer experienced no vision in the right monitor of the system surgeon console. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with a monitor in the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected hrsv. The procedure was converted to open surgical techniques to complete the planned procedure, and there was no injury to the patient. As of august 15, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 877601 | 899334 | 861634 | 2955842200700263 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-08-17 | 2007-08-16 | Yes | Yes | Health Professional | IS2000 A51P2 | No | No | 2007-07-18 | Device Not Retu | No | Reuse | 2007
420179XX | 2007-07-02 | Malfunction |  | It was reported that during a prostatectomy procedure, the surgical staff noticed that the monopolar curved scissors instrument's tip cover accessory had "charring" where the clear material interfaces with the gray material on the tip cover. There were approximately 2 "pin point" holes on the interface and an additional 3 "pin point" holes in the clear section. No sharp edges could be located by the surgical staff. Upon inspection. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda as related events from this site. 2955842-2007-00257. | The instrument and/or instrument tip cover has not yet been received for analysis. Once analysis has been performed, a follow up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 878600 | 893406 | 855714 | 2955842200700256 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 420179XX | No | No | 2007-07-12 | No | No | Unkown | 2007
40017905 |  | Malfunction |  | It was reported that during testing, prior to a procedure, the monopolar curved scissors instrument began smoking from its tip. The instrument was exchanged for an instrument of the same type to successfully complete the procedure without significant delay. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per customer reported complainant engineering observed that a hole was burnt through the distal end of the instrument's main tube and tube extension. Material around the hole has been removed and burn marks were also observed on the hypotubes. A crack was found on the main tube which likely acted as a path for arcing. Per the case notes, the arcing occured prior a procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 875903 | 893407 | 855715 | 2955842200700260 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 40017905 | Yes | No | 2007-07-13 | Yes | No | Initial | 2007
42009305 |  | Malfunction |  | It was reported that the release tabs of the endowrist prograsp instrument fell out of the housing. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument and detached levers were returned and evaluated. Per the customer reported complaint, engineering observed that the two housing pegs (next to the snaps) on one side of the housing were broken, indicating that the instrument was dropped. Engineering also observed that the distal end of the instrument main tube had material removed on one side, which may have been due to a defective cannula accessory. No other damage was found. The site will be contacted to request the return of all their cannula accessories. If one or more cannula are found to be defective in a way that could contribute to the removal of material from an instrument, additional reports will be submitted. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 875847 | 893409 | 855717 | 2955842200700255 | 1 | NAY | Manufacturer | User facility | Physician | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 42009305 | Yes | Yes | 2007-07-12 | Yes | No | Reuse | 2007
42020702 |  | Malfunction |  | It was reported that the shaft was of the teneculum forceps instrument was splintering. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube to have various deep scratches where material has been removed. Deep scratches are not aligned with the longitudinal axis of the tube, suggesting a defective cannula was not the cause of damage. Scratches may be due to collisions with other instruments or mishandling during cleaning. No other damage found. | TENECULUM FORCEPS INSTRUMENTS | ENDOSCOPIC INSTRUMENT | 931449 | 893410 | 855718 | 2955842200700254 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | No | 2007-07-11 | Yes | No | Reuse | 2007
420179XX | 2007-06-27 | Malfunction |  | It was reported that during a prostatectomy procedure the surgeon noticed a hole near the tip of the monopolar curved scissors instrument tip cover accessory. Both the tip cover and the monopolar curved scissors instrument were removed and the procedure was successfully completed by replacing the tip cover. The procedure was not significantly lengthened. No pt harm, adverse outcome or injury was reported. | The instrument was not returned, however, the tip cover accessory used with the instrument was returned and evaluated. Engineering evaluation found the tip cover to be damaged in two locations at the silicone overmold. There is a hole at the midpoint of the silicone, located near one of the silicone parting lines. The hole has a crater-like appearance, with a smaller diameter on the inner surface of the overmold, and material removed on the outer surface. Hold appears to be energy/heat related. Engineering also observed a tear and an adjacent oblong shaped hole roughly 180 degrees away and. 125" below the hole. The tear does not penetrate through the silicone and appears to have been created by a collision with a sharp object. The oblong hole has a similar damage characteristics (crater-like appearance) as the round hole, suggesting heat/electrical energy caused the hole. Likely that as cautery was activated, electrical energy found a path to escape through the silicone, burning a hole. Site previously returned a tip cover with a hole burned in it. (see mdr 2955842-2007-00192). | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890307 | 893411 | 855719 | 2955842200700258 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 420179XX | No | No | 2007-07-13 | Device Not Retu | No | Unkown | 2007
420179XX | 2007-07-06 | Malfunction |  | It was reported that while removing the monopolar curved scissors instruments from use during a prostatectomy procedure, the surgeon noticed a burn hole had formed on the side of the instrument's tip cover accessory. The procedure had been successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda as related events from this site. 2955842-2007-00256. | The instrument and/or instrument tip cover has not yet been received for analyses. Once analyses has been performed a follow up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 905765 | 893412 | 855720 | 2955842200700257 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 420179XX | No | No | 2007-07-12 | No | No | Unkown | 2007
42017907 |  | Malfunction |  | It was reported that during a prostatectomy procedure the surgeon noticed a burning smell and a black mark with smoke at the distal end of the monopolar curved scissors instrument, around the instrument's tip cover. This occurred while coagulating tissue. The instrument was removed and replaced with an instrument of the same type. The procedure was successfully completed without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned. Per the customer reported complaint, engineering observed that the tube extension was broken resulting in the proximal clevis to be dislodged. A broken tube extension may have created a path for arcing, however, there are no signs of arcing through the main tube or tube extension. The conductor wire is not damaged. There is a black mark on the proximal clevis that lines up with the hole on the tip cover. This may possibly be where the instrument arced. The tube extension may have been damaged further during shipping, as broken pieces were found loose in packing. Electrical continuity passed. No other damage found. The tip cover accessory was also returned. Engineering evaluation found the tip cover to have a hole at the silicone to sleeve interface. The plastic sleeve has charring below the hole. The hole is located approximately halfway between the parting lines of the silicone and the sleeve on either side of the hole, which may have created a path for arcing. There are also various tears in the silicone that are not heat/energy related. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890303 | 893413 | 855721 | 2955842200700259 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-10 | 2007-08-10 | No | Yes | Health Professional | 42017907 | Yes | No | 2007-07-13 | Device Not Retu | No | Unkown | 2007
40021403 | 2007-07-09 | Malfunction |  | It was reported that after an hour of use during a hysterectomy procedure, the pk dissecting forceps instrument was difficult to remove due to a bent tip. The surgical staff applied force to remove the instrument causing the instrument to break. A piece of the broken insulation fell inside the patient and was successfully removed using a grasper. The procedure was completed using another instrument of the same type and without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a break at the proximal clevis to main tube interface. This resulted in the clevis to become dislodged from the main tube. The proximal clevis and main tube are missing pieces. The clevis is missing a piece below one ear. Engineering concluded that the instrument may have been overloaded at the tip. Engineering also observed the main tube to have a couple of deep scratches at the distal end. One conductor wire is broken at the yaw pulley exit. No other damage was found. Per the case notes, the broken insulation that fell inside the patient was successfully retrieved. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 878792 | 892334 | 854681 | 2955842200700253 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-08 | 2007-08-08 | No | Yes | Health Professional | 40021403 | Device Returned To Manufacturer | No | 2007-07-09 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument was not working. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr. Report #2955842-2006-00223, 2955842-2006-00224, 2955842-2007-00225. | The instrument was returned and evaluated. Engineering conducted performance testing and observed that the instrument moved intuitively with full range of motion in all directions. All instrument components were found to be functioning properly and the customer reported complaint could not be confirmed. Engineering also observed severe scraping on one side of the instrument main tube, two inches from the wrist. It was determined that the scraping may have been caused by a defective cannula accessory. No other damage was found. The site will be contacted to request the return of all their cannula accessories. If one or more cannula are found to be defective in a way that could contribute to the removal of material from an instrument, additional reports will be submitted. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 878784 | 892335 | 854682 | 2955842200700252 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-08 | 2007-08-08 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-07-09 | Yes | No | Reuse | 2007
42020503 |  | Malfunction |  | It was reported that after sterilization, it was noticed that one of the prongs of the cord of the fenestrated bipolar forceps instrument was pushed all the way down inside of the instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument top plug riser pin was pushed into the back end and the chassis feature which retains the pins was broken. Engineering also observed the distal end of the main tube had sections with material removed. This damage appears to be consistent with the use of a potentially defective cannula accessory. No other damage found. The site has returned defective cannula accessories. Additional medwatch reports will be submitted for the defective cannula accessories. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890293 | 891936 | 854290 | 2955842200700249 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-06 | 2007-08-07 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-07-07 | Yes | No | Reuse | 2007
40000620 | 2007-05-15 | Malfunction |  | It was reported by a distributor that during a prostatectomy procedure, the fine tip insert of the large needle driver instrument became dislodged during the surgery. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with one tungsten carbide insert missing from one grip. Engineering evaluation observed that the brazing surface shows small traces of filler metal. This discovery leads engineering to conclude that the carbide insert may have been inadequately brazed onto the grip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 905427 | 891958 | 854312 | 2955842200700251 | 1 | NAY | Manufacturer | ForeignDistributor | UNKNOWN | Initial | 2007-08-07 | 2007-08-07 | No | Yes | Health Professional | 40000620 | Yes | No | 2007-06-08 | Yes | No | Reuse | 2007
40017206 | 2007-07-03 | Malfunction | Other | It was reported while cauterizing the seminal vesicles during the prostatectomy procedure it was noticed that the wires were broken at the distal tip of the maryland bipolar forceps instrument and the surrounding tissue were being cauterized by the wires. The planned surgical procedure was delayed approx. 20 minutes. The procedure was completed and no patient injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one broken conductor wire at the yaw pulley exit, which is pointing away from the yaw pulley. The conductor insulation was not melted and no signs of arcing were present. Engineering concluded that the broken wire and grip together with the unbroken wire may have been touching tissue when the cautery was activated, thus causing the unexpected burning of the tissue. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 873563 | 889900 | 852270 | 2955842200700245 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-07-05 | 2007-08-03 | Yes | Yes | Health Professional | 40017206 | Device Returned To Manufacturer | Yes | 2007-07-05 | Yes | No | Reuse | 2007
42017904 | 2007-06-28 | Malfunction |  | It was reported that during a prostatectomy procedure, the monopolar curved scissors instrument stopped working. The surgeon found a loose cable at the tip of the instrument. The instrument was replaced with another instrument of the same type and the procedure was successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the grip close cable to be broken at the distal idlers resulting in the pulley spinning freely. It was determined that the cable broke under tensile loading. Other cables at the wrist are not damaged. Electrical continuity passed. Engineering also observed a section on the main tube located 4 inches from the wrist with light scraping all around the tube's outer diameter. Some material has been removed, as surface finish is rough. Damage may be due to a defective cannula. Cannulae returned from this site have been evaluated and determined to be within specification. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 904209 | 889901 | 855719 | 2955842200700250 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2007-07-07 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional information has been provided. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Per the reported complaint failure analyses observed the instrument while placed on a system for a latex cut test. The scissors do not cut cleanly through. 006" latex. The latex gets pinched at scissor tips. The blades are not damaged. The blade edges are slightly rounded at tips, negatively affecting cut performance. Engineering also observed the distal end of main tube to have various small, deep scratches concentrated on one side of the tube. Damage does not appear to be cannula related. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 904225 | 889902 | 852272 | 2955842200700248 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-07-07 | Yes | No | Reuse | 2007
42009305 |  | Malfunction |  | It was reported that the wrist of the prograsp forceps instrument broke between the shaft and the tip. It was reported that no components had fallen off into a patient. No additional information has been provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analysis observed the instrument to be broken at the proximal clevis to main tube interface. The clevis is not dislodged from the main tube but it is missing a piece below one of the ears. The main tube is not missing pieces, but the stiffness of interface wall has been compromised. Engineering also observed the main tube to have a section with material removed on one side, located 1. 5 inches from the wrist. The surface finish is rough, possibly damaged by a defective cannula. No other damage found. This site has returned cannulae for evaluation. If one or more cannulae are found to be defective in a way that would contribute to the removal of material from an instrument, additional reports will be submitted. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 904237 | 889903 | 852273 | 2955842200700247 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-07-07 | Yes | No | Reuse | 2007
42020503 |  | Malfunction |  | It was reported that the tips of the fenestrated bipolar forceps instrument did not meet. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed one grip to be bent, causing misalignment of the grips. The yaw pulley of the bent grip is not damaged. The grip likely bent due to excessive loading. Engineering also observed that the distal end of main tube has material removed all around the outer diameter. Damage appears to be a combination of deep scratches and scraping. Tube is gray and feels rough in the damaged area. No other damage found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 904245 | 889904 | 852274 | 2955842200700244 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-07-04 | Yes | No | Reuse | 2007
42009305 | 2007-07-06 | Malfunction |  | It was reported that during the middle of the total hysterectomy procedure, the tip of the endowrist prograsp instrument fell inside the patient and was retrieved. The planned surgical procedure was completed using a replacement instrument of the same type and without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the wrist detached from the main tube. Per the customer reported complaint, engineering evaluation observed that the instrument proximal clevis was intact, but the main tube interface wall was collapsed, indicating a break at the proximal clevis to main tube interface. The cables had been cut at the distal end of the hypotubes, which leads engineering to conclude that the instrument wrist was most likely cut off to aid in removal of the instrument from the cannula. Per the case notes, the broken instrument component was successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 869292 | 889905 | 852275 | 2955842200700246 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-07-06 | Yes | No | Reuse | 2007
42009305 |  | Malfunction |  | It was reported that the jaws of the endowrist prograsp instrument would not open properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and the instrument moved intuitively with full range of motion. The grips opened and closed properly. Engineering also observed the distal end of the main tube to have two narrow sections, located roughly 180 degrees apart, with material removed. The tube surface feels rough in the damaged areas. Damage may possibly be caused by a defective cannula. No another damage found. The site has been requested to return all cannulae. If one or more cannulae are found to be defective in a way that would contribute to the removal of material from an instrument, additional reports will be submitted. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 869385 | 889913 | 852283 | 2955842200700243 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2007-08-03 | 2007-08-03 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-07-04 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-07-03 | Other | Required Intervention | It was reported that during the surgical procedure the customer experienced a recurring 20013 system error code which could not be overridden. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering discovered that the motor encoder assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of august 2, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 891735 | 889934 | 852304 | 2955842200700240 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2007-08-02 | 2007-08-02 | Yes | Yes | Health Professional | IS2000 A51P2 | Yes | No | 2007-07-03 | Yes | No | Reuse | 2007
42017907 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the tube extension was dislodged from the instrument main tube and the keys on the tube extension had broken off. Engineering also discovered a section at the distal end of the instrument main tube where material had been removed. It was determined that this failure mode is consistent with the use of a potentially defective cannula accessory which may remove material from the instrument during use. The site will be contacted to request the return of their cannulae for evaluation. If the cannulae are found to be defective, additional mdrs will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 873583 | 889845 | 852215 | 2955842200700239 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-08-02 | 2007-08-01 | No | Yes | Health Professional | 42017907 | Yes | No | 2007-07-03 | Yes | No | Reuse | 2007
42017204 | 2007-06-27 | Malfunction |  | It was reported that during a lumboaortic scraping surgical procedure, the black plastic part on the grip of the maryland bipolar forceps instrument broke. Broken pieces fell inside of the pt and were retrieved. The instrument was removed, and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube. Both the proximal clevis and main tube were missing pieces. Per the case notes, the broken instrument components were successfully retrieved from the pt. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 931457 | 886632 | 849003 | 2955842200700236 | 1 | GEI | Manufacturer | ForeignUser facility | Other | Initial | 2007-07-27 | 2007-07-27 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-06-27 | Yes | No | Reuse | 2007
IS1200 | 2006-05-01 | Other | Other | During a surgical procedure, the da vinci surgical system presented a fault code to the user. The da vinci safety systems generated the fault code in response to a differential change in the angular position of one or more robotic joints as measured by a primary and secondary sensor being out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The error was produced while the surgeon was working between two of the pt's ribs. The surgeon decided to convert the planned procedure to non-robotic surgical techniques. Whether the procedure was converted to open or laparoscopic surgical techniques was not recorded. No pt harm, adverse outcome or injury was reported. No additional pt info was provided. | Isi field service engineering was unable to identify a malfunction of the da vinci surgical system. A review of the system error logs indicated a potential issue with the master tool manipulator (mtm) assembly. To conservatively address the hospital's concern, the master tool manipulator (mtm) of the da vinci surgical system was replaced and returned to isi for further assessment. Through extensive testing, the isi return material service technician was also unable to find a malfunction with the master tool manipulator. Issue was concluded as likely caused by poor surgical case setup. Inappropriately located surgical instrument cannula ports may have applied back pressure on the robotic joints and produced aberrant sensor signals. The da vinci safety systems functioned as designed by transitioning to a "recoverable safe state". | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 882774 | 886690 | 849061 | 2955842200600200 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2006-05-01 | Yes | No | Reuse | 2007
IS1200 | 2006-05-03 | Other | Other | During a surgical procedure with the da vinci surgical system, the surgeon reported that he could not move the instrument down the insertion axis. The surgeon decided to convert the planned procedure to non-robotic surgical techniques. Whether the procedure was converted to open or laparoscopic surgical techniques was not recorded. No pt harm, adverse outcome or injury was reported. No additional pt info was provided. | Isi field service engineer was unable to find a malfunction with the da vinci surgical system. A review of the system error logs produced no indication of a system malfunction. Issue was concluded as likely caused by poor surgical case setup. Inappropriately located surgical instrument cannulae ports may have applied back pressure on the robotic joints thus restricting movement of the instrument down the insertion axis. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 883277 | 886691 | 849061 | 2955842200600201 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2006-05-03 | Yes | No | Reuse | 2007
IS2000 | 2006-04-18 | Other | Other | During the beginning of a surgical procedure, the da vinci surgical system faulted. The da vinci safety systems generated the fault code in response to detecting a communication error. Upon determining this condition, the safety systems put da vinci in a "non-recoverable safe state". The system fault was detected after the system had been docked to the pt, but before any surgical tasks had been performed with the system. No pt harm, adverse outcome or injury was reported. No additional pt information was provided. | Isi field service engineer determined through a review of the system error logs, that the da vinci safety systems generated the fault code in response to detecting a communication error, between the surgeon's side console and the pt side console of the da vinci surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 883287 | 886693 | 849064 | 2955842200600202 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS2000 | Yes | No | 2006-04-18 | Yes | No | Reuse | 2007
IS1200 |  | Other | Other | Before any surgical tasks had been performed with the da vinci surgical system and while attempting to dock the system to the pt, the system faulted. The da vinci safety systems generated the fault code in response to a differential change in the angular position of one or more robotic joints as measured by a primary and secondary sensor being out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". No pt harm, adverse outcome or injury was reported. | The isi field service engineer evaluated the system and determined that a drivetrain cable had been dislodged. The noted damage can only occur as a result of abusive forces being applied to the da vinci surgical system. These abusive forces do not occur during normal system use and may have resulted prior to surgery when the da vinci surgical system was subjected to a collision (e. G. , during movement between hospital rooms). | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 869593 | 886694 | 849065 | 2955842200600203 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2006-05-18 | Yes | No | Reuse | 2007
IS1200 |  | Other | Other | During a surgical procedure, the da vinci surgical system began to present a fault code to the user. The da vinci safety systems generated the fault code in response to a differential change in the angular position of one or more robotic joints as measured by a primary and secondary sensor being out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The surgeon decided to convert the procedure to open surgical techniques to complete the planned procedure. No pt harm, adverse outcome or injury was reported. | The isi field service engineer identified the issue as originating from the endoscopic camera manipulator arm (ecm). He repaired the system by replacing the ecm. The ecm was returned to isi for additional failure analysis. An isi return material service technician performed additional electrical tests on the returned ecm and determined the issue was likely caused by degraded potentiometer (i. E. Secondary sensor) producing a noisy signal. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 882442 | 886695 | 849066 | 2955842200600204 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Device Returned To Manufacturer | No | 2006-11-06 | Yes | No | Reuse | 2007
IS1200 | 2006-03-20 | Other | Other | During a single vessel small thoracotomy procedure, the da vinci surgical system presented a fault code to the user and it was reported that the "system locked up". The site did not record the fault code. The procedure was converted to an open sternotomy to complete the planned procedure. No pt harm, adverse outcome or injury was reported. No additional pt info was provided. | The isi field service engineer was unable to determine why the da vinci safety systems transitioned the system to a "safe state". A review of the system errors logs produced no indication of a system malfunction. The cause of the customer complaint could not be identified. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 869629 | 886697 | 849068 | 2955842200600199 | 1 | NAY | Manufacturer | User facility |  | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2006-03-20 | Yes | No | Reuse | 2007
IS1200 |  | Other | Other | During a surgical procedure with the da vinci surgical system, the master tool manipulator (mtm) malfunctioned. The surgeon decided to convert the procedure to open surgical techniques to completed the planned procedure. No pt harm, adverse outcome or injury was reported. | The isi field service engineer repaired the system by replacing the mtm. The mtm was returned to isi for additional failure analysis. An isi return material service technician performed an additional eval on the returned mtm and determined that the roll pin on the mtm had sheared. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 869581 | 886698 | 849069 | 2955842200700237 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2007-02-28 | Yes | No | Reuse | 2007
42020702 | 2007-03-30 | Other | Other | During a hysterectomy procedure, an isi tenaculum forceps instrument was used as a retraction device. While under load, the instrument shaft cracked and the co2 insufflation gas started escaping from the pt's abdomen. It was reported that no pieces of the instrument had fallen into the pt. The damaged instrument was removed from the pt and the planned surgical procedure was completed using the da vinci surgical system. The surgery was lengthened approximately 30 minutes to 1 hour due to this incident. No pt harm, adverse outcome or injury was reported. | The instrument was returned with the wrist section cut off at the distal end of the main tube. The instrument wrist was not returned with the instrument. Splintering found at the distal end is due to cutting off of the wrist and cables. The engineering evaluation determined that the failure mode experienced by the customer was most likely a break in the proximal clevis at the main tube interface. No other damage found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882432 | 886699 | 849070 | 2955842200700238 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-07-27 | 2007-07-27 | Yes | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | No | 2007-04-03 | Yes | No | Initial | 2007
40217206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument had a broken cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument pitch down cable was broken at the distal clevis and the pitch up cable was frayed at the distal clevis. Both the distal and proximal clevis were free of burrs. Engineering also observed that material was removed at the distal end of the instrument main tube. The cause of the damage cannot be determined. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890471 | 893007 | 855323 | 2955842200700241 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-08-03 | 2007-07-27 | No | Yes | Health Professional | 40217206 | Yes | No | 2007-07-02 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that during a surgical procedure, the maryland bipolar forceps instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's main tube had a large section with material removed all around the outer diameter. The damaged area is gray in color and has a rough surface finish. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory which may remove material from the instrument during use. This account has recently returned a defective cannula. Please reference the following mdr's: 2955842-2006-00114, 2955842-2006-00116, 2955842-2006-00118, 2955842-2006-00119, 2955842-2006-00120, 2955842-2006-00121. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890296 | 886620 | 848991 | 2955842200700234 | 1 | GEI | Manufacturer | Unknown | UNKNOWN | Initial | 2007-07-26 | 2007-07-26 | No | Yes | Other | 42017204 | Yes | No | 2007-06-27 | Yes | No | Reuse | 2007
42018404 |  | Malfunction |  | It was reported that during a total hysterectomy procedure, the coating of the permanent cautery spatula instrument began chipping. It is unknown if fragments of the instrument had fallen into the patient. The instrument was exchanged with another permanent cautery spatula instrument and the procedure was successfully completed without significant delay. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned and evaluated. Engineering evaluation observed that the distal most part of the ceramic sleeve is missing, exposing bare metal. The ceramic sleeve broke where it transitions to a smaller diameter. Both sides of spatula have fairly deep scratches which extend down to the bare metal exposed by the missing ceramic sleeve. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890300 | 886621 | 848992 | 2955842200700233 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-07-26 | 2007-07-26 | No | Yes | Health Professional | 42018404 | Yes | No | 2007-06-26 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument did not grasp well. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering tested the instrument and found it to move intuitively with a full range of motion. No problem with grasping was noticed. Engineering also observed deep scratches at the distal end of the main tube, causing white marks to appear on the tube. Some scratches were deep enough to have removed material. The damage is likely caused by contact from other instruments. No other damage found. The damage of the main tube does not appear to be caused by a defective cannula. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 879777 | 886015 | 848403 | 2955842200700235 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-25 | 2007-07-25 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-06-26 | Yes | No | Reuse | 2007
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument had loose cables. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and the instrument moved intuitively with full range of motion. No loose cables were observed. Engineering also discovered deep scratches on one side of the distal end of the main tube. One scratch was deep enough to have removed a small amount of material. The scratches do not appear to have been caused by a defective cannula. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 879774 | 886017 | 848405 | 2955842200700230 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-25 | 2007-07-25 | No | Yes | Health Professional | 42009307 | Yes | No | 2007-06-25 | Yes | No | Reuse | 2007
40009306 | 2007-06-26 | Malfunction |  | It was reported that the surgical staff had difficulties removing the endowrist prograsp instrument after it had been in use for 2 hrs while performing a da vinci prostatectomy. The tip of the instrument was bent at a 90 degree angle and became caught on the trocar. A piece of the distal clevis broke off and it is assumed to have fallen inside of the pt. The surgical staff did not notice any pieces falling in the pt and did not take an x-ray as the plastic piece is not radiopaque. The case was successfully completed and no pt harm has been reported. No additional surgical procedure will be performed and this event did not significantly lengthen the procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface, and a piece is missing below one of the ears. The clevis was not dislodged from the main tube. The main tube does not appear to be missing pieces, but the section exposed by the missing clevis piece has been damaged. The instrument was likely overloaded at the tip. The main tube has various small scratches at distal end. No other damage found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882757 | 886018 | 848406 | 2955842200700232 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-07-25 | 2007-07-25 | No | Yes | Health Professional | 40009306 | Device Returned To Manufacturer | No | 2007-06-27 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that during sterilization, splinters formed on the shaft of the maryland bipolar forceps instrument. The instrument was discarded by the site. No pt harm, adverse outcome or injury was reported. | Visual inspection of the instrument was performed by an isi representative prior to the site discarding the instrument. The inspection suggests that the splintering was a by-product of a s scratch/ scrape on the exterior coating of the instrument. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882702 | 883974 | 846389 | 2955842200700140 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42017204 | No | No | 2007-06-20 | Yes | No | Unkown | 2007
42014204 |  | Malfunction |  | It was reported that the shaft of the 5mm monopolar cautery instrument was peeling off. No additional information was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr report #2955842-2006-00223, mfr report # 2955842-2006-00225. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section with missing insulation located 2 inches from the snake wrist on the main tube. The metal part of the tube is exposed. The damaged area has scratch marks at an angle to the tube axis. Inspection of the tube under magnification shows various small scratch marks below the damaged area. Electrical continuity passed. It was concluded that the instrument damage may have been caused by the 5mm cannula accessory. No other damage was found. This account has recently returned two damaged cannula accessories. Please see mdr's 2955842-2007-00226 and 2955842-2007-00227. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882710 | 883975 | 846390 | 2955842200700224 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42014204 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
42017907 | 2007-06-20 | Malfunction |  | It was reported that during a prostatectomy procedure, the tip of the monopolar curved scissors instrument broke off and fell into the pt. The separated tip was successfully removed with a grasper instrument and discarded by the site. The procedure was completed with a replacement instrument of the same type without injury or significantly lengthening the procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated with one blade missing. Per the customer reported complaint, engineering observed the blade break at the cross-section of the grip cable crimp. The returned blade was sent to an independent test lab for analysis. Testing results confirmed that stress corrosion cracking was the cause of failure, as large amounts of chlorine were found on the fracture surface of the broken blade. Per the case notes, the broken instrument component was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882736 | 883976 | 846391 | 2955842200700229 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42017907 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
42001101 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: mfr report # 2955842-2007-00227. | The cannula accessory was evaluated. Engineering observed a burr on the inner diameter of the cannula near the distal tip. It was determined that the burr had the potential to cause scraping in certain loading conditions. Site has previously returned scraped instruments, see mfr report numbers: 2955842-2007-00223, 2955842-2007-00224, and 2955842-2007-00225. | CANNULA 5MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882722 | 883977 | 846392 | 2955842200700226 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42001101 | Yes | Yes | 2007-07-18 | Yes | No | Reuse | 2007
42014204 |  | Malfunction |  | It was reported that the shaft of the 5mm monopolar cautery instrument was peeling off. No additional information was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related compliants used at the same facility. Mfr report #2955842-2006-00224, mfr report # 2955842-2006-00225. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a large section with missing insulation located 5 inches from the snake wrist on the main tube. The metal part of the tube is exposed. Inspection of the tube under magnification shows various small scratch marks below the damaged area. Electrical continuity passed. It was concluded that the instrument damage may have been caused by the 5mm cannula accessory. No other damage was found. This account has recently returned two damaged cannula accessories. Please see mdr's 2955842-2007-00226 and 2955842-2007-00227. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882717 | 883978 | 846393 | 2955842200700223 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42014204 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
42001101 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: mfr report: # 2955842-2007-00226. | The cannula accessory was evaluated. Engineering observed a dent on the lip of the cannula on the long side of the distal tip. The dent has created a flat spot on the lip, displacing material towards the inner diameter and creating a small bump on the inner surface. It was determined that the bump had the potential to cause scraping in certain loading conditions. Site has previously returned scraped instruments, see mfr report numbers: 2955842-2007-00223, 2955842-2007-00224, and 2955842-2007-00225. | CANNULA 5MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882707 | 883979 | 846394 | 2955842200700227 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42001101 | Yes | Yes | 2007-07-18 | Yes | No | Reuse | 2007
42011704 |  | Malfunction |  | It was reported that the shaft of the 5mm needle driver instrument wsa peeling off. No additional information was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments with related complaints used at the same facility. Mfr report #2955842-2006-00223, mfr report # 2955842-2006-00224. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a small, narrow section with missing insulation located 2 inches from the snake wrist on the main tube. The metal part of the tube is exposed. Inspection of the tube under magnification shows long scratches next to the damaged area. Electrical continuity passed. It was concluded that the instrument damage may have been caused by the 5mm cannula accessory. No other damage was found. This account has recently returned two damaged cannula accessories. Please see mdr's 2955842-2007-00226 and 2955842-2007-00227. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882731 | 883980 | 846395 | 2955842200700225 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42011704 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
42009305 | 2007-06-20 | Malfunction |  | It was reported that the endowrist prograsp instrument was in use for approx 4 hrs during a da vinci prostatectomy procedure when the tip became offset from the shaft following the lymph node dissection. While trying to remove the instrument from the cannula, the instrument broke and a piece fell inside of the pt and was retrieved with a needle holder. The instrument and cannula were removed concurrently from the pt and the planned surgical procedure was completed without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube. The clevis is intact, however, one side of the main tube interface wall has collapsed. Per the case notes, the broken instrument component was successfully retrieved from the pt. The cannula accessory was also returned and no damage was found. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882725 | 883981 | 846396 | 2955842200700222 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
42020503 |  | Malfunction |  | It was reported that the shaft of the fenestrated bipolar forceps instrument was frayed. No pieces of the instrument fell inside of the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a section of the instrument main tube with scraping, three inches from the wrist. It was determined that this failure mode is consistent with the use of a potentially defective cannula accessory which may remove material from the instrument during use. The site will be contacted to return their cannula accessories for evaluation. If the cannula are found to be defective, additional mdrs will be submitted. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 864057 | 883982 | 846397 | 2955842200700228 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-20 | 2007-07-20 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2007-06-21 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument would not grasp. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing, and confirmed that material becomes pinched at the tips when trying to cut. Visual inspection revealed the blades to be undamaged. Engineering also discovered a section towards the distal end of the instrument main tube where material had been removed all around the outer diameter. It was determined that this failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. The site has been requested to return their cannulae for evaluation. If the cannulae are found to be defective, additional mdrs will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 865108 | 883930 | 846345 | 2955842200700220 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-19 | 2007-07-19 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-06-19 | Yes | No | Reuse | 2007
42017204 | 2007-04-11 | Malfunction |  | It was reported that toward the end of a prostatectomy procedure, the wrist of the maryland bipolar forceps instrument broke. No pieces of the broken instrument fell into the pt. The cannula and the instrument were removed from the pt, and the instrument was replaced with an instrument of the same type. The procedure was delayed approximately 40 minutes as the account did not have a sterilized back-up instrument ready for use. The procedure was successfully completed and no pt harm was reported. | The instrument was returned for evaluation with the wrist detached. Per the customer reported complaint, engineering observed the cables and conductors at the distal end to be cut. This is likely due to the instrument breaking at the proximal clevis to main tube interface. Wrist had been cut to remove from instrument cannula. Both main tube and proximal clevis are missing pieces. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 865099 | 883932 | 846347 | 2955842200700221 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2007-07-19 | 2007-07-19 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-06-20 | Yes | No | Reuse | 2007
40017907 |  | Malfunction |  | It was reported that during a prostatectomy procedure, shavings from the monopolar curved scissors instrument fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various small, light scratches at the distal end of the main tube. The surface finish of the tube feels fairly smooth, as if little to no material has been removed. Tube damage is not consistent with cannula related scraping. Additionally, engineering observed the tube extension to have a triangle shaped burn mark at the interface with the metal proximal clevis indicating that arcing had occurred. The burn mark is located roughly 45 degrees radially from one of the keys and is around. 100" wide. The tube extension has a hairline crack at the other key. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 867450 | 882912 | 845328 | 2955842200700219 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2007-07-18 | 2007-07-18 | No | Yes | Health Professional | 40017907 | Yes | No | 2007-07-18 | Yes | No | Initial | 2007
40018302 |  | Malfunction |  | It was reported that the tip of the permanent cautery hook instruement was cracked and without much insulation. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found the yaw pulley, distal clevis, and proximal clevis to show severe charring, indicating that arcing had occurred. Material has been removed from the yaw pulley due to burning. Part of the distal clevis is broken off. Arcing may have originated at the cable crimp, as the yaw pulley is the most severely damaged component. Conductor wire is not damaged. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 867456 | 882913 | 845329 | 2955842200700218 | 1 | GEI | Manufacturer | ForeignDistributor | Unknown | Initial | 2007-07-18 | 2007-07-18 | No | Yes | Health Professional | 40018302 | Yes | No | 2007-06-18 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-06-18 | Other | Other | It was reported that during a da vinci prostectomy procedure, the customer experienced a mechanical problem with a pt side manipulator of the da vinci surgical system. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Field service engineering repaired the system by replacing the affected patient side manipulator (psm). The psm was returned to isi for failure analysis investigation. Engineering discovered two cables for insertion were off their pulleys. As of june 18, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 890317 | 883048 | 843491 | 2955842200700216 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-07-18 | 2007-07-18 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-06-18 | Yes | No | Reuse | 2007
IS1200 A43P6 | 2007-06-18 | Other | Other | It was reported that during a da vinci proctectomy procedure, the customer experienced 21013 and 20013 system error codes during homing. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 18, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 890306 | 883050 | 843491 | 2955842200700217 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2007-07-18 | 2007-07-18 | Yes | Yes | Health Professional | IS1200 A43P6 | Yes | No | 2007-06-18 | Yes | No | Reuse | 2007
4200201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. 2955842-2007-00210, 2955842-2007-00211, 2955842-2007-00213. | The cannula accessory was evaluated. Engineering observed the cannula to be visibly bent at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Inspection under magnification reveals a deformation at the lip of the distal tip, creating a raised edge which has the potential to cause scraping in certain loading conditions. Site has returned an instrument with damaged main tube before. See mdr 2955842-2007-00185. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882687 | 881048 | 843473 | 2955842200700212 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 4200201 | Device Returned To Manufacturer | No | 2007-06-14 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr. Report #2955842-2007-00208. Mfr. Report #2955842-2007-00209. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface of the cannula. The ledge is present on half of the inner diameter. The ledge has the potential to cause scraping in certain loading condition. Site has previously returned scraped instruments, see mfr. Report #2955842-2007-00155. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882662 | 881050 | 843475 | 2955842200700207 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | Yes | 2007-06-13 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. 2955842-2007-00211, 2955842-2007-00212, 2955842-2007-00213. | The cannula accessory was evaluated. Engineering observed the cannula to be visibly bent at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter, making the cannula out of specification due to a deformation. The deformation has the potential to cause scraping in certain loading condition. Site has returned an instrument with damaged main tube before. See mdr 2955842-2007-00185. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882692 | 881051 | 843476 | 2955842200700210 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-06-14 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr. Report #2955842-2007-00207. Mfr. Report #2955842-2007-00209. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface of the cannula. The ledge is present on half of the inner diameter. The ledge has the potential to cause scraping in certain loading condition. Site has previously returned scraped instruments, see mfr. Report #2955842-2007-00155. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882678 | 881053 | 843478 | 2955842200700208 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | Yes | 2007-06-13 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. Mfr. Report #2955842-2007-00207. Mfr. Report #2955842-2007-00208. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface of the cannula. The ledge is present on half of the inner diameter. The ledge has the potential to cause scraping in certain loading condition. Site has previously returned scraped instruments, see mfr. Report #2955842-2007-00155. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882670 | 881055 | 843480 | 2955842200700209 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | Yes | 2007-06-13 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site. 2955842-2007-00210, 2955842-2007-00211, 2955842-2007-00212. | The cannula accessory was evaluated. Engineering observed the cannula to be visibly bent at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. Inspection under magnification reveals a small deformation at the lip of the distal tip, creating a raised edge which has the potential to cause scraping in certain loading conditions. Site has returned an instrument with damaged main tube before. See mdr 2955842-2007-00185. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 890279 | 881056 | 843481 | 2955842200700213 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-06-14 | Yes | No | Unkown | 2007
42021403 |  | Malfunction |  | A pk dissecting forceps instrument was returned without any info. No pt harm, adverse outcome or injury was reported. | The instrument was evaluated. Engineering observed a burn mark on one yaw pulley near the cable crimp, with some material removed/deformed due to burning. Engineering also discovered a section towards the distal end of the instrument main tube with scraping and where material had been removed. It was determined that this failure mode is consistent with the use of a potentially defective cannula accessory which may remove material from the instrument during use. The site has been contacted to return their cannula for evaluation. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882744 | 881058 | 843483 | 2955842200700214 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42021403 | Yes | No | 2007-06-14 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00210, 2955842-2007-00212, 2955842-2007-00213. | The cannula accessory was evaluated. Engineering observed the cannula to be visibly bent at the distal tip. A. 342 gage pin does not fit through the distal tip inner diameter. The tip of the cannula has a flat spot, possibly due to mishandling. Flat spot has the potential to cause scraping in certain loading conditions. Site has returned an instrument with a damaged main tube before. See mdr 2955842-2007-00185. No other damage found. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 890282 | 881060 | 843485 | 2955842200700211 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-07-13 | 2007-07-13 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-06-14 | Yes | No | Unkown | 2007
IS1200 A43P8 | 2007-06-15 | Other | Other | It was reported that during a da vinci prostectomy procedure, the customer experienced an unrecoverable 21013 system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered two defective potentiometer assemblies within the psm, thus generating the system error experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 13, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 890278 | 881066 | 843491 | 2955842200700215 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-07-13 | 2007-07-13 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-06-15 | Yes | No | Reuse | 2007
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on the system and determined the motion to be intuitive. The distal end of main tube has a section with light scraping. Material has been removed from one side of the tube. This is consistent with the use of a defective cannula. Electrical continuity passed. No other damage found. This site has been requested to return all cannulae. A follow up mdr will be filed upon receipt and evaluation of cannulae. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 907811 | 880980 | 843405 | 2955842200700206 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-12 | 2007-07-12 | No | Yes | Health Professional | 42021403 | Yes | No | 2007-06-04 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00197. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface of the cannula. The ledge is present on half of the inner diameter. The ledge has the potential to cause scraping in certain loading conditions. Site has previously returned scraped instruments. See mdr 2955842-2007-00071. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882696 | 878762 | 841208 | 2955842200700198 | 1 | GCJ | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 42000201 | Yes | Yes | 2007-06-07 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00198. | The accessory was returned and evaluated. Engineering observed a small ledge at the bowl/tube interface of the cannula. The ledge is present on half of the inner diameter. The ledge has the potential to cause scraping in certain loading conditions. Site has previously returned scraped instruments. See mdr 2955842-2007-00071. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882689 | 878763 | 841209 | 2955842200700197 | 1 | GCJ | Manufacturer | User facilityForeign | UNKNOWN | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | Yes | 2007-06-07 | Yes | No | Reuse | 2007
31146404 | 2007-04-24 | Malfunction | Other | It was reported that during a gastric bypass one lens of the endoscope fell into the patient. The location of the missing lens was not confirmed and the physicians opted not to perform open surgery to attempt to locate the missing lens as the patient is obese. No post operative complication have been reported by this account. | The endoscope has been returned to the manufacturer but failure analysis has not been completed yet. A follow up mdr will be submitted upon the conclusion of this investigation. | ZERO DEGREE SCHOELLY STEREO ENDOSCOPE | ENDOSCOPE | 882674 | 878764 | 841210 | 2955842200700200 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2007-07-06 | 2007-07-06 | Yes | Yes | Health Professional | 31146404 | Yes | No | 2007-04-24 | No | No | Reuse | 2007
42020702 | 2007-06-07 | Malfunction |  | It was reported that 1/2 hour into the total hysterectomy procedure, the tenaculum forceps instrument broke during the mobilization of the uterus. The instrument wrist was at its maximum angle when the surgeon tried pushing the uterus further and the fracture occurred. A broken piece fell inside of the patient and was retrieved with a laparoscopic grasper. The planned surgical procedure was completed without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube. Both the proximal clevis and main tube were missing pieces. No other damage was found. Per the case notes, the broken instrument component was successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882672 | 878765 | 841211 | 2955842200700203 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 42020702 | Yes | No | 2007-06-07 | Yes | No | Reuse | 2007
42020703 | 2007-06-06 | Malfunction |  | It was reported that during a myomectomy procedure, while grasping the myoma, the proximal clevis of the tenaculum forceps instrument broke into pieces. Pieces fell inside of the patient and were retrieved. The instrument was replaced with a cobra grasper instrument to successfully complete the procedure without significant delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube. Both the proximal clevis and main tube were missing pieces. No other damage was found. Per the case notes, the broken instrument component was successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882697 | 878766 | 841211 | 2955842200700196 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-07-05 | 2007-07-06 | No | Yes | Health Professional | 42020703 | Yes | No | 2007-06-06 | Yes | No | Reuse | 2007
42004904 |  | Malfunction |  | It was reported that one hour into a mediastinal mass resection surgery, a cable on the wrist of the cadiere forceps instrument became disconnected. The procedure was completed without incident. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip open cable is frayed at the grip cable groove. The other grip open cable has slight birdcaging near the grip exit. Both grips still open and close properly. Engineering also observed scraping and deep scratches on one side of the distal end of the main tube. The affected area is approximately 3 inches long. Material has been removed in this area. It was concluded that the instrument damage may have been caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. No other damage was found. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882691 | 878767 | 841213 | 2955842200700199 | 1 | GEI | Manufacturer | User facilityDistributor | Other | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 42004904 | Yes | No | 2007-06-07 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that after sterilization, it was noticed that the serrated portion on one side of the tip of the large needle driver instrument was missing. It cannot be determined if the component separated from the instrument during a surgical procedure or during handling after a procedure. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with the carbide insert missing from one grip. Engineering evaluation observed that the brazing surface shows very little presence of filler metal. This discovery leads engineering to conclude that the carbide insert may have not been properly brazed onto the grip, thus causing the carbide insert to loosen and fall off. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882684 | 878768 | 841214 | 2955842200700204 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-06-08 | Yes | No | Reuse | 2007
40000620 | 2007-06-08 | Malfunction |  | It was reported that during the surgical procedure, the inner part of the jaw of the large needle driver instrument fell inside of the patient and was retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with the carbide insert missing from one grip. Engineering evaluation observed that the brazing surface shows very little presence of filler metal. This discovery leads engineering to conclude that the carbide insert may have not been properly brazed onto the grip, thus causing the carbide insert to loosen and fall off. As indicated in the case notes, the carbide insert was successfully retrieved from the patient. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882679 | 878769 | 841215 | 2955842200700205 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-06 | 2007-07-06 | No | Yes | Health Professional | 40000620 | Yes | No | 2007-06-08 | Yes | No | Reuse | 2007
420179XX | 2007-06-04 | Injury | Required Intervention | It was reported that during the prostatectomy procedure, the surgeon noticed that the patient's iliac vein was bleeding. The monopolar curved scissors instrument was removed and a hole on the tip cover accessory was observed. The planned surgical procedure was completed, however, the case was lengthened by thirty minutes as the surgeon had to repair the bleeding iliac vein. No patient harm, adverse outcome or injury was reported. | The instrument was not returned, however, the tip cover accessory used with the instrument was returned and evaluated. Engineering evaluation found a hole in the silicone around 1/16" in diameter at the proximal end of the tip cover. The hole does not appear to come from a puncture, as material is missing from the outer surface of the silicone overmold. A video of the procedure was returned. During review of the video, it was noticed that the tip cover accessory was pressing on the tissue, when the iliac vein began to bleed. It may have been possible that the cautery activation with the tip cover pressing on the tissue caused energy to transfer from the instrument, through the silicon, and to the iliac vein, causing the bleeding experienced by the patient. Engineering also observed a tear on the distal end of the tip cover that was not likely caused by heat. The silicone is thinner at the proximal end (where the hole is) than at the distal end (where the tear is), making it more susceptible to being damaged. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 858014 | 875685 | 838161 | 2955842200700192 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-03 | 2007-07-03 | Yes | Yes | Health Professional | 420179XX | No | No | 2007-06-04 | Device Not Retu | No | Unkown | 2007
42009305 | 2007-06-04 | Malfunction |  | It was reported that after sterilization of the instrument, the hospital staff noticed the shaft of the prograsp forceps instrument was peeling off. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed severe scraping on the main tube. Material has been removed. Damage may be caused by a defective cannula. This account has previously returned three defective cannulae, all with dents at the tip likely caused by mishandling. See mdr's 2955842-2007-00021, 2955842-2007-00022 and 2955842-2007-00023. This account has been contacted and a request has been made to return the cannulae used with this instrument for evaluation. After the cannulae evaluation, an assessment will be made to determine if product handling re-training should be administered to this account. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890322 | 875777 | 838253 | 2955842200700195 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-03 | 2007-07-03 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-06-04 | Yes | No | Reuse | 2007
42000604 | 2007-06-04 | Malfunction |  | It was reported that after sterilization of the instrument, the hospital staff noticed the shaft of the large needle driver instrument was peeling off. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed deep scratches on main tube, 1 inch from wrist. One scratch has material removed. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890313 | 875778 | 838254 | 2955842200700194 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-07-03 | 2007-07-03 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-06-04 | Yes | No | Reuse | 2007
42021405 | 2007-06-04 | Malfunction |  | It was reported that during a total hysterectomy procedure, the pk dissecting forceps instrument had been in use for five minutes when it was noticed that the tip was grasping, but had no energy. A small wire sticking out of the instrument tip was observed. The instrument was removed and a small black component was seen inside of the patient and retrieved. The surgical staff did not know source of the black component. The planned surgical procedure was completed using a replacement instrument of the same type and the procedure was not significantly lengthened. The customer indicated that the black component would be returned with the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated; however, no components were returned with the instrument. Per the customer reported complaint, engineering observed one conductor wire broken at the yaw pulley exit. The grip cables on the same side of the conductor break are derailed at the distal idlers. Engineering found the proximal clevis and main tube to be intact. The conductor insulation was not torn and no fragments were missing from the instrument. The source of the black component is unknown. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890311 | 875779 | 838255 | 2955842200700193 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-07-02 | 2007-07-03 | No | Yes | Health Professional | 42021405 | Yes | No | 2007-06-04 | Yes | No | Reuse | 2007
40021405 | 2007-05-29 | Malfunction |  | It was reported that during a total hysterectomy procedure, that the proximal clevis of a pk dissecting forceps instrument broke and fragments fell into the patient. The fragments were retrieved. It was also reported that during the procedure, fragments from the shaft of this instrument came off and fell into the patient. Most of these fragments were removed, but some small fragments were unable to be retrieved. The procedure was completed using a replacement instrument. The procedure was not significantly lengthened. As of 2007, there had been no complaints from this patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various a section 4 inches from the wrist with light scraping. A small amount of material is removed. Damage is localized, only on one side of the tube. Damaged area is gray in color and has rough surface finish. Engineering also observed that the proximal clevis is intact, there is no damage to any of the components other than the main tube. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882894 | 874114 | 836595 | 2955842200700191 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-29 | 2007-06-29 | No | Yes | Health Professional | 40021405 | Device Returned To Manufacturer | No | 2007-06-01 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the wires were exposed and frayed on the maryland bipolar forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire had the insulation cut, exposing the wires, which have frayed slightly. No signs of charring or arcing on the yaw pulley. Did not observe any burs or sharp edges on the distal clevis which may have damaged conductor. Engineering also observed scraping on the distal end of the main tube were material had been removed. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882897 | 874116 | 836597 | 2955842200700190 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-06-29 | 2007-06-29 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-05-31 | Yes | No | Reuse | 2007
40017904 | 2007-05-03 | Malfunction |  | "the pt underwent a robitic assisted total laparoscopic hysterectomy with bilateral salpingo-oophorectomy for fibroid tumors. During the procedure, one of the blades on the robotic scissors broke off. Device usage problem: device failed (e. G. , broke, couldn't get it to work or stop working)". During isi's investigation it was determined that the broken blade was removed from the pt and the procedure was successfully completed without significantly lengthening the procedure. | The device has not been returned for eval yet. It is expected back. Once the device has been rec'd and evaluated a follow up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 841419 | 973625 | 819671 | 2955842200700189 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2007-06-29 | 2007-06-29 | No | Yes | Health Professional | 40017904 | Yes | No | 2007-05-30 | No | No | Unkown | 2007
42017204 |  | Malfunction |  | It was reported that the tips of the bipolar maryland forceps instrument do not meet. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one of the instrument grips was bent, causing misalignment of the grips. The tips were offset and the yaw pulley and conductor cap were separated at the glue joint due to the bending force. A burn mark was observed on a yaw pulley and conductor wire near the yaw pulley exit. Engineering also discovered a three inch section of scraping on one side of the main tube with material removed, located three inches from the wrist. The instrument scraping does not appear to be related to a defective cannula. No other damage was found. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882892 | 874049 | 836530 | 2955842200700188 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-27 | 2007-06-28 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-05-29 | Yes | No | Reuse | 2007
IS1200 A43P8 | 2007-05-29 | Other | Other | It was reported that during the surgical procedure, the customer experienced "power fluctuation" errors. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with a remote interface adaptor (ria) pca. The system was repaired by replacing the affected ria pca. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 26, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 862785 | 874133 | 836614 | 2955842200700187 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-26 | 2007-06-28 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-05-29 | No | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the shaft of the maryland bipolar forceps instrument had splintered. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various small scratches on the distal end of main tube. Small amounts of material had been removed in this area. Scratches are only found on one side of the main tube. Electrical continuity passed. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 852609 | 870225 | 832781 | 2955842200700186 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-06-22 | 2007-06-22 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-05-24 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. | The cannula accessory was evaluated. Engineering observed chipped sections at the lip of the distal end of the cannula, which may create the potential to remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2006-00168 and # 2955842-2006-00169. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 852622 | 870226 | 832782 | 2955842200700184 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-20 | 2007-06-22 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-25 | Yes | No | Unkown | 2007
42000202 | 2007-05-25 | Malfunction |  | It was reported that during the surgical procedure, the surgeon could not move the instrument arm in and out of the patient. The bipolar maryland forceps instrument was checked, and it was noticed that plastic shavings and tissue had built up inside of the cannula, thus prohibiting the movement of the instrument. The instrument arm was removed from the patient which allowed the surgical staff to remove the instrument and cannula concurrently. The planned surgical procedure was delayed by one hour and completed with a replacement instrument of the same type. The surgical staff did not observe any instrument shavings fall inside of the patient. No patient harm, adverse outcome or injury was reported. Mfr. Report # 2955842-2007-00183 was created for the bipolar maryland forceps instrument. | The cannula accessory was returned and evaluated by engineering. No damage was found at the distal tip, and the scraping could not be replicated using a representative instrument. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882663 | 870227 | 832783 | 2955842200700182 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-22 | 2007-06-22 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-05-25 | Yes | No | Reuse | 2007
42017204 | 2007-05-25 | Malfunction |  | It was reported that during the surgical procedure, the surgeon could not move the instrument arm in and out of the patient. The bipolar maryland forceps instrument was checked and it was noticed that plastic shavings and tissue had built up inside of the cannula, thus prohibiting the movement of the instrument. The instrument arm was removed from the patient which allowed the surgical staff to remove the instrument and cannula concurrently. The planned surgical procedure was delayed by one hour and completed with a replacement instrument of the same type. The surgical staff did not observe any instrument shavings fall inside of the patient. No patient harm, adverse outcome or injury was reported. Mfr. Report # 2955842-2007-00182 was created for the cannula accessory. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument main tube had two distinct sections with scraping. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory which may remove material from the instrument during use; however, the cannula accessory returned with the instrument showed no signs of damage at the tip. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882666 | 870228 | 832784 | 2955842200700183 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-22 | 2007-06-22 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-05-25 | Yes | No | Reuse | 2007
37137802 |  | Malfunction |  | It was reported that during a prostatectomy procedure the back of an unspecified instrument was not working on the conmed machine. The cable was unscrewed and screwed back in. The nurse noticed that the cable was frayed. The cable worked again and was in use for 3 to 4 hours. The case was completed, the pt was not injured, but there was a "burn like" area around the port site on the pt. No add'l pt harm, adverse outcome or injury was reported. | The electrical cable was returned. Failure analysis is anticipated, but not yet completed. The event is currently under investigation and a follow up mdr will be submitted once all investigations are complete. | CONMED CAUTERY EXTENSION CABLE | CAUTERY EXTENSION CABLE | 852669 | 870229 | 832785 | 2955842200700185 | 1 | JOS | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-22 | 2007-06-22 | Yes | No | Health Professional | 37137802 |  | No | 2007-05-24 | No | No | Initial | 2007
IS1200 A43P8 | 2007-05-23 | Other | Required Intervention | It was reported that during a da vinci prostectomy procedure, the customer experienced a system error "20008" fault. The hospital opted to discontinue the use of the da vinci surgical system to perform the planned surgery. No pt harm was reported. The procedure was converted to traditional open surgical techniques to successfully complete the surgery. | The investigation conducted by field service engineering indicated the pt side manipulator (psm) had malfunctioned. The system was repaired by replacing the affected psm. The psm was returned to isi for further failure analysis investigation. During failure analysis of the returned psm concluded that the root cause for the system fault was a defective axis-4 motor/encoder assembly. A replacement axis-4 motor/encoder assembly was installed to resolve this problem. No related complaints have been reported from this account as of 2007. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 861299 | 870336 | 832892 | 2955842200700181 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-21 | 2007-06-22 | Yes | Yes | Health Professional | IS1200 A43P8 | Yes | No | 2007-05-23 | Yes | No | Reuse | 2007
IS2000 A50P5 | 2007-05-21 | Other | Required Intervention | It was reported that prior to a da vinci prostectomy procedure the customer experienced a system error "2" fault. The hospital opted to not use the da vinci s surgical system to perform the planned surgery. No surgical tasks were performed with the system. No pt harm was reported. The access ports had been created and the procedure was converted to traditional open surgical techniques to successfully finish the planned surgery. | The investigation conducted by field service engineering indicated the pt cart controller (pcc) had malfunctioned. The system was repaired by replacing the affected pcc. No related complaints have been reported from this account as of june, 20, 2007. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 853556 | 869497 | 832053 | 2955842200700180 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2007-06-20 | 2007-06-21 | Yes | Yes | Health Professional | IS2000 A50P5 | Yes | No | 2007-05-21 | No | No | Reuse | 2007
40017905 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was defective. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Electrical and mechanical testing was performed. The instrument was able to move intuitively with full range of motion in all directions, electrical continuity passed, and the instrument was able to cut according to specifications. The customer reported complaint cannot be confirmed. During evaluation, engineering observed deep scratches on the distal end of the main tube. All of the scratches were found on the same side of the tube. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 853608 | 869145 | 831708 | 2955842200700178 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-19 | 2007-06-20 | No | Yes | Health Professional | 40017905 | Yes | No | 2007-03-26 | Yes | No | Reuse | 2007
40017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument stopped working. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Performance testing was conducted and the instrument was found to move intuitively with full range of motion in all directions. The customer reported complaint cannot be confirmed. During evaluation, engineering observed deep scratches and light scraping on the distal end of the main tube. The scratches and scrapes were mainly concentrated on side of the main tube. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 852216 | 869146 | 831709 | 2955842200700177 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-19 | 2007-06-20 | No | Yes | Health Professional | 40017206 | Device Returned To Manufacturer | No | 2007-02-20 | Yes | No | Reuse | 2007
40000620 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument is cracked and/or flaking. No add'l info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the main tube was not cracked, however, deep scratches were found on the distal end of the main tube, causing white marks to appear on the tube. The main tube also has various small scratches concentrated in one area, making the area gray in color. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 852220 | 869147 | 831710 | 2955842200700176 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-19 | 2007-06-20 | No | Yes | Health Professional | 40000620 | Device Returned To Manufacturer | No | 2007-01-09 | Yes | No | Reuse | 2007
42020501 | 2007-05-21 | Malfunction |  | It was reported that during the surgical procedure, the shaft of the fenestrated bipolar forceps instrument broke when it collided with another instrument. A broken piece fell inside of the pt and was retrieved. The planned surgical procedure was completed using a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube and the cables at the wrist were slack. Both the proximal clevis and main tube were missing pieces. Per the case notes, the broken instrument component was successfully retrieved from the patient. As indicated in the case notes, the instrument breakage was the result of a collision with another instrument. The endowrist instruments instructions for use specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 852240 | 869148 | 831711 | 2955842200700179 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-19 | 2007-06-20 | No | Yes | Health Professional | 42020501 | Device Returned To Manufacturer | No | 2007-05-21 | Yes | No | Reuse | 2007
40017206 |  | Malfunction |  | It was reported that the shaft of the bipolar maryland forceps instrument is cracked/flaking. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the main tube was not cracked, however, deep scratches were found on the distal end of the main tube, causing white marks to appear on the tube. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 852295 | 869158 | 831721 | 2955842200700175 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-19 | 2007-06-20 | No | Yes | Health Professional | 40017206 | Device Returned To Manufacturer | No | 2007-01-09 | Yes | No | Reuse | 2007
40016906 | 2007-05-16 | Malfunction |  | It was reported that the harmonic curved shears insert broke and pieces fell inside of the patient. All broken pieces were retrieved and the procedure was completed with a replacement insert. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned with the clamp arm and curved blade broken off. Both broken components were returned with the insert. Engineering observed that the clamp arm attachment joints were bent on both sides and the clamp arm had a section of deformed teflon near the hinge. The curved blade was broken near the base and had a small chipped section at the break point. Visual evidence suggests that the components were subjected to excessive force, causing the clamp arm and curved blade to break off. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 854577 | 866750 | 829375 | 2955842200700172 | 1 | KOG | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-15 | No | Yes | Health Professional | 40016906 | Yes | No | 2007-05-16 | Yes | Yes | Initial | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00167, #2955842-2007-00168, #2955842-2007-00169, #2955842-2007-00170. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report # 2955842-2007-00166. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 854416 | 866751 | 829376 | 2955842200700171 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-15 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-16 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00167, #2955842-2007-00168, #2955842-2007-00169, #2955842-2007-00171. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mrf. Report # 2955842-2007-00166. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 854428 | 866752 | 829377 | 2955842200700170 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-15 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-16 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accesory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00167, #2955842-2007-00168, #2955842-2007-00170, #2955842-2007-00171. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00166. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 854436 | 866753 | 829378 | 2955842200700169 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial |  | 2007-06-15 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-16 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00168, #2955842-2007-00169, #2955842-2007-00170, #2955842-2007-00171. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip with two dents on opposite sites of the tip, thus creating raised ledges which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr report# 2955842-2007-00166. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 854475 | 866754 | 829379 | 2955842200700167 | 1 | KOG | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-15 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-16 | Yes | No | Unkown | 2007
42017904 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tube extension was broken at the main tube interface. The instrument wrist was also covered in a sticky residue, making the grips difficult to open and close. During evaluation, engineering also observed that the distal end of the instrument main tube had a 5" section of material removed. It was determined that the failure mode is consistent with the use of potentially defective cannula accessory, which may remove material from the instrument during use. The following defective cannula medwatch reports listed below were also sent to the fda. These reports only pertain to defective cannula returned from this specific site. Mfr. Report #2955842-2007-00167, mfr. Report #2955842-2007-00168, mfr. Report #2955842-2007-00169,mfr. Report #2955842-2007-00170, mfr. Report #2955842-2007-00171,. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 854532 | 866755 | 829380 | 2955842200700166 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-22 | 2007-06-15 | No | Yes | Health Professional | 42017904 | Yes | No | 2006-12-05 | Yes | No | Reuse | 2007
IS2000 A51P2 | 2007-05-16 | Other | Required Intervention | It was reported that during the surgical procedure, the surgeon was unable to continue with the case as the site did not have the correct instrument available. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Per the case notes, it was determined that the instrument that the site had in their inventory was for their da vinci surgical system, and is incompatible with their da vinci s surgical system. The site did not have the correct instrument for their da vinci s surgical system in their inventory in order for them to complete the planned surgical procedure. The da vinci and da vinci s surgical system instruments differ in the color, and fit of their housing and are not interchangeable. The instruments are clearly marked indicating which system they should be used on. The da vinci s surgical system user manual explicitly states that "environmental or equipment failures may cause the da vinci s system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. ". | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 854461 | 866792 | 829416 | 2955842200700173 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-15 | 2007-06-15 | No | Yes | Health Professional | IS2000 A51P2 | No | No | 2007-05-16 | Device Not Retu | No | Reuse | 2007
IS1200 A43P7 | 2007-05-18 | Other | Required Intervention | It was reported that during a da vinci prostectomy procedure, the customer experienced a mechanical problem with a master tool manipulator (mtm). No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering indicated mechanical damage on the mtm. The spring went out on the grip. The system was repaired by replacing the affected right mtm. The mtm was returned to isi for further failure analysis investigation. Failure analysis of the returned mtm concluded that the root cause for the mechanical damage was that the grip levers were missing. New design grip parts were installed to resolve this problem. No related complaints have been reported from this account as of june 15, 2007. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INTRUMENT CONTROL S | 854456 | 866793 | 829417 | 2955842200700174 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2007-06-15 | 2007-06-15 | Yes | Yes | Health Professional | IS1200 A43P7 | Device Returned To Manufacturer | No | 2007-05-18 | Yes | No | Reuse | 2007
42015703 | 2007-05-14 | Malfunction | Required Intervention | It was reported that during a beating heart totally endoscopic coronary artery bypass, the blade of the snap-fit instrument fell into the patient's anatomy. The blade was retrieved and the procedure was successfully completed without being significantly lengthening. No patient harm, adverse outcome of injury was reported. | The instrument was returned for evaluation. Per the reported complaint, engineering observed the following: instrument was returned with the snap fit insert in a separate container. Instrument shows no signs of damage. Pocket of yaw pulley was checked with a go/no go gage and found to be in specification. Insert does not show any damage to the blade or the legs. Insert was installed on instrument and driven. Insert did not fall off while moving the instrument in all directions with varying speed. It took a collision which deformed the tip and one leg in order for the insert to fall off. Likely that insert was loaded incorrectly, legs did not snap into yaw pulley. No other damage found. | SNAPFIT INSTRUMENT | ENDOSCOPIC INSTRUMENT | 849977 | 865879 | 828530 | 2955842200700162 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-06-13 | 2007-06-14 | Yes | Yes | Health Professional | 42015703 | Yes | No | 2007-05-14 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00071, #2955842-2007-00072, #2955842-2007-00164. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip. There is a flat spot at the tip, reducing the clearance for the main tube. Deformation has the potential to cause scrapping in certain loading conditions. No other damage found. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00071 and 2955842-2007-00072. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882577 | 866736 | 829361 | 2955842200700163 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-14 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-05-15 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00071, #2955842-2007-00072, #2955842-2007-00163. | The cannula accessory was evaluated. Engineering observed that the cannula was not visibly bent at the distal tip. A. 342 gage pin does not fit through the distal tip inner-diameter. With the gage pin installed, the tip looks out of round, possibly due to a slight deformation. A deformation has the potential to cause scraping in certain loading conditions. No other damage found. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00071 and #2955842-2007-00072. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882585 | 866739 | 829364 | 2955842200700164 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-14 | 2007-06-14 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-15 | Yes | No | Unkown | 2007
42000619 |  | Malfunction |  | It was reported that the large needle driver instrument did not articulate properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a da vinci s system and drove it. Movement in yaw is non-intuitive due to a derailment of a cable. In addition, engineering also observed the distal end of main tube to have a scraped section 2. 5 inches from the wrist. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882556 | 866740 | 829365 | 2955842200700165 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-14 | 2007-06-14 | No | Yes | Health Professional | 42000619 | Device Returned To Manufacturer | No | 2007-05-15 | Yes | No | Reuse | 2007
42019003 |  | Malfunction |  | It was reported that the cobra grasper instrument would not open and close. It was also reported that the shaft of the instrument had been damaged. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the grips to open and close properly on the system. Movement in all axes were found to be intuitive with a full range of motion. The distal end of the main tube has various deep scratches, causing white marks on the tube. No other damage found. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 860065 | 865233 | 827889 | 2955842200700161 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-06-12 | 2007-06-12 | No | Yes | Health Professional | 42019003 | Yes | No | 2007-02-09 | Yes | No | Reuse | 2007
42004903 |  | Malfunction |  | It was reported that the shaft on the cadiere forceps instrument was scraped. No additional information was provided. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Engineering found the distal end of the main tube to have various deep scratches, causing white marks on the tube. Scratching is mainly at least 3 inches of the main tube's distal end. No other damage found. | CADIERE FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 860071 | 865234 | 827890 | 2955842200700160 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-06-12 | 2007-06-12 | No | Yes | Health Professional | 42004903 | Yes | No | 2007-01-22 | Yes | No | Reuse | 2007
42009303 |  | Malfunction |  | It was reported that the jaws on the prograsp forceps instrument do not close completely. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the grips to open and close properly on the system. Movement in all axes were intuitive with a full range of motion. The distal end of the main tube had various deep scratches, causing white marks on the tube. No other damage was found. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 860076 | 865235 | 827891 | 2955842200700159 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-06-12 | 2007-06-12 | No | Yes | Health Professional | 42009303 | Yes | No | 2007-02-09 | Yes | No | Reuse | 2007
42020702 |  | Malfunction |  | It was reported that a piece of the black shaft was "lifting off" from the teneculum forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found no pieces of the main tube or proximal clevis broken or "lifting off". All components are intact. The distal end of the main tube has deep scratches. No other damage found. | TENECULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 860086 | 865236 | 827892 | 2955842200700158 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-12 | 2007-06-12 | No | Yes | Health Professional | 42020702 | Yes | No | 2007-03-22 | Yes | No | Reuse | 2007
42009305 | 2007-05-07 | Malfunction |  | It was reported that during the surgical procedure, when the customer attempted to release tissue, the jaws of the endowrist prograsp instrument stopped working. The customer was able to pull the tissue out of the instrument jaws. When attempting to remove the instrument, the instrument became stuck in the cannula. As a result, the instrument and trocar were removed simultaneously. The planned surgical procedure was completed after a 15-20 minute delay. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the main tube to proximal clevis interface. The clevis was dislodged from the main tube as a result of the breakage and the main tube interface had collapsed inwards. The proximal clevis was found to be intact and pieces of the main tube were missing. Per the case notes, there are no indications from the hospital that the missing instrument components fell inside of the patient during the surgical procedure. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 883797 | 862452 | 826104 | 2955842200700144 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-06-06 | 2007-06-07 | No | Yes | Health Professional | 42009305 | Device Returned To Manufacturer | No | 2007-05-07 | Yes | No | Reuse | 2007
42003604 |  | Malfunction |  | It was reported that the debakey forceps instrument experienced flaking on the outer shaft. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customers reported complaint, engineering observed material removed from the distal end of main tube, right above the wrist. Damaged area is roughly 1. 5 inches long and only on one side of the tube. Damage may be a result of a defective cannula. No other damage found. | DEBAKEY FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882858 | 863446 | 826101 | 2955842200700155 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42003604 | Yes | No | 2007-05-11 | Yes | No | Reuse | 2007
42009305 | 2007-05-08 | Malfunction |  | It was reported that during the surgical procedure, the tip of the endowrist prograsp instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed using a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. As a result of the breakage, the clevis was dislodged from the main tube and the cables at the wrist were slack. Both the proximal clevis and main tube were missing pieces. Per the case notes, the broken instrument component was successfully retrieved from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882863 | 863449 | 826104 | 2955842200700150 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42009305 | Yes | No | 2007-05-08 | Yes | No | Reuse | 2007
42004804 |  | Malfunction |  | It was reported that the long tip forceps instrument experienced flaking on the outer shaft. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customers reported complaint, engineering observed material removed from distal end of main tube, right above the wrist. Damaged area is roughly 1. 5 inches long and only on one side of the tube. Damage may be a result of a defective cannula. No other damage found. | LONG TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882559 | 863450 | 826105 | 2955842200700156 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42004804 | Device Returned To Manufacturer | No | 2007-05-11 | Yes | No | Reuse | 2007
40000619 |  | Malfunction |  | It was reported that the large needle driver instrument did not work properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the main tube with localized deep scratches, located 2 inches from the wrist. When tested, the instrument moved intuitively with a full range of motion in all directions. No components were broken. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882842 | 863452 | 826107 | 2955842200700152 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-06 | 2007-06-07 | No | Yes | Health Professional | 40000619 | Device Returned To Manufacturer | No | 2007-05-09 | Yes | No | Reuse | 2007
40017205 |  | Malfunction |  | It was reported that the bipolar forceps instrument did not work properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the main tube with localized deep scratches, located 2 inches from the wrist. When tested, the instrument moved intuitively with a full range of motion in all directions. No components were broken. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882848 | 863453 | 826108 | 2955842200700151 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-06 | 2007-06-07 | No | Yes | Health Professional | 40017205 | Yes | No | 2007-05-09 | Yes | No | Reuse | 2007
42004904 |  | Malfunction |  | It was reported that the grip of the cadiere forceps instrument do not answer correctly any more [sic]. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument moved intuitively with full range of motion in all directions. The instrument cable tension was slightly lower than normal; however, this is typical of an instrument with heavy usage. Engineering also observed deep scratches on the distal end of the main tube, causing white/gray marks to appear on tube. No other damage was found. The scratches on the main tube do not appear to be caused by a defective cannula. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882869 | 863454 | 826109 | 2955842200700148 | 1 | GEI | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-06-06 | 2007-06-07 | No | Yes | Health Professional | 42004904 | Yes | No | 2007-05-08 | Yes | No | Reuse | 2007
42017204 |  | Malfunction |  | It was reported that the tip of the bipolar forceps instrument broke. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section on the main tube with localized deep scratches, located 1 inch from the wrist. When tested, the instrument moved intuitively with a full range of motion in all directions. No components were broken. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882856 | 863455 | 826110 | 2955842200700153 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-06-06 | 2007-06-07 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-05-10 | Yes | No | Reuse | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome of injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00145, #2955842-2007-00146, #2955842-2007-00147. | The cannula accessory was evaluated. Engineering observed that the cannula was out of specification as the tip opening was smaller than the approved acceptable size. The smaller opening may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00112. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 883775 | 863456 | 826111 | 2955842200700154 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-05-08 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome of injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00145, #2955842-2007-00146, #2955842-2007-00154. | The cannula accessory was evaluated. Engineering observed a small deformation at the lip of the distal tip, thus creating a small raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00112. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 883779 | 863457 | 826112 | 2955842200700147 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-05-08 | Yes | No | Unkown | 2007
40017907 |  | Malfunction | Other | It was reported that during an unknown procedure and while the surgeon was cauterizing tissue, small burn marks were noticed on the patient's uterus. The monopolar curved scissors instrument was removed, cleaned, and the instrument was found to have a small crack in the black carbon fiber. Another monopolar curved scissors instrument was used to successfully complete the procedure without significant delay. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a dislodged tube extension. Keys that mate with the main tube are broken and the tube extension can rotate. Main tube has various deep scratches at the distal end. Did not observe any cracks on the main tube or any holes burned through the main tube. Instrument was placed on system and cautery was perform. Cautery passed, with arcing occurring at the tube extension/main tube interface. A gap is created at the interface due to the dislodgement, the insulation is compromised, and energy escapes from the hypotubes through the gap, burning tissue. This suggests the tube extension breakage occurred before the alleged arcing. Latex cut test passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 883785 | 863458 | 826113 | 2955842200700157 | 1 | GEI | Manufacturer | User facilityCompany Representative | Other | Initial | 2007-04-12 | 2007-06-07 | Yes | Yes | Health Professional | 40017907 | Device Returned To Manufacturer | No | 2007-05-11 | Yes | No | Initial | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome of injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00145, #2955842-2007-00147, #2955842-2007-00154. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00112. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 883806 | 863459 | 826114 | 2955842200700146 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-08 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome of injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00146, #2955842-2007-00147, #2955842-2007-00154. | The cannula accessory was evaluated. Engineering observed that the cannula was visibly deformed at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. This cannula may have been associated with previously reported mfr. Report #2955842-2007-00112. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 883811 | 863460 | 826115 | 2955842200700145 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-06-07 | 2007-06-07 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-05-08 | Yes | No | Unkown | 2007
IS1200 A43P8 | 2007-05-08 | Other | Required Intervention | It was reported that during the surgical procedure, the customer experienced a recurring 20013 system error code which could not be overridden. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering discovered that a potentiometer and motor encoder assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 882850 | 863694 | 826349 | 2955842200700149 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-06 | 2007-06-07 | Yes | Yes | Health Professional | IS1200 A43P8 | No | No | 2007-05-08 | Yes | No | Reuse | 2007
40000620 |  | Malfunction |  | It was reported that the surface of the large needle driver instrument was rough. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the distal end of he main tube to have a deep scratch one inch from the wrist. There are also a few other scratches in the same area. The surface of the tube is rougher in this damage area, but is not severe as in cannula related scraping. Did not feel any rough or sharp edges on the metal components at the wrist. Instrument jaws open and close properly. Motion is smooth in all directions. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 883818 | 861810 | 824463 | 2955842200700141 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-06-05 | 2007-06-05 | No | Yes | Health Professional | 40000620 | Device Returned To Manufacturer | No | 2007-04-23 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument did not work properly. No additional information was provided. No patient harm, adverse outcomes or injury was reported. | The instrument was returned and evaluated. Engineering observed deep scratches on once side of the distal end of the main tube, causing white marks to appear on tube. No other damage found. The damage to the main tube does not appear to be caused by a defective cannula. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 883791 | 861811 | 824464 | 2955842200700143 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2007-06-05 | 2007-06-05 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-05-04 | Yes | No | Reuse | 2007
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had a "broken wire". No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the grip close cable to be broken at the distal idlers. Engineering also observed scraping on one side of the main tube. This account has returned cannulae which were determined not be damaged. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 883787 | 861812 | 824465 | 2955842200700142 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-06-05 | 2007-06-05 | No | Yes | Health Professional | 42021403 | Device Returned To Manufacturer | No | 2007-04-25 | Yes | No | Reuse | 2007
42014204 | 2007-05-04 | Malfunction |  | It was reported that the insulation of the 5mm monopolar cautery instrument sheared off when the instrument rubbed against the trocar. No pieces fell inside of the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been received. A follow-up medwatch report will be sent upon return and evaluation of the instrument. Per the case notes, no instrument components fell inside of the patient. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890328 | 859568 | 822219 | 2955842200700134 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-01 | 2007-06-01 | No | Yes | Health Professional | 42014204 | No | No | 2007-05-04 | Device Not Retu | No | Unkown | 2007
42013904 | 2007-05-04 | Malfunction |  | It was reported that the insulation of the 5mm fenestrated schertel grasper instrument sheared off when the instrument rubbed against the trocar. No pieces fell inside of the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been received. A follow-up medwatch report will be sent upon return and evaluation of the instrument. Per the case notes, no instrument components fell inside of the patient. | 5MM FENESTRATED SCHERTEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890331 | 859569 | 822220 | 2955842200700135 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-06-01 | 2007-06-01 | No | Yes | Health Professional | 42013904 | No | No | 2007-05-04 | Device Not Retu | No | Unkown | 2007
42017204 | 2007-04-18 | Malfunction |  | It was reported that during the surgical procedure, the bipolar maryland tip forceps instrument no longer moved. The surgical staff was unable to remove the instrument from the trochar and a small instrument fragment fell into the patient. The fragmented component was successfully retrieved. The instrument and trochar were removed from the patient simultaneously, and the surgical staff was then able to remove the instrument from the trochar. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the main tube to proximal clevis interface. The clevis was dislodged from the main tube as a result of the breakage. Both the clevis and main tube were missing pieces. As indicated in the complaint notes, the fragmented component was successfully retrieved. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890339 | 859571 | 822222 | 2955842200700136 | 1 | GEI | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2007-06-01 | 2007-06-01 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-05-04 | Yes | No | Reuse | 2007
40000620 | 2007-03-29 | Malfunction |  | It was reported that during the surgical procedure, the insert of one of the jaws of the large needle driver instrument fell inside of the patient. The insert was retrieved and discarded and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with the carbide insert missing from one grip. Engineering evaluation observed that the brazing surface shows very little presence of filler metal. This discovery leads engineering to conclude that the carbide insert may have not been properly brazed onto the grip, thus causing the carbide insert to loosen and fall off. As indicated in the case notes, the carbide insert was successfully retrieved from the patient. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890351 | 859573 | 869587 | 2955842200700137 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2007-06-01 | 2007-06-01 | No | Yes | Health Professional | 40000620 | Yes | No | 2007-05-04 | Yes | No | Reuse | 2007
42009304 | 2007-05-02 | Malfunction |  | It was reported that during a prostatectomy procedure, the wrist of the endowrist prograsp instrument broke and the instrument could not be removed from the trocar. The instrument was cut to remove it from the trocar. It is unknown if any pieces had fallen inside of the patient. Another instrument of the same kind was used to successfully complete the procedure. No patient harm, adverse outcome or injury was reported during or after this procedure. | The instrument has not yet returned from this account for evaluation. A follow up medwatch will be sent once the instrument has been received and evaluated. | ENDOWRIST PROGRASP INSTRUMENT | ENDOWRIST INSTRUMENT | 890356 | 859575 | 822227 | 2955842200700138 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-06-01 | 2007-06-01 | No | Yes | Health Professional | 42009304 | No | No | 2007-05-02 | No | No | Unkown | 2007
42009304 | 2007-05-02 | Malfunction | Required Intervention | It was reported that during a hysterectomy procedure, the shaft of the endowrist prograsp instrument broke. Subsequently, the shaft was missing pieces. The surgeon decided to irrigate and suction the potentially effected area as well as take x-rays as a precautionary measure. No fragments were located in the pt's anatomy. The broken instrument was manually removed from the system and a second instrument of the same kind was used to successfully complete the procedure. No pt harm, adverse outcome or injury has been reported. | The instrument was returned and per the customer reported complaint, engineering observed a break at the prox clevis to main tube interface. Clevis is dislodged from main tube. Both clevis and main tube are missing pieces. Distal end of main tube has scraping, likely caused by attempting to remove instrument from cannula. Roll cables have derailed from clamping pulleys and one of the crimps has slipped off of a roll cable. No other damage found. As indicated in the complaint notes, no broken pieces were found in the pt's anatomy. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 890274 | 859576 | 822227 | 2955842200700139 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2007-06-01 | 2007-06-01 | Yes | Yes | Health Professional | 42009304 | Yes | No | 2007-05-03 | Yes | No | Reuse | 2007
IS1200 A43P7 | 2007-04-26 | Other | Required Intervention | It was reported that during a da vinci proctectomy procedure the customer experienced an intermittent 20008 fault error mode. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering was unable to duplicate the fault, but concluded that the system issue experienced by the customer was likely associated with the left master tool manipulator (mtm) and/or the pt side manipulator (psm). The system was repaired by replacing the affected mtm and psm. Both the mtm and psm were returned to isi for further failure analysis investigation. Failure analysis of the returned components concluded that the root cause of the intermittent errors was a noisy potentiometer on the psm. Results from performance testing and visual inspection concluded that there was no failure of the mtm component associated with the intermittent errors. As designed, the system alarm (system generated fault codes) functioned correctly during the surgery and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of may 30, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INTRUMENT CONTROL S | 943164 | 859463 | 822114 | 2955842200700115 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-30 | 2007-05-30 | Yes | Yes | Health Professional | IS1200 A43P7 | Device Returned To Manufacturer | No | 2007-04-24 | Yes | No | Reuse | 2007
42009304 | 2007-04-18 | Malfunction |  | It was reported that during the surgical procedure, a black chip from the shaft of the endowrist prograsp instrument fell inside of the pt and was retrieved. The planned surgical procedure was completed using a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for eval, therefore, the root cause of the customer reported complaint cannot be determined. As indicated in the complaint notes, the broken piece was successfully retrieved. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 846282 | 856693 | 819369 | 2955842200700133 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-22 | 2007-05-24 | No | Yes | Health Professional | 42009304 | No | No | 2007-04-26 | Device Not Retu | No | Unkown | 2007
42021403 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument had splinters approx three to four inches from the wrist. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument main tube had some light scraping approx three inches from the wrist, all around the outer diameter of the tube. There was also a section on one side of the tube where material had been removed, located beneath the scraped area. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. The site has previously returned damaged cannula. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 845700 | 856694 | 819370 | 2955842200700130 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-24 | 2007-05-24 | No | Yes | Health Professional | 42021403 | Yes | No | 2007-04-24 | Yes | No | Reuse | 2007
40017206 |  | Malfunction |  | It was reported that the bipolar maryland forceps instrument had a frayed cable. No add'l info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found all instrument cables to be undamaged and intact. During eval, engineering observed that the distal end of the instrument main tube had a 5" section of material removed in various places around the outer diameter. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1404489 | 856695 | 819371 | 2955842200700132 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-05-22 | 2007-05-24 | No | Yes | Health Professional | 40017206 | Yes | No | 2007-04-26 | Yes | No | Reuse | 2007
IS1200 A33 | 2007-04-24 | Other | Required Intervention | It was reported that during the beginning of the surgical procedure the insertion axis of the system endoscopic camera manipulator became stuck and would not move. No patient harm was reported. The patient incisions were closed and the planned surgical procedure was aborted. | The investigation conducted by field service engineering concluded that the system endoscopic camera manipulator (ecm) was beyond it's normal upper range. The system was repaired by running the system in maintenance mode. As of may 23, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 846223 | 856778 | 819454 | 2955842200700131 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-23 | 2007-05-24 | Yes | Yes | Health Professional | IS1200 A33 | No | No | 2007-04-24 | Yes | No | Reuse | 2007
42004904 | 2007-04-27 | Malfunction |  | It was reported that after sterilization of the cadiere forceps instrument the account noticed that a piece of the distal end of the main tube had broken off. It is undetermined as to when the piece broke off. A myomectomy was performed 7 days prior, but no patient harm, adverse outcome or injury was reported. | The instrument was returned for eval. Per the customer reported complaint, engineering observed that the instrument was missing a piece from the "ear" of the proximal clevis. In addition to the missing piece ("ear") of the proximal clevis, there was a crack on the proximal clevis at the main shaft interface underneath the ear and the proximal pin that is swaged into the ear was bent. This add'l damage suggests that the proximal clevis broke due to a high force impact on the tip of the instrument. We can not determine whether this occurred during use or during handling subsequent to use. If this damage occurred during use, operation of the instrument would be impaired and likely observable with the endoscopic vision provided by the system. | CADIERE FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 846880 | 853270 | 815967 | 2955842200700122 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42004904 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Reuse | 2007
31146404 | 2007-04-18 | Malfunction | Required Intervention | It was reported that after 30 minutes of performing a colectomy procedure, the surgeon could not see in one eye. The scope was removed and it was determined that a lens was missing. The lens could not be located through open surgery or by x-ray. It has not been determined if the lens fell into the patient's anatomy. No post operative complication have been reported by this account. | The endoscope has been returned to the manufacturer, but failure analysis has not been completed yet. A follow up mdr will be submitted upon the conclusion of this investigation. | 0 DEGREE SCHOELLY STEREO ENDOSCOPE | ENDOSCOPE | 882899 | 853271 | 815968 | 2955842200700108 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | Yes | Yes | Health Professional | 31146404 | Device Returned To Manufacturer | No | 2007-04-19 | No | No | Initial | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00123, 2955842-2006-00124, 2955842-2006-00125,2955842-2006-00126, 2955842-2006-00128. | The cannula accessory was evaluated. Engineering observed the cannula accessory to be bent at the distal tip, thus potentially causing scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882901 | 853272 | 815969 | 2955842200700127 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00124, 2955842-2006-00125, 2955842-2006-00126,2955842-2006-00127, 2955842-2006-00128. | The cannula accessory was evaluated. Engineering observed the cannula accessory to be bent at the distal tip, thus potentially causing scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882903 | 853273 | 815970 | 2955842200700123 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00123, 2955842-2006-00125, 2955842-2006-00126,2955842-2006-00127, 2955842-2006-00128. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a small deformation at the lip of the distal tip, thus creating a raised edge that could cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882904 | 853274 | 815971 | 2955842200700124 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00123, 00124, 00126, 00127, 00128. | The cannula accessory was evaluated. Engineering observed the cannula accessory to be bent at the distal tip, thus potentially causing scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882636 | 853275 | 815972 | 2955842200700125 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00123, 00124, 00125, 00127, 00128. | The cannula accessory was evaluated. Engineering observed the cannula accessory to be bent at the distal tip, thus potentially causing scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882642 | 853276 | 815973 | 2955842200700126 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00123, 00124, 00125, 00126, 00127. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a small deformation at the lip of the distal tip, thus creating a raised edge that could cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS200 | ENDOSCOPIC ACCESSORY | 882647 | 853277 | 815974 | 2955842200700128 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-26 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 00117, 00118, 00119, 00120, 00121. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the distal tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882658 | 853281 | 815978 | 2955842200700116 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-18 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 00116, 00118, 00119, 00120, 00121. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the distal tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882682 | 853282 | 815979 | 2955842200700117 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-04-18 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 00116, 00117, 00119, 00120, 00121. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a small deformation at the distal tip, which created a raised ledge that may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882693 | 853284 | 815981 | 2955842200700118 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-04-18 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 00116, 00117, 00118, 00120, 00121. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a small deformation at the distal tip, which created a raised ledge that may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882704 | 853285 | 815982 | 2955842200700119 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-04-18 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 2955842-2006-00116, 2955842-2006-00117, 2955842-2006-00118, 2955842-2006-00119, 2955842-2006-00121. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the distal tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882563 | 853286 | 815983 | 2955842200700120 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000401 | Yes | No | 2007-04-18 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00114, 2955842-2006-00116, 2955842-2006-00117, 2955842-2006-00118, 2955842-2006-00119, 2955842-2006-00120. | The cannula accessory was evaluated. Engineering observed that the inner diameter of the cannula accessory has a small ledge at the distal tip, which creates a raised ledge that may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882566 | 853287 | 815984 | 2955842200700121 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-05-17 | 2007-05-18 | No | Yes | Health Professional | 42000401 | Device Returned To Manufacturer | No | 2007-04-18 | Yes | No | Unkown | 2007
IS2000 A51P2 | 2007-04-24 | Other | Required Intervention | It was reported that during the surgical procedure, the customer experienced a recurring 23015 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering discovered that a motor encoder assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of may 18, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 882571 | 853304 | 816001 | 2955842200700129 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-05-18 | 2007-05-18 | Yes | Yes | Health Professional | IS2000 A51P2 | No | No | 2007-04-24 | Yes | No | Reuse | 2007
|  | Malfunction |  |  | The instrument was returned and eval. Per the customer reported complaint, engineering observed a 5 inch long section on the distal end of the main tube with scraping. This site has previously returned cannulae with damage at the distal tip which may have been used in conjuction with the instrument. The following medwatch reports listed below will also be sent to the fda. These reports only pertain to the cannulae returned from this specific site: 2955842-2007-00116, 00117, 00118, 00119, 00120, 00121. |  |  | 946286 | 978414 | 803300 | 2955842200700114 | 1 | GCJ | Manufacturer | Other |  | Initial |  | 2007-05-17 | No |  | Invalid Data |  |  |  | 2007-04-17 | Yes | No | Unkown | 2007
IS2000 A50P5 | 2007-04-14 | Malfunction |  | It was reported that before a radical prostatectomy procedure, the physician was unable to balance the left or right camera controller unit (ccu). While under anesthesia, the patient's incisions were closed. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right camera controller unit (ccu). The system was repaired by replacing the affected ccu. As of may 14, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC IINSTRUMENT CONTROL | 838301 | 850952 | 813648 | 2955842200700113 | 1 | KOG | Manufacturer | User facility | Nurse | Initial | 2007-05-14 | 2007-05-15 | No | Yes | Health Professional | IS2000 A50P5 | No | Yes | 2007-04-14 | Yes | No | Reuse | 2007
| 2007-02-20 | Malfunction |  | A surgeon was performing a laparoscopic robotically assisted myomectomy on patient. The case was performed using the davinci robot. During the procedure, a piece of the orange tube extension broke off of the monopolar curved scissor. Another monopolar curved scissor was given to the surgeon to use and the orange tube extension on that instrument broke into pieces. The silicone tip cover was not used with either of the monopolar curved scissors. The silicone tip cover is used to insulate the exposed metal at the tip of the instrument except for the scissors themselves. The surgeon was using the 8mm monopolar curved scissors (part# 400179, lot # 0112061 835 ver-05 and 0112061 878 ver-05). Two weeks later, we spoke to an instrument engineer at intuitive surgical. Our initial thought was that the damage was caused when the electrosurgical unit was activated without the silicone tip cover. The intuitive engineer informed us that the tip cover adds support to the end of the tube extension. When the instrument is cutting it puts additional stress on the end of the tube extension. The engineer also informed us of the importance of using the silicone tip cover to prevent alternate site burns. Staff that were trained by intuitive clearly understood that the silicone tip cover provided insulation to prevent burns when used for cauterizing but did not have any knowledge of the silicone tip cover providing support for the monopolar curved scissors; which may have given the impression that the tip cover was not required if the cautery function was not being utilized. Our evaluation of the monopolar curved scissors is that the failure was due to human error and the design of the instrument. The monopolar curved scissors were returned to intuitive surgical for evaluation. Our recommendation for future designs would be for intuitive to develop a better product that takes into consideration the possibility of failure. |  | MONOPOLAR SCISSOR DAVINCI | ROBOTIC DEVICE ACCESSORY | 841269 | 856777 | 819453 | 856777 | 1 | GEI | User Facility |  |  | Initial | 2007-05-02 | 2007-05-15 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2007
| 2007-05-03 | Malfunction |  | The patient underwent robotic-assisted total laparoscopic hysterectomy with bilateral salpingo-oophorectomy for fibroid tumors. During the procedure, one of the blades on the robotic scissors broke off. |  | DA VINCI S ENDOWRIST | ROBOTIC SURGICAL INSTRUMENT CU | 841419 | 856995 | 819671 | 856995 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2007-05-14 | 2007-05-15 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2007
37094602 | 2007-05-14 | Malfunction |  | Surgeon using robot for robot assisted total abdominal hysterectomy/bilateral salpingo-oophorectomy (tah/bso) when he began having trouble with robotic 3rd arm. A surgery technician discovered that the cannula holder had broken. The instrument was removed, the cannula holder was replaced and the case proceeded without further incident. The patient was not harmed. |  | CANNULA HOLDER TWISTLOCK | ROBOTIC DEVICE ACCESSORY | 858143 | 878900 | 841343 | 878900 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-05-15 | 2007-05-15 | No | Yes | Physician | 37094602 | Yes | No Answer Provided |  |  |  |  | 2007
DA VINCI S | 2007-03-29 | No Answer Provided |  | One of the metal connectors from the bifurcated light guide cables became disconnected from the endoscope as the physician was handling the endoscope. The light guide cable landed on the surgical drape that was covering the patient on the operative table. A 1/8-inch burn hole was noted on the top layer of the drape. The lower layer of the drape was intact. No injury to patient. The concern is that the unit becomes very hot, hot enough to burn drapes quickly. For these types of urologic procedures, the surgery staff sets up the camera, light cable and endoscope on a separately draped table. There is a metal tray placed on this table for the placement of the light cable/endoscope when not in use to avoid the drapes being burned. The surgeons are to place the removed scope and scope assembly on this metal tray on this separately draped table prepared exclusively for this purpose. The surgeons are not to place the scope and scope assembly on the drapes covering or next to the patient or the operating table. There is no room available to place a metal tray on the patient and or the operating table. The heat from the scope is very intense, because the surgical tech's visual awareness is obscured due to the size, draping, staff placement. The surgery tech can not see when the surgeon is taking the scope from the trocar and where exactly the surgeon is setting it down. The manufacturer's literature describes no safety procedures to prevent the risk. Also, no equipment is provided for the placement of device between steps in the procedure. The use of the metal pan was improvised by surgery staff as a safety measure. Surgery staff believe that the risk of fire is high. There have been two other similar cases involving drape burns with no patient injuries at our facility. |  | DA VINCI S | ROBOTIC DEVICE | 850948 | 869814 | 832370 | 869814 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-04-24 | 2007-05-10 | No | No | Physician | DA VINCI S | Yes | No Answer Provided |  |  |  |  | 2007
|  | Malfunction |  |  | The instrument was not returned for evaluation, however, it was determined that the reported complaint is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. This site has previously returned two cannulae with a raised edge which may have been used in conjunction with the instrument. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 851715 | 870950 | 807145 | 2955842200700112 | 1 | GCJ | Manufacturer | Other |  | Initial | 2007-05-10 | 2007-05-10 | [5 empty] | No | 2007-04-10 | Device Not Retu | No | Unkown | 2007
42000202 |  | Malfunction |  | It was reported that during a surgical procedure, a cannula accessory came apart at the hub creating an air leak. Additional insufflation and anesthesia were needed to successfully complete the scheduled procedure. No patient harm, adverse outcome or injury was reported. Three cannula's will be returning for investigation, two of which have been reported as breaking at the hub. 2955842-20074-00109, 2955842-2007-00111. | The cannula accessory has not yet returned for evaluation. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882101 | 907584 | 869818 | 2955842200700110 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-05-09 | 2007-05-09 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-04-09 | Device Not Retu | No | Unkown | 2007
42000202 |  | Malfunction |  | It was reported that during a surgical procedure, a cannula accessory came apart at the hub creating an air leak. Additional insufflation and anesthesia were needed to successfully complete the scheduled procedure. No patient harm, adverse outcome or injury was reported. Three cannula's will be returning for investigation, two of which have been reported as breaking at the hub. 2955842-20074-00110, 2955842-2007-00111. | The cannula accessory has not yet returned for evaluation. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882103 | 907585 | 869819 | 2955842200700109 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-05-09 | 2007-05-09 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-04-09 | Device Not Retu | No | Unkown | 2007
42000202 |  | Malfunction |  | It was reported that during a surgical procedure, a cannula accessory came apart at the hub creating an air leak. Additional insufflation and anesthesia were needed to successfully complete the scheduled procedure. No patient harm, adverse outcome or injury was reported. Three cannula's will be returning for investigation, two of which have been reported as breaking at the hub. 2955842-20074-00109, 2955842-2007-00110. | The cannula accessory has not yet returned for evaluation. | CANNULA 8M REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882112 | 907588 | 869822 | 2955842200700111 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-05-09 | 2007-05-09 | No | Yes | Health Professional | 42000202 | Yes | No | 2007-04-09 | Device Not Retu | No | Unkown | 2007
40011008 |  | Malfunction |  | It was reported that the tip of the precise bipolar pyramid tip instrument was bent. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint as no damage was observed on the instrument tip and the grips were properly aligned. During failure analysis, engineering also observed deep scratches at the distal end of the main tube, including a scratch where material had been removed. The damage to the main tube does appear to be caused by a defective cannula. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882160 | 907599 | 869833 | 2955842200700092 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-27 | 2007-04-27 | No | Yes | Health Professional | 40011008 | Device Returned To Manufacturer | No | 2007-04-10 | Yes | No | Reuse | 2007
40018302 |  | Malfunction |  | It was reported that the permanent cautery hook instrument was not cauterizing. No additional details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found signs or arcing on the instrument yaw pulley. Engineering also observed that the ceramic sleeve was missing from the instrument hook, which most likely created a path for the arcing to occur. This complaint is being reported due to the evidence of arcing found, which has led engineering to conclude that the instrument may have arced during a previous surgical procedure, thus contributing to the malfunction experienced by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882173 | 907601 | 869835 | 2955842200700103 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-04-20 | 2007-04-27 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | No | 2007-03-30 | Yes | No | Reuse | 2007
40017905 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument began splintering. It was also reported that no components had fallen off into a pt. No additional info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint failure analyses observed a four inch long section of scraping on the outer diameter of the main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881974 | 907369 | 869603 | 2955842200700107 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 40017905 | Yes | No | 2007-04-13 | Yes | No | Reuse | 2007
42017204 | 2007-04-11 | Malfunction |  | It was reported that during a prostatectomy procedure, while the maryland bipolar forceps instrument was being removed, the tip broke and fragments of the instrument fell inside of the pt. All the fragments were retrieved and the procedure was successfully completed using a replacement maryland bipolar forceps instrument. No pt harm was reported. | The instrument was returned for eval. Per the customer reported complaint, engineering observed that the instrument was broken at the interface between the proximal clevis and main tube. Both the proximal clevis and the main tube were missing pieces. The clevis is dislodged from the tube as a result of the breakage. As indicated in the complaint notes, all broken pieces were successfully retrieved. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 882019 | 907419 | 869653 | 2955842200700101 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | Yes | 2007-04-11 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00189, 2955842-2007-00104, 2955842-2007-00105. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a misalignment at the bowl/tube interface, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882022 | 907420 | 869654 | 2955842200700106 | 1 | KOG | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-12 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00189, 2955842-2007-00105, 2955842-2007-00106. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a misalignment at the bowl/tube interface, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882043 | 907507 | 869741 | 2955842200700104 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-12 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00189, 2955842-2007-00104, 2955842-2007-00106. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a misalignment at the bowl/tube interface, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882050 | 907520 | 869754 | 2955842200700105 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-12 | Yes | No | Unkown | 2007
42009304 | 2007-03-28 | Malfunction |  | It was reported that during the surgical procedure the tip of the endowrist prograsp instrument broke. A broken piece of the instrument fell inside of the pt and was retrieved. When attempting to remove the instrument, the trocar became stuck on the instrument. As a result the instrument and trocar were removed from the pt and the customer had to cut the tip of the instrument off to remove it from the trocar. The planned procedure was completed using a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned with the main tube and wrist detached. Per the customer reported complaint, engineering observed a loss of rigidity on a section of the distal end of the main tube, at the proximal clevis interface. This observation indicates that the failure mode was most likely a break at the proximal clevis to main tube interface, with the main tube absorbing the majority of the damage. As indicated in the complaint notes, the broken piece was successfully retrieved. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882063 | 907527 | 869761 | 2955842200700102 | 1 | KOG | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2007-04-25 | 2007-04-25 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-03-28 | Yes | No | Reuse | 2007
IS2000 A50P5 | 2007-03-22 | Other | Other |  | In 2007: the field service engineering was unable to duplicate the customer's reported issue. The root cause of the customer's difficulties were not identified. On 3/26/07: the field service engineer placed a follow up visit to this account to monitor a case. The customer reported issue was not duplicated and the root cause of the customers difficulties were not identified. As of 4/23/07, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 824622 | 844818 | 800210 | 2955842200700082 | 1 | NAY | Manufacturer | User facilityCompany Representative | BIOMEDICAL ENGINEER | Initial | 2007-04-24 | 2007-04-24 | Yes | No | Health Professional | IS2000 A50P5 | Yes | No | 2007-03-22 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00058, #2955842-2007-00097, #2955842-2007-00098 #2955842-2007-00099. | The cannula accessory was evaluated. Engineering determined that the taper angle of the cannula accessory may possibly be too large at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM LONG IS2000 | ENDOSCOPIC ACCESSORY | 881913 | 907351 | 869585 | 2955842200700100 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00058, #2955842-2007-00097, #2955842-2007-00098 #2955842-2007-00100. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was slightly bent at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM LONG IS2000 | ENDOSCOPIC ACCESSORY | 881916 | 907352 | 869586 | 2955842200700099 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000604 | 2007-03-30 | Malfunction |  | It was reported that the shaft of the large needle driver instrument was splintering. No add'l info was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed some light scraping on one side of the distal end of the instrument main tube. Engineering also observed some scratches on the outer surface of the instrument grips. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881917 | 907353 | 869587 | 2955842200700081 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-24 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-03-30 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00058, #2955842-2007-00097, #2955842-2007-00099 #2955842-2007-00100. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was misaligned at the bowl/tube interface,thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 881919 | 907354 | 869588 | 2955842200700098 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00058, #2955842-2007-00098, #2955842-2007-00099 #2955842-2007-00100. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was misaligned at the bowl/tube interface,thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 881924 | 907355 | 869589 | 2955842200700097 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -00179, 2955842-2007-00085, -00086, -00087, -00088, -00089, -00090, -00091. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 881925 | 907356 | 869590 | 2955842200700084 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000401 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00055, #2955842-2007-00093, #2955842-2007-00094 #2955842-2007-00095. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was misaligned at the distal tip,thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM LONG IS2000 | ENDOSCOPIC ACCESSORY | 881927 | 907357 | 869591 | 2955842200700096 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000401 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -00179, 2955842-2007-00084, -00085, -00087, -00088, -00089, -00090, -00091. | Conclusions: the cannula accessory was evaluated. Engineering observed that the cannula accessory has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 881929 | 907358 | 869592 | 2955842200700086 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42011003 | 2007-03-19 | Malfunction |  | It was reported that during a hysterectomy procedure the precise bipolar forceps instrument had broken and pieces of the instrument fell inside of the pt. All broken pieces were retrieved using a grasper through the accessory port. The planned procedure was completed using a replacement precise bipolar forceps instrument. No pt harm was reported. | The instrument was returned for eval. Per the customer reported complaint, engineering observed that the instrument was broken at the interface between the proximal clevis and main tube. Both the proximal clevis and the main tube were missing pieces. As indicated in the complaint notes, all broken pieces were successfully retrieved. Engineering also observed scraping at the distal end of the main tube. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881968 | 907367 | 869601 | 2955842200700079 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-04-20 | 2007-04-24 | No | Yes | Health Professional | 42011003 | Device Returned To Manufacturer | Yes | 2007-03-29 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00067, #2955842-2007-00075, #2955842-2007-00076, # 2955842-2007-00077. | Conclusions: - the cannula accessory was evaluated. Engineering observed a small area at the distal tip with material removed. It was determined that the damaged area could create a rough surface and may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPY ACCESSORY | 881990 | 907373 | 869607 | 2955842200700078 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-03-30 | Yes | No | Unkown | 2007
4200021 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00067, #2955842-2007-00075, #2955842-2007-00076, # 2955842-2007-00078. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 881998 | 907394 | 869628 | 2955842200700077 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-24 | No | Yes | Health Professional | 4200021 | Yes | No | 2007-03-30 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00067, #2955842-2007-00075, #2955842-2007-00077, # 2955842-2007-00078. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the distal tip, thus creating a raised ledge which may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882011 | 907417 | 869651 | 2955842200700076 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-03-30 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  |  | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a bent edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882071 | 907564 | 869798 | 2955842200700085 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: #2955842-2007-00067, #2955842-2007-00076, #2955842-2007-00077, # 2955842-2007-00078. | The cannula accessory was evaluated. Engineering observed that the cannula accessory had a burr located approximately two inches from the distal tip. It was determined that the burr may potentially remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882184 | 907603 | 869837 | 2955842200700075 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-03-30 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -00179, 2955842-2007-00084, -00085, -00086, -00088, -00089, -00090, -00091. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a bent edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882428 | 907840 | 870076 | 2955842200700087 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -00179, 2955842-2007-00084, -00085, -00086, -00087, -00089, -00090, -00091. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882438 | 907843 | 870079 | 2955842200700088 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, 00177, -00178, -00179, 2955842-2007-00084, -00085, -00086, -00087, -00088, -00090,-00091. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882448 | 907848 | 870084 | 2955842200700089 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -00179, 2955842-2007-00084, -00085, -00086, -00087, -00088, -00089, -00091. | The cannula accessory was evaluated. Engineering observed that the cannula accessory has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882455 | 907852 | 870088 | 2955842200700090 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00154, -00176, -00177, -00178, -000179, 2955842-2007-00084, -00085, -00086, -00087, -00088, -00089, -00090. | The cannula accessory was evaluated. Engineering observed that the cannula has a deformation on the lip of the distal tip, thus creating a raised edge which may cause scraping of an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882518 | 908026 | 870262 | 2955842200700091 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00055, -00094, -00095, -00096. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was misaligned at the bowl/tube interface, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882524 | 908048 | 870284 | 2955842200700093 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00055, -00093, -00095, -00096. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was misaligned at the bowl/tube ineterface, thus creating a raised ledge which may remove material from an insturment during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882526 | 908049 | 870285 | 2955842200700094 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No pt harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2007-00055, -00093, -00094, -00096. | The cannula accessory was evaluated. Engineering observed that the cannula accesory was damaged at the distal tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882535 | 908051 | 870287 | 2955842200700095 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-23 | 2007-04-24 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-04-11 | Yes | No | Unkown | 2007
IS1200 A43P7 | 2007-04-09 | Malfunction |  | It was reported that during the beginning of a prostatectomy procedure, the surgeon experienced a recurring 23008 system error code. No pt harm was reported. The pt's incisions were closed and the planned surgical procedure was rescheduled. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered a defective potentiometer assembly within the mtmr, thus generating the system error experienced by the customer. The system alarm (the 23008 system error code) functioned as designed and there was no injury to the pt. As of (b)(4) 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2333820 | 2304480 | 2201376 | 2955842200700083 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | IS1200 A43P7 | Yes | No | 2007-04-09 | Yes | No | Reuse | 2007
42017204 | 2007-02-15 | Malfunction |  | It was reported that during a radical prostatectomy the (b)(4) bipolar forceps instrument "did not hold grip". No add'l info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint failure analyses observed the grips to open and close properly. Also observed was a 4 inch long section with scraping at distal end of main tube. Damage is similar to that caused by a defective cannula. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2333891 | 2304494 | 2201390 | 2955842200700072 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-04-24 | 2007-04-24 | No | Yes | Health Professional | 42017204 | Yes | Yes | 2007-03-28 | Yes | No | Reuse | 2007
42000604 | 2007-03-30 | Malfunction |  | It was reported that the shaft of the large needle driver instrument was splintering. No add'l info was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a combination of scraping and deep scratches on the distal end of the instrument main tube. Engineering also observed that material had been removed in the damaged area. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. Which may remove material from the instrument during use. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882416 | 907839 | 870075 | 2955842200700080 | 1 | KOG | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-04-19 | 2007-04-23 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-03-30 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00181, 2955842-2006-00161, -00162, -00163, -00164, -00182, -00183, -00184, -00185, -00186, -00187, -00188, -00192, -00070, -00074. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 827428 | 840006 | 802880 | 2955842200700071 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-13 | 2007-04-19 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-03-23 | Yes | No | Unkown | 2007
IS1200 A43P7 | 2007-03-20 | Other |  | It was reported that during a da vinci prostatectomy, an error 7 occurred causing the account to convert to an open prostatectomy to complete the planned surgical procedure. No patient harm was reported. | The fse evaluated the system on site and confirmed the system error. The system was repaired by replacing a pca in the system. The repaired system passed all tests and was returned to service. As of april 19, 2007, there have been no reported recurrences of this failure mode. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 881959 | 907364 | 869598 | 2955842200700069 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-19 | 2007-04-19 | No | Yes | Health Professional | IS1200 A43P7 | Device Returned To Manufacturer | No | 2007-03-20 | Yes | No | Reuse | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00181, 2955842-2006-00161, 2955842-2006-00162, 2955842-2006-00163, 2955842-2006-00164, 2955842-2006-00182, 2955842-2006-00183, 2955842-2006-00184, 2955842-2006-00185, 2955842-2006-00186, 2955842-2006-00187, 2955842-2006-00188, 2955842-2006-00192, 2955842-2007-00071, 2955842-2007-00074. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882147 | 907595 | 869829 | 2955842200700070 | 1 | KOG | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-13 | 2007-04-19 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-03-23 | Yes | No | Unkown | 2007
42000201 |  | Malfunction |  | The cannula accessory was returned as part of a field removal action. No patient harm, adverse outcome or injury was reported. The following medwatch reports listed below were also sent to the fda. These reports only pertain to the field removal action from this specific site: 2955842-2006-00181, 2955842-2006-00161, 2955842-2006-00162, 2955842-2006-00163, 2955842-2006-00164, 2955842-2006-00182, 2955842-2006-00183, 2955842-2006-00184, 2955842-2006-00185, 2955842-2006-00186, 2955842-2006-00187, 2955842-2006-00188, 2955842-2006-00192, 2955842-2007-00070, 2955842-2007-00071. | The cannula accessory was evaluated. Engineering observed that the cannula accessory was bent at the tip, thus creating a raised ledge which may remove material from an instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 882155 | 907598 | 869832 | 2955842200700074 | 1 | KOG | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-18 | 2007-04-19 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-04-11 | Yes | No | Unkown | 2007
42000604 | 2007-03-12 | Malfunction |  | It was reported that during a radical prostatectomy the tip of the large needle driver instrument broke off and fell inside of the patient. The physician chose to leave the piece in the pt and completed the procedure using a replacement large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one entire jaw broken off. The broken piece was not returned. The jaw is broken off at the swage, partially exposing the cable crimp. The other grip has a crack in the same location as one of the fracture surfaces. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882220 | 907628 | 869862 | 2955842200700073 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-04-19 | 2007-04-19 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | Yes | 2007-03-26 | Yes | No | Reuse | 2007
| 2007-03-12 | Malfunction |  | The patient was in the or suite undergoing a robot assisted laparoscopic radical prostatectomy. The large needle driver broke during the procedure and a small piece of the tip was left in the patient per the surgeon. It was felt that it would do more harm to retrieve it. |  | DA VINCI | ROBOTIC DEVICE NEEDLE | 856559 | 877278 | 839734 | 877278 | 1 | GCJ | User Facility |  |  | Initial | 2007-04-11 | 2007-04-11 | No | Yes | Invalid Data |  | No | No Answer Provided |  |  |  |  | 2007
42017204 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument experienced "scraping of black coating". It is unknown if the event occurred during a procedure. No additional information has been provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint failure analyses observed a very thin section with scraping and various scratches on the distal end of the main tube. This site has previously returned bent cannulas which were the likely cause of scraping. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881962 | 907366 | 869600 | 2955842200700067 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-11 | 2007-04-11 | No | Yes | Health Professional | 42017204 | Yes | No | 2007-03-15 | Yes | No | Reuse | 2007
42000604 | 2007-03-12 | Malfunction |  | It was reported that during the surgical procedure, the wrist of the large needle driver instrument collided against the cannula accessory. Towards the end of the surgical procedure, it was noticed that the instrument was broken and a piece of the instrument had fallen inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. The proximal clevis was dislodged from the main tube as a result of the breakage. Both the proximal clevis and main tube were missing pieces. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis at the main tube interface. The cannula accessory was also evaluated and no damage was found. As indicated in the complaints notes, the broken piece was successfully retrieved from the patient. | LARGE NEEDLE DRIVE INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881972 | 907368 | 869602 | 2955842200700066 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-04-10 | 2007-04-11 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-03-12 | Yes | No | Reuse | 2007
420172 | 2007-03-29 | Malfunction |  | While in a patient's abdomen, the maryland bipolar endo forceps looked like it had shavings or particulate matter on it. The device was removed and replaced. It remains unknown whether these shavings entered the patient. |  | DA VINCI MARYLAND BIPOLAR ENDO FORCEPS | SURGICAL ROBOT ACCESSORY | 827728 | 840431 | 803300 | 840431 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-04-06 | 2007-04-06 | No | Yes | Physician | 420172 | Yes | No Answer Provided |  |  |  |  | 2007
40018301 |  | Malfunction |  | It was reported that the permanent cautery hook instrument stopped functioning. No further info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering eval observed that the instrument proximal clevis and yaw pulley were charred on one side. Engineering also discovered that electrical continuity fails with the instrument wrist at maximum yaw positions, indicating damage to the conductor wire. These findings may have led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881888 | 907345 | 869579 | 2955842200700056 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-04-02 | 2007-04-06 | No | Yes | Health Professional | 40018301 | Yes | Yes | 2007-03-08 | Yes | No | Reuse | 2007
40009306 |  | Malfunction |  | It was reported that the shaft of the prograsp forceps began splintering. It was reported that no components had fallen off into a pt. No add'l info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analyses observed surface scraping on one side of the tube. The surface finish is rough and the color of the tube is gray at the damaged area. This site has previously returned bent cannulas which were the likely cause of scraping. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881897 | 907347 | 869581 | 2955842200700059 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 40009306 | Yes | No | 2007-03-09 | Yes | No | Reuse | 2007
40009304 |  | Malfunction |  | It was reported that the shaft of the prograsp forceps began splintering. It was reported that no components had fallen off into a pt. No add'l info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analyses observed deep scratches at the distal end of main tube causing white marks. This damage to tube does not seem to be caused by a defective cannula. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881902 | 907348 | 869582 | 2955842200700060 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 40009304 | Yes | No | 2007-03-09 | Yes | No | Reuse | 2007
42000604 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument began splintering. It was reported that no components had fallen off into a pt. No add'l info has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analyses observed a 1 inch long section of light scraping around the outer diameter. This site has previously returned bent cannulas which were the likely cause of scraping. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 881905 | 907349 | 869583 | 2955842200700061 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-03-09 | Yes | No | Reuse | 2007
42021403 |  | Malfunction |  | It was reported that the shaft of the pk dissecting forceps instrument began splintering. It was also reported that no components had fallen off into a pt. No add'l info has been provided. No pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the reported complaint failure analyses observed severe scraping on the outer diameter of the main tube. This site has previously returned bent cannulas which were the likely cause of scraping. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881911 | 907350 | 869584 | 2955842200700062 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 42021403 | Device Returned To Manufacturer | No | 2007-03-09 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument began splintering. It was also reported that no components had fallen off into a patient. No additional information has been provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analyses observed a two inch long section of scraping on the outer diameter of the main tube. This site has previously returned bent cannulas which were the likely cause of scraping. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881947 | 907362 | 869596 | 2955842200700063 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-03-09 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that the shaft of the monopolar curved scissors instrument began splintering. It was also reported that no components had fallen off into a patient. No additional information has been provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, failure analyses observed a two inch long section of scraping on the outer diameter of the main tube. This site has previously returned bent cannulas which were the likely cause of scraping. | MONOPOLARCURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881952 | 907363 | 869597 | 2955842200700064 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-05 | 2007-04-06 | No | Yes | Health Professional | 42017904 | Yes | No | 2007-03-09 | Yes | No | Reuse | 2007
IS1200 A43P7 | 2007-03-08 | Malfunction | Life Threatening | It was reported that when changing the instrument during the surgical procedure, the instrument advanced into the patient without being guided. The planned surgical procedure was completed and no patient harm or injury was reported. | Investigation was conducted by field service engineering. Testing was performed on the system and the customer reported complaint could not be replicated. The patient side manipulator (psm) was replaced as a precaution. The psm was returned to isi for failure analysis investigation. Engineering was also unable to replicate the customer reported complaint and found no issues with the psm. Engineering believes that the issue experienced by the customer may have been a normal operation of the guided tool change, a feature of the system that allows the instrument to return to the exact position it was before, after an instrument exchange, however, no definitive conclusion can be drawn. As of april 5, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 881980 | 907370 | 869604 | 2955842200700057 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-04-05 | 2007-04-06 | Yes | Yes | Health Professional | IS1200 A43P7 | No | Yes | 2007-03-08 | Device Not Retu | No | Reuse | 2007
42009304 |  | Malfunction |  | It was reported that the is2000 endowrist prograsp instrument did not register on the sys. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was tested for recognition and passed. The instrument had full range of motion in all directions and engineering was unable to replicate the customer reported complaint. During eval, engineering observed a scrape on the distal end of the instrument main tube. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | IS2000 ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881884 | 907344 | 869578 | 2955842200700055 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-04-02 | 2007-04-04 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | Yes | 2007-03-05 | Yes | No | Reuse | 2007
42017904 |  | Malfunction |  | The monopolar curved scissors instrument was returned as an unk failure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was tested and found not to cut clearly, due to damage of the cutting edge of the scissors. During eval, engineering observed major wear on the main tube starting at midpoint of the tube and spreading distally. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 881894 | 907346 | 869580 | 2955842200700058 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-04-03 | 2007-04-04 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2007-04-03 | Yes | No | Reuse | 2007
42009304 | 2007-02-27 | Malfunction |  | It was reported that during the surgical procedure, the surgeon collided the prograsp forceps instrument with another instrument. As a result, the instrument tip snapped and the proximal clevis shattered and pieces fell inside of the pt. All visible pieces were removed and the planned surgical procedure was completed using a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. The proximal clevis was dislodged from the main tube as a result of the breakage. Both the proximal clevis and the main tube were missing pieces. As indicated in the case notes, all visible broken pieces were successfully retrieved. The case notes also indicate that the surgeon collided the instrument with another instrument, thus causing the instrument breakage. Based on this info, it was determined that user error caused the event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 881874 | 907342 | 869576 | 2955842200700053 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-03-28 | 2007-03-30 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-02-27 | Yes | No | Initial | 2007
IS2000 A50P5 | 2007-03-01 | Other | Other | It was reported that during the surgical procedure, the sys arms were moving without prompting. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Investigation was conducted by field svc engineering. Extensive troubleshooting was performed and the sys was found to be functioning properly. Add'l info provided by the site indicates that the surgeon was moving the sys arms erratically, which may have led to the "non-intuitive" motion experienced by the surgeon. As of march 29, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 881880 | 907343 | 869577 | 2955842200700054 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-03-29 | 2007-03-30 | Yes | Yes | Health Professional | IS2000 A50P5 | No | Yes | 2007-03-01 | No Answer Provi | No | Reuse | 2007
42009304 | 2007-01-17 | Malfunction |  | It was reported that prograsp forceps instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014; mfr. Report #2955842-2007-00015; mfr. Report # 2955842-2007-00016; mfr. Report # 2955842-2007-00017; mfr. Report #2955842-2007-00018, mfr. Report # 2955842-2007-00019, mfr. Report #2955842-2007-00020, mfr. Report #2955842-2007-00021; mfr. Report #2955842-2007-00022; mfr. Report #2955842-2007-00023. | The instrument was returned and evaluated. Engineering observed that the distal end of the main tube has a few deep scratches. Per the customer reported complaint, it was determined that the failure mode is not consistent with the use of a potentially defective cannula accessory. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 882194 | 907623 | 869857 | 2955842200700051 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-03-28 | 2007-03-28 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-02-26 | Yes | No | Reuse | 2007
40009306 | 2007-02-26 | Malfunction |  | It was reported that the black shaft of the prograsp forceps instrument felt "splintery". This was determined during inventory and not during a procedure. No additional information has been provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of main tube has various long, narrow scrape marks. Material has been removed. The damage is not similar to that caused by a bent cannula. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 882209 | 907625 | 869859 | 2955842200700052 | 1 | KOG | Manufacturer | User facility | UNKNOWN | Initial | 2007-03-28 | 2007-03-28 | No | Yes | Health Professional | 40009306 | Device Returned To Manufacturer | No | 2007-02-26 | Yes | No | Reuse | 2007
DA VINCI | 2007-03-22 | Injury | Required InterventionOther | During a cardiac robotic surgical procedure using a davinci surgical system, the robot's 3rd instrument arm (arm4) stopped functioning. The surgeon removed the instrument from the affected arm and placed it on arm1 after which the instrument tested okay. The surgeon then tried unsuccessfully to reattach the instrument to arm4 three different times, as instructed by the manufacturer's clinical representative, who was present during this surgical procedure. The surgeon had to make another incision in the patient to attach an atrium retraction instrument to complete the procedure. The manufacturer representative finally placed arm4 in the complete stow position and redeployed arm to a cold reboot status. After the reboot of arm4, the system was then able to detect the instrument and the surgeon completed the case without having to use the additional incision site, and without further robotic system failure. |  | DAVINCI SURGICAL SYSTEM | SURGICAL ROBOT | 824622 | 837314 | 800210 | 837314 | 1 | GCJ | User Facility |  |  | Initial | 2007-03-26 | 2007-03-26 | Yes |  | Physician | DA VINCI | Yes | No Answer Provided |  |  |  |  | 2007
|  | Other | Other | A myomectomy was successfully completed using the da vinci system. A few days after the procedure, the pt returned complaining of abdominal pain. It was determined that the pt had a perforation in the ileum portion of her intestines. The pt's current condition is unk. No add'l info has been provided. | The instruments used for the procedure were not returned for failure analysis. Root cause is undetermined. There have been no related complaints from this account since 2/20/07. |  |  | 820519 | 833211 | 796167 | 2955842200700048 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-03-22 | 2007-03-23 | Yes | No | Health Professional |  | No | No | 2007-02-20 | Device Not Retu | No | Unkown | 2007
40017905 | 2007-02-23 | Malfunction |  | It was reported that during the surgical procedure the monopolar curved scissors instrument broke and a piece of the orange cover fell inside of the pt. The broken piece was retrieved. Prior to the instrument breaking, another monopolar curved scissors instrument being used during the surgical procedure broke and a small piece fell inside of the pt. The broken piece was retrieved. Medwatch mfr report # 2955842-2007-00049 was created to report the first instrument breakage. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tube extension was severely broken and pieces were missing at the proximal clevis interface. Add'l info was provided by an isi rep. It was discovered that the customer did not install the tip cover accessory on either instrument prior to use. It was determined that if the tip cover accessory had been properly installed on the instruments, the broken piece would have been prevented from falling into the pt. Based on this info, it was concluded that user error directly contributed to the event. Per the case notes, the broken piece was successfully retrieved from the pt. The monopolar curved scissors instrument instructions for use supplement specifically states: general precautions and warnings: the endowrist monopolar curved scissors instrument must always be used in conjunction with the tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890169 | 917372 | 879602 | 2955842200700050 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-02-23 | 2007-03-23 | Yes | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | Yes | 2007-02-23 | Yes | No | Reuse | 2007
40017905 | 2007-02-23 | Malfunction |  | It was reported that during the surgical procedure the orange cover of the monopolar curved scissors instrument broke and a small piece fell inside of the pt. The broken piece was retrieved. It was also reported that another monopolar curved scissors instrument used during the surgical procedure broke and a piece of the orange cover fell inside of the pt. The broken piece was retrieved. Medwatch mfr report #2955842-2007-00050 was created to report the second instrument breakage. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tube extension was broken and a piece was missing at the proximal clevis interface. Add'l info was provided by an isi rep. It was discovered that the customer did not install the tip cover accessory on either instrument prior to use. It was determined that if the tip cover accessory had been properly installed on the instruments, the broken piece would have been prevented from falling into the pt. Based on this info, it was concluded that user error directly contributed to the event. Per the case notes, the broken piece was successfully retrieved from the pt. The monopolar curved scissors instrument instructions for use supplement specifically states: general precautions and warnings: the endowrist monopolar curved scissors instrument must always be used in conjunction with the tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 890211 | 917488 | 879718 | 2955842200700049 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-02-23 | 2007-03-23 | Yes | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | Yes | 2007-02-23 | Yes | No | Reuse | 2007
| 2007-01-26 | Malfunction |  | At the end of a laparoscopic radical prostatectomy with robotic assistance, it was noted that the coating had rubbed off of the prograsp instrument. The patient's abdomen was irrigated thoroughly and no noted particles remained. Staff is aware that the coating can rub off from continual friction because of the positioning of the instrument. Because of draping, the field is hard to visualize. The davinci surgical team was made aware of the potential for this problem and will monitor the instruments while in use. Note: the prograsp instrument is tracked on number of uses. It was not documented as to how many times (out of ten), this particular prograsp had been used at the time of occurrence. |  | ENDOWRIST PROGRASP FORCEPS | SURGICAL ROBOT FORCEPS | 831534 | 844349 | 807145 | 844349 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-03-21 | 2007-03-21 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2007
40016906 | 2007-02-20 | Malfunction |  | It was reported that during a gastrectomy, the tip of the harmonic curved shears insert broke off and had fallen in the pt. The broken piece was retrieved without incident. The planned procedure was successfully completed using another harmonic curved shears instrument. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was missing from the harmonic shears insert. One side of clamp arm attachment area was bent outward, creating a large gap. Visual evidence suggests this joint was subjected to excessive forces, causing the clamp arm to break off. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 843102 | 859887 | 822538 | 2955842200700046 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2007-03-21 | 2007-03-21 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | No | 2007-02-20 | Yes | Yes | Unkown | 2007
42009304 | 2007-02-21 | Malfunction |  | It was reported that during a procedure, the physician lost control of the prograsp forceps instrument. The physician chose to replace the instrument, but had difficulties removing it from the trocar. This resulted in the removal of the trocar. While removing the trocar the proximal clevis of the instrument broke and pieces of the instrument fell inside of the pt. All visible pieces were removed and the planned surgical procedure was successfully completed using another prograsp forceps instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the wrist missing. The wrist was cut off at the distal end of the main tube. Engineering's evaluation determined that the failure mode experienced by the customer was a break at the proximal clevis to main tube interface. As indicated in the complaints notes, all visible broken pieces were successfully retrieved. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 843660 | 860563 | 823219 | 2955842200700047 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-03-21 | 2007-03-21 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-02-21 | Yes | No | Reuse | 2007
40000620 | 2007-02-14 | Malfunction |  | It was reported that towards the end of the surgical procedure, it was noticed that parts on each side of the jaw of the large needle driver instrument were missing. A fragmented peice was retrieved from the pt, however, the other missing piece was not located. The planned surgical procedure was completed with a slight delay. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported complaint cannot be determined. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 851525 | 870612 | 833168 | 2955842200700045 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-03-15 | 2007-03-16 | No | Yes | Health Professional | 40000620 | No | No | 2007-02-16 | Device Not Retu | No | Unkown | 2007
42000604 | 2007-02-14 | Malfunction |  | It was reported that during the surgical procedure, part of the jaw of the large needle driver instrument fell inside of the patient. The fragmented piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with the carbide insert missing from one grip. Engineering evaluation observed discolored areas on the brazing surface of the grip and area where the bare grip is visible. This discover leads engineering to conclude that the carbide insert may have not been properly brazed onto the grip, thus causing the carbide insert to loosen and fall off. As indicated in the case notes, the carbide insert was successfully retrieved from the patient. No other instruments from this manufactured lot have exhibited this failure. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 852104 | 871450 | 833993 | 2955842200700044 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-03-13 | 2007-03-16 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-02-16 | Yes | No | Reuse | 2007
40020705 | 2007-01-24 | Malfunction |  | On 2/22/07, medwatch uf/importer report was received. The event description is as follows: "the tip of the 5mm tenaculum instrument broke off in the pt, which required additional procedures to be done. X-ray was taken and hand assisted gel port was used to remove the broken instrument tip. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working)". | Investigation performed indicates that the instrument used during the surgical procedure was an 8mm tenaculum forceps instrument and not a "5 mm myomectomy tenaculum" instrument as referenced on the customer's medwatch report. Further investigation revealed that a myomectomy procedure was being performed at the time of the incident, which most likely caused the customer to incorrectly refer to the instrument as a "5 mm myomectomy tenaculum". The instrument was discarded by the site, therefore, the root cause of the customer reported complaint cannot be determined. As indicated in the event description, the broken piece was successfully retrieved from the pt. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 805946 | 974680 | 781771 | 2955842200700039 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2007-03-12 | 2007-03-16 | Yes | Yes | Health Professional | 40020705 | No | Yes | 2007-02-22 | Device Not Retu | No | Unkown | 2007
IS1200 A43P7 | 2007-02-13 | Other | Other | It was reported that during the surgical procedure, the customer experienced blurry vision. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering found that the light output being emitted from the system illuminators was below specification. The low light output was most likely interpreted by the customer as "blurry vision". It was determined that the low light output was caused by the illuminator bulbs. Replacement bulbs were ordered to repair the system. As of march 14, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 821902 | 834590 | 797512 | 2955842200700043 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-02-13 | 2007-03-15 | Yes | Yes | Health Professional | IS1200 A43P7 | No | Yes | 2007-02-13 | Device Not Retu | No | Reuse | 2007
| 2007-02-13 | Other | Other | It was reported that during a mitra valve repair procedure, the customer experienced difficulties in straightening the wrist of an instrument. A regional support specialist attempted to troubleshoot the issue over the phone. Subsequently, the surgeon chose to convert to an open surgical technique to finish the planned procedure. No pt harm was reported. No add'l info was provided. | The field service engineering was unable to duplicate the customer's reported issue. The root cause of the customer's difficulties were not identified. As a precaution, an isi cardiac clinical sales representative was scheduled to attend the next case. As of march 15, 2007, there have been no reported recurrences of the issue at this hospital. |  | DA VINCI INSTRUMENT | 821910 | 834598 | 797520 | 2955842200700041 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-03-15 | 2007-03-15 | Yes | Yes | Health Professional |  | No | No | 2007-02-13 | No Answer Provi | No | Unkown | 2007
42009304 | 2007-02-07 | Malfunction |  | It was reported during a myomectomy procedure, small fragments from the prograsp forceps instrument's shaft fell into the pt. The fragments were removed. The procedure was successfully completed using a second prograsp forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was not returned and the reported complaint could not be confirmed. | PROGRASP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 820703 | 833395 | 796351 | 2955842200700042 | 1 | GEI | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2007-03-13 | 2007-03-14 | No | Yes | Health Professional | 42009304 | No | No | 2007-02-14 | Device Not Retu | No | Unkown | 2007
42017204 |  | Malfunction |  | It was reported that clips on the side of the maryland bipolar forceps instrument had been depressed inward. No additional info was provided. The hosp did not identify this issue as having occurred during a surgical procedure. | The instrument was returned and evaluated. Engineering did not notice a problem with levers (clips) pressed inward. Engineering did observed areas on the instrument's main tube where material had been removed. It was determined that the failure mode is consistent with the use of the potentially defective cannula accessory, which may removed material from the instrument during use. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 820709 | 833402 | 796358 | 2955842200700040 | 1 | GEI | Manufacturer | User facility | Physician | Initial | 2007-03-14 | 2007-03-14 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | Yes | 2007-02-12 | Yes | No | Reuse | 2007
42021403 | 2007-02-06 | Malfunction |  | It was reported that during the surgical procedure, a portion of the tip of the pk dissecting forceps instrument broke and fell into the patient's peritoneum. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was dislodged from the main tube interface and missing pieces. Engineering determined that the failure mode experienced by the customer was a result of damage to the proximal clevis at the main tube interface. As indicated in the complaints notes, the broken piece was successfully retrieved from the patient. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 818152 | 830808 | 793820 | 2955842200700036 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-03-06 | 2007-03-07 | No | Yes | Health Professional | 42021403 | Device Returned To Manufacturer | No | 2007-02-06 | Yes | No | Reuse | 2007
40017103 | 2007-01-31 | Malfunction |  | It was reported that during the surgical procedure, the tip of the micro bipolar forceps instrument broke and a piece fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken at the main tube interface and a small piece of the proximal clevis was missing. Engineering determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. As indicated in the complaints notes, the broken piece was successfully retrieved from the patient. | MICRO BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 818163 | 830819 | 793831 | 2955842200700035 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2007-03-06 | 2007-03-07 | No | Yes | Health Professional | 40017103 | Device Returned To Manufacturer | Yes | 2007-02-05 | Yes | No | Reuse | 2007
40018301 | 2007-01-25 | Malfunction |  | It was reported that the housing on the back end of the permanent cautery hook instrument had broken. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found evidence that the instrument may have arced during a surgical procedure. Part of the yaw pulley, conducter cap and the distal clevis have been charred. The customer's complaint that the instrument's "housing was broken" was not confirmed by failure analyses. This complaint is being reported via an mdr since the instrument may have arced during a surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 818167 | 830823 | 793835 | 2955842200700037 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-03-07 | 2007-03-07 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2007-02-09 | Yes | No | Reuse | 2007
400179 | 2007-02-08 | Other | Other | It was reported that during an unknown pediatric procedure, a monopolar curved scissors instrument induced a voltage into the maryland bipolar forceps instrument causing inadvertent charring of tissue. The patient's condition is unknown. The following medwatch reports was created for the maryland bipolar forceps instrument. Mfr report #2955842-2007-00033. | Failure analysis was performed on site for the da vinci system and similar instruments as the initial instruments were unavailable. The field service engineer was unable to replicate the event. The account has been advised to record future cases for ongoing evaluation to determine if the cause is due to a malfunction or user error. If additional information becomes available, a follow up mdr will be submitted. | MONOPOLAR CURVED SCISSORS | ELECTROSURGICAL INSTRUMENT | 810836 | 823433 | 786620 | 2955842200700032 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-03-02 | 2007-03-02 | Yes | Yes | Health Professional | 400179 | No | No | 2007-02-08 | No | No | Unkown | 2007
400172 | 2007-02-08 | Other | Other | It was reported that during an unk pediatric procedure, a monopolar curved scissors instrument induced a voltage into the maryland bipolar forceps instrument causing inadvertent charring of tissue. The pt's condition is unk. The following medwatch reports was created for the monopolar curved scissors instrument. Mfr. Report #2955842-2007-00032. | Failure analysis was performed on site for the da vinci system and similiar instruments as the initial instruments were unavailable. The field service engineer was unable to replicate the event. The account has been advised to record future cases for ongoing eval to determine if cause is due to a malfunction or user error. If add'l info becomes available, a follow up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 810846 | 823443 | 786630 | 2955842200700033 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-03-02 | 2007-03-02 | Yes | Yes | Health Professional | 400172 | No | No | 2007-02-08 | No | No | Unkown | 2007
| 2006-12-20 | Injury | Hospitalization | Physician was performing laparoscopic radical prostatectomy with robotic assistance- intuitive surgical da vinci device- when during the procedure, a light bulb on the control tower exploded. Glass from the explosion ended up on to the wall of the o. R. The control tower went dead and the procedure could go no further. The robot was moved away from the pt and the procedure was converted to an open radical prostatectomy. The pt received three blood transfusions during the surgery. His heart rate went up to 190. After surgery, he spent two days in intensive care and another two days on the cardiac care floor. |  | INTUITIVE SURGICAL | DA VINCI ROBOTIC DEVICE | 811731 | 824328 | 787508 | MW1042102 | 1 | GCJ | Voluntary |  | Patient FAMILY MEMBER OR FRIEND | Initial | 2007-03-02 | 2007-03-02 | Yes | Yes | Invalid Data |  | No Answer Provided | No |  |  |  |  | 2007
42009304 | 2006-11-28 | Malfunction |  |  | Conclusion: the instrument was returned with the cannula accessory still attached. Engineering eval observed that no pieces were missing from the instrument proximal clevis, however, the distal end of the main tube had material removed. Engineering also noticed that the cannula accessory was visibly dented at the distal end. This discovery leads engineering to conclude that the dented cannula accessory removed material from the main tube of the instrument during use. The report notes indicate that the instrument fragments were successfully retrieived. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 780408 | 828220 | 756577 | 2955842200700028 | 1 | KOG | Manufacturer | Other | NOT APPLICABLE | Initial | 2007-02-26 | 2007-03-02 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-02-05 | Yes | No | Reuse | 2007
40017905 | 2007-02-01 | Malfunction |  | It was reported that small metal shaving came off of the monopolar curved scissors instrument during sterilization. At this time, the account has not reported an incident of particulate coming off on an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument accessories related complaints from the same account. Mfr report #2955842-2007-00030, mfr report # 2955842-2007-00031. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that part of the main tube had scratches and material removed at the distal end. The cause of the instrument wear is still under investigation. If add'l info becomes available a f/u mdr will be submitted. | MONOPOLAR CURVED SCISSORS | ELECTROSURGICAL INSTRUMENT | 831906 | 844710 | 807499 | 2955842200700029 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-03-01 | 2007-03-02 | No | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | No | 2007-02-01 | Yes | No | Reuse | 2007
40017905 | 2007-02-01 | Malfunction |  | It was reported that small metal shaving came off of the monopolar curved scissors instrument during sterilization. At this time, the account has not reported an incident of particulate coming off on an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument accessories related complaints from the same account. Mfr report #2955842-2007-00029, mfr report # 2955842-2007-00031. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that part of the main tube had scratches and material removed at the distal end. The cause of the instrument wear is still under investigation. If add'l info becomes available a f/u mdr will be submitted. | MONOPOLAR CURVED SCISSORS | ELECTROSURGICAL INSTRUMENT | 831908 | 844712 | 807501 | 2955842200700030 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-03-01 | 2007-03-02 | No | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | No | 2007-02-01 | Yes | No | Reuse | 2007
40000620 | 2007-02-01 | Malfunction |  | It was reported that small metal shaving came off of the large needle driver instrument during sterilization. At this time, the account has not reported an incident of particulate coming off on an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument accessories related complaints from the same account. Mfr report #2955842-2007-00029, mfr report # 2955842-2007-00030. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that part of the main tube had scratches and material removed at the distal end. The cause of the instrument wear is still under investigation. If add'l info becomes available, a f/u mdr will be submitted. | LARGE NEEDLE DRIVER | ELECTROSURGICAL INSTRUMENT | 831915 | 844719 | 807508 | 2955842200700031 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-03-01 | 2007-03-02 | No | Yes | Health Professional | 40000620 | Device Returned To Manufacturer | No | 2007-02-01 | Yes | No | Reuse | 2007
42000201 | 2006-12-20 | Malfunction |  | It was reported that instrument shavings fell inside of the pt. The instrument shavings were not retrieved. The planned surgical procedrue was completed and no add'l procedure was required. No pt harm, adverse outcome or injury was reported. | The instrument and cannula accessory were not returned for eval. The event description is consistent with the use of a potentially defective cannula, which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 832063 | 844879 | 807653 | 2955842200600192 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-03-01 | 2007-03-02 | No | Yes | Health Professional | 42000201 | No | No | 2006-12-20 | Device Not Retu | No | Unkown | 2007
40000620 | 2007-01-31 | Malfunction |  | It was reported that "one of the teeth" of the large needle driver instrument fell inside of the pt and was retrieved. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for eval, therefore, the root cause of the customer reported complaint cannot be determined. The "teeth" referenced by the customer is most likely the grasper portion of the instrument. As indicated in the case notes, the broken piece was successfully retrieved. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 832069 | 844885 | 807659 | 2955842200700034 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-03-02 | 2007-03-02 | No | Yes | Health Professional | 40000620 | No | No | 2007-02-02 | Device Not Retu | No | Unkown | 2007
31146404 | 2007-01-22 | Other |  | It was reported that during the surgical procedure the surgeon could not see in the left side of the high resolution system viewer. When the system endoscope was removed from the patient, it was noticed that the lens of the left eye was missing. No patient harm, adverse outcome or injury was reported. | The affected endoscope was replaced and returned for repair. It cannot be determined if the missing lens fell inside of the patient during the surgical procedure. | 0 DEGREE SCHOELLY STEREO ENDOSCOPE | ENDOSCOPE | 813863 | 826476 | 789614 | 2955842200700025 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2007-02-22 | 2007-02-23 | Yes | Yes | Health Professional | 31146404 | No | No | 2007-01-22 | Device Not Retu | No | Reuse | 2007
40000620 | 2006-12-06 | Malfunction |  | It was reported that metal shavings came off the large needle driver instrument during sterilization. Particulate could have been created during a previous procedure. To date, the account has not reported any incidences of particulate coming off of an instrument during a procedure. No patient harm, adverse outcome or injury was reported. No additional information was provided. The following medwatch report was created for the maryland bipolar forceps instrument that is related to this event. Mfr. Report # 2955842-2006-00191. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed scratches and scraping on the main tube with material that has been removed. Based on investigation performed on instruments returned for a similar failure, it was concluded that the instrument damage may have been caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825459 | 838156 | 801052 | 2955842200600190 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 40000620 | Yes | Yes | 2006-12-07 | Yes | No | Reuse | 2007
40017206 | 2006-12-07 | Malfunction |  | It was reported that metal shavings came off the maryland bipolar forceps instrument during sterilization. Particulate could have been created during a previous procedure. To date, the account has not reported any incidences of particulate coming off of an instrument during a procedure. No patient harm, adverse outcome or injury was reported. No additional information was provided. The following medwatch report was created for the large needle driver instrument that is related to this event. Mfr. Report # 2955842-2006-00190. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed scratches and scraping on the main tube. Based on investigation performed on instruments returned for a similar failure, it was concluded that the instrument damage may have been caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825463 | 838160 | 801056 | 2955842200600191 | 1 | GCJ | Manufacturer | User facility | Nurse | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 40017206 | Yes | Yes | 2006-12-07 | Yes | No | Reuse | 2007
42000604 | 2007-01-17 | Malfunction |  | It was reported that large needle driver instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014, 00015, 00016, 00018, 00019, 00020, 00021, 00022, 00023. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that part of the main tube had small scratches. This is considered to be normal wear and tear of the instrument. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 825544 | 838240 | 801136 | 2955842200700017 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-01-26 | Yes | No | Reuse | 2007
42017904 | 2007-01-17 | Malfunction |  | It was reported that monopolar curved scissors instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Reports #2955842-2007-00014, 00015, 00016, 00017, 00019, 00020, 00021, 00022, and 00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had scraping around the tube. The damaged area is about 6 inches long, ending at the distal end. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | MONOPOLAR CURVED SCISSORS | ELECTROSURGICAL INSTRUMENT | 825553 | 838249 | 801145 | 2955842200700018 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2007-01-26 | Yes | No | Reuse | 2007
42021402 | 2007-01-17 | Malfunction |  | It was reported that pk dissecting forceps instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014, 00015, 00016, 00017, 00018, 00020, 00021, 00022, 00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had deep scratches on the main tube. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | PK DISSECTING FORCEPS | ELECTROSURGICAL INSTRUMENT | 825559 | 838255 | 801151 | 2955842200700019 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42021402 | Yes | No | 2007-01-26 | Yes | No | Reuse | 2007
42000604 | 2007-01-17 | Malfunction |  | It was reported that large needle driver instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014, 00015, 00016, 00017, 00018, 00019, 00020, 00021, 00022, 00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had one deep scratch near the distal end, one small area with scraping located 5 inches from the wrist, and the removal of material on the main tube. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 825563 | 838259 | 801155 | 2955842200700020 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42000604 | Yes | No | 2007-01-26 | Yes | No | Reuse | 2007
42000201 | 2007-01-17 | Malfunction |  | It was reported that cannula instrument was bent and may have contributed to instruments splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014, 00015, 00016, 00017, 00018, 00019, 00020, 00022, 00023. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the cannula to be bent inward at the distal tip thus making the opening smaller likely to scrape and/or remove material from the instrument during use. | CANNULA INSTRUMENT | ENDOSCOPIC ACCESSORY | 825569 | 838265 | 801161 | 2955842200700021 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-01-26 | Yes | No | Reuse | 2007
42000201 | 2007-01-17 | Malfunction |  | It was reported that cannula instrument was bent and may have contributed to instruments splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014, 00015, 00016, 00017, 00018, 00019, 00020, 00021, 00023. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment and the cannula to be bent inward at the distal tip thus making the opening smaller likely to scrape and/or remove material from the instrument during use. | CANNULA INSTRUMENT | ENDOSCOPIC ACCESSORY | 825574 | 838270 | 801166 | 2955842200700022 | 1 | GCJ | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-01-26 | Yes | No | Reuse | 2007
40016906 | 2007-01-24 | Malfunction |  | It was reported that during a laparoscopic assisted vaginal hysterectomy, the harmonic curved shears insert broke. Broken pieces fell inside the patient. All pieces were retrieved except for the hinge. The procedure was delayed approximately 20 minutes and completed with another isi instrument. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was missing. Engineering also observed a gap between the dual hinged feature that was slightly wider than normal. Biocompatibility testing was performed on the materials of the hinge. Testing results found the materials to be non-cytotoxic, non-sensitizing, non-hemolytic and non-pyrogenic. Testing also showed no evidence of systemic toxicity and no evidence of acute intracutaneous reactivity. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 825582 | 838278 | 801174 | 2955842200700026 | 1 | GDW | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-23 | 2007-02-23 | No | Yes | Health Professional | 40016906 | Yes | No | 2007-01-24 | Yes | Yes | Unkown | 2007
IS1200 A43P7 | 2006-01-24 | Malfunction |  | It was reported that during a cardiac procedure, the customer experienced low light output from their two light sources. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The instrument was evaluated on site. Per the customer reported complaint, engineering obsered bad light output of both light sources. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 825587 | 838283 | 801179 | 2955842200700027 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-23 | 2007-02-23 | No | Yes | Health Professional | IS1200 A43P7 | No | No | 2007-01-24 | Yes | No | Unkown | 2007
42017204 | 2007-01-17 | Malfunction |  | It was reported that maryland bipolar forceps instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00015; mfr. Report #2955842-2007-00016; mfr. Report # 2955842-2007-00017; mfr. Report # 2955842-2007-00018; mfr. Report #2955842-2007-00019, mfr. Report # 2955842-2007-00020, mfr. Report #2955842-2007-00021, mfr. Report #2955842-2007-00022; mfr. Report #2955842-2007-00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had scratches and material removed toward the distal end. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | MARYLAND BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 826096 | 838792 | 801688 | 2955842200700014 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2007-01-26 | Yes | No | Reuse | 2007
42009304 | 2007-01-17 | Malfunction |  | It was reported that prograsp forceps instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014; mfr. Report #2955842-2007-00015; mfr. Report # 2955842-2007-00017; mfr. Report # 2955842-2007-00018; mfr. Report #2955842-2007-00019, mfr. Report # 2955842-2007-00020, mfr. Report #2955842-2007-00021, mfr. Report #2955842-2007-00022; mfr. Report #2955842-2007-00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had small scratches toward the distal end. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 826117 | 838813 | 801709 | 2955842200700016 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2007-01-26 | Yes | No | Reuse | 2007
42000604 | 2007-01-17 | Malfunction |  | It was reported that large needle driver instrument splintered. This was discovered prior to a surgical procedure and removed from use. At this time, the account has not reported an incident of particulate coming off of an instrument during a procedure. No add'l info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for instrument and cannula accessories related complaints from the same account. Mfr. Report #2955842-2007-00014; mfr. Report #2955842-2007-00016; mfr. Report # 2955842-2007-00017; mfr. Report # 2955842-2007-00018; mfr. Report #2955842-2007-00019, mfr. Report # 2955842-2007-00020, mfr. Report #2955842-2007-00021, mfr. Report #2955842-2007-00022; mfr. Report #2955842-2007-00023. | The instrument was returned and evaluated. Engineering observed that part of the main tube had deep scratches toward the distal end. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found that cannulas were bent at the tip, thus likely causing the scraping and the removal of material from the instrument. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 826130 | 838827 | 801723 | 2955842200700015 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2007-02-22 | 2007-02-23 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2007-01-26 | Yes | No | Reuse | 2007
IS1200 A4 3P7 | 2007-01-24 | Other | Other | It was reported that during the surgical procedure, the surgical field was too dark. No patient hram was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering revealed that, the glass on the bifurcated light guide cable was cracked, thus causing the low light output experienced by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 827140 | 839854 | 802728 | 2955842200700024 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-02-22 | 2007-02-23 | Yes | Yes | Health Professional | IS1200 A4 3P7 | No | Yes | 2007-01-24 | Device Not Retu | No | Reuse | 2007
42000201 | 2006-07-26 | Malfunction |  | It was reported that during the surgical procedure, the is2000 monopolar curved scissors instrument was rubbing against the trocar and residue fell inside of the pt. The residue was retrieved and the planned procedure was completed with no delay. No add'l procedure was required. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments and cannula accessories returned from this site. Mfr report# 2955842-2006-00160, 00161, 00162, 00163, 00164, 00181, 00182, 00183, 00184, 00185, 00186, 00187, 00188. | The device was not returned for evaluation. The event description is consistent with the use of a potentially defective cannula. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 825019 | 837716 | 800612 | 2955842200600180 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-14 | 2007-02-22 | No | Yes | Health Professional | 42000201 | No | No | 2006-07-26 | Device Not Retu | No | Unkown | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instruments was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr report# 2955842-2006-00173, 00174, 00175, 00009, 00010, 00011, 00012, 00013. | Conclusions: the cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found a dent at the distal end of the cannula, which may remove material from the instrument during use. It was determined that the dent was most likely caused by being dropped or mishandling. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825031 | 837728 | 800624 | 2955842200700008 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Reuse | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instrument was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr report #2955842-2006-00173, mfr report #2955842-2006-00174, mfr report #2955842-2006-175, mfr report #2955842-2007-00008, mfr report #2955842-2007-00009, mfr report #2955842-2007-00010, mfr report #2955842-2007-00011, mfr report #2955842-2007-00012. | Conclusion: the cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment and a burr on the distal inner diameter, which may remove material from the instrument during use. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825662 | 838358 | 801254 | 2955842200700013 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Reuse | 2007
42004903 | 2006-12-01 | Malfunction |  | It was reported that during a surgical procedure, residue from the cadiere forceps instrument fell into the pt. No additional information was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr. Report #2955842-2006-00174, mfr. Report #2955842-2006-00175, mfr. Report #2955842-2007-00008, mfr. Report #2955842-2007-00009, mfr. Report #2955842-2007-00010, mfr. Report #2955842-2007-00011, mfr. Report #2955842-2007-00012, mfr. Report #2955842-2007-00013. | The device was not returned. The event description is consistent with the use of potentially defective cannula. | CADIERE FORCEPS | ELECTROSURGICAL INSTRUMENT | 825678 | 838373 | 801269 | 2955842200600173 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42004903 | No | No | 2006-12-01 | Device Not Retu | No | Unkown | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instrument was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr report #2955842-2006-00173, mfr report #2955842-2006-00174, mfr report #2955842-2006-175, mfr report #2955842-2007-00008, mfr report #2955842-2007-00009, mfr report #2955842-2007-00010, mfr report #2955842-2007-00011, mfr report #2955842-2007-00013. | Conclusion: the cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to cause abrasions on an instrument's main tube. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825687 | 838382 | 801278 | 2955842200700012 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Reuse | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instrument was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr report #2955842-2006-00173, mfr report #2955842-2006-00174, mfr report #2955842-2006-175, mfr report #2955842-2007-00008, mfr report #2955842-2007-00009, mfr report #2955842-2007-00010, mfr report #2955842-2007-00012, mfr report #2955842-2007-00013. | Conclusion: the cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to cause abrasions on an instrument's main tube. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825695 | 838390 | 801286 | 2955842200700011 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Reuse | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instrument was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr report #2955842-2006-00173, mfr report #2955842-2006-00174, mfr report #2955842-2006-175, mfr report #2955842-2007-00008, mfr report #2955842-2007-00009, mfr report #2955842-2007-00011, mfr report #2955842-2007-00012, mfr report #2955842-2007-00013. | Conclusion: the cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, which may remove material from the instrument during use. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825703 | 838398 | 801294 | 2955842200700010 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Reuse | 2007
42009304 | 2006-12-18 | Malfunction |  | It was reported that during a surgical procedure or during processing the prograsp forceps, instrument underwent damage to the fiberglass with dents and chips missing. No additional info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for related instrument complaints at the same facility. Mfr. Reports #2955842-2006-00154, #2955842-2006-00176, #2955842-2006-00178, and #2955842-2006-00179. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument's main tube where deep scratches were found at the distal end. The instrument damage does not correlate with use of a defective cannula. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 825754 | 838449 | 801345 | 2955842200600177 | 1 | HTD | Manufacturer | Company Representative | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42009304 | Yes | No | 2006-12-18 | Yes | No | Reuse | 2007
42017204 | 2006-12-18 | Malfunction |  | It was reported that during a surgical procedure or during processing the maryland bipolar forceps, instrument underwent damage to the fiberglass with dents and chips missing. No additional info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for related instrument complaints at the same facility. Mfr. Reports #2955842-2006-00154, #2955842-2006-00176, #2955842-2006-00177, and #2955842-2006-00178. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument's main tube where scratches were found at the distal end. The instrument damage does not correlate with use of a defective cannula. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC INSTRUMENT | 825758 | 838453 | 801349 | 2955842200600179 | 1 | HTD | Manufacturer | Company Representative | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42017204 | Yes | No | 2006-12-18 | Yes | No | Reuse | 2007
42017904 | 2006-12-18 | Malfunction |  | It was reported that during a surgical procedure or during processing the monopolar curved scissors, instrument underwent damage to the fiberglass with dents and chips missing. No additional info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for related instrument complaints at the same facility. Mfr. Reports #2955842-2006-00154, #2955842-2006-00176, #2955842-2006-00177, and #2955842-2006-00179. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument's main tube where a scratch was found at the distal end. The instrument damage does not correlate with use of a defective cannula. | MONOPOLAR CURVED SCISSORS | ENDOSCOPIC INSTRUMENT | 825763 | 838458 | 801354 | 2955842200600178 | 1 | LRW | Manufacturer | Company Representative | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42017904 | Yes | No | 2006-12-18 | Yes | No | Reuse | 2007
42017203 | 2006-12-01 | Malfunction |  | It was reported that during a surgical procedure, residue from the maryland bipolar instrument fell into the pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr reports #2955842-2006-00173, #2955842-2006-00175, #2955842-2007-00008, #2955842-2007-00009, #2955842-2007-00010, #2955842-2007-00011, #2955842-2007-00012, and #2955842-2007-00013. | The device was not returned. The event description is consistent with the use of a potentially defective cannula. | MARYLAND BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 825768 | 838463 | 801359 | 2955842200600174 | 1 | HTD | Manufacturer | Company Representative | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42017203 | No | No | 2006-12-01 | Device Not Retu | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories returned from this site. Mfr reports #2955842-2006-00160, #2955842-2006-00161, #2955842-2006-00162, #2955842-2006-00163, #2955842-2007-00164, #2955842-2007-00180, #2955842-2007-00181, #2955842-2007-00183, #2955842-2007-00184, #2955842-2007-00185, #2955842-2006-00186, #29558242-2006-00187, and #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 825771 | 838466 | 801362 | 2955842200600182 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-08-17 | Yes | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories returned from this site. Mfr reports #2955842-2006-00160, #2955842-2006-00161, #2955842-2006-00162, #2955842-2006-00163, #2955842-2007-00164, #2955842-2007-00180, #2955842-2007-00181, #2955842-2007-00182, #2955842-2007-00184, #2955842-2007-00185, #2955842-2006-00186, #29558242-2006-00187, and #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 825774 | 838469 | 801365 | 2955842200600183 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Yes | No | 2006-08-17 | Yes | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories returned from this site. Mfr reports #2955842-2006-00160, #2955842-2006-00161, #2955842-2006-00162, #2955842-2006-00163, #2955842-2007-00164, #2955842-2007-00180, #2955842-2007-00181, #2955842-2007-00182, #2955842-2007-00183, #2955842-2007-00184, #2955842-2006-00186, #29558242-2006-00187, and #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 825777 | 838472 | 801368 | 2955842200600185 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Yes | No | 2006-08-17 | Yes | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories returned from this site. Mfr reports #2955842-2006-00160, #2955842-2006-00161, #2955842-2006-00162, #2955842-2006-00163, #2955842-2007-00164, #2955842-2007-00180, #2955842-2007-00181, #2955842-2007-00182, #2955842-2007-00183, #2955842-2007-00185, #2955842-2006-00186, #29558242-2006-00187, and #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 825779 | 838474 | 801370 | 2955842200600184 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Yes | No | 2006-08-17 | Yes | No | Unkown | 2007
42017904 | 2006-07-31 | Malfunction |  | It was reported that during the surgical procedure, the customer noticed that the is2000 monopolar curved scissors instrument was rubbing against the trocar and metal residue fell inside of the pt. The instrument was removed and replaced with an instrument of the same type. The procedure was completed with no delay and no additional procedure was required. No pt harm, adverse outcome or injury was report. The following medwatch reports were created for the instruments and cannula accessories returned from this site. Mfr reports #2955842-2006-00160, #2955842-2006-00161, #2955842-2006-00162, #2955842-2006-00163, #2955842-2007-00164, #2955842-2007-00180, #2955842-2007-00182, #2955842-2007-00183, #2955842-2007-00184, #2955842-2007-00185, #2955842-2006-00186, #29558242-2006-00187, and #2955842-2006-00188. | The instrument was not returned for evaluation. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula. | IS2000 MONOPOLAR SCISSORS CURVED INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 825781 | 838476 | 801372 | 2955842200600181 | 1 | LRW | Manufacturer | Company Representative | Other | Initial | 2007-02-13 | 2007-02-22 | No | Yes | Health Professional | 42017904 | No | No | 2006-07-31 | Device Not Retu | No | Unkown | 2007
42017903 | 2006-12-01 | Malfunction |  | It was reported that during a surgical procedure, residue from the monopolar curved instrument fell into the pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr reports #2955842-2006-00173, #2955842-2006-00174, #2955842-2007-00008, #2955842-2007-00009, #2955842-2007-00010, #2955842-2007-00011, #2955842-2007-00012, and #2955842-2007-00013. | The device was not returned. The event description is consistent with the use of a potentially defective cannula. | MONOPOLAR CURVED SCISSORS | ELECTROSURGICAL INSTRUMENT | 825784 | 838479 | 801375 | 2955842200600175 | 1 | LRW | Manufacturer | User facility | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42017903 | No | No | 2006-12-01 | Device Not Retu | No | Unkown | 2007
42000201 | 2007-01-03 | Malfunction |  | It was reported that the cannula instrument was being returned as it is believed it may have damaged one or more instruments. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the cannula accessories and instruments used at the same facility. Mfr. Reports #2955842-2006-00173, #2955842-2006-00174, #2955842-2006-00175, #2955842-2007-00008, #2955842-2007-00010, #2955842-2007-00011, #2955842-2007-00012, and #2955842-2007-00013. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found a dent at the distal end of the cannula, which may remove material from the instrument during use. It was determined that the dent was most likely caused by being dropped or mishandling. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 825785 | 838480 | 801376 | 2955842200700009 | 1 | KOG | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Yes | No | 2007-01-03 | Yes | No | Reuse | 2007
42000604 | 2006-12-18 | Malfunction |  | It was reported that during a surgical procedure or during processing the the large needle driver instrument underwent damage to the fiberglass with dents and chips missing. No additional info was provided. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for related instrument complaints at the same facility. Mfr. Reports #2955842-2006-00154, #2955842-2006-00177, #2955842-2006-00178, and #2955842-2006-00179. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed scratches at distal end, more likely due to normal wear and tear of the instrument. The instrument damage does not correlate with use of a defective cannula. | IS2000 LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 826111 | 838807 | 801703 | 2955842200600176 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-21 | 2007-02-22 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2006-12-18 | Yes | No | Reuse | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments and cannula accessories returned from this site. Mfr. Report #2955842-2006-00160, mfr. Report #2955842-2006-00161, mfr. Report #2955842-2006-00162, mfr. Report #2955842-2006-00163, mfr. Report #2955842-2006-00164, mfr. Report #2955842-2006-00180, mfr. Report #2955842-2006-00181, mfr. Report #2955842-2006-00182, mfr. Report #2955842-2006-00183, mfr. Report #2955842-2006-00184, mfr. Report #2955842-2006-00185, mfr. Report #2955842-2006-00186, mfr. Report #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 827423 | 840139 | 803012 | 2955842200600187 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-08-17 | Yes | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments and cannula accessories returned from this site. Mfr. Report #2955842-2006-00160, mfr. Report #2955842-2006-00161, mfr. Report #2955842-2006-00162, mfr. Report #2955842-2006-00163, mfr. Report #2955842-2006-00164, mfr. Report #2955842-2006-00180, mfr. Report #2955842-2006-00181, mfr. Report #2955842-2006-00182, mfr. Report #2955842-2006-00183, mfr. Report #2955842-2006-00184, mfr. Report #2955842-2006-00185, mfr. Report #2955842-2006-00187, mfr. Report #2955842-2006-00188. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 827432 | 840147 | 803020 | 2955842200600186 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-08-17 | Yes | No | Unkown | 2007
42000201 | 2006-08-17 | Malfunction |  | It was reported that the instrument was dropping metal shavings. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the instruments and cannula accessories returned from this site. Mfr. Report #2955842-2006-00160, mfr. Report #2955842-2006-00161, mfr. Report #2955842-2006-00162, mfr. Report #2955842-2006-00163, mfr. Report #2955842-2006-00164, mfr. Report #2955842-2006-00180, mfr. Report #2955842-2006-00181, mfr. Report #2955842-2006-00182, mfr. Report #2955842-2006-00183, mfr. Report #2955842-2006-00184, mfr. Report #2955842-2006-00185, mfr. Report #2955842-2006-00186, mfr. Report #2955842-2006-00187. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 827435 | 840150 | 803023 | 2955842200600188 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-22 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-08-17 | Yes | No | Unkown | 2007
42017904 |  | Malfunction |  | It was reported that the is2000 monopolar curved scissors instrument was not working properly. No additional info was provided. The hospital did not identify this issue as having occurred during a surgical procedure. | The instrument was returned and evaluated. Engineering observed areas on the instrument main tube, where material had been removed. It was determined that the failure mode is consistent with the use of a potentially defective cannula accessory, which may remove material from the instrument during use. | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 827437 | 840152 | 803025 | 2955842200600189 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-22 | 2007-02-22 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-11-27 | Yes | No | Reuse | 2007
400183020 | 2007-01-09 | Malfunction |  | It was reported that the permanent cautery hook intstrument did not dissect properly. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found evidence that the instrument may have arced during a surgical procedure. Part of the yaw pulley and the distal clevis have been charred. The customer's complaint that the instrument "did not dissect properly" may have been due to a decline in performance as the insulation began to degrade. This complaint is being reported via an mdr since the instrument may have arced during a surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 824407 | 837104 | 800000 | 2955842200700007 | 1 | GEI | Manufacturer | User facility | Nurse | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 400183020 | Device Returned To Manufacturer | Yes | 2007-01-18 | Yes | No | Reuse | 2007
40017905 | 2006-01-17 | Malfunction |  | It was reported that metal shavings came off the monopolar curved scissors instrument during sterilization. Particulate could have been created during a previous procedure. To date, the account has not reported any incidences of particulate coming off of an instrument during a procedure. No patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where the tube extension is broken at the proximal end. Keyed sections of the tube extension which mate with the main tube are broken off. The distal end of the main tube is scratched on one side. Based on investigation performed on instruments returned for a similar failure, it was concluded that the instrument damage was likely caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 824418 | 837115 | 800011 | 2955842200700005 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | No | 2007-01-17 | Yes | No | Reuse | 2007
42017904 | 2006-10-31 | Malfunction |  |  |  | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 824425 | 837122 | 800018 | 2955842200700168 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-31 | 2007-02-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-10-31 | Yes | No | Reuse | 2007
42017904 | 2006-10-31 | Malfunction |  | It was reported that during the beginning of the surgical procedure, the is2000 monopolar curved scissors instrument got scraped inside of the trochar, causing fragments to fall inside of the patient. The instrument fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. Medwatch mfr. Report #2955842-2006-00169 was created for the cannula accessory used during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where material had been removed. It was determined that the failure mode is consistent with the use of a potentially defective cannula. As indicated in the complaints notes, the instrument fragments were successfully retrieved. | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 824432 | 837129 | 800025 | 2955842200600168 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-31 | 2007-02-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-10-31 | Yes | No | Reuse | 2007
42000201 | 2006-10-31 | Malfunction |  | It was reported that during the beginning of the surgical procedure, the is2000 monopolar curved scissors instrument got scraped inside of the trochar, causing fragments to fall inside of the patient. The instrument fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. Medwatch mfr. Report #2955842-2006-00168 was created for the is2000 monopolar curved scissors instrument used during the surgical procedure. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found a dent at the distal end of the cannula, which may remove material from the instrument during use. It was determined that the dent was most likely caused by user handling. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 824436 | 837133 | 800029 | 2955842200600169 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2006-10-31 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-10-31 | Yes | No | Reuse | 2007
40009306 | 2006-01-17 | Malfunction |  | It was reported that metal shavings came off the prograsp forceps instrument during sterilization. Particulate could have been created during a previous procedure. To date, the account has not reported any incidences of particulate coming off of an instrument during a procedure. No patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where the main tube is dislodged from the roll shaft and cracked at the proximal end. The tube is bent at a 15-degree angle at the roll shaft. The distal end of the tube is scratched. Based on investigation performed on instruments returned for a similar failure, it was concluded that the instrument damage was likely caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 824445 | 837142 | 800038 | 2955842200700006 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 40009306 | Device Returned To Manufacturer | No | 2007-01-17 | Yes | No | Reuse | 2007
42000201 | 2006-07-05 | Malfunction |  | It was reported that during the surgical procedure, the trocar left metal paint residue on the patient. The procedure was completed with no delay and no additional procedure was required. No patient harm, adverse outcome or injury was reported. | Additional information provided by an isi representative confirmed that the particulate was not paint residue as initially reported, but residue from an instrument. The device was not returned for evaluation. The event description is consistent with the use of a potentially defective cannula. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 824452 | 837149 | 800045 | 2955842200600160 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-12 | 2007-02-21 | No | Yes | Health Professional | 42000201 | No | No | 2006-07-05 | Device Not Retu | No | Unkown | 2007
42017904 | 2006-07-07 | Malfunction |  | It was reported that during the surgical procedure, the customer noticed that the is2000 monopolar curved scissors instrument was rubbing against the trocar and metal residue fell inside the patient. The procedure was completed with no delay and no additional procedure was required. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the three cannula accessories used during the surgical procedure. Mfr. Report # 2955842-2006-00162, mfr report # 2955842-2006-00163, and mfr report # 2955842-2006-00164. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument main tube had material removed. It was determined that the failure mode is consistent with the use of a potentially defective cannula. | IS2000 MONOPOLAR SCISSORS CURVED INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 824473 | 837169 | 800065 | 2955842200600161 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-02-13 | 2007-02-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-07-07 | Device Not Retu | No | Unkown | 2007
42000201 | 2006-07-07 | Malfunction |  | It was reported that during the surgical procedure, the customer noticed that the is2000 monopolar curved scissors instrument was rubbing against the trocar and metal residue fell inside the patient. The procedure was completed with no delay and no additional procedure was required. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the is2000 monopolar curved scissors instrument and the add'l two cannula accessories used during the surgical procedure. Mfr. Report # 2955842-2006-00161, mfr report # 2955842-2006-00163 and mfr report # 2955842-2006-00164. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 824481 | 837177 | 800073 | 2955842200600162 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-13 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-07-07 | Device Not Retu | No | Unkown | 2007
42000201 | 2006-07-07 | Malfunction |  | It was reported that during the surgical procedure, the customer noticed that the is2000 monopolar curved scissors instrument was rubbing against the trocar and metal residue fell inside the patient. The procedure was completed with no delay and no additional procedure was required. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the is2000 monopolar curved scissors instrument and the add'l two cannula accessories used during the surgical procedure. Mfr. Report # 2955842-2006-00161, mfr report # 2955842-2006-00162 and mfr report # 2955842-2006-00164. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 824486 | 837182 | 800078 | 2955842200600163 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-13 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-07-07 | Device Not Retu | No | Unkown | 2007
42000201 | 2006-07-07 | Malfunction |  | It was reported that during the surgical procedure, the customer noticed that the is2000 monopolar curved scissors instrument was rubbing against the trocar and metal residue fell inside the patient. The procedure was completed with no delay and no additional procedure was required. No patient harm, adverse outcome or injury was reported. The following medwatch reports were created for the is2000 monopolar curved scissors instrument and the add'l two cannula accessories used during the surgical procedure. Mfr. Report # 2955842-2006-00161, mfr report # 2955842-2006-00162 and mfr report # 2955842-2006-00163. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 824492 | 837188 | 800084 | 2955842200600164 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2007-02-13 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-07-07 | Device Not Retu | No | Unkown | 2007
42017203 | 2006-11-22 | Malfunction |  | It was reported that during a surgical procedure, the is2000 maryland bipolar forceps instrument became scraped by the inside of the cannula causing fragments to fall inside of the patient. No patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where material had been removed. Based on investigation performed on instruments returned for a similar failure, it was concluded that the instrument damage was likely caused by a defective instrument cannula accessory, which scraped off material as the instrument was moved. | MARYLAND BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 824495 | 837190 | 800086 | 2955842200600170 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-13 | 2007-02-21 | No | Yes | Health Professional | 42017203 | Device Returned To Manufacturer | No | 2006-11-22 | Yes | No | Reuse | 2007
42020501 | 2006-11-13 | Malfunction |  | It was reported that two instruments had scratches on the shaft. The complaint record states, "fragments did not fall inside of patient. " no patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where material had been removed. The cause of the scratches on the instrument shaft could not be conclusively determined. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC INSTRUMENT | 824505 | 837200 | 800096 | 2955842200600171 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-16 | 2007-02-21 | No | Yes | Health Professional | 42020501 | Device Returned To Manufacturer | No | 2006-11-13 | Yes | No | Reuse | 2007
42020501 | 2006-11-13 | Malfunction |  | It was reported that two instruments had scratches on the shaft. The complaint record states, "fragments did not fall inside of patient. " no patient harm, adverse outcome or injury was reported. No additional information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed areas on the instrument main tube where material had been removed. The cause of the scratches on the instrument shaft could not be conclusively determined. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC INSTRUMENT | 824510 | 837205 | 800101 | 2955842200600172 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-16 | 2007-02-21 | No | Yes | Health Professional | 42020501 | Device Returned To Manufacturer | No | 2006-11-13 | Yes | No | Reuse | 2007
42017904 | 2006-09-27 | Malfunction |  | It was reported that residue from the monopolar curved scissors instrument fell inside of the pt. The pt wound was irrigated and the residue was suctioned out. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the two cannula accessories used during the surgical procedure. Mfr report #2955842-2006-166. Mfr report #2955842-2006-167. | The instrument was returned and evaluated. Engineering observed that part of the main tube had material removed toward the distal end. Per the customer reported complaint, it was determined that the failure mode is consistent with the use of a potentially defective cannula accessory. The cannula accessories used during the surgical procedure were also returned and evaluated. Engineering found the inner surface of the cannulas to be out of alignment, thus creating a raised ledge, which may remove material from the instrument during use. | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 943167 | 974958 | 742884 | 2955842200600165 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-09-27 | Yes | No | Reuse | 2007
42000201 | 2006-09-27 | Malfunction |  | It was reported that residue from the monopolar curved scissors instrument fell inside of the pt. The pt wound was irrigated and the residue was suctioned out. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the monopolar curved scissors instrument and additional cannula accessory used during the surgical procedure. Mfr report #2955842-2006-00165. Mfr report #2955842-2006-00166. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus creating a raised ledge, which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 943188 | 974981 | 742884 | 2955842200600167 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-09-27 | Yes | No | Reuse | 2007
42000201 | 2006-09-27 | Malfunction |  | It was reported that residue from the monopolar curved scissors instrument fell inside of the pt. The pt wound was irrigated and the residue was suctioned out. No pt harm, adverse outcome or injury was reported. The following medwatch reports were created for the monopolar curved scissors instrument and additional cannula accessory used during the surgical procedure. Mfr. Report #2955842-2006-00165. Mfr report #2955842-2006-00167. | The cannula accessory was returned and evaluated. Per the customer reported complaint, engineering found the inner surface of the cannula to be out of alignment, thus, creating a raised ledge, which may remove material from the instrument during use. | CANNULA 8MM REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 943268 | 975128 | 742884 | 2955842200600166 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-02-15 | 2007-02-21 | No | Yes | Health Professional | 42000201 | Device Returned To Manufacturer | No | 2006-09-27 | Yes | No | Reuse | 2007
PS 2000 | 2007-01-03 | Injury | Required Intervention | Perforated terminal ileum with intra-abdominal free air, peritonitis following a myomectomy performed robotically (laparoscopically). |  | DA VINCI ROBOT | DA VINCI S SURGICAL SYSTEM | 2706094 | 2671300 | 2567916 | 2671300 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2007-02-15 | 2007-02-21 | Yes | No | Service Personnel | PS 2000 | Yes | No |  | No Answer Provi |  | Invalid Data | 2007
| 2007-02-01 | Malfunction |  | Patient to undergo laparoscopic lysis of adhesions and retropubic radical prostatectomy utilizing robotic assistance with the davinci system. At beginning of the procedure, there was a malfunction on robotic arm #2. One of the plastic wings used to tighten down the arm broke off while surgeon was attaching arm #2 to cannula. Because of the robotic sensors, it would not allow the tool to be utilized. At this point, a company representative was called and alternatives were discussed. (this was not a repairable piece of the robot during this surgery). Patient's spouse was notified and options discussed and agreed upon if robotic arm remained nonfunctional. Surgeon was able to get the second arm functional and it seemed to work, but as soon as one tool was changed out for another, it became nonfunctional again. At this point surgeon and staff believed it would have been imprudent to proceed with the robotic assisted procedure. The robot was undocked and all of the ports were closed in the usual fashion. Patient prepped again, redrapped and underwent open radial prostatectomy. The day after the event the service representative from intuitive surgical came to site and installed a new mount (broken cannula mount) on the psm2 (patient-side manipulator), tested system and system passed all tests. |  | DAVINCI | ROBOTIC SURGICAL SYSTEM | 810026 | 822627 | 785829 | 822627 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2007-02-15 | 2007-02-15 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2007
40020705 | 2007-01-15 | Malfunction |  | It was reported that during the surgical procedure, the tip of the tenaculum forceps instrument broke and a piece fell inside the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the wrist section cut off at the distal end of the main tube. The wrist was also returned with the instrument. Per the customer reported complaint, engineering observed that the instrument wrist had broken proximal clevis and pieces of the proximal clevis were missing. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis at the main tube interface. As indicated in the complaints notes, the broken piece was successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 824579 | 837273 | 800169 | 2955842200700004 | 1 | GDW | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-02-14 | 2007-02-15 | No | Yes | Health Professional | 40020705 | Yes | No | 2007-01-16 | Yes | No | Reuse | 2007
| 2007-01-24 | Malfunction |  | The tip of the 5 mm tenaculum instrument broke off in the patient, which required additional procedures to be done. X-ray was taken and hand assisted gel port was used to remove the broken instrument tip. |  | 5 MM MYOMECTOMY TENACULUM | ACCESSORY SURGICAL ROBOT | 805946 | 818463 | 781771 | 818463 | 1 | KOG | User Facility |  |  | Initial | 2007-02-08 | 2007-02-08 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2007
IS1200 A43P7 |  | Other | Other | It was reported that during a da vinci prostatectomy, the customer experienced a 20013 system error code followed by a 20127 non-recoverable fault. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering, concluded that the system faults experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Based on the system logs and check sensor test, the axis 3 potentiometer was replaced. The system alarm (system generated fault codes) functioned as designed and there was injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 805601 | 818117 | 781425 | 2955842200700003 | 1 | KOG | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-07 | 2007-02-07 | Yes | Yes | Health Professional | IS1200 A43P7 | Device Returned To Manufacturer | No | 2007-01-09 | Yes | No | Reuse | 2007
IS1200 A43P7 | 2006-12-20 | Other | Other | It was reported that the customer experienced vision loss during the planned surgical procedure. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering found that the bulb installed in the system illuminator was the incorrect size, thereby, causing the bulb to overheat and explode. This lead to the loss of vision experienced by the customer. The system was repaired by installing the correct size bulb. As of 1/31/07, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 802098 | 814616 | 777944 | 2955842200600158 | 1 | GDW | Manufacturer | User facility | Other | Initial | 2007-01-30 | 2007-02-01 | Yes | Yes | Health Professional | IS1200 A43P7 | No | No | 2006-12-20 | Device Not Retu | No | Reuse | 2007
42017904 |  | Malfunction |  | It was reported that during the surgical procedure, the shaft was heating up and a burning smell came from the monopolar curved scissors instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint of heating up and a burning smell came from the instrument, engineering observed that a hole was burnt through the instrument's main tube and part of the tube extension, indicating that arcing had occurred. A hairline crack was also found on the distal end of the main tube, which may have been caused by excessive side loading on the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 805273 | 817793 | 781097 | 2955842200700002 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-30 | 2007-02-01 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-01-03 | Yes | No | Reuse | 2007
42021403 | 2006-12-26 | Malfunction |  | It was reported that during the surgical procedure, the tip of the forcep broke inside of the pt. All broken pieces were retrieved and the planned surgical procedure was completed. During instrument removal, the instrument became stuck on the cannula accessory and the instrument tip was severed in order to remove the instrument from the cannula. No pt harm, adverse outcome or injury was reported. | The instrument was returned with the wrist section cut off at the distal end of the main tube. The instrument wrist was not returned with the instrument. Per the customer reported complaint symptoms, engineering evaluation determined that the failure mode experienced by the customer was most likely a break in the proximal clevis at the main tube interface. | IS2000 PK DESSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 805388 | 817908 | 781212 | 2955842200700001 | 1 | GCJ | Manufacturer | User facility | NOT APPLICABLE | Initial | 2007-01-30 | 2007-02-01 | No | Yes | Health Professional | 42021403 | Device Returned To Manufacturer | No | 2007-01-03 | Yes | No | Initial | 2007
IS1200 A43P7 | 2006-12-14 | Other | Other | It was reported that during the surgical procedure, the scope was intermittently fogging. Due to the vision issue experienced by the customer and patient condition, the procedure was converted to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | Information obtained from the site by an isi representative indicates that the seal around the instrument cannula accessory was no longer tight, thereby contributing to the fogging experienced by the customer. Investigation conducted by isi field service engineering determined that a faulty camera and endoscope may have also contributed to the vision issue experienced by the customer. Per the case notes, patient condition was also a factor in converting to traditional open surgical techniques to finish the planned surgery. As of january 31, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 813727 | 826335 | 789480 | 2955842200600157 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2007-01-30 | 2007-02-01 | Yes | Yes | Health Professional | IS1200 A43P7 | No | No | 2006-12-14 | Device Not Retu | No | Reuse | 2007
|  | Malfunction |  | It was reported that during the surgical procedure, it was noticed that the shaft of the instrument began flaking. The patient's body cavity was irrigated and the instrument fragments were removed. It is unknown which instrument was flaking. No further information was provided. No patient harm, adverse outcome or injury was reported. | The affected instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. Per the case notes, the instrument fragments were successfully retrieved. | Unknown Brand Name | UNKNOWN | 814309 | 826914 | 790052 | 2955842200600159 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-01-31 | 2007-01-31 | No | Yes | Health Professional |  | No | Yes | 2006-11-29 | Device Not Retu | No | Unkown | 2007
42017904 |  | Malfunction |  | It was reported that during the surgical procedure, a hole burnt through the insulation of the monopolar curved scissors instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a hole was burnt through the instrument main tube and part of the tube extension, indicating that arcing had occurred. A hairline crack was also found on the distal end of the main tube, which may have been caused by excessive side loading on the instrument. Based on this discovery, engineering concluded that the instrument arced during the surgical procedure and the source of the arcing most likely originated from the crack found in the main tube. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 801813 | 814323 | 777656 | 2955842200600156 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2007-01-25 | 2007-01-26 | No | Yes | Health Professional | 42017904 | Device Returned To Manufacturer | No | 2006-12-27 | Yes | No | Reuse | 2007
40016906 | 2006-12-06 | Other | Other | Sterile product from manufacturer has chemical indicator that didn't change color when eo sterilized. Contacted company, stated product ok, indicator was bad. Requested in writing. |  | INTUITIVE SURGICAL | HARMONIC CURVED SHEARS INSERT | 795863 | 808380 | 771854 | MW1041724 | 1 | LFL | Voluntary |  | RISK MANAGER | Initial | 2006-12-06 | 2007-01-24 | No | Yes | Invalid Data | 40016906 | No Answer Provided | No |  |  |  |  | 2007
42011003 | 2006-12-20 | Malfunction |  | It was reported that during the surgical procedure, the is2000 precise bipolar forceps instrument became stuck in the cannula accessory and could not be pulled out. The instrument fractured and pieces fell inside of the patient. All broken pieces were retrieved and the instrument and cannula accessory were removed from the patient. It was also reported that due to vision issues related to the patient's anatomy, the procedure was converted to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The instrument was returned with the cannula accessory attached. Per the customer reported complaint, engineering evaluation observed that the instrument was broken at the interface between the proximal clevis and main tube. Both the proximal clevis and the main tube were missing pieces. It was determined that the failure mode experienced by the customer was a result of damage to the proximal clevis and/or main tube. As indicated in the complaint notes, all broken pieces were successfully retrieved. | IS2000 PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 800608 | 813123 | 776469 | 2955842200600155 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-01-18 | 2007-01-19 | No | Yes | Health Professional | 42011003 | Device Returned To Manufacturer | Yes | 2006-12-20 | Yes | No | Reuse | 2007
42000604 | 2006-12-18 | Malfunction |  | It was reported that during the surgical procedure, when removing the is2000 large needle driver instrument from the pt, a 4" piece of the instrument sheared off when it came in contact with the trocar and fell inside of the patient. The sheared off piece was retrieved and the planned surgical procedure was completed. The sheared off piece will not be returning for evaluation. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed a severe crack in the main tube, approximately five inches from the distal end of the instrument. It appears that an unusually high force was applied to the main tube, causing it to break. A five inch sliver of the main tube, starting at the crack and ending at the distal end, was missing. Engineering was unable to determine whether the cause of the abrasion may have been due to the cannula accessory. Since the cannula used in conjunction with the instrument was not returned for evaluation, no definitive conclusion can be drawn. The missing piece was not returned with the instrument, however, the complaint notes indicate that the missing piece was successfully retrieved from the patient. | IS2000 LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 809681 | 822277 | 785479 | 2955842200600154 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-01-17 | 2007-01-17 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2006-12-18 | Yes | No | Reuse | 2007
40019002 | 2006-12-15 | Malfunction |  | It was reported that during the surgical procedure, the cobra grasper instrument, the bipolar forceps instrument and the monopolar curved scissors instrument were arcing. Additional information provided indicates that the cobra grasper instrument (a non-energy instrument) was being used for retraction when the arcing was observed. The instrument did not come into contact with the monopolar curved scissors instrument and the port had not been used with any energized instruments prior to insertion of the cobra grasper instrument. The hospital did not provide information regarding which instrument(s) arced, or if all three arced. No patient harm, adverse outcome or injury was reported. The alleged arcing with the bipolar forceps instrument is being reported on mfr report #2955842-2006-00152. The alleged arcing with the monopolar curved instrument is being reported on mfr's report # 2955842-2006-00153. | The instrument was returned and evaluated. The cobra grasper instrument is not an electrosurgical instrument. No physical damage was found. It cannot be determined if other instruments arced to this instrument. The customer reported failure mode cannot be confirmed. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 798079 | 810622 | 773995 | 2955842200600151 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-01-11 | 2007-01-12 | No | Yes | Health Professional | 40019002 | Yes | Yes | 2006-12-15 | Yes | No | Reuse | 2007
40017206 | 2006-12-15 | Malfunction |  | It was reported that during the surgical procedure, the cobra grasper instrument, the bipolar forceps instrument and the monopolar curved scissors instrument were arcing. Additional information provided indicates that the cobra grasper instrument (a non-energy instrument) was being used for retraction when the arcing was observed. The instrument did not come into contact with the monopolar curved scissors instrument and the port had not been used with any energized instruments prior to insertion of the cobra grasper instrument. The hospital did not provide information regarding which instrument(s) arced, or if all three arced. No patient harm, adverse outcome or injury was reported. The alleged arcing with the corba grasper instrument is being reported on mfr's report # 2955842-2006-00151. The alleged arcing with the monopolar curved instrument is being reported on mfr's report # 2955842-2006-00153. | The instrument was returned and evaluated. Engineering evaluation found no evidence of the instrument having previously arced. The instrument was tested for electrical continuity and passed. It cannot be determined if other instruments arced to this instrument. The customer reported failure mode cannot be confirmed. | BIPOLAR MARYLAND FROCEPS INSTRUMENTS | ELECTROSURGICAL INSTRUMENT | 798088 | 810631 | 774004 | 2955842200600152 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-01-11 | 2007-01-12 | No | Yes | Health Professional | 40017206 | Yes | Yes | 2006-12-15 | Yes | No | Reuse | 2007
40017905 | 2006-12-15 | Malfunction |  | It was reported that during the surgical procedure the cobra grasper instrument, the bipolar forceps instrument and the monopolar curved scissors instrument were arcing. Additional information provided indicates that the cobra grasper instrument (a non-energy instrument) was being used for retraction when the arcing was observed. The instrument did not come into contact with the monopolar curved scissors instrument and the port had not been used with any energized instruments prior to insertion of the cobra grasper instrument. The hospital did not provide information regarding which instrument(s) arced, or if all three arced. No patient harm, adverse outcome or injury was reported. The alleged arcing with the cobra grasper instrument is being reported on mfr's report # 2955842-2006-00151. The alleged arcing with the monopolar curved instrument is being reported on mfr's report # 2955842-2006-00153. | The instrument was returned and evaluated. Engineering evaluation found no evidence of the instrument having previously arced. The instrument tip cover, which provides electrical insulation, was not returned. It is possible that damage to the tip cover may have compromised the electrical insulation. Since the tip cover was not returned, the customer reported failure mode cannot be confirmed. It also cannot be determined if other instruments arced to this instrument. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 798100 | 810643 | 774016 | 2955842200600153 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2007-01-11 | 2007-01-12 | No | Yes | Health Professional | 40017905 | Yes | Yes | 2006-12-15 | Yes | No | Reuse | 2007
IS1200 A43P7 | 2006-12-13 | Other | Other | It was reported that during the surgical procedure, the customer experienced a recurring 20013 system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded, that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered a defective potentiometer assembly within the mtmr, thus generating the system error experienced by the customer. The system alarm (the 20013 system error code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of 01/11/07, there have been no reported recurrences of the issue at the hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 799895 | 812411 | 775759 | 2955842200600150 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-01-11 | 2007-01-12 | Yes | Yes | Health Professional | IS1200 A43P7 | No | Yes | 2006-12-13 | Device Not Retu | No | Reuse | 2007
40016906 | 2006-12-06 | Malfunction |  | It was reported that during the surgical procedure, the harmonic curved shears insert broke. A broken piece fell inside of the patient and could not be located. The insert was removed and the planned surgical procedure was completed with another instrument. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, engineering observed that one of the clamp arms had broken off. Engineering also found that one side of the insert where the clamp arm attaches is bent outwards and the insert is also bent outwards on the same side, opposite to where the clamp arm attaches. This discovery indicates that high force was applied at the joint where the clamp arm attaches. Biocompatibility testing was performed on the materials of the clamp arm. Testing results found the materials to be non-cytotoxic, non-sensitizing, non-hemolytic and non-pyrogenic. Testing also showed no evidence of systemic toxicity and no evidence of acute intracutaneous reactivity. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 793871 | 806358 | 769904 | 2955842200600149 | 1 | LFL | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2007-01-03 | 2007-01-04 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | Yes | 2006-12-06 | Yes | Yes | Initial | 2007
IS1200 A43P7 | 2006-12-05 | Other | Other | It was reported that during the surgical procedure, the customer experienced a recurring 212 system error code, which caused the instrument arms to "softlock" and release the patient tissue held by the instruments. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Based on the system logs, a motor encoder and potentiometer were replaced. As of january 3, 2007, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 794152 | 806638 | 770185 | 2955842200600146 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2007-01-03 | 2007-01-04 | Yes | Yes | Health Professional | IS1200 A43P7 | No | Yes | 2006-12-05 | Device Not Retu | No | Reuse | 2007
IS1200 A43P6 | 2006-11-29 | Other | Other | It was reported that during the surgical procedure, the customer experienced low light output from their two light sources. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Per the customer reported complaint, field service engineering conducted testing on the customer's two illuminator lamps and found them to be below specification. The customer indicated that replacement lamps had been ordered and were going to be installed by their biomedical engineer. The da vinci surgical system user manual explicitly states that "environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. ". | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 794156 | 806642 | 770189 | 2955842200600145 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2006-12-26 | 2007-01-04 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-11-29 | Device Not Retu | No | Reuse | 2007
IS2000 A5 0P5 | 2006-11-28 | Other | Other | It was reported that during the surgical procedure, the is2000 fenestrated bipolar forceps instrument stopped cauterizing and would no longer function. The customer did not have any additional is2000 fenestrated bipolar forceps instruments, therefore, the case was completed with three other instruments of a diffrent type. Because the three replacement instruments are not usually used during the planned surgical procedure, the case was lengthened by three hours and the patient had to receive an additional amount of anesthetic. No patient harm was reported, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed loss of cautery functionality on one of the instrument grips due to a broken conductor wire near the instrument proximal clevis wire hole. The customer indicated that because they did not have a replacement instrument of the same type available, the planned surgical procedure was prolonged due to the use of three replacement instruments which are not usually utilized during the planned surgical. The da vinci s surgical system user manual explicitly states that "environmental or equipment failures may cause the da vinci s system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. ". | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 794532 | 807028 | 770563 | 2955842200600144 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2006-12-26 | 2007-01-04 | Yes | Yes | Health Professional | IS2000 A5 0P5 | No | No | 2006-11-28 | Device Not Retu | No | Reuse | 2007
40009306 | 2006-11-01 | Malfunction |  | On november 2, 2006, it was reported that during the surgical procedure, the cable broke at the wrist of the endowrist prograsp instrument, thereby breaking the housing. The instrument was removed and replaced with an instrument of a different type. Later in the procedure, the console surgeon decided to convert to an open procedure due to pt condition and not due to the da vinci system. No pt harm was reported. The following additional info was provided: during procedure, endowrist prograsp instrument broke while engaged and holding on the fibroid attached to the pt's uterus. While under laparoscopic observation, it was noted that one piece of the instrument had broken off. The piece was recovered laparoscopically and removed. While removing the instrument from the pt's abdominal cavity, additional pieces of the instrument broke off. The procedure was converted to an open procedure. The abdomen was irrigated and explored by direct vision and feel. The procedure carried on without further incident. | The insturment was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 935891 | 966334 | 744934 | 2955842200600147 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2007-01-03 | 2007-01-04 | No | Yes | Health Professional | 40009306 | No | No | 2006-12-05 | Device Not Retu | No | Unkown | 2007
400172XX | 2006-10-20 | Other |  | During a laparoscopic assisted vaginal hysterectomy, the surgeon placed the maryland bipolar shears on the pt's abdomen. He inadvertently stepped on the foot pedal, causing a 2 mm superficial burn to the pt's abdomen. A contributing factor was that the gyn laparoscopic pack was not available; this pack contains an instrument pouch which would have held the shears. The or was also out of the alternative instrument pouch. | The instrument was not returned for evaluation. The info provided indicates that, the surgeon inadvertently stepped on the cautery foot pedal while the instrument was lying dormant on the pt's abdomen. Based on this info, it has been determined that the pt cautery burn was caused by user error. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 775239 | 966367 | 751467 | 2955842200600148 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2007-01-03 | 2007-01-04 | No | Yes | Health Professional | 400172XX | No | Yes | 2006-12-05 | Device Not Retu | No | Unkown | 2007
40018401 | 2006-11-03 | Malfunction |  | It was reported that during the surgical procedure, the isolation portion of the permanent cautery spatula instrument tip broke and a piece fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument ceramic sleeve was broken and a small piece was missing. Engineering evaluation determined that the damage to the instrument ceramic sleeve may have been caused by a collision of the instrument tip and a hard object. As indicated in the complaint notes, the broken piece was successfully retrieved. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 786911 | 799393 | 763067 | 2955842200600143 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-12-15 | 2006-12-15 | No | Yes | Health Professional | 40018401 | Yes | Yes | 2006-11-16 | Yes | No | Reuse | 2006
40000619 | 2006-11-17 | Malfunction |  | It was reported that during the surgical procedure, it was noticed that one part of the large needle drive instrument jaw was missing. It is unknown whether the missing piece separated from the instrument during the surgical procedure or during sterilization and handling prior to the surgical procedure. The instrument was removed and replaced and the planned surgical procedure was completed. The missing piece was not found. No patient harm, adverse outcome or injury was reported. | Per the customer reported complaint, the instrument was returned with one of the carbide tips missing. Engineering evaluation observed trace amounts of residual brazing on the instrument grip surface. This observation leads engineering to conclude that the grip surface may not have been sufficiently cleaned prior to the brazing being applied. This contamination may have caused a weakened braze. Engineering also found gouge marks on the outer surface of the damaged grip where the material is removed. This discovery indicates that excessive force was applied on the grip multiple times. Engineering determined that a combination of a weakened braze and excessive force may have led to the instrument carbide tip falling off. No other instruments from this manufactured lot have exhibited this failure. | LARGE NEEDLE DRIVE INSTRUMENT | ENDOSCOPIC INSTRUMENT | 786977 | 799459 | 763133 | 2955842200600142 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2006-12-15 | 2006-12-15 | No | Yes | Health Professional | 40000619 | Yes | Yes | 2006-11-17 | Yes | No | Reuse | 2006
42017204 | 2006-11-08 | Malfunction |  | It was reported that during a surgical procedure when the instrument was moved "backward, the wrist broke off and fell inside of the pt. " all pieces were retrieved and removed from the pt. The planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was not return for evaluation, therefore, the root cause of the customer reported complaint cannot be determined. | ENDOWRIST MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 782949 | 795364 | 759106 | 2955842200600139 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-11-08 | 2006-12-08 | No | Yes | Health Professional | 42017204 | No | No | 2006-11-08 | Device Not Retu | No | Unkown | 2006
42020702 | 2006-11-09 | Malfunction |  | It was reported that during the surgical procedure, the tip of the is2000 tenaculum forceps instrument broke and pieces fell inside of the patient. All broken pieces were retrieved and the planned surgical procedure was completed. During instrument removal, the instrument became stuck on the cannula accessory and the instrument tip was severed in order to remove the instrument from the cannula. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the wrist section cut off at the distal end of the main tube. The instrument wrist was not returned with the instrument. Per the customer reported complaint symptoms, engineering evaluation determined that the failure mode experienced by the customer was most likely a break in the proximal clevis at the main tube interface. | IS2000 TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 783720 | 796168 | 759875 | 2955842200600140 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-11-09 | 2006-12-08 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | Yes | 2006-11-09 | Yes | No | Initial | 2006
40018301 | 2005-03-24 | Malfunction |  | It was reported that the thermal protector on the instrument had slipped down. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the electrical insulation had become slightly dislodged from its proper location. Engineering also found evidence that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 783774 | 796222 | 759929 | 2955842200600141 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-11-09 | 2006-12-08 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2006-11-09 | Yes | No | Reuse | 2006
| 2006-11-07 | Death | Death | Elderly male had prostatectomy using the davinci robotic surgery equipment. Returned to hospital on three times with hemoperitoneum. Coded in bathroom 27 days postoperatively and died. |  | DAVINCI | ROBOTIC DEVICE SURGICAL | 787272 | 799747 | 763421 | 799747 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-12-07 | 2006-12-07 | Yes | No | Physician |  | Yes | No Answer Provided |  |  |  |  | 2006
| 2006-11-28 | Malfunction |  | During the robotic-assisted surgical procedure (dissection of the prostate), one of the grasping (prograsp forceps) instruments broke where the specialized jaw end of the instrument is hinged to the shaft. There were three small fragments of black colored material that were noted intraperitoneally and these were retrieved uneventfully. The entire docking port was taken off the patient and was replaced with another long trocar. The remainder of the procedure went well and the patient left the operating room in excellent condition. This was the 5th use out of 10. For our tracking purposes, a piece of tape is placed on the instrument with numbers 1 - 10 written on it. We cross off the number after each use. However, if we were to try to use instrument on the 11th time, the davinci system will not allow it and it has to be replaced. The manufacturer was notified regarding this failure and we are retaining the device in the event the manufacturer requests to evaluate it. We have not experienced any similar problems with this particular instrument. |  | PROGRASP FORCEPS | SURGICAL ROBOT FORCEPS | 780408 | 792767 | 756577 | 792767 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-11-29 | 2006-12-05 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2006
42009304 | 2006-10-30 | Malfunction |  | It was reported that during the surgical procedure the shaft of the is2000 endowrist prograsp instrument broke and two pieces fell inside of the pt. Both broken pieces were retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | IS2000 ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 780919 | 793301 | 757081 | 2955842200600138 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-30 | 2006-11-29 | No | Yes | Health Professional | 42009304 | No | No | 2006-10-30 | Device Not Retu | No | Unkown | 2006
42018404 | 2006-10-24 | Malfunction |  | It was reported that during the surgical procedure the proximal clevis of the is2000 permanent cautery spatula instrument broke into 2 pieces and fell inside of the pt. The broken pieces were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering eval observed that a portion of the instrument distal clevis was broken off at the base and that the conductor cap was also missing. The missing pieces were not returned with the instrument; however, the complaint notes indicate that all broken pieces were successfully retrieved from the pt. | IS2000 PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 778445 | 790789 | 754631 | 2955842200600133 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-10-24 | 2006-11-22 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | Yes | 2006-10-24 | Yes | No | Initial | 2006
40018301 |  | Malfunction |  | It was reported that during the surgical procedure the cable of the permanent cautery hook instrument broke. No further details were provided. No pt harm, adverse outcome or injury was reported. | H6: the instrument was returned and evaluated. Per the customer reported complaint, engineering eval observed that the instrument conductor wire was broken at the wrist near the distal idler pulley. Burn marks were also found on the proximal clevis, distal clevis, and yaw pulley. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 778463 | 790807 | 754649 | 2955842200600131 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2006-10-23 | 2006-11-22 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2006-10-23 | Yes | No | Reuse | 2006
40020502 | 2006-10-24 | Malfunction |  | It was reported that during the surgical procedure, it was noticed that the accessory port was very close to the left system arm and instrument port. The surgeon was advised to reposition the accessory port, but declined. When the surgeon used the fenestrated bipolar forceps instrument, the instrument fractured and pieces fell inside of the patient. The instrument fragments were retrieved and the fenestrated bipolar forceps instrument was replaced with a tenaculum forceps instrument to continue with the planned surgical procedure. When using the tenaculum forceps instrument, the surgeon again declined to reposition the accessory port. The instrument fractured inside of the patient and pieces fell inside of the patient. The instrument fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of a different type. Mfr report #2955842-2006-00135 was submitted for the tenaculum forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, however, per the information provided in the case notes, it was determined that the instrument fracture experienced by the customer was due to a collision between an isi instrument and a non-isi instrument accessory. The instrument and accessory collision may be associated with the location of the surgical port placements. As indicated in the complaints notes, the fragmented components were successfully retrieved. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 778649 | 791002 | 754833 | 2955842200600134 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-24 | 2006-11-22 | No | Yes | Health Professional | 40020502 | No | No | 2006-10-24 | Device Not Retu | No | Unkown | 2006
40020705 | 2006-10-24 | Malfunction |  | It was reported that during the surgical procedure, it was noticed that the accessory port was very close to the left system arm and instrument port. The surgeon was advised to reposition the accessory port, but declined. When the surgeon used the fenestrated bipolar forceps instrument, the instrument fractured and pieces fell inside of the patient. The instrument fragments were retrieved and the fenestrated bipolar forceps instrument was replaced with a tenaculum forceps instrument to continue with the planned surgical procedure. Mfr report #2955842-2006-00134 was submitted for the fenestrated bipolar forceps instrument. When using the tenaculum forceps instrument, the surgeon again declined to reposition the accessory port. The instrument fractured inside of the patient and pieces fell inside of the patient. The instrument fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of a different type. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, however, per the information provided in the case notes, it was determined that the instrument fracture experienced by the customer was due to a collision between an isi instrument and a non-isi instrument accessory. The instrument and accessory collision may be associated with the location of the surgical port placements. As indicated in the complaints notes, the fragmented components were successfully retrieved. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 778671 | 791025 | 754855 | 2955842200600135 | 1 | GDW | Manufacturer | Company Representative | Other | Initial | 2006-10-24 | 2006-11-22 | No | Yes | Health Professional | 40020705 | No | No | 2006-10-24 | Device Not Retu | No | Unkown | 2006
40016906 | 2006-10-18 | Malfunction |  | It was reported that during the surgical procedure, the harmonic shears insert broke and fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement insert. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 778676 | 791030 | 754860 | 2955842200600132 | 1 | GDW | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-10-24 | 2006-11-22 | No | Yes | Health Professional | 40016906 | No | Yes | 2006-10-24 | Device Not Retu | Yes | Initial | 2006
40017904 | 2006-10-25 | Malfunction |  | It was reported that during the surgical procedure the surgeon noticed an unusual burning smell coming from the monopolar curved scissors instrument. The surgeon performed a visual inspection of the instrument and a small hole in the shaft of the instrument was observed. The instrument was removed and the planned surgical procedure was completed with a replacement instrument of a different type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a hole was burnt through the instrument main tube and part of the tube extension, indicating that arcing had occurred. A hairline crack was also found on the distal end of the main tube. Based on this discovery, engineering concluded that the instrument arced during the surgical procedure and the source of the arcing most likely originated from the crack found in the main tube. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 778679 | 791034 | 754863 | 2955842200600137 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-25 | 2006-11-22 | No | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | No | 2006-10-25 | Yes | No | Reuse | 2006
IS2000 A50P5 | 2006-10-25 | Other | Other | It was reported that during the beginning of the surgical procedure the surgeon experienced a recurring 23008 system error code when he rolled his wrist. No patient harm was reported. The patient's incisions were closed and the planned surgical procedure was rescheduled. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Engineering discovered a defective potentiometer assembly within the mtml, thus generating the system error experienced by the customer. The system alarm (the 23008 system error code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of november 20, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 785018 | 797485 | 761171 | 2955842200600136 | 1 | KOG | Manufacturer | Company Representative | Other | Initial |  | 2006-11-22 | Yes | Yes | Health Professional | IS2000 A50P5 | No | No | 2006-10-25 | Device Not Retu | No | Reuse | 2006
| 2006-10-18 | Malfunction |  | Surgeon was using the harmonic curved shears during a robotic-assisted laparoscopic radical prostatectomy with preservation of the vascular bundles on both sides when the movable cutting tip broke. The tip was retrieved and removed. |  | HARMONIC CURVED SHEARS INSERT | ROBOTIC HARMONIC CURVED SHEAR | 774979 | 787214 | 751213 | 787214 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-11-20 | 2006-11-20 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2006
IS1200 A43P7 | 2006-10-18 | Other | Other | It was reported that during the surgical procedure, the system shut down and the customer experienced a low battery error code when restarting the system. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering discovered that the surgeon's console power plug was plugged into a multi-plug power distribution unit which was not connected correctly, thus causing the system to switch to battery power multiple times during the surgical procedure and consequently depleting the battery power and generating the error code experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 775898 | 788157 | 752117 | 2955842200600128 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-10-18 | 2006-11-17 | Yes | Yes | Health Professional | IS1200 A43P7 | No | No | 2006-10-18 | Device Not Retu | No | Reuse | 2006
IS2000 A5 0P5 | 2006-10-20 | Other | Other | It was reported that during the surgical procedure the customer experienced a recurring 23008 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered a defective potentiometer assembly within the mtmr, thus generating the system error experienced by the customer. The system alarm (the 23008 system error code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 775915 | 788174 | 752134 | 2955842200600129 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-10-20 | 2006-11-17 | Yes | Yes | Health Professional | IS2000 A5 0P5 | No | No | 2006-10-20 | Device Not Retu | No | Reuse | 2006
40017206 |  | Malfunction |  | It was reported that during the surgical procedure, the wrist of the bipolar maryland forceps instrument broke and fragments fell inside of the pt. The fragments were retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken proximal clevis and main tube were missing. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. As indicated in the complaints notes, the fragmented components were successfully retrieved. | BIOPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 776680 | 788994 | 752886 | 2955842200600130 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-10-20 | 2006-11-17 | No | Yes | Health Professional | 40017206 | Device Returned To Manufacturer | Yes | 2006-10-20 | Yes | No | Reuse | 2006
IS1200 A4 3P5 | 2006-09-28 | Other | Other | It was reported that during the surgical procedure, the customer experienced a nonrecoverable 20008 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered excessive jitter noise emanating from a potentiometer assembly within the mtmr, which consequently generated the system error experienced by the customer. The system alarm (20008 system error code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 769014 | 781234 | 745314 | 2955842200600127 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2006-09-28 | 2006-11-10 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | Yes | 2006-09-28 | Device Not Retu | No | Reuse | 2006
| 2006-11-01 | Malfunction |  | During procedure endowrist prograsp instrument broke while engaged and holding on to the fibroid attached to the patient's uterus. While under laparoscopic observation, it was noted that one piece of the instrument had broken off. The piece was recovered laparoscopically and removed. While removing the instrument from the patient's abdominal cavity additional pieces of the instrument broke off. The procedure was converted to an open procedure. The abdomen was irrigated and explored by direct vision and feel. The procedure carried on without further incident. |  | DAVINCI ENDOWRIST PROGRASP | ROBOTIC ASSISTED DEVICE | 768633 | 780854 | 744934 | 780854 | 1 | GEI | User Facility |  |  | Initial | 2006-11-09 | 2006-11-09 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2006
40017204 |  | Malfunction |  | It was reported that during the surgical procedure, it was noticed that the tip of the bipolar maryland tip forceps instrument was broken and a part of the shaft was missing. The planned surgical procedure was completed with another instrument. No further details were provided. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken at the main tube interface and that a piece was missing from the proximal clevis and main tube. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTOSURGICAL INSTRUMENT | 770054 | 782257 | 746337 | 2955842200600126 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2006-10-09 | 2006-11-08 | No | Yes | Health Professional | 40017204 | Yes | Yes | 2006-10-09 | Yes | No | Reuse | 2006
| 2006-10-20 | No Answer Provided |  | During a laparoscopic assisted vaginal hysterectomy, the surgeon placed the maryland bipolar shears on the patient's abdomen. He inadvertently stepped on the foot pedal, causing a 2 mm superficial burn to the patient's abdomen. A contributing factor was that the gyn laparoscopic pack was not available. This pack contains an instrument pouch which would have held the shears. The or was also out of the alternative instrument pouch. |  | DAVINCI MARYLAND | SHEARS BIPOOLAR | 775239 | 787471 | 751467 | 787471 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-11-03 | 2006-11-03 | No | No | Physician |  | No | No Answer Provided |  |  |  |  | 2006
42009304 | 2006-10-04 | Malfunction |  | It was reported that during the surgical procedure the wrist of the is2000 endowrist prograsp instrument separated from the shaft and broke inside of the patient. All broken pieces were retrieved and discarded. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken at the main tube interface and that a piece of the proximal clevis was missing. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. As indicated in the complaints notes, all broken pieces were successfully retrieved. | IS2000 ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 766334 | 778559 | 742649 | 2955842200600125 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-04 | 2006-11-02 | No | Yes | Health Professional | 42009304 | Device Returned To Manufacturer | No | 2006-10-04 | Yes | No | Reuse | 2006
42017204 | 2006-09-29 | Malfunction |  | It was reported that during the surgical procedure the shaft of the is2000 bipolar maryland tip forceps instrument broke and a piece fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | H10: - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken at the main tube interface. A piece of the proximal clevis was missing and the main tube was found to be intact. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. As indicated in the complaint notes, the broken piece was successfully retrieved. | IS2000 BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 766344 | 778569 | 742659 | 2955842200600124 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-10-03 | 2006-11-02 | No | Yes | Health Professional | 42017204 | Device Returned To Manufacturer | No | 2006-10-03 | Yes | No | Reuse | 2006
IS2000 A50P3 | 2006-10-03 | Other | Other | It was reported that during the surgical procedure, the system faulted and the system video became lost and could not be restored. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Conclusions: the investigation conducted by field service engineering concluded that the system fault and loss of video experienced by the customer was caused by a non-isi monitor which was loading the system down. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of november 1, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 768813 | 781034 | 745114 | 2955842200600123 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-10-03 | 2006-11-02 | Yes | Yes | Health Professional | IS2000 A50P3 | No | No | 2006-10-03 | Device Not Retu | No | Reuse | 2006
40017206 | 2006-09-10 | Malfunction |  | It was reported that during the surgical procedure, the shaft of the bipolar maryland tip forceps instrument shattered inside of the patient. All broken pieces were believed to be retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken at the main tube interface. The proximal clevis and main tube were both missing pieces. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis. As indicated in the complaints notes, it is believed that all broken pieces were retrieved. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 766054 | 778284 | 742375 | 2955842200600122 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-10-02 | 2006-11-01 | No | Yes | Health Professional | 40017206 | Device Returned To Manufacturer | Yes | 2006-10-02 | Yes | No | Reuse | 2006
IS2000 A50P5 | 2006-09-27 | Other | Other | It was reported that during the surgical procedure, the customer experienced multiple system error codes which could not be overridden. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Based on the system logs, the motor encoder assemblies and potentiometers were replaced. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of october 27, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURIGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 764045 | 776267 | 740367 | 2955842200600119 | 1 | KOG | Manufacturer | Company Representative | Other Health Care Professional | Initial | 2006-09-27 | 2006-10-27 | Yes | Yes | Health Professional | IS2000 A50P5 | No | Yes | 2006-09-27 | Device Not Retu | No | Reuse | 2006
40018301 |  | Malfunction |  | A permanent cautery hook instrument was returned for evaluation. The reason for the instrument return was not provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation observed that the instrument conductor was broken at the wrist near the proximal clevis pulley. Burn marks were also found on the proximal clevis, distal clevis, and distal idler pulley, near the location where the conductor broke. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 799143 | 811656 | 775020 | 2955842200600121 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-29 | 2006-10-27 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-09-29 | Yes | No | Reuse | 2006
40018301 |  | Malfunction |  | It was reported that during the surgical procedure, the permanent cautery hook instrument would start to work and then stop. The planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | H6: the instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation observed that the instrument hook was partially pulled out from the ceramic sleeve, preventing it from making electrical contact with the conductor. Engineering also found signs of arcing on the instrument yaw pulley and proximal clevis. This complaint is being reported due to the evidence of arcing found, which has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 799259 | 811772 | 775134 | 2955842200600118 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-28 | 2006-10-27 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-09-28 | Yes | No | Reuse | 2006
40000620 | 2006-09-28 | Malfunction |  | It was reported that the tip of the large needle driver instrument broke off and fell inside of the pt. The broken tip was not retrieved. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering eval observed that one of the instrument grips was completely broken off and missing. The grip had broken off at the cable crimp hole, leaving the cable crimp partially exposed. Biocompatibility testing was previously performed on a rep instrument that contained stainless steel and the tungsten carbide materials. The testing met the requirements of the fda and iso 10993-1 standard for externally communicating devices having indirect blood path contact for limited contact duration. A list of the biocompatibility tests performed on the test instrument and results are as follows: test: cytoxicity study. Results: passed - materials were determined to be non-cytotoxic. Tests: irritation test: results: passed - no evidence of causing delayed dermal contact sensitization. Test: acute systemic toxicity. Results: passed - no mortality or evidence of systemic toxicity. Test: hemolysis study. Results: passed. Test: pyrogen study. Results: passed - materials were determined to be non-pyrogenic. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 799569 | 812086 | 775443 | 2955842200600120 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-28 | 2006-10-27 | No | Yes | Health Professional | 40000620 | Device Returned To Manufacturer | Yes | 2006-09-28 | Yes | No | Reuse | 2006
42017903 |  | Malfunction |  | It was reported that electrical failure was experienced with the is2000 monopolar curved scissors instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that a hole was burnt through the instrument main tube and part of the tube extension, indicating that arcing had occurred. A hairline crack was also found on the distal end of the main tube. Based on this discovery, engineering concluded that the instrument arced during a previous surgical procedure and the source of the arcing most likely originated from the crack found in the main tube. No other damage was found. | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 815369 | 827964 | 791086 | 2955842200600117 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-25 | 2006-10-27 | No | Yes | Health Professional | 42017903 | Device Returned To Manufacturer | Yes | 2006-09-25 | Yes | No | Reuse | 2006
42011003 | 2006-09-20 | Malfunction |  | It was reported that during the surgical procedure, the cable at the end of the precise bipolar pyramid tip forceps instrument broke. Also, a small "metal ball" was noticed inside of the pt. The "metal ball" was retrieved and the planned surgical procedure was completed with another instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one of the instrument conductor wires was broken at the location where it enters the yaw pulley and connects to the grip. The distal clevis also had a small amount of material removed at the edge. It was concluded that the "metal ball" referenced by the customer was actually a piece of the distal clevis, as all of the other metal components in the instrument wrist were found to be intact. Engineering eval determined that the failure mode experienced by the customer was a result of damage to the instrument distal clevis. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 820192 | 832873 | 795837 | 2955842200600116 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-09-20 | 2006-10-20 | No | Yes | Health Professional | 42011003 | Device Returned To Manufacturer | No | 2006-09-20 | Yes | No | Reuse | 2006
40016904 | 2006-09-18 | Malfunction |  | It was reported that during the surgical procedure, the white portion of the harmonic shears insert broke off inside the trocar. No components fell into the pt. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. As indicated in the case notes, no components fell into the pt. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 817053 | 829670 | 792758 | 2955842200600115 | 1 | LFL | Manufacturer | Company Representative | Other | Initial | 2006-09-18 | 2006-10-18 | No | Yes | Health Professional | 40016904 | No | Yes | 2006-09-18 | Device Not Retu | Yes | Initial | 2006
| 2006-09-26 | Malfunction |  | Physician was performing laparoscopic radical prostatectomy with robotic assistance when during the procedure, a fault occurred with one of the instrument arms. Despite several attempts at restarting the system and with telephone contact to the company, the system fault could not be overridden and the procedure could not go further in its current status. The robot was then moved away from the patient and the procedure was converted to an open radical prostatectomy. The patient tolerated the procedure well and was transferred to the recovery room in excellent condition. The patient's expected outcome changed with the failure of the robotic assisted surgery, as patient went from a minimally invasive procedure to a more invasive one (open). Analysis of the device identified a "3rd arm fault. " error code 25504 was identified, which basically indicated there was a "mismatch" code, meaning that there is a problem with the arm lining up perfectly. Mismatch can occur for a variety of reasons, from part failure to a dirty sensor. It was determined that the psm (patient-side manipulator) was suspect. The service representative ordered and installed a new psm; tested system, system passed all tests. The removed part was sent to the home office for analysis. |  | DA VINCI S SURGICAL SYSTEM | SURGICAL ROBOT | 762595 | 774813 | 738913 | 774813 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2006-09-22 | 2006-10-17 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2006
IS1200 A43P6 | 2006-09-14 | Other | Other | It was reported that during the surgical procedure, the customer experienced a nonrecoverable 20008 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Based on the system logs, a motor encoder and potentiometer were replaced. As of october 10, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 758202 | 770372 | 734566 | 2955842200600114 | 1 | KOG | Manufacturer | Distributor | Other | Initial | 2006-09-14 | 2006-10-12 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-09-14 | Device Not Retu | No | Reuse | 2006
40018401 | 2006-07-05 | Malfunction |  | A sterile processing staff member noticed that a wire was broken on the permanent cautery spatula instrument and a piece of the instrument was missing. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for eval, therefore, the cause of the customer reported complaint cannot be determined nor can it be determined if any components separated from the instrument during a surgical procedure or during handling after the procedure. The hosp has not returned requests for add'l info. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 745732 | 970507 | 722431 | 2955842200600113 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2006-09-12 | 2006-10-12 | No | Yes | Health Professional | 40018401 | No | Yes | 2006-09-12 | Device Not Retu | No | Unkown | 2006
40018301 | 2006-08-16 | Malfunction |  | It was originally reported that the permanent cautery hook instrument malfunctioned and there was no report of components having separated from the instrument. On 09/12/06, additional info was rec'd. It was reported that during the surgical procedure, the permanent cautery hook instrument malfunctioned. The instrument was removed from the patient and upon closer inspection, the working end of the instrument appeared frayed and spider web cracks were noted on the shaft. Two of three pieces of plastic material were manually removed from the pt's abdomen and more pieces were found inside the laparoscopic trocar. The pt's abdomen was suctioned and irrigated with one liter of saline. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a break at the insturment proximal clevis and main tube interface. The proximal clevis was found to be intact, however, the main tube was missing pieces. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage in the main tube. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 744897 | 973637 | 721592 | 2955842200600112 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-12 | 2006-10-12 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-09-12 | Yes | No | Reuse | 2006
40016906 | 2006-09-08 | Malfunction |  | It was reported that during the surgical procedure the tip of the harmonic shears insert broke off and fell into the pt. The tip was retrieved and the planned surgical procedure was completed using a replacement insert. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, engineering evaluation found one of the arms to be broken off. The section of the shaft where the clamp arm inserts is bent inwards, indicating that breakage most likely occurred when high force was applied at that location. Per the complaints notes, the detached clamp arm was successfully retrieved from the pt. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 841974 | 857964 | 820620 | 2955842200600110 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-09-08 | 2006-10-06 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | Yes | 2006-09-08 | Yes | Yes | Initial | 2006
40018401 | 2006-09-06 | Malfunction |  | It was reported that small pieces of the permanent cautery spatula instrument broke off and fell inside of the pt. All fragmented pieces were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument ceramic sleeve was cracked and missing pieces. The conductor wire was also kinked near the proximal clevis pulley. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the ceramic sleeve of the instrument. As indicated in the complaint notes, all fragmented pieces were retrieved. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 841977 | 857965 | 820621 | 2955842200600109 | 1 | GEI | Manufacturer | Company Representative | Other Health Care Professional | Initial | 2006-09-06 | 2006-10-06 | No | Yes | Health Professional | 40018401 | Device Returned To Manufacturer | Yes | 2006-09-06 | Yes | No | Reuse | 2006
40016906 | 2006-09-08 | Malfunction |  | It was reported that during the surgical procedure the tip of the harmonic shears insert broke off and fell into the pt. The tip was retrieved and the planned surgical procedure was completed using a replacement insert. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, engineering evaluation found a section of one of the clamping arms to be broken off. Engineering evaluation determined that the breakage most likely occurred when high force was applied to the clamping arm. No other damaged was found. As indicated in the complaint notes, the broken piece was successfully retrieved from the pt. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 842023 | 858151 | 820807 | 2955842200600111 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-09-08 | 2006-10-06 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | Yes | 2006-09-08 | Yes | Yes | Initial | 2006
40012608 |  | Malfunction |  | It was reported that the permanent cautery hook instrument snapped. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation observed the instrument conductor wire had been damaged. Engineering also found evidence that the instrument had previously arced due to the damaged conductor wire. It was determined that due to the burnt material in the instrument yaw pulley, the cable crimp had slipped and became exposed, thus creating the "snapped" appearance reported by the customer. The instrument was returned with no missing components. Based on the evidence of previous arcing, engineering concluded that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 842009 | 858143 | 820799 | 2955842200600108 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-09-05 | 2006-10-05 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2006-09-05 | Yes | No | Reuse | 2006
| 2006-09-27 | Malfunction |  | During procedure, some small black particles were noted as debris from davinci's monopolar curved scissor. Wound irrigated and black particles suctioned out of wound by surgeon. |  | DA VINCI SURGICAL SYSTEM | INSTRUMENT ROBOTIC SURGICAL MO | 766571 | 778794 | 742884 | 778794 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-10-03 | 2006-10-03 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2006
IS1200 A43P6 | 2006-08-25 | Other | Other | It was reported that during a surgical procedure, the customer experienced a 20008 system error code which could not be overridden. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered a broken carriage clamp screw on one of the axes, thus generating the system error experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of september 21, 2006 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 753091 | 765238 | 729626 | 2955842200600107 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-08-25 | 2006-09-22 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-08-25 | Device Not Retu | No | Reuse | 2006
40017905 |  | Malfunction |  | It was reported that the housing of the monopolar curved scissors instrument was broken. No further details were provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that a hole was burnt through the instrument main tube and part of the tube extension, which indicates that arcing had occurred. A hairline crack was also found on the distal end of the main tube, below the hole. Based on this discovery, engineering concluded that the instrument arced during a previous surgical procedure, and the source of the arcing most likely originated from the crack found in the main tube. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 848892 | 867378 | 829994 | 2955842200600105 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-08-24 | 2006-09-22 | No | Yes | Health Professional | 40017905 | Yes | Yes | 2006-08-24 | Yes | No | Initial | 2006
40012607 | 2006-08-25 | Malfunction |  | It was reported that during the surgical procedure, the permanent cautery hook instrument was arcing and smoking at the tip. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found evidence of arcing at the instrument yaw pulley, conductor cap, and on one side of the distal clevis. This discovery has led engineering to conclude that the instrument arced during the surgical procedure, thus causing the smoke experienced by the customer. This instrument was manufactured in january 2004. The instrument has since been redesigned to improve performance. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 848903 | 867431 | 830047 | 2955842200600106 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-08-25 | 2006-09-22 | No | Yes | Health Professional | 40012607 | Yes | Yes | 2006-08-25 | Yes | No | Reuse | 2006
40018301 | 2006-06-30 | Malfunction |  | It was reported that the permanent cautery hook instrument was not getting an electrical charge. No further details were provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found that the instrument conductor was broken as it exits the proximal clevis and that part of the distal clevis was burnt where the conductor broke. Evidence of insulation melting is also visible. Based on this discovery, engineering concluded that the instrument may have arced during a previous surgical procedure, thus contributing to the malfunction experienced by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 848888 | 867354 | 829970 | 2955842200600104 | 1 | GEI | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2006-08-22 | 2006-09-21 | No | Yes | Health Professional | 40018301 | Yes | Yes | 2006-08-22 | Yes | No | Reuse | 2006
IS1200 A43P6 | 2006-08-21 | Other | Other | It was reported that during the surgical procedure, the customer experienced a 20008 system error code when trying to attach a sterile adapter onto 1 of the system arms. The spare arm was used in its place, and although the spare arm was lifted to it's maximum capacity, the surgeon experienced low visibility and had difficulty distinguishing planes. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field svc engineering concluded that the 20008 system fault experienced by the customer was associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered a broken clamp screw on 1 of the axes which allowed the shaft to spin freely, thus generating the system error experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of september 18, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 752325 | 764468 | 728864 | 2955842200600103 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2006-08-21 | 2006-09-20 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-08-21 | Device Not Retu | No | Reuse | 2006
40017904 | 2006-07-21 | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No further details were provided. No pt harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Performance testing was conducted and engineering found the instrument to perform as intended. Engineering also observed that an area of the instrument tube extension was burnt at the distal end and that part of the instrument proximal clevis was also visible in the burnt area. Based on this discovery, engineering concluded that the instrument may have arced during a previous surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 851703 | 870882 | 833430 | 2955842200600102 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-08-21 | 2006-09-20 | No | Yes | Health Professional | 40017904 | Yes | Yes | 2006-08-21 | Yes | No | Reuse | 2006
40012608 |  | Malfunction |  | It was reported that the permanent cautery hook instrument was not being recognized by the system. No further details were provided. No pt harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Engineering eval found evidence of arc tracking located at the instrument gooseneck and yaw pulley. Based on this discovery, engineering concluded that the instrument may have arced during a previous surgical procedure, thus contributing to the system recognition issue experienced by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 851764 | 871093 | 833637 | 2955842200600101 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2006-08-21 | 2006-09-20 | No | Yes | Invalid Data | 40012608 | Device Returned To Manufacturer | No | 2006-08-21 | Yes | No | Reuse | 2006
IS1200 A43P6 | 2006-05-11 | Other | Other | It was reported that during the surgical procedure, the system shifted to battery power. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation concluded that the system issue experienced was associated with the battery pack module (bpm). The system was repaired by replacing the affected bpm. The bpm was returned to isi for failure analysis investigation. It was determined that all four batteries of the bpm were out of spec and needed to be replaced. As of 9/20/06, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 853870 | 873694 | 836175 | 2955842200600100 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-05-11 | 2006-09-20 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-05-11 | Device Not Retu | No | Reuse | 2006
400172 | 2006-09-11 | Other | Other | Patinet was having laparoscopic robot assisted radical prostatectomy. A small piece of plastic broke from the proximal clevis of the instrument. Piece was retrieved by surgeon during the procedure. No adverse effect on patient. |  | INTUITIVE DEVINCI ROBOT | MARYLAND BIPOLAR ARM | 750471 | 762572 | 727046 | MW1040453 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2006-09-13 | 2006-09-19 | No | Yes | Health Professional | 400172 | Yes | Yes |  |  |  |  | 2006
AESOP 3000 | 2006-08-22 | Malfunction |  | When bovie in use aesop command conflicts between microphone connected, microphone disconnected making the aesop inoperable. There were two such incidents; no harm to either patient. Biomed assessment: possible interference between valleylab esu units and the aesop units. Recommendation is that the or use only a certain esu. When using aesop units this particular esu does not seem to cause the problem. All three aesop units have had related issues. Recommendation is that a certain esu only be used when the aesop units are being used. Biomed is also following up with the manufacturer to determine if they are aware of the problem. The site has not received a response. |  | ROBOTIC LAPARASCOPE HOLDER | ROBOTIC SYSTEM | 750555 | 762655 | 727129 | 762655 | 1 | KOG | User Facility |  |  | Initial | 2006-09-15 | 2006-09-15 | No | Yes | Physician | AESOP 3000 | Yes | No Answer Provided |  |  |  |  | 2006
IS2000 A50P3 | 2006-08-18 | Other | Other | It was reported that during the surgical procedure, the customer experienced an unrecoverable system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the s5vp pca. The system was repaired by replacing the affected s5vp pca. As of 09/12/06, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 751333 | 763448 | 727890 | 2955842200600099 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-08-18 | 2006-09-15 | Yes | Yes | Health Professional | IS2000 A50P3 | No | No | 2006-08-18 | Device Not Retu | No | Reuse | 2006
40017904 | 2006-08-17 | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and the instrument performed as intended. Engineering also observed that the instrument conductor was burnt just after entering the idler block. The area where the burn was located was not covered by clear plastic insulation, which indicates that arcing most likely occurred. Based on this discovery, engineering concluded that the instrument may have arced a previous surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 853716 | 873430 | 835911 | 2955842200600098 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-08-17 | 2006-09-15 | No | Yes | Health Professional | 40017904 | Yes | No | 2006-08-17 | Yes | No | Reuse | 2006
IS2000 A50P3 | 2006-08-14 | Other | Other | It was reported that during the surgical procedure, the customer experienced multiple 23013 and 23008 recoverable system error codes. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field svc engineering concluded that the system faults experienced by the customer were associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Performance and functional tests were performed and engineering was unable to replicate the customer reported complaint. As of september 12, 2006, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 749832 | 761941 | 726428 | 2955842200600097 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2006-08-14 | 2006-09-13 | Yes | Yes | Health Professional | IS2000 A50P3 | No | Yes | 2006-08-14 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2006-08-09 | Other | Other | It was reported that during the surgical procedure, the da vinci surgical system would not recognize the surgical instruments, therefore, the system could not be used to complete the planned surgical procedure. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system was evaluated by an isi field svc engineer (fse). The isi fse was unable to replicate the customer reported symptom. Per additional engineering investigation performed on similar failure modes, it was determined that the instrument engagement issues experienced by the customer were related to the instrument chassis not properly engaging with the instrument sterile adapter. As of september 7, 2006, there have been no reported recurrence of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 747193 | 759276 | 723817 | 2955842200600095 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-08-09 | 2006-09-08 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-08-09 | Yes | No | Reuse | 2006
40018402 | 2006-08-09 | Malfunction |  | It was reported that during the surgical procedure, the permanent cautery spatula instrument was arcing at the tip and the wire in the tip broke. No pt harm, adverse outcome or injury was reported. | Conclusion: the instrument was returned and evaluated. Per the customer reported complaint, engineering found one broken pitch cable near the cable crimp. Engineering also observed that the instrument conductor wire was burnt and the insulation was removed near one distal idler pulley. Part of the pulley had been burned of, thus exposing the conductor wire. This discovery has led engineering to conclude that the instrument arced during the surgical procedure and that instrument arcing most likely occurred due to the exposed conductor wire. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 858459 | 879202 | 841643 | 2955842200600096 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-08-09 | 2006-09-08 | No | Yes | Health Professional | 40018402 | Device Returned To Manufacturer | No | 2006-08-09 | Yes | No | Reuse | 2006
400169XX | 2006-08-07 | Malfunction | Other | It was reported that during the surgical procedure, the tip of the harmonic curved shears insert broke and fell inside of the pt. The tip was retrieved, the planned surgical procedure was completed. It was also reported that the site disposed of the insert. No pt harm, adverse outcome or injury was reported. | Because the harmonic curved shears insert was disposed of the by the site, evaluation could not be performed and no conclusion can be drawn regarding the cause of the insert breakage. | HARMONIC CURVED INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 764978 | 777198 | 741297 | 2955842200600094 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-08-07 | 2006-09-06 | No | Yes | Health Professional | 400169XX | No | No | 2006-08-07 | Device Not Retu | Yes | Initial | 2006
| 2006-07-05 | Malfunction |  | Near miss event, after a 2nd robotic surgery of the day, sterile processing staff identified that a wire was broken and a piece of the device was missing. |  | ENDOWRIST | ROBOTIC PERMANENT CAUTERY SPAT | 745732 | 757832 | 722431 | 757832 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2006-08-30 | 2006-08-30 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2006
40017905 | 2006-07-27 | Malfunction |  | It was reported that during the surgical procedure a bad smell was emanating from the monopolar curved scissors instrument when pressing the coag footswitch. The instrument was removed and replaced with a different type of instrument to complete the planned surgical procedure. During cleaning, a burn near the protective cover of the instrument was noticed and that the wires inside of the shaft were exposed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering discovered a hairline crack on the distal end of the main tube of the instrument, below the hole. It was determined that as the instrument was energized, arcing originated from the crack, which then spread distally and compromised the overmolded seal at the tube extension. It was concluded that mechanical damage to the instrument shaft near the tip of the instrument is the root cause of the hairline crack. It is understood that the instrument was used without any issues during the first use, therefore, it is believed that the damage most likely occurred prior to the second use of the instrument. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 770195 | 782397 | 746477 | 2955842200600093 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2006-07-31 | 2006-08-30 | No | Yes | Health Professional | 40017905 | Device Returned To Manufacturer | No | 2006-07-31 | Yes | No | Reuse | 2006
4001833 | 2006-08-16 | Malfunction |  | Surgeon was using the 8mm davinci hook cautery, the instrument malfunctioned. It was removed from the patient and upon closer inspection the working end of the instrument appeared frayed and spider webbed cracks were noted on the shaft. Two or three pieces of plastic material were manually removed from the patient's abdomen by the surgeon, and more pieces were found inside the laparoscopic trocar. The patient's abdomen was suctioned and irrigated with one liter of saline. |  | DAVINCI | SURGICAL ROBOT CAUTERY HOOK | 744897 | 756994 | 721592 | 756994 | 1 | GEI | User Facility |  |  | Initial | 2006-08-29 | 2006-08-29 | No | Yes | Physician | 4001833 | No | No Answer Provided |  |  |  |  | 2006
IS1200 A43P6 | 2006-07-27 | Other | Other | It was reported that during the beginning of the surgical procedure, when the first instruments were inserted into the system, the customer experienced a nonrecoverable 21008 system fault. Although the patient port incisions were already made, no patient harm was reported and the planned surgical procedure was cancelled and rescheduled for a later date. | Fourty seven - the investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered two pulley cables had dislodged, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of august 24, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 745264 | 757366 | 721965 | 2955842200600091 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-07-27 | 2006-08-25 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-07-27 | Device Not Retu | No | Reuse | 2006
SS2000 | 2006-07-11 | No Answer Provided |  | The event occurred during a robotic thymectomy. The surgeon was operating with the robot and we lost power to the robot. Everything went black. The instruments were removed from the patient. The sales rep was present. We began to inspect the plugs & power strips to the tower. Biomed came to the or. Before biomed arrived we had gotten the system back up. We unplugged and replugged in the tower until it repowered. The robot was reactivated and surgery continued. The sales rep discovered that a plug to the davinci was changed and placed in one of the power strips which caused too much power to be drawn. The light source was to have its own outlet. Biomed corrected the problem by getting us a new cord which was replaced after the procedure. |  | DA VINCI | ROBOTIC SURGICAL DEVICE | 749427 | 761520 | 726022 | 761520 | 1 | KOG | User Facility |  | RISK MANAGER | Initial | 2006-08-25 | 2006-08-25 | No |  | Physician | SS2000 | Yes | No Answer Provided |  |  |  |  | 2006
40020704 | 2006-07-27 | Malfunction |  | It was reported that during the surgical procedure, when the surgeon was attempting to remove the myomas from the patient's uterus, the tip of the tenaculm forceps instrument broke and pieces of the instrument fell inside of the patient. As a result of the instrument breakage, the instrument could not be removed from the cannula and the surgical staff had to remove the instrument and cannula simultaneously from the patient. It was reported that all of the instrument fragments were retrieved and the planned surgical procedure was completed with a different type of instrument. No patient harm, adverse outcome or injury was reported. | The affected instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determned. As stated in the complaint notes, all instrument fragments were retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 764903 | 777123 | 741222 | 2955842200600092 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-07-27 | 2006-08-25 | No | Yes | Health Professional | 40020704 | No | No | 2006-07-27 | Device Not Retu | No | Initial | 2006
42009304 | 2006-07-24 | Other | Other | It was reported that during the surgical procedure, while the surgeon was attempting to reposition the trocar and retract the pt's bladder, the surgeon collided the is2000 endowrist prograsp forceps instrument with another instrument while inside of the pt. As a result, the area of the instrument where the wrist and shaft meet broke. No instrument components fell inside of the pt. The instrument was removed and replaced with the same type of instrument. While using the replacement instrument, the surgeon collided the instrument with the pt's pelvis or abdominal wall, which resulted in the portion of the instrument, where the shaft and wrist meet broke. No instrument components fell inside of the pt. Due to the delay in the procedure and the pt's anatomy, the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. The regulatory reporting info for the initial is2000 endowrist prograsp forceps instrument used during the surgical procedure can be found in mfr #2955842-2006-00089. | Although, the instrument was not returned for eval, it was determined that user error contributed to the malfunction of the instrument when the surgeon collided the affected instrument with the pt's pelvis or abdominal wall (unsure of which one). The instruments instructions for use specifically states: contraindications: this instrument may only be used on soft tissue. Do not use it on cartilage, bone or hard objects. Doing so may damage the instrument and make it impossible to remove it from the cannula. General precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the complaint notes, no instrument components fell into the pt and the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery due to the delay in the procedure caused by the instrument malfunctions and the pt's anatomy. | IS2000 ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 745200 | 757296 | 721886 | 2955842200600090 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-07-24 | 2006-08-23 | Yes | Yes | Health Professional | 42009304 | No | No | 2006-07-24 | Device Not Retu | No | Unkown | 2006
42009303 | 2006-07-24 | Other | Other | It was reported that during the surgical procedure, while the surgeon was attempting to reposition the trocar and retract the pt's bladder, the surgeon collided the is2000 endowrist prograsp forceps instrument with another instrument while inside of the pt. As a result, the area of the instrument where the wrist and shaft meet broke. No instrument components fell inside of the pt. The instrument was removed and replaced with the same type of instrument. While using the replacement instrument, the surgeon collided the instrument with the pt's pelvis or abdominal wall, which resulted in the portion of the instrument, where the shaft and wrist meet broke. No instrument components fell inside of the pt. Due to the delay in the procedure and the pt's anatomy, the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery. No pt harm was reported. The regulatory reporting info for the initial is2000 endowrist prograsp forceps instrument used during the surgical procedure can be found in mfr #2955842-2006-00090. | Although, the instrument was not returned for eval, it was determined that user error contributed to the malfunction of the instrument when the surgeon collided the affected instrument with another instrument. The instruments instructions for use specifically states: contraindications: this instrument may only be used on soft tissue. Do not use it on cartilage, bone or hard objects. Doing so may damage the instrument and make it impossible to remove it from the cannula. General precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the complaint notes, no instrument components fell into the pt and the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery, due to the delay in the procedure caused by the instrument malfunctions and the pt's anatomy. | IS2000 ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 745202 | 757298 | 721888 | 2955842200600089 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-07-24 | 2006-08-23 | Yes | Yes | Health Professional | 42009303 | No | No | 2006-07-24 | Device Not Retu | No | Unkown | 2006
IS1200 A43P6 | 2006-07-18 | Other | Other | It was reported that during the beginning of the surgical procedure, the customer experienced a 20008 system error code. The patient incisions were closed and the planned surgical procedure was rescheduled for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the customer was associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering found two insertion cables out of the pulley, thus generating the system error code experienced by the customer. The system alarm (system error codes) functioned as designed and there was no injury to the patient. The planned surgical procedure was aborted and rescheduled to a later date. As of august 14, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 743219 | 755392 | 720029 | 2955842200600086 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-07-18 | 2006-08-17 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-07-18 | Device Not Retu | No | Reuse | 2006
IS2000 A50P3 | 2006-07-18 | Other | Other | It was reported that during the surgical procedure the customer experienced 20008 and 20013 system error codes which could not be overridden. The system was powered down and restarted, and upon restart, a 26004 system error code was generated. The system was again powered down and restarted and was in use for approximately 30 minutes when the customer experienced another 26004 error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered that a motor encoder assembly had malfunctioned, thus generating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of august 14, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 743225 | 755398 | 720035 | 2955842200600088 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-07-18 | 2006-08-17 | Yes | Yes | Health Professional | IS2000 A50P3 | No | Yes | 2006-07-18 | Device Not Retu | No | Reuse | 2006
40018301 | 2006-05-31 | Malfunction |  | It was reported that the permanent cautery hook instrument had a broken cable at the tip of the instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the conductor wire was burnt at the pulley interface. Engineering also found evidence of arcing at the distal clevis and pulley. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure and that instrument arcing most likely occurred due to the damaged conductor wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 761060 | 773271 | 737377 | 2955842200600085 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-05-31 | 2006-08-17 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2006-07-18 | Yes | No | Reuse | 2006
40000319 | 2006-07-10 | Malfunction | Other | It was reported that during the surgical procedure, the tip of the clip applier instrument broke when the customer attempted to move the stabilizer with the instrument. The broken tip was not retrieved and the planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one of the instrument tips was broken off. It was determined that user error contributed the reported failure mode when the user attempted to move the stabilizer with the instrument. The instructions for use specifically states: caution: endoscopic instruments are designed and manufactured for a specific surgical function. Use of an instrument for task other than the instrument's designed use may result in a damaged or broken instrument. | CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 761074 | 773285 | 737391 | 2955842200600087 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2006-07-18 | 2006-08-17 | No | Yes | Health Professional | 40000319 | Device Returned To Manufacturer | Yes | 2006-07-18 | Yes | No | Reuse | 2006
40003519 | 2006-07-12 | Malfunction | Other | It was reported that during the surgical procedure, part of the tip of the round tooth forceps instrument broke off and fell inside the patient. The tip was retrieved using a magnet and the planned surgical procedure was completed with a different type of instrument. No patient harm, adverse outcome or injury was reported. | The instrument and detached tip were returned and evaluated. Per the customer reported complaint, engineering observed that the grip of the instrument had broken off slightly above where the cable crimp enters the grip. Engineering determined that the failure mode experienced by the customer was a result of damage to the grip of the instrument. As stated in the complaint report, the detached instrument tip was retrieved and no patient harm was reported. | ROUND TOOTH FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 760745 | 772960 | 737067 | 2955842200600083 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-07-17 | 2006-08-16 | No | Yes | Health Professional | 40003519 | Device Returned To Manufacturer | Yes | 2006-07-17 | Yes | No | Reuse | 2006
40017904 | 2006-07-17 | Malfunction |  | It was reported that during dissection of the seminal vesicles, the surgeon noticed arcing marks on the tissue. Upon inspection of the monopolar curved scissors instrument, it appeared as though a hole had been burnt into instrument tip cover accessory. The instrument tip cover accessory was removed and replaced and the same issue occurred. The instrument tip cover accessory was again removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned without the affected tip cover accessories. Functional testing was conducted on the instrument with a new tip cover accessory, and the instrument performed as designed and no arcing was observed. The new tip cover accessory used for testing did not have any holes or burn marks after removal from the instrument. Based on the testing results, engineering concluded that the arcing experienced by the customer was most likely due to defective instrument tip cover accessories. The cause of the damages to the tip cover accessories is inconclusive. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 760861 | 773074 | 737181 | 2955842200600084 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-07-17 | 2006-08-16 | No | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | No | 2006-07-17 | Yes | No | Reuse | 2006
400179XX | 2006-07-17 | Malfunction |  | It was reported that during the surgical procedure, arcing was observed. Upon further inspection of the monopolar curved scissors instrument, a split was discovered on the instrument tip cover accessory. The instrument tip cover accessory was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument and tip cover accessory were not returned for evaluation, therefore, the cause of the split on the instrument tip cover accessory cannot be determined. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 760883 | 773096 | 737202 | 2955842200600082 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-07-17 | 2006-08-16 | No | Yes | Health Professional | 400179XX | No | No | 2006-07-17 | Device Not Retu | No | Unkown | 2006
IS1200 A43P6 | 2006-06-22 | Other | Other | It was reported that during the beginning of the surgical procedure, the customer experienced multiple system power on faults. No patient harm was reported. The patient incisions were closed and the surgical procedure was rescheduled for a later date. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with an iod pca and msd pca. The system was repaired by replacing the affected iod pca and msd pca. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The patient incisions were closed and the surgical procedure was rescheduled for a later date. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 727634 | 739747 | 704961 | 2955842200600079 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-06-22 | 2006-07-21 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-06-22 | Device Not Retu | No | Reuse | 2006
400169XX | 2006-06-16 | Malfunction |  | It was reported that during the surgical procedure, the "blade" of the harmonic shears insert fell into the patient. The "blade" was retrieved and the planned surgical procedure was completed using a replacement instrument and insert. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears instrument and insert accessory were not returned for evaluation. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps. " per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. Because the instrument and insert were not returned for evaluation, the cause of the customer reported complaint cannot be determined. Per the complaints notes, the detached clamp arm was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT 5 PACK | ENDOSCOPIC INSTRUMENT | 730509 | 742681 | 707735 | 2955842200600080 | 1 | GDW | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-06-23 | 2006-07-21 | No | Yes | Health Professional | 400169XX | No | Yes | 2006-06-23 | Device Not Retu | Yes | Initial | 2006
IS1200 A43P6 | 2006-07-05 | Other | Other | It was reported that during the surgical procedure, the customer experienced multiple 20008 and 20013 system error codes. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered a broken shaft clamp screw on one of the axes which allowed the shaft to spin freely, thus generating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of july 19, 2006, there have been no reported recurrence of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 726921 | 739040 | 704243 | 2955842200600077 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-07-05 | 2006-07-19 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-07-05 | Device Not Retu | No | Reuse | 2006
IS1200 A43P6 | 2006-06-27 | Other | Other | It was reported that during the surgical procedure, the customer experienced multiple 20008 and 20013 system error codes. No pt harm was reported. The procedure was converted to traditional open surgical technique to finish the planned surgery. | The investigation conducted by field servide engineering concluded that the system faults experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure anlaysis investigation. Engineering discovered that a potentiometer had malfunctioned, thus denerating the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical technique to complete the planned procedure. As of july 19, 2006, there have been no reported recurrence of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 726934 | 739053 | 704256 | 2955842200600076 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-06-27 | 2006-07-19 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-06-27 | Device Not Retu | No | Reuse | 2006
IS2000 A5 0P3 | 2006-07-07 | Other | Other | It was reported that during the surgical procedure, the customer experienced unrecoverable 23015 system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was eapaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering discovered that the motor encoder assemblies had malfunctioned on two axes, thus generating the system errors experienced by the customer. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of july 19, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 726947 | 739066 | 704269 | 2955842200600078 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-07-07 | 2006-07-19 | Yes | Yes | Health Professional | IS2000 A5 0P3 | No | Yes | 2006-07-07 | Device Not Retu | No | Reuse | 2006
IS1200 A4 3P6 | 2006-06-21 | Other | Other | It was reported that towards the end of the surgical procedure, the customer experienced an unrecoverable 20008 system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system fault experienced by the customer was associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering determined that the malfunction was due to a potentiometer with a failed solder joint. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of july 19, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 726960 | 739079 | 704281 | 2955842200600075 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-06-21 | 2006-07-19 | Yes | Yes | Health Professional | IS1200 A4 3P6 | No | Yes | 2006-06-21 | Device Not Retu | No | Reuse | 2006
40018301 | 2005-06-12 | Malfunction |  | It was reported that during the surgical procedure, the permanent cautery hook instrument began to smoke/while inside the patient. The instrument was removed and replaced, and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed evidence of arcing at the base of the hook-pulley interface. The conductor wire was found to be damaged, which leads engineering to conclude that instrument arcing most likely occurred due to the damaged conductor wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 727115 | 739234 | 704436 | 2955842200600074 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2005-06-12 | 2006-07-14 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-06-12 | Yes | No | Reuse | 2006
40018301 | 2006-04-03 | Malfunction |  | It was reported that during the surgical procedure the cable of the permanent cautery hook instrument broke. The instrument was removed and the planned surgical procedure was completed with a different instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the conductor wire had separated from the instrument wrist. Engineering also observed evidence of charred plastic on the instrument wrist. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 725129 | 737228 | 702466 | 2955842200600071 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2006-04-03 | 2006-07-07 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-06-07 | Yes | No | Reuse | 2006
IS1200 A43P6 | 2006-06-08 | Other | Other | It was reported that during the surgical procedure, the surgeon experienced instrument engagement issues. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system was evaluated by an isi field service engineer (fse). The isi fse found no issues with the system, however, decided to replace the extended psm as a precaution. The extended psm was returned to isi for failure analysis investigation. Engineering performed functional tests on the extended psm and was unable to replicate the customer reported symptom. Per additional engineering investigation performed on similar failure modes, it was determined that the instrument engagement issues experienced by the customer were related to the instrument chassis not properly engaging with the instrument sterile adapter. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 725196 | 737295 | 702533 | 2955842200600072 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-06-08 | 2006-07-07 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-06-08 | Yes | No | Reuse | 2006
40011008 | 2006-06-09 | Malfunction |  | It was reported that during the surgical procedure smoke and fire came out of the precise bipolar pyramid tip forceps instrument. The instrument was removed and the planned surgical procedure was completed with a different instrument. No pt harm, adverse outcome or injury was reported. | The affected instrument was not returned. Isi cannot confirm or deny the reported symptom. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 725203 | 737302 | 702540 | 2955842200600073 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-06-09 | 2006-07-07 | No | Yes | Health Professional | 40011008 | No | No | 2006-06-09 | Yes | No | Initial | 2006
40018301 | 2005-06-05 | Malfunction |  | It was reported that during the begining of the surgical procedure, there was electrical arcing from the joint of the permanent cautery hook instrument directly to the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed evidence of arcing, which appears to be localized at the instrument conductor wire. A kink in the conductor wire was found, which leads engineering concluded that instrument arcing most likely occurred due to the damaged conductor wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 725739 | 737851 | 703086 | 2955842200600068 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2005-06-05 | 2006-07-05 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-06-05 | Yes | No | Reuse | 2006
40018301 | 2005-06-05 | Malfunction |  | It was reported that during the begining of the surgical procedure, there was electrical arcing from the joint of the permanent cautery hook instrument directly to the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found evidence of arc tracking at the base of the instrument cautery hook, pulley and proximal clevis. Engineering also observed that the instrument ceramic sleeve was dislodged, indicating that the pulley-hook interface may have been damaged, thus allowing arcing to occur. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 725746 | 737858 | 703093 | 2955842200600069 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2005-06-05 | 2006-07-05 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-06-05 | Yes | No | Reuse | 2006
40018301 | 2006-05-31 | Malfunction | Other | It was reported that during the surgical procedure, the surgeon noticed some carbon buildup on the tip of the permanent cautery hook instrument. As a result, the surgeon used another instrument to scrape the residue off the tip and while doing so; the permanent cautery hook instrument tip fell inside the patient. The tip was retrieved and the planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation. Although the instrument was not returned for evaluation, the instructions for use specifically state: general precautions and warnings. Do not use an instrument to clean debris off another instrument intraoperatively. This may result in damage to the instrument. | PERMANENT CAUTERY HOOD INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 725765 | 737877 | 703112 | 2955842200600070 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-06-05 | 2006-07-05 | No | Yes | Health Professional | 40018301 | No | Yes | 2006-06-05 | Device Not Retu | No | Reuse | 2006
IS1200 A43P6 | 2006-06-05 | Other | Other | It was reported that during the surgical procedure, the customer was unable to focus the wide angle camera head. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system camera had malfunctioned. The system was repaired by replacing the affected camera. As of june 26, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 723900 | 736010 | 701256 | 2955842200600066 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-06-05 | 2006-06-29 | Yes | Yes | Health Professional | IS1200 A43P6 | No | Yes | 2006-06-05 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2006-05-31 | Other | Other | It was reported the customer experienced recurring 20008 and 21008 faults during the surgical procedure. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering found broken clamp screws on two of the axis gear assemblies, thus generating the system faults experienced by the customer. The system alarms (the 20008 and 21008 system error codes) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 723914 | 736024 | 701270 | 2955842200600065 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-05-31 | 2006-06-29 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-05-31 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2006-06-06 | Other | Other | It was reported that during the surgical procedure the customer experienced several recoverable system faults. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered two axis potentiometers to be out of specification, which consequently generated the system errors experienced by the customer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 26, 2006, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 725409 | 737508 | 702746 | 2955842200600067 | 1 | KOG | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2006-06-06 | 2006-06-29 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-06-06 | Device Not Retu | No | Reuse | 2006
IS1200 A43P6 | 2006-05-22 | Other | Other | It was reported that during the surgical procedure, the customer experienced multiple system error codes. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | H10: 423, 923, 47 - the investigation conducted by field service engineering concluded that the multiple system error codes experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering found an axis cable out of its pulley, which subsequently generated the system error codes experienced by the customer. The system alarm (system error codes) functioned as desigend and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 20, 2006, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 718565 | 730274 | 695565 | 2955842200600064 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-05-22 | 2006-06-21 | Yes | Yes | Health Professional | IS1200 A43P6 | No | No | 2006-05-22 | Device Not Retu | No | Reuse | 2006
40015403 | 2005-05-16 | Malfunction |  | It was reported that during the surgical procedure the permanent cautery spatula instrument was arcing. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | H10: 67 the instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot to determined. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 719033 | 730745 | 696038 | 2955842200600062 | 1 | GEI | Manufacturer | Health Professional | Other Health Care Professional | Initial | 2005-05-16 | 2006-06-15 | No | Yes | Health Professional | 40015403 | No | Yes | 2006-05-16 | Device Not Retu | No | Unkown | 2006
40018301 | 2005-05-16 | Malfunction |  | It was reported that during the surgical procedure a piece of the permanent cautery hook instrujent tip broke off and fell inside the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, no patient harm, adverse outcome or injury was reported. | H10: 67 the instrument was not returned for evaluation, therefore, the cause of the custome reported complaint cannot be determined. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 719085 | 730797 | 696092 | 2955842200600061 | 1 | GEI | Manufacturer | Health Professional | Other Health Care Professional | Initial | 2005-05-16 | 2006-06-15 | No | Yes | Health Professional | 40018301 | No | Yes | 2006-05-16 | Device Not Retu | No | Unkown | 2006
40017904 | 2006-05-16 | Malfunction |  | It was reported that during the surgical procedure the tip cover accessory of the monopolar curved scissors instrument fell inside the patient and was retrieved. Traces of orange paint from the instrument warning label were also seen inside the patient and were removed. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | H10: 400, 47, 66 the instrument was returned and evaluated. Engineering concluded that the tip cover accessory had slipped off of the distal end of the instrument due to a reduction in friction at the tube extension surface. Engineering also determined that between the tip cover and the tube extension during installation and removal of the tip cover accessory. The continous instalation and removal of the tip cover accessory eventually caused the ink to abrade, thus decreasing the friction between the tube extension and tip cover. The reduction in friction subsequently contributed to the tip cover accessory slipping. Biocompatibility testing was previously performed on a representative instrument sample containing the orange ink. Test results demonstrate the sample to be non-cytotoxic, non-hemolytic. Non-pyrogenic, and non-sensitizing. In addition, no acute toxicity or intracutaneous reactivity were identified. It was determined that patient health and safety were not compromised. This is a unusual event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 719086 | 730798 | 696093 | 2955842200600063 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-05-16 | 2006-06-15 | No | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | No | 2006-05-16 | Yes | No | Reuse | 2006
IS1200 A4 3P5 | 2006-05-12 | Other | Other | It was reported that during the surgical procedure, the customer experienced a recurring 20013 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the recurring system error experienced by the customer was associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered excessive jitter noise emanating from a potentiometer assembly within the extended psm, which consequently generated the system errors experienced by the customer. The system alarm (the 20013 system error code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 8, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 790480 | 802976 | 766605 | 2955842200600060 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-05-12 | 2006-06-09 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | Yes | 2006-05-12 | Device Not Retu | No | Reuse | 2006
42017903 | 2006-05-11 | Malfunction |  | It was reported that during the surgical procedure, a blade of the is2000 monopolar curved scissors instrument fell inside the pt. The blade was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Visual examination confirmed that one blade of the is2000 monopolar curve scissors instrument had separated from the instrument. The fragment was returned with the instrument. Engineering evaluation concluded that the failure mode experienced by the customer appears to be a result of a crack at the cross section of the instrument cable crimp. The customer reported failure mode is determined to be a unique and unusual event. | IS2000 MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 800473 | 812989 | 776335 | 2955842200600059 | 1 | GEI | Manufacturer | User facility | Nurse | Initial | 2006-05-11 | 2006-06-09 | No | Yes | Health Professional | 42017903 | Device Returned To Manufacturer | Yes |  | Yes | No | Reuse | 2006
40011709 | 2006-05-10 | Malfunction |  | It was reported that during the surgical procedure, one side of the jaw of the 5mm needle driver instrument broke while suturing. A broken piece fell inside of the pt and was retrieved. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument and broken jaw piece were returned and evaluated. Engineering evaluation found evidence of porosity within the instrument jaw. The porous condition allowed for contamination through the cross section of the jaw, which consequently weakened the jaw and compromised the jaw strength when high forces were applied. This is a unique event. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 800487 | 813003 | 776349 | 2955842200600058 | 1 | KOG | Manufacturer | User facility | Nurse | Initial | 2006-05-10 | 2006-06-09 | No | Yes | Health Professional | 40011709 | Device Returned To Manufacturer | Yes | 2006-05-10 | Yes | No | Reuse | 2006
400117 | 2006-05-06 | Malfunction |  | Robotic 5mm needle driver was being used to suture. The instrument was holding a 2/0 silk and being used to tie a knot when one side of the needle driver tip broke off. The broken needle driver was not holding the suture needle. The driver was removed from the patient's abdomen and taken off the system. The piece was easily visualized and retrieved. |  | DAVINCI ENDOWRIST | ROBOTIC SURGICAL INSTRUMENT NE | 715325 | 726884 | 692278 | 726884 | 1 | KOG | User Facility |  |  | Initial | 2006-06-08 | 2006-06-08 | No | Yes | Physician | 400117 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2006
40018401 | 2006-05-08 | Malfunction |  |  |  | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 713662 | 724782 | 690196 | 2955842200600056 | 1 | GEI | Manufacturer | Health Professional | RISK MANAGER | Initial | 2006-05-08 | 2006-06-07 | No | Yes | Health Professional | 40018401 | Device Returned To Manufacturer | Yes | 2006-05-08 | Yes | No | Reuse | 2006
ISI200 A4 3P5 | 2006-05-08 | Other | Other | It was reported that during system homing the customer experienced an irrecoverable 20008 system error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | H10: 47 - the investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering evaluation determined that a potentiometer of the mtmr had malfunctioned, thus causing the system error experienced by the customer. The system alarms functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of june 6, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 713666 | 724784 | 690198 | 2955842200600055 | 1 | KOG | Manufacturer | Company Representative |  | Initial | 2006-05-08 | 2006-06-07 | Yes | Yes | Health Professional | ISI200 A4 3P5 | No | No | 2006-05-08 | Device Not Retu | No | Reuse | 2006
ISI200 A4 3P5 | 2006-05-05 | Other | Other | During the surgical procedure the customer experienced intermittent 212 system error codes. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The intermittent 212 system error that occurred during the procedure is a recoverable error. An investigation was conducted by a field service engineer (fse) and the fse was unable to replicate the customer reported complaint. The fse replaced the patient side manipulator (psm) and msd pca as a precaution. The psm and msd pca were returned to isi for further failure analysis investigation. Visual inspection and function testing were performed on both units. The engineering department was also unable to replicate the issue experienced by the customer. The intermittent 212 system error may have been caused by oxidation on the system cable connectors. It is possible that the manipulation of system cables performed by the fse while exchanging components, may have removed any problematic oxidation buildup from the connectors. As of june 5, 2006, there have been no reported recurrence of the reported issue as this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 714449 | 726187 | 691586 | 2955842200600054 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-05-05 | 2006-06-05 | Yes | Yes | Health Professional | ISI200 A4 3P5 | No | No | 2006-05-05 | Device Not Retu | No | Reuse | 2006
IS1200 A43P4 | 2006-05-03 | Other | Other | During the surgical procedure the system experienced intermittent faults. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system errors experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Although engineering was unable to replicate the customer reported complaint, a visual inspection revealed that the potentiometer grease was contaminated. The contamination may have prevented proper electrical contact between the rotor and potentiometr resistive element. The system alarms functioned as designed and there was no injury to the pt. The procedure was converted to open surgical techniques to complete the planned procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 712374 | 723296 | 688724 | 2955842200600053 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-05-03 | 2006-06-02 | Yes | Yes | Health Professional | IS1200 A43P4 | No | No | 2006-05-03 | Yes | No | Reuse | 2006
400169 | 2006-02-10 | Malfunction |  | During the surgical procedure the tip of the harmonic curved shears insert broke off inside the patient and was retrieved. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 683766 | 694637 | 661041 | 2955842200600031 | 1 | GDW | Manufacturer | User facility | Nurse | Initial | 2006-02-27 | 2006-05-29 | No | Yes | Health Professional | 400169 | No | Yes | 2006-02-27 | Device Not Retu | Yes | Initial | 2006
40018301 | 2005-04-26 | Malfunction |  | The permanent cautery hook instrument has a burnt insulator, however, no electrical arcing during the surgical procedure was observed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found signs of localized arc tracking at the ultem pulley. Engineering also observed that part of the instrument pulley and clevis has been charred, thus damaging the conductor wire. Based on the evidence of arc tracking found, it was concluded that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 710492 | 721424 | 686896 | 2955842200600049 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2005-04-26 | 2006-05-26 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-04-26 | Yes | No | Reuse | 2006
40012608 | 2005-04-26 | Malfunction |  | There was electricity dispersion on the tip of permanent cautery hook instrument, however, no electrical arcing during the surgical procedure was observed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the electrical conductor going into the instrument hook had been damaged. Signs of an arc track at the ultem pulley were discovered, which appear to be localized at the conductor wire. Engineering determined that the damage to the electrical conductor was due to arcing, therefore, it has been concluded that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 710497 | 721429 | 686901 | 2955842200600048 | 1 | GEY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2005-04-26 | 2006-05-26 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2006-04-26 | Yes | No | Reuse | 2006
40018301 | 2005-04-26 | Malfunction |  | During the surgical procedure it was noticed that the tip of the permanent cautery hook instrument was burnt. No electrical arcing during the surgical procedure was observed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found area of arc tracking at the instrument distal and proximal clevises. The charring appears to be localized at the distal clevis and pulley, however, there are small arc tracks at the proximal pin at the proximal clevis. Engineering also observed that the instrument electrical conductor was damaged. Based on the evidence of arc tracking found, it was concluded that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 710503 | 721435 | 686906 | 2955842200600050 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2005-04-26 | 2006-05-26 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-04-26 | Yes | No | Reuse | 2006
42009303 | 2006-04-28 | Malfunction |  | Towards the end of the surgical procedure a crack in the wrist of the is2000 endowrist prograsp instrument was noticed and pieces of the instrument wrist fell inside the pt. The instrument had been in use for a few minutes when the incident occured. It was reported the all fragmented pieces were retrieved with a laparoscopic grasper. The customer had difficulty removing the instrument from the cannula and as a result, had to break the instrument wrist to remove it. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the instrument wrist fracture cannot be determined. As indicated in the customer reported complaint, all fragmented pieces were retrieved with a laparoscopic grasper. | IS2000 ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 710515 | 721447 | 686918 | 2955842200600051 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-04-28 | 2006-05-26 | No | Yes | Health Professional | 42009303 | No | No | 2006-04-28 | Device Not Retu | No | Unkown | 2006
IS1200 A4 3P5 | 2006-04-19 | Other | Other | During the surgical procedure the surgeon experienced tool insertion/recognition issues. The surgical staff suggested trying a different sterile adapter and instrument, however, the surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery. No pt harm reported. | The system was evaluated by an isi field service engineer (fse). The isi fse performed a complete preventative maintenance of the system and found no problems. As of march 2006 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 710521 | 721453 | 686924 | 2955842200600052 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-04-20 | 2006-05-26 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | No | 2006-04-20 | Yes | No | Reuse | 2006
40017904 | 2006-04-18 | Malfunction |  | Towards the end of the surgical procedure arcing around the tip of the monopolar curved scissors instrument was observed. A hole in the instrument tip cover was also noticed. The instrument and tip cover accessory wereremoved and replaced and the planned surgical procedure was completed. The instrument was sued for over an hour prior to the incident. Cautery burns on the right lateral wall of the pt's pelvis were reported, bu no pt harm, adverse outcome or injury was reported. | The instrument and tip cover accessory were not returned for evaluation, therefore, no conclusion can be drawn regarding the cause of the instrument arcing and hole in the monopolar curved scissors instrument tip over. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 707359 | 718290 | 683852 | 2955842200600047 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-04-18 | 2006-05-18 | No | Yes | Health Professional | 40017904 | No | No | 2006-04-18 | Device Not Retu | No | Unkown | 2006
400179XX | 2006-04-18 | Malfunction |  | Two hours into the surgical procedure, the surgeon noticed an arc of energy near the bowel area of the pt while using the monopolar cirved scissors instrument. The instrument was immediately removed and upon removal, a hole in the instrument tip cover was observed. The instrument tip cover weas removed and replaced and the same instrument was used to complete the planned surgical procedure. The surgeon indicated that it appeared that the hole in the instrument tip cover was caused by the tip of another instrument rubbing against it. An arching mark near the pt's bowel was reported. But no pt harem, adverse outcome or injury was rpeorted. | The instrument and tip cover accessory were not returned for evaluation. Per the customer reported complaint, the surgeon indicated that it appeared that the hole in the instrument tip cover was caused by the tip of another instrument rubbing against it, however, this theory cannot be confirmed or denied. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 707365 | 718297 | 683858 | 2955842200600046 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-04-18 | 2006-05-18 | No | Yes | Health Professional | 400179XX | No | No | 2006-04-18 | Device Not Retu | Yes | Unkown | 2006
IS1200 A43P5 | 2006-04-17 | Other | Other | During the surgical preocedure, the right arm of the da vinvi surgical system was not recognizing instruments. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the instrument recognition issue experienced by the customer was associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering evaluation found the extended psm to be working properly and no issues were found. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 709165 | 720096 | 685604 | 2955842200600045 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2006-04-17 | 2006-05-17 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-04-17 | Device Not Retu | No | Reuse | 2006
42009303 | 2006-04-03 | Malfunction |  | During the surgical procedure the shaft of the isi2000 endowrist prograsp instrument broke while the doctor was "making a turn with the instrument". A broken piece of the instrument fell inside the patient and was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument and grip were returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis had broken at the main tube interface. Engineering determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. As stated in the complaint report, the broken instrument piece was retrieved. | 1S2000 ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 709177 | 720108 | 685616 | 2955842200600044 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-05-10 | 2006-05-17 | No | Yes | Health Professional | 42009303 | Device Returned To Manufacturer | Yes | 2006-05-10 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2006-04-20 | Other | Other | During the surgical procedure the customer experienced a 20008 system error code which they could not recover from. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered that the axis 6 clamp screw had broken off from the hub clamp, thus generating the system fault experienced by the customer. The system alarm (the 20008 system error code) functioned as designed and there was no injury to the patient. The procedure was converted to open surgical techniques to complete the planned procedure. As of may 12, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 704181 | 715115 | 680744 | 2955842200600043 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-04-20 | 2006-05-12 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-04-20 | Device Not Retu | No | Reuse | 2006
40011006 | 2006-03-22 | Malfunction |  | During the surgical procedure the black isolation part of the precise bipolar pyramid tip forceps instrument broke into two pieces and fell inside the pt. All components were retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. Per the customer reported complaint, all components were retrieved from the pt. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 704241 | 715175 | 680801 | 2955842200600042 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2006-04-14 | 2006-05-12 | No | Yes | Health Professional | 40011006 | No | Yes | 2006-04-14 | Device Not Retu | No | Unkown | 2006
40000619 | 2006-05-10 | Malfunction |  | It was reported that during the surgical procedure, the "pad" of the large needle driver instrument fell inside of the pt. The pad was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned with one of the carbide tips (which the customer referred to as the "pad" of the instrument) missing. As indicated in the complaint notes, the carbide tip was retrieved from the pt. Engineering evaluation observed trace amounts of residual brazing on the instrument grip surface. This observation leads engineering to conclude that the grip surface was not sufficiently cleaned prior to the brazing being applied. This contamination may have caused a weakened braze, thus allowing the carbide tip to fall off. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 800495 | 813011 | 776357 | 2955842200600057 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2006-05-10 | 2006-05-10 | No | Yes | Health Professional | 40000619 | Device Returned To Manufacturer | No | 2006-05-10 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2006-04-04 | Other | Other | It was reported that at the end of the surgical procedure, the system generated a 20008 error code which the customer was unable to override. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system issues experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering discovered that the axis 3 pot screw had broken off from the hub clamp, which prevented the pot from moving through it's full range of motion, thus causing the system fault experienced by the customer. As of may 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 698765 | 709721 | 675459 | 2955842200600041 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-04-04 | 2006-05-04 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-04-04 | Device Not Retu | No | Reuse | 2006
40017205 | 2006-03-27 | Malfunction | Other | During the beginning of the surgical procedure the tip of the bipolar maryland tip forceps instrument broke off and fell inside the pt. The broken tip was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument and detached tip were returned and evaluated. Per the customer reported complaint, engineering observed that all of the cables and both conductor wires were broken at the distal end of the hypotubes. Engineering concluded that the cable breakages may have been a result of material degradation. It has been theorized that a site specific processing issue may have precipitated this failure. As indicated in the complaint notes, the broken tip was retrieved. No other damage to the instrument was observed. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 696359 | 707293 | 673208 | 2955842200600039 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-03-28 | 2006-05-01 | No | Yes | Health Professional | 40017205 | Device Returned To Manufacturer | Yes | 2006-03-28 | Yes | No | Reuse | 2006
40011008 | 2006-03-14 | Malfunction |  | It was reported that during the surgical procedure it was noticed that the shaft of the precise bipolar pyramid tip forceps instrument was cracked. The customer was unable to straighten the instrument tip to remove it from the cannula, therefore, they severed the tip to enable removal. The severed tip fell into the pt and was immediately retrieved and removed. No pt harm, adverse outcome or injury was reported. | The instrument and tip were returned and evaluated. Per the customer reported complaint, engineering observed that the instrument proximal clevis was broken and there was damage to the main tube interface. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. As indicated in the complaint notes, the severed tip of the instrument was retrieved. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 696701 | 707649 | 673541 | 2955842200600040 | 1 | GEI | Manufacturer | User facility | Physician | Initial | 2006-03-31 | 2006-04-28 | No | Yes | Health Professional | 40011008 | Device Returned To Manufacturer | Yes | 2006-03-31 | Yes | No | Reuse | 2006
IS2000 A5 0P2 | 2006-04-04 | Other | Other | During the surgical procedure the customer experienced 25504 error codes on the patient side manipulator. The system was rebooted and upon start-up, the system generated a 26004 error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system error codes experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering evaluation determined that the axis 1 encoder of the psm had malfunctioned, thus causing the system error codes experienced by the customer. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 692465 | 703344 | 669416 | 2955842200600038 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-04-04 | 2006-04-19 | Yes | Yes | Health Professional | IS2000 A5 0P2 | No | No | 2006-04-04 | Device Not Retu | No | Reuse | 2006
40017203 | 2006-01-11 | Malfunction | Other | During the surgical procedure the customer experienced difficulty removing the bipolar tip forceps instrument. A visual inspection of the insturment after removal did not show any abnormalities. The instrument was reinserted into the pt and the tip of the instrument broke while grasping tissue. Upon breakage, instrument components fell into the pt. The components were retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 692949 | 703848 | 669884 | 2955842200600037 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-03-20 | 2006-04-19 | No | Yes | Health Professional | 40017203 | No | Yes | 2006-03-20 | Device Not Retu | No | Unkown | 2006
40018401 | 2006-03-15 | Malfunction |  | During the surgical procedure, the permanent cautery spatula instrument was arcing. No further details were provided. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 690530 | 701420 | 667589 | 2955842200600036 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-03-15 | 2006-04-14 | No | Yes | Health Professional | 40018401 | No | Yes | 2006-03-15 | Device Not Retu | No | Unkown | 2006
40016906 | 2006-03-13 | Malfunction |  | It was reported, that during the surgical procedure, when attempting to remove the harmonic curved shears insert from the trocar, the insert broke and a piece of the insert fell inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 689786 | 700698 | 666891 | 2955842200600035 | 1 | LFL | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-03-13 | 2006-04-12 | No | Yes | Health Professional | 40016906 | No | Yes | 2006-03-13 | Device Not Retu | Yes | Initial | 2006
IS1200 A42 | 2006-03-09 | Other | Other | During the surgical procedure the right pt side manipulator was not accepting instruments. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the instrument engagement and movement issues experienced by the customer were associated with the patient side manipulator (psm). The system was repaired by replacing the affected extended psm. As of april 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 687432 | 698343 | 664596 | 2955842200600034 | 1 | KOG | Manufacturer | User facility | Physician | Initial | 2006-03-09 | 2006-04-07 | Yes | Yes | Health Professional | IS1200 A42 | No | Yes | 2006-03-09 | Device Not Retu | No | Reuse | 2006
5000073 | 2006-03-01 | Malfunction |  | During the surgical procedure, the arm of the aesop system progressively lowered until it came to rest on the pt's rib cage. The endoscope was not inserted when the event occured. No serious injury or adverse outcome was reported. | H10: the aesop system was evaluated by an isi field service engineer (fse). The fse was unable to reproduce the customer reported complaint. During evalution, the fse was informed that the surgeon did not set the lower limit when starting the aesop system. The aesop user's guide specifically states: set the lower limit. As a safety feature, the aesop system requires that you set the lowest point the positioning arm shall go for each case. Warning: disregarding this information may prove hazardous to the safety of a healthcare professional or pt. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 686129 | 697020 | 663337 | 2955842200600033 | 1 | GCJ | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2006-03-02 | 2006-04-03 | No | Yes | Health Professional | 5000073 | No | Yes | 2006-03-02 | Device Not Retu | No | Reuse | 2006
40009305 | 2006-02-27 | Malfunction |  | During the surgical procedure the endowrist prograsp instrument was not working properly. At the end of the surgical procedure it was noticed that a piece of the proximal clevis was missing. No pt harm, adverse outcome or injury was reported. The customer indicated that they would not be returning the instrument to isi for evaluation. | Although the instrument was not returned for evaluation, an isi clinical sales rep was able to obtain photographs of the affected instrument. Based on a review of the photographs, engineering determined that the instrument issues experienced by the customer were a result of damage to the proximal clevis of the instrument. The location of the missing portion of the proximal clevis cannot be determined. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 684442 | 695310 | 661693 | 2955842200600032 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-02-27 | 2006-03-31 | No | Yes | Health Professional | 40009305 | No | No | 2006-02-27 | Device Not Retu | No | Unkown | 2006
IS1200 A43P5 | 2006-03-01 | Other | Other | The customer experienced 20008 faults when trying to home the system. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system faults experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering found a broken clamp screw on one of the axis gear assemblies, which caused the system faults experienced by the customer. As of march 27, 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 683805 | 694676 | 661079 | 2955842200600030 | 1 | GDW | Manufacturer | Company Representative | Other | Initial | 2006-03-01 | 2006-03-29 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-03-01 | Device Not Retu | No | Reuse | 2006
40009305 | 2006-02-10 | Malfunction |  | During the surgical procedure the tip of the endowrist prograsp instrument seperated from the shaft and the instrument was not inside the pt at the time the malfunction occurred. On march 15, 2006 additional information was received which reported that when the distal shaft of the instrument broke, it shifted the instrument and the surgeon was unable to remove the instrument from the cannula. To remove the instrument, the surgeon cut the instrument wires and upon doing so the tip of the instrument fell inside the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering determined that the instrument malfunction experienced by the customer was a result of damage to the proximal clevis of the instrument. | ENDOWRIST PROGRASP INTRUMENT | ENDOSCOPIC INSTRUMENT | 684032 | 694903 | 661295 | 2955842200600029 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-02-10 | 2006-03-29 | No | Yes | Invalid Data | 40009305 | Device Returned To Manufacturer | No | 2006-03-15 | Yes | No | Reuse | 2006
| 2006-02-27 | Malfunction |  | A patient was in the or undergoing robotic gyn surgery (laparoscopic hysterectomy) when the tip of the endowrist prograsp broke while in the patient. The instrument was removed and examined by the physician as was the patient. It was noted that a small plastic piece was missing, however, it could not be determined if the missing piece was in the patient and it could not be detected by xray. |  | ENDOWRIST PROGRASP 8MM | GRASPER | 684011 | 694882 | 661275 | 694882 | 1 | KOG | User Facility |  | RISK MANAGER | Initial | 2006-03-27 | 2006-03-27 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2006
IS1200 A43P5 | 2006-02-24 | Other | Other | It was reported that during the surgical procedure the system would home but fail as soon as a sterile adapter was installed. Upon system failure, the system would generate a 20008 error code. No pt harm was reported. The procedure was converted to tradational open surgical techniques to finish the planned surgery. | The investigation conducted by field svc engineering concluded that the system issues experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering found the extended psm carriage pot assemblies to be damaged, thus causing the system to fault during sterile adapter engagement. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 682577 | 693449 | 659893 | 2955842200600027 | 1 | NAY | Manufacturer | Company Representative | Other | Initial |  | 2006-03-24 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-02-24 | Device Not Retu | No | Reuse | 2006
40000619 | 2006-02-24 | Malfunction |  | It was reported that during the surgical procedure fibers from the large needle driver instrument fell inside of the pt. All of the fibers were retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument grip cable had broken and frayed at the distal idlers. The fibers that the customer was referring to are from the tungstent cable. As stated in the complaint notes, all of the fibers were retrieved from the pt. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 682579 | 693451 | 659895 | 2955842200600028 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-24 | 2006-03-24 | No | Yes | Health Professional | 40000619 | Device Returned To Manufacturer | Yes | 2006-02-24 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2006-02-21 | Other | Other | During the surgical procedure the customer experienced a recurring 20008 error code. No pt harm was rpeorted. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system was evaluated. A review of the system error logs was performed and engineering found indication that the iod to ria cable was not connected, which caused the recurring error experienced by the customer. The system alarm (i. E. Error code and audible alarm) functioned as designed. User error during system setup directly caused the event. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 681989 | 692833 | 659331 | 2955842200600025 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-02-21 | 2006-03-23 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-02-21 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2006-02-20 | Other | Other | During the surgical procedure the customer experienced multiple systems errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system errors experienced by the customer were associated with the left master tool manipulator (mtml). The sytem was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Engineering evaluation determined that a potentiometer of the mtml had malfunctioned, thus causing the system errors experienced by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 682079 | 692933 | 659419 | 2955842200600024 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-02-21 | 2006-03-23 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2006-02-21 | Device Not Retu | No | Reuse | 2006
IS1200 A43P4 | 2006-03-13 | Other | Other | During the surgical procedure the customer experienced a recurring 20008 system error code. During the evaluation performed by the hospital biomedical engineer, a broken clamp screw was found on the extended patient side manipulator. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The hospital biomedical engineer repaired the system by replacing the affected extended patient side manipulator (psm). The extended psm was returned to isi for failure analysis investigation and per the reported complaint, engineering confirmed that the fastener on the carriage hub clamp had broken, which caused the system errors experienced by the customer. As of 03/ 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 682979 | 693852 | 660284 | 2955842200600026 | 1 | KOG | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2006-03-13 | 2006-03-23 | Yes | Yes | Health Professional | IS1200 A43P4 | No | No | 2006-03-13 | Device Not Retu | No | Reuse | 2006
40016906 | 2006-02-16 | Malfunction |  | During the surgical procedure the harmonic shears insert broke and a piece of the insert fell into the pt. The broken piece was retrieved adn the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 679731 | 690583 | 657170 | 2955842200600021 | 1 | LFL | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-16 | 2006-03-17 | No | Yes | Health Professional | 40016906 | No | Yes | 2006-02-16 | Device Not Retu | Yes | Initial | 2006
40000319 | 2006-02-17 | Malfunction | Other | During the surgical procedure, the clips applied to the pt vessel were approximating, yet appeared to be closed. Nevertheless, the surgeons continued with the planned surgical procedure. After the robotic portion of the procedure, a clip re-opened or slipped off the vessel and the pt began bleeding. An unplanned non-robotic surgical procedure was performed to control the pt bleeding. Two clip applier instruments were used during the original planned surgical procedure. The model number, lot number, and failure analysis results of the other clip applier instrument used can be found in mdr report #22955842-2005-00023. | H10: the instrument was returned and evaluated. Engineering evaluation found no instrument damage passed all performance tests. Engineering also reviewed a video of the surgical procedure. Engineering analysis of the procedure video observed that there appeared to be an instance where the clip was not properly compressed. The clip takes on a diamond shape, as opposed to being fully closed. The clip did not appear to have any deformation prior to the clipping. In another instance, two clips were used to stop a slight hemorrhage, however, one of the clips appears to not have clipped entirely and this particular clip detached and the bleeding cont'd. No problem could be found with the instrument and its functionality. | CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 679746 | 690598 | 657184 | 2955842200600022 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-17 | 2006-03-17 | Yes | Yes | Health Professional | 40000319 | Device Returned To Manufacturer | Yes | 2006-02-17 | Yes | No | Reuse | 2006
40000319 | 2006-02-17 | Malfunction | Other | During the surgical procedure, the clips applied to the pt vessel weren't approximating, yet appeared to be closed. Nevertheless, the surgeons cont'd with the planned surgical procedure. After the robotic portion of the surgical procedure, a clip re-opened or slipped off the vessel and the pt began bleeding. An unplanned non-robotic surgical procedure was performed to control the pt bleeding. Two clip applier instruments were used during the original planned surgical procedure. The model number, the lot number, and failure analysis results of the othe clip applier instrument used can be found in mdr report #2955842-2005-00022. | H10: the instrument was returned and evaluated. Engineering evaluation found no instrument damage or abnormalities and the instrument passed all performance test. Engineering also reviewed a video of the surgical procedure. Engineering analysis of the procedure video observed that there appeared to be an instance where the clip was not properly compressed. The clip takes on a diamond shape, as opposed to being fully closed. The clip did not appear to any deformation prior to the clipping. In another instance, two clips were used to stop a slight hemmorhage, however, one of the clips appears not to have clipped entirely and this paticular clip detached and the bleeding cont'd. No problem could be found with the instrument and its functionality. | CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 679768 | 690620 | 657205 | 2955842200600023 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-17 | 2006-03-17 | Yes | Yes | Health Professional | 40000319 | Device Returned To Manufacturer | Yes | 2006-02-17 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2006-02-13 | Other | Other | During the surgical procedure the customer was unable to use the right arm of the system after the tool was inserted. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Field service engineering tested the system's extended patient side manipulator (psm), which the customer referred to as the "right arm", and was unable to reproduce the customer reported complaint. The field service engineer decided to replace the affected extended psm as a precaution. The extended psm was returned to isi for failure analysis investigation. Engineering performed functional tests on the extended psm and was also unable to replicate the customer reported symptom. As of march 2006, there have been no reported recurrences of the tissue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 678807 | 689630 | 656267 | 2955842200600020 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-13 | 2006-03-15 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-02-13 | Device Not Retu | No | Reuse | 2006
IS1200 A4 3P5 | 2006-02-06 | Other | Other | It was reported that during the beginning of the surgical procedure the customer experienced nonrecoverable 20000 and 21000 internal error codes. No patient harm was reported. The patient incisions were closed and the surgicla procedure will be performed at a later date. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the left master tool manipulator (mtml). The system was also repaired by replacing the affected mtml. The mtml was returned to isi for failure anslysis investigation. Engineering discovered that the axis-6 motor gear had become loose and disengaged due to a broken clamp screw, thus causing the mtml to malfunction. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 674995 | 685776 | 652573 | 2955842200600018 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2006-02-06 | 2006-03-08 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | No | 2006-02-06 | Device Not Retu | No | Reuse | 2006
40016906 | 2006-01-23 | Malfunction |  | It was reported that towards the end of the surgical procedure the "blade" of the harmonic shears insert fell into the pt. The "blade" was retrieved and the planned surgical procedure was completed using a replacement instrument and insert. The "blade" was disposed of by the hospital. No pt harm, adverse outcome or injury was reported. | There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. Because the harmonic curved shears insert is not being returned for eval, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 2821155 | 2783739 | 2680335 | 2955842200500128 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-06 | 2006-03-08 | No | Yes | Health Professional | 40016906 | No | Yes | 2006-02-06 | Device Not Retu | Yes | Initial | 2006
40011008 | 2006-02-03 | Malfunction |  | During the surgical procedure the black plastic piece which holds the jaws together of the precise bipolar pyramid tip forceps instrument fell inside the pt. The procedure was completed laparoscopically, however, the fragmented component was not retrieved. No pt harm, adverse outcome or injury was reported. The customer indicated that they would not be returning the instrument to isi for evaluation. | Although the instrument was not returned for evaluation, an isi account manager was able to obtain photographs of the affected instrument and a video of the surgical procedure. Based on a review of the photographs and procedure video, engineering determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 672953 | 683705 | 650574 | 2955842200600015 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-02-03 | 2006-03-03 | No | Yes | Health Professional | 40011008 | Device Returned To Manufacturer | Yes | 2006-02-03 | Device Not Retu | No | Reuse | 2006
40016906 | 2006-02-03 | Malfunction |  | It was reported that during the surgical procedure the harmonic shears insert inside of the pt. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 680512 | 691364 | 657918 | 2955842200600017 | 1 | LFL | Manufacturer | Company Representative | Other | Initial | 2006-02-03 | 2006-03-03 | No | Yes | Health Professional | 40016906 | No | No | 2006-02-03 | Device Not Retu | Yes | Initial | 2006
40016906 | 2006-01-27 | Malfunction |  | It was reported that during the surgical procedure the harmonic shears insert fell inside of the pt. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The harmonic curved shears insert was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 2832162 | 2783756 | 2680352 | 2955842200500040 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2006-02-03 | 2006-03-03 | No | Yes | Health Professional | 40016906 | No | No | 2006-02-03 | Device Not Retu | Yes | Initial | 2006
IS1200 A43P5 | 2006-01-30 | Other | Other | It was reported that during the surgical procedure the right patient side manipulator was not accepting instruments. The customer also experienced wrist and grip movement problems when using multiple instruments. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the instrument engagement and movement issues experienced by the customer were associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering discovered a broken cannula piece stuck inside the elbow bowl wirings, which most likely prevented the instruments from properly engaging on the extended psm. As of march 2006, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 671073 | 681803 | 648704 | 2955842200600014 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-01-31 | 2006-03-01 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-01-31 | Device Not Retu | No | Reuse | 2006
IS1200 A4 3P4 | 2006-01-30 | Other | Other | None | The investigation conducted by field service engineering concluded that the system camera had malfunctioned. The system was repaired by replacing the affected camera. As of 2/06, there have been no reported recurrences of the issue at this hosp. | DE VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 671079 | 681809 | 648710 | 2955842200600013 | 1 | KOG | Manufacturer | Health Professional | BIOMEDICAL ENGINEER | Initial | 2006-01-30 | 2006-03-01 | Yes | Yes | Health Professional | IS1200 A4 3P4 | No | No | 2006-01-30 | Device Not Retu | No | Reuse | 2006
37094601 | 2006-01-26 | Other | Other | During the surgical procedure, the instrument twist-lock cannula mount (a non-patient contact instrument accessory) broke in half. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery since the hospital did not have a backup instrument cannula mount. | The instrument twist-lock cannula mount was not returned for evaluation, therefore, the cause of the reported complaint cannot be determined. | INSTRUMENT TWISTLOCK CANNULA MOUNT | INSTRUMENT ACCESSORY | 668665 | 679397 | 646392 | 2955842200600011 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2006-01-26 | 2006-02-24 | Yes | Yes | Health Professional | 37094601 | No | No | 2006-01-26 | Device Not Retu | No | Unkown | 2006
40000417 | 2006-01-25 | Malfunction |  | During the surgical procedure the scalpel cautery instrument tip broke and became detached. No instrument components fell into the patient and planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the customer reported complaint cannot be determined. This instrument was manufactured in november 2001. The affected instrument model is no longer distributed by intuitive surgical, inc. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 670574 | 681308 | 648245 | 2955842200600010 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2006-01-25 | 2006-02-24 | No | Yes | Health Professional | 40000417 | No | No | 2006-01-25 | Device Not Retu | No | Unkown | 2006
40009305 | 2006-01-25 | Malfunction |  | During the middle of the surgical procedure the surgeon attempted to remove the enodwrist prograsp instrumenbt from the patient, however, the instrument would not fit through the trocar because the pin holding the jaws of the instrument and dislodged, leaving the jaws to hang free in the patient. The surgeon was able to push the pin back into place and the instrument was successfully removed from the patient. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation determined that the grip axis pin was improperly swaged, thus allowing the pin to become loose and subsequently causing the instrument to lose its functionality. As indicated in the complaint notes, the surgeon was able to push the pin back into place and the instrument was successfully removed from the patient. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 670704 | 681437 | 648367 | 2955842200600012 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-01-27 | 2006-02-24 | No | Yes | Invalid Data | 40009305 | Device Returned To Manufacturer | Yes | 2006-01-27 | Yes | No | Initial | 2006
40009305 | 2006-01-20 | Malfunction |  | During the middle of the surgical procedure, the pin of the prograsp endowrist instrument came out of the instrument jaws and fell into the patient. The surgeon was able to retrieve the pin. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the grip axis pin missing. Engineering evaluation of the instrument determined that the grip axis pin was improperly swaged, thus allowing the pin to separate and dislodge from the instrument. As indicated in the complaint notes, the surgeon was able to retrieve the pin from the patient. | PROGRASP ENDOWRIST INSTRUMENT | ENDOSCOPIC INSTRUMENT | 675123 | 685912 | 652689 | 2955842200600009 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-01-20 | 2006-02-24 | No | Yes | Invalid Data | 40009305 | Device Returned To Manufacturer | No | 2006-01-20 | Yes | No | Reuse | 2006
40011007 | 2006-01-17 | Malfunction |  | During the surgical procedure, the tip of the precise bipolar pyramid tip forceps instrument broke and fell into the patient. The tip was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument proximal clevis had broken at the main tube interface. As indicated in the complaint notes, the fragmented component of the instrument was retrieved. No other damage to the instrument was observed. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 668448 | 679183 | 646184 | 2955842200600008 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial |  | 2006-02-16 | No | Yes | Health Professional | 40011007 | Device Returned To Manufacturer | Yes | 2006-01-17 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2006-01-16 | Other | Other | It was reported that during the surgical procedure the right eye of the high resolution stereo viewer froze. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was associagted with the system synchronizer. The system was repaired by replacing the affected synchronizer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 666903 | 677633 | 644682 | 2955842200600006 | 1 | NAY | Manufacturer | User facility | Physician | Initial | 2006-01-16 | 2006-02-15 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-01-16 | Device Not Retu | No | Reuse | 2006
PRECISE BIPOLAR | 2006-02-03 | Injury | Hospitalization | A laparoscopic hysterectomy was being performed at the hosp in 2006. The equipment being used was an intuitive robot and a part# 400110 pricise bipolar accessory. The instrument had been in use for 2 hours and was being used to cauterize the uterine arteries when the device failed. The bipolar accessory broke at the pincher tip. The device was removed and the hysterectomy was completed vaginally. |  | PRECISE BIPOLAR | BIOPOLAR FORCEP | 671867 | 682599 | 649484 | MW1038017 | 1 | GEI | Voluntary |  | BIOMEDICAL ENGINEER | Initial | 2006-02-13 | 2006-02-13 | Yes | Yes | Health Professional | PRECISE BIPOLAR | Yes | Yes |  |  |  |  | 2006
IS1200 A43P5 | 2006-01-13 | Other | Other | It was reported that during the surgical procedure the surgeon experienced poor/blurry vision. No patient harm was reported and the planned surgical procedure was aborted. | The investigation conducted by field service engineering concluded that the system camera cable had frayed internally, thus causing the poor/blurry vision experienced by the customer. The system was repaired by replacing the affected camera. As of 2006, there have been no reported recurrence of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 663511 | 674232 | 641401 | 2955842200600004 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-01-13 | 2006-02-10 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-01-13 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2006-01-23 | Other | Other | It was reported that during the surgical procedure the customer experienced a recurring 20008 error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering evaluation determined that the pot element of the extended psm hadmalfunctioned, thus causing the 20008 errors experienced by the customer. As of 2006, there have been no reported recurrence of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 663518 | 674239 | 641407 | 2955842200600005 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-01-23 | 2006-02-10 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2006-01-23 | Device Not Retu | No | Reuse | 2006
40018301 | 2005-11-17 | Malfunction |  | There was a problem with the insulation of the permanent cautery hook instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the electrical conductor going into the instrument hook had been damaged due to some localized charring at the base of the hook assembly. The damage had caused the instrument to lose its cautery functions. Engineering determined that the localized charring was due to minor arcing, therefore, it has been concluded that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 665190 | 675907 | 643061 | 2955842200600003 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-11-17 | 2006-02-10 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-02-02 | Yes | No | Reuse | 2006
40011008 | 2006-01-06 | Malfunction | Other | During the middle of the surgical procedure, the precise bipolar pyramid tip forceps instrument became stuck to the cannula. When attempting to remove the instrument out of the cannula, a small piece of the instrument fell into the pt and another small piece fell into the cannula. The surgeon retrieved both instrument pieces and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument proximal clevis had broken and fragmented at the main tube interface. As indicated in the complaint notes, the fragmented components of the instrument were retrieved. No other damage to the instrument was observed. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 661976 | 672653 | 639917 | 2955842200600001 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2006-01-06 | 2006-02-03 | No | Yes | Health Professional | 40011008 | Device Returned To Manufacturer | Yes | 2006-01-06 | Yes | No | Reuse | 2006
40018301 | 2006-01-06 | Malfunction |  | During the surgical procedure, the permanent cautery hook instrument was arcing. The planned surgical procedure was completed and no further details were provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument passed electrical testing and engineering found no signs of electrical arcing during visual inspection. The customer reported complaint cannot be confirmed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 661982 | 672659 | 639923 | 2955842200600002 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2006-01-06 | 2006-02-03 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2006-01-06 | Yes | No | Reuse | 2006
40018301 | 2005-11-22 | Malfunction |  | It was reported that the permanent cautery hook instrument no longer functioned. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the electrical conductor had charred at the base of the instrument hook, which led to electrical continuity failure, thus causing the loss of functionality report by the customer. Engineering determined that the charring was due to localized arcing, therefore it has been concluded that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 659201 | 669876 | 637243 | 2955842200500126 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-11-22 | 2006-01-26 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2006-01-20 | Yes | No | Reuse | 2006
IS1200 A43P5 | 2005-12-27 | Other | Other | It was reported that during the surgical procedure the customer experienced a 20002 internal error code. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering observed poor soldering on the hall effect sensor which led to an intermittent loss of connection, thus causing the mtmr to malfunction. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INTRUMENT CONTROL S | 659247 | 669922 | 637289 | 2955842200500127 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-12-27 | 2006-01-26 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2005-12-27 | Device Not Retu | No | Reuse | 2006
IS1200 A43P5 | 2005-12-21 | Other | Other | It was reported that during the beginning of the surgical procedure the surgeon experienced a "moving prism on one of the eye" of the system. No pt harm was reported ahd the planned surgical procedure was cancelled. | The investigation conducted by field service engineer concluded that the system endoscope had been damaged, therefore causing the "moving prism" issue experienced by the customer. The system was repaired by replacing the affected endoscopic. The vision issue with the endoscope should have been discovered by the hospital during equipment set-up, prior to the start of the surgical procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 655978 | 666596 | 634103 | 2955842200500124 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2005-12-21 | 2006-01-20 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2005-12-21 | Device Not Retu | No | Reuse | 2006
40018301 | 2005-12-21 | Malfunction | Other | During the surgical procedure it was noticed that the wire was broken on the wrist of the permanent cautery hook instrument and that a small plastic piece of the instrument wrist fell into the patient and was retrieved. The instrument would not come out of the cannula, therefore, the customer had to break the instrument to remove it from the cannula. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, the instrument was returned with the main tube and distal-proximal clevis separated from the rest of the instrument as a result of the customer having to break apart the instrument to remove it from the cannula. As indicated in the complaint notes, the fragmented component of the instrument wrist was retrieved. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 657138 | 667780 | 635221 | 2955842200500125 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-12-21 | 2006-01-20 | No | Yes | Invalid Data | 40018301 | Device Returned To Manufacturer | No | 2005-12-21 | Yes | No | Reuse | 2006
40000619 | 2005-11-21 | Malfunction |  | It was reported that during the later part of a pyeloplasty surgical procedure the tip of a large needle driver instrument broke off inside of the patient. The broken piece was retrieved and the surgical procedure was completed with a replacement instrument. The procedure was not significantly lengthened due to the incident. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one of the large needle driver tips had broken in the radial direction. The broken piece was retrieved and returned with the instrument. Upon microscopic analysis, a possible failure mechanism was determined to be a weakened carbide-brazing interface. Engineering was able to reproduce this failure mode in the laboratory and found that the planes of the fractures were consistent with the device return from the field. | LARGE NEEDLE DRIVER ASSEMBLY | SURGICAL INSTRUMENT | 646558 | 657075 | 625064 | 2955842200500115 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2005-12-19 | 2006-01-06 | No | Yes | Health Professional | 40000619 | Device Returned To Manufacturer | Yes | 2005-11-22 | Yes | No | Reuse | 2006
40017204 | 2005-10-06 | Malfunction |  | It was reported that during the surgical procedure the "bolt" at the instrument tip became loose. The instrument and cannula had to be removed together due to the loose pin. The procedure was completed after replacing the defective instrument and the cannula. The loose pin was reported to have fallen on the floor of the or. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The distal clevis pin, which holds the grips in place, was not returned with the device. It is likely that the "bolt" referenced by the customer was the distal clevis pin. There was no visible damage to the instrument. The fact that the distal clevis pin was missing in the returned device led engineering to conclude that if this failure mode were to recur, the pin could potentially separate from the instrument during the surgical procedure. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 651371 | 661961 | 629694 | 2955842200500123 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2005-12-14 | 2006-01-06 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | Yes | 2005-10-06 | Yes | No | Unkown | 2006
40017205 | 2005-11-18 | Malfunction |  | It was inttially reported that during the surgical procedure the bipolar maryland tip forceps instrument was "catching fire in the wires. " the instrument was removed and replaced and the planned surgical procedure was completed. It was later reported that the instrument was not "catching fire at the wires. " the complaint was modified to state that smoking was observed. No patient harm, adverse outcome or injury was reported. |  | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 653298 | 663879 | 631550 | 2955842200500122 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-12-06 | 2006-01-05 | No | Yes | Health Professional | 40017205 | No | No | 2005-12-06 | Device Not Retu | No | Unkown | 2006
IS1200 A4 3P5 | 2005-11-30 | Other | Other | It was reported that during the surgical procedure the 4th arm was not raising up and down. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Field service engineering determined that the extended pt side manipulator (psm), which the customer referred to as the "4th arm", was out of calibration. The system was repaired by calibrating the extended psm. As of 2005, no additional recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 648353 | 658912 | 626747 | 2955842200500119 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-11-30 | 2005-12-29 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | No | 2005-11-30 | Device Not Retu | No | Reuse | 2005
IS1200 A4 3P5 | 2005-11-30 | Other | Other | During the surgical procedure the customer experienced a critical fault 212 error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. As of 2005, no additional recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 648359 | 658919 | 626753 | 2955842200500120 | 1 | KOG | Manufacturer | Distributor | Other | Initial | 2005-11-30 | 2005-12-29 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | No | 2005-11-30 | Device Not Retu | No | Reuse | 2005
40018301 | 2005-11-26 | Malfunction |  | Event Description: --/-- It was reported that during the surgical procedure the permanent cautery hook instrument disarticulated itself and could not be straightened. The customer had to disengage the cannula from the pt in order to remove the instrument. After the instrument and cannula were removed from the pt cavity, the customer had to cut the instrument to remove it from the cannula. No instrument pieces fell into the pt. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- H10: 67 - the instrument was returned and evaluated. Per the customer reported complaint engineering found that the instrument proximal clevis had malfunctioned at the main tube interface. The malfunction caused the instrument wrist to become misaligned ( i. E. The disartivulation noted by the customer). The misalignment of the wrist prevented the normal removal of the instrument through the cannula. As indicated in the complaint notes, no instrument components fell into the pt. --/-- Additional Manufacturer Narrative: --/-- The instrument was not returned for evaluation. Therefore, no conclusion can be drawn regarding the cause of the malfunction of the device. The hosp indicated that all device components were retrieved and no pt injury occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 649819 | 660382 | 628180 | 2955842200500121 | 1 | GEI | Manufacturer | User facility | Other | InitialFollowup | 2005-12-22 | 2005-12-29 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-11-28 | Yes | No |  | 2005
IS1200 A43P5 | 2005-11-29 | Other | Other | It was reported that during the surgical procedure the customer experienced a "2008" system error. No patient harm was reported. The procedure was aborted and rescheduled for later in december. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. As 0f 2005, there have been no reported recurrences of the 20008 system error since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 649911 | 660473 | 628268 | 2955842200500118 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-12-20 | 2005-12-29 | No | Yes | Health Professional | IS1200 A43P5 | No | No | 2005-11-29 | Device Not Retu | No | Reuse | 2005
40009305 | 2005-11-08 | Malfunction |  | It was reported that during a surgical procedure the endowrist prograsp instrument snagged and tore tissue. The physician stopped using the device and finished the procedure with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The endowrist prograsp instrument was returned for eval. The instrument passed both visual inspection and performance testing. Engineering has requested add'l case info. | ENDOWRIST PROGRASP | ENDOSCOPIC INSTRUMENT | 646554 | 657070 | 625060 | 2955842200500116 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2005-12-20 | 2005-12-22 | No | Yes | Health Professional | 40009305 | Device Returned To Manufacturer | Yes | 2005-11-23 | Yes | No | Unkown | 2005
40011007 | 2005-11-14 | Malfunction | Other | The instrument was not returned for evaluation. Therefore, no conclusion can be drawn regarding the cause of the fractured device. The hospital indicated that all components were retrieved from the patient and no patient injury occurred. | Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument.It was reported that during a prostatectomy surgical procedure, the neck of the precise bipolar pyramid tip forceps instrument broke into four pieces. All of the pieces fell inside of the patient and were retrieved. The procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 646817 | 657329 | 625321 | 2955842200500117 | 1 | GEI | Manufacturer | Company Representative | Other | InitialFollowup | 2005-12-20 | 2005-12-22 | No | Yes | Health Professional | 40011007 | Device Returned To Manufacturer | No | 2005-11-28 | Yes | No | Reuse | 2005
IS1200 A43P5 | 2005-11-16 | Other | Other | It was reported that during the surgical procedure the customer experienced "arm not free to move" errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering concluded that the encoder of the extended psm had malfunctioned. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 643633 | 654165 | 622274 | 2955842200500109 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-12-12 | 2005-12-15 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2005-11-16 | Device Not Retu | No | Reuse | 2005
IS1200 A43P5 | 2005-11-21 | Other | Other | During a procedure, the surgeon experienced a vision problem in the left eye of the stereo viewer. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation by isi field engineering concluded that the left monitor in the high resolution stereo viewer (hrsv) was not powering up. The isi field engineer repaired the hospital's system by replacing the hrsv. The hospital's da vinci vision system passed testing after the repair was performed. As of 2005, no recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 643641 | 654173 | 622282 | 2955842200500110 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-12-12 | 2005-12-15 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes | 2005-11-21 | Device Not Retu | No | Reuse | 2005
40016905 | 2005-11-17 | Malfunction |  | During the surgical procedure the harmonic shears insert broke and a "non-ceramic" piece of the insert fell into the patient cavity. As a result, the surgeon had to convert to a laparoscopic procedure. No pt harm, or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed nor denied. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 643965 | 654497 | 622583 | 2955842200500112 | 1 | LFL | Manufacturer | User facility | Other | Initial | 2005-12-14 | 2005-12-15 | No | Yes | Health Professional | 40016905 | No | Yes | 2005-11-17 | Device Not Retu | Yes | Initial | 2005
400179XX | 2005-11-18 | Malfunction |  | During the surgical procedure the tip of the monopolar curved scissors instrument burned a hole in the tip cover and that the instrument arced back to itself. Procedure was completed after replacing the defective part. No patient harm, adverse outcome or injury was reported. | The instrument and tip cover accessory were not returned for evaluation. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 643967 | 654499 | 622585 | 2955842200500113 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2005-12-14 | 2005-12-15 | No | Yes | Health Professional | 400179XX | No | Yes | 2005-11-18 | Device Not Retu | No | Unkown | 2005
400179XX | 2005-11-18 | Malfunction |  | During the surgical procedure the tip of the monopolar curved scissors instrument burned a hole in the tip cover and the instrument arced back to itself. Procedure was completed after replacing the defective part. No patient harm, adverse outcome or injury was reported. | The instrument and tip cover accessory were not returned for evaluation. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 643968 | 654500 | 622586 | 2955842200500114 | 1 | GEI | Manufacturer | User facility | Other | InitialFollowup | 2005-12-14 | 2005-12-15 | No | Yes | Health Professional | 400179XX | No | Yes | 2005-11-18 | Device Not Retu | No | Unkown | 2005
40017904 | 2005-11-15 | Malfunction |  | Event Description: --/-- It was reported that during the surgical procedure the tip of the monopolar curved scissors instrument burned a hole in the tip cover. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. The burning appears to have been a result of arcing through two holes/tears in the instrument tip cover. The cause of the damages to the tip cover is inconclusive. --/-- Additional Manufacturer Narrative: --/-- The instrument was not returned for evaluation, therefore, no conclusion can be drawn regarding the cause of the hole in the monopolar curved scissors instrument tip cover. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 644836 | 655351 | 623398 | 2955842200500111 | 1 | GEI | Manufacturer | User facility | Other | InitialFollowup | 2005-12-14 | 2005-12-15 | No | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | Yes | 2005-11-15 | Yes | No |  | 2005
DA VINCI | 2005-09-22 | Injury | Required InterventionHospitalization | During a procedure to repair the mitral valve regurgitation involving use of the da vinci robotic surgical system, the aortic valve leaflet was perforated during placement of the cardioplegia catheter. This necessitated initiation of cardiopulmonary bypass to repair the leaflet and complete the procedure. Upon removal from bypass, the patient entered cardiogenic shock due to biventricular failure. Additionally the patient went into renal failure. The patient remains hospitalized with partial recovery of left ventricular function. |  | DA VINCI | TELEMANIPULATION SYSTEM SURGIC | 649511 | 660083 | 627898 | 660083 | 1 | KOG | User Facility |  |  | Initial | 2005-11-14 | 2005-12-05 | Yes | No | Physician | DA VINCI | No | No Answer Provided |  |  |  |  | 2005
DA VINCI | 2005-09-22 | Injury | Required InterventionHospitalization | During a procedure to repair the mitral valve regurgitation involving use of the da vinci robotic surgical system, the aortic valve leaflet was perforated during placement of the cardioplegia catheter. This necessitated initiation of cardiopulmonary bypass to repair the leaflet and complete the procedure. Upon removal from bypass, the patient entered cardiogenic shock due to biventricular failure. Additionally the patient went into renal failure. The patient remains hospitalized with partial recovery of left ventricular function. |  | DA VINCI | TELEMANIPULATION SYSTEM SURGIC | 649511 | 660083 | 627898 | 660083 | 1 | KOG | User Facility |  |  | Initial | 2005-11-14 | 2005-12-05 | Yes | No | Physician | DA VINCI | No | No Answer Provided |  |  |  |  | 2005
DA VINCI | 2005-09-08 | Injury | Hospitalization | During a procedure employing the use of the da vinci robotic surgical system to treat atrial fibrillation via rf ablation of ectopic foci, intraoperative bleeding was experienced, that required initiation of cardiopulmonary bypass. The bleeding of the atrial wall was caused by the atricure ablation electrode (isolator transpolar pen), which had been unhooked from the robotic arm and was being manipulated manually. Cardiopulmonary bypass was initiated to rectify and complete the procedure. Secondary complications included a stroke and aortic dissection. The patient has been discharged. |  | DA VINCI | TELEMANIPULATION SYSTEM SURGIC | 649575 | 660137 | 627952 | 660137 | 1 | KOG | User Facility |  |  | Initial | 2005-10-24 | 2005-12-05 | Yes | No | Physician | DA VINCI | No | No Answer Provided |  |  |  |  | 2005
DA VINCI | 2005-09-08 | Injury | Hospitalization | During a procedure employing the use of the da vinci robotic surgical system to treat atrial fibrillation via rf ablation of ectopic foci, intraoperative bleeding was experienced, that required initiation of cardiopulmonary bypass. The bleeding of the atrial wall was caused by the atricure ablation electrode (isolator transpolar pen), which had been unhooked from the robotic arm and was being manipulated manually. Cardiopulmonary bypass was initiated to rectify and complete the procedure. Secondary complications included a stroke and aortic dissection. The patient has been discharged. |  | DA VINCI | TELEMANIPULATION SYSTEM SURGIC | 649575 | 660137 | 627952 | 660137 | 1 | KOG | User Facility |  |  | Initial | 2005-10-24 | 2005-12-05 | Yes | No | Physician | DA VINCI | No | No Answer Provided |  |  |  |  | 2005
40018301 | 2005-09-21 | Malfunction |  | It was reported that there was a crack in the insulation at the distal tip of the permanent cautery hook instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that both the distal and proximal clevises are melted in several locations. The distal clevis has excessive charring in the area where the conductor enters the hook and the proximal clevis has an arc track that runs laterally between the clevis pin and the main tube. The customer reported "crack" in the insulation was concluded by isi engineering to be the arc track in the insulation. No components were reported to have been missing from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 639471 | 649986 | 618273 | 2955842200500106 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-09-21 | 2005-12-02 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2005-09-21 | Yes | No | Reuse | 2005
40016905 | 2005-11-04 | Malfunction |  | It was reported that during the surgical procedure the customer experienced difficulty removing the harmonic curved shears instrument from the sterile adapter. As a result, they had to disengage the sterile adapter to remove the instrument. Upon disengagement, the harmonic shears insert broke and the non-ceramic piece of the insert fell onto a drape, outside of the patient cavity. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed nor denied. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 639479 | 649994 | 618280 | 2955842200500107 | 1 | LFL | Manufacturer | Company Representative | Other | Initial | 2005-11-04 | 2005-12-02 | No | Yes | Health Professional | 40016905 | No | No | 2005-11-04 | Device Not Retu | Yes | Initial | 2005
IS1200 A4 3P5 | 2005-10-11 | Other | Other | It was reported that during the surgical procedure the customer experienced a 20008 error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. Additional info provided by the isi account manager indicates that the surgeon found the pt's anatomy to be very challenging, therefore, the decision to convert the procedure to open surgical techniques was not solely due to the error experienced by the customer. | The extended psm was returned to isi for failure analysis investigation. Engineering observed that the rolling loop was damaged and the cables were off the pulley, thus causing the extended psm to malfunction. Based on the additional information provided by the isi field engineer, it was determined that product malfunction and patient anatomy may have contributed to the surgeon's decision to convert the procedure to open surgical techniques. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 639494 | 650011 | 618294 | 2955842200500108 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-10-11 | 2005-12-02 | Yes | Yes | Health Professional | IS1200 A4 3P5 | No | Yes | 2005-11-21 | Device Not Retu | No | Reuse | 2005
IS1200 A43P5 | 2005-10-24 | Other | Other | During the surgical procedure the customer experienced "temporary motion sensor data mismatch" errors. No pt harm was reported. The procedure was coverted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the extended patient side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. Engineering concluded that the potentiometer of the extended psm had malfunctioned. As of 2o05, no additional recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 638705 | 649212 | 617541 | 2955842200500103 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-10-24 | 2005-11-23 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2005-10-24 | Device Not Retu | No | Reuse | 2005
40017904 | 2005-10-26 | Malfunction |  | During the surgical procedure the tip of the monopolar curved scissors instrument burned a hole in the tip cover. No patient harm, adverse outcome or injury was reported. Additional info later provided by the isi account manager indicated that there was a frayed wire in the instrument wrist. | The instrument was not returned for evaluation, therefore, no conclusion can be drawn regarding the cause of the hole in the monopolar curved scissiors instrument tip cover. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 693991 | 704923 | 670915 | 2955842200500105 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-10-26 | 2005-11-23 | No | Yes | Health Professional | 40017904 | No | No | 2005-10-26 | Device Not Retu | No | Unkown | 2005
40018301 | 2005-10-26 | Malfunction |  | Halfway through the surgical procedure some of the insulation of the permanent cautery hook instrument broke off and fell inside the patient's pelvic cavity. The broken pieces were retrieved and the planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noticed signs of damage to the ultem plastic. Engineering also observed that the wire to the instrument hook had been charred. This discovery has led engineering to conclude that the instrument likely arced during the surgical procedure, which may have caused the ultem to deform and fracture. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 694005 | 704937 | 670928 | 2955842200500104 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-10-26 | 2005-11-23 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-10-26 | Yes | No | Reuse | 2005
40017904 | 2005-10-21 | Malfunction |  | During the surgical procedure there was a tear on the protective tip cover of the monopolar curved scissors instrument which was not immediately noticed. As a result, there was slight arcing to the iliac artery which caused bleeding. The bleeding was coagulated and sutured without any complications and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering concluded that the tip cover was most likely damaged when the instrument was being retracted out of the cannula. This issue can only occur if the instrument's wrist was not straightened prior to removal. It pulled with enough force, the edge of the cannula can catch against the tip cover, causing the tear seen by the customer and allowing arcing to occur. Engineering also observed scratch-like marks at the distal end of the instrument main tube. This discovery indicates that the main tube was damaged by the cannula during insertion or retraction of the instrument. The instruments for use specifically states under general precautions and warnings, "to prevent damage to the tip cover accessory insulation, always straighten the instrument wrist under endoscopic visualization prior to removing the instrument through the cannula. ". | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 694027 | 704959 | 670949 | 2955842200500102 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-10-24 | 2005-11-23 | Yes | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | No | 2005-10-24 | Yes | No | Reuse | 2005
40018301 | 2005-10-18 | Malfunction |  | During the surdical procedure, the permanant cautery hook instrument was arcing. No further details were provided. No patient harm, asverse outcome was reported. | The instrument was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 701768 | 712700 | 678395 | 2955842200500101 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-10-18 | 2005-11-16 | No | Yes | Health Professional | 40018301 | No | Yes | 2005-10-18 | Device Not Retu | No | Unkown | 2005
40016905 | 2005-10-17 | Malfunction | Other | During the surgical procedure, the harmonic curved shears insert broke and a part fell into the patient. The broken part was retrieved and the harmonic curved shears instrument and insert were removed and replaced. The planned surgical procedure was completed, however, the incident lengthened the procedure by 30 minutes. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned along with the detached clamping arm. Inspection revealed no signs of significant damage or deformation to the clamping arm and the outer and inner tube shafts. The detached clamping arm was re-attached onto the harmonic curved shears insert and inserted into the harmonic curved shears instrument so that a functionality test could be performed. The instrument and insert functioned normally. Based on the evaluation results, no conclusion can be drawn regarding the cause of the clamping arm breakage. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 718596 | 730305 | 695596 | 2955842200500100 | 1 | LFL | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-10-17 | 2005-11-16 | No | Yes | Health Professional | 40016905 | Device Returned To Manufacturer | Yes | 2005-10-17 | Yes | Yes | Initial | 2005
IS1200 A43P5 | 2005-10-11 | Other | Other | It was reported that during the surgical procedure, the customer experienced intermittent system alarms with a 20008 error code. No patietn harm was reported. The procedure wa converted to traditional open surgical techniques to finish the planned surgery. Additional information provided by the isi field enginer idicated that the patient was informed prior to surgery with the da vinci surgical system, that the procedure may have to be converted to open surgical techniques, due to his pathology. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the left master tool manipulator (mtmr). The system was repaired by replacing the affected mtml. The mtmr was returned to isi for failure analysis investigation. Engineering concluded that the axis 6 potentiometer of the mtml was malfunctioning. Based on the additional information provided by the isi field engineer, it was determined that product malfunction and patient pathology may have contributed to the surgeon's decision to convert the procedure to open surgical techniques. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 749266 | 761359 | 725861 | 2955842200500099 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-10-11 | 2005-11-10 | Yes | Yes | Health Professional | IS1200 A43P5 | No | Yes |  | Device Not Retu | No | Reuse | 2005
40008305 | 2005-09-12 | Malfunction |  | Reported that during the surgical procedure, the tip of the ultrasonic shears instrument broke and two fragments fell inside the pt. The fragments were retrieved from the pt and thrown away. No pt harm, adverse outcome or injury was reported. | The ultrasonic shears instrument was returned without the ultrasonic probe accessory. Failure analysis of the ultrasonic shears instrument revealed that the tip of the instrument was intact; however, excessive charring and teflon deformation was observed. The observed instrument damage is known to occur when the instrument is energized and there is no tissue between the grip and the probe; therefore, it was concluded that the instrument was likely misused. Because the ultrasonic probe was not returned for failure analysis evaluation, no conclusion can be drawn regarding the cause of the probe breakage. | ULTRASONIC SHEARS INSTRUMENT | ULTRASONIC SHEARS INSTRUMENT | 724405 | 736516 | 701761 | 2955842200500098 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2005-09-12 | 2005-11-04 | No | Yes | Invalid Data | 40008305 | Device Returned To Manufacturer | No |  | Yes | No | Reuse | 2005
40016905 | 2005-09-13 | Malfunction | Other | It was reported that during the surgical procedure, part of the jaw of the harmonic curved shears insert broke off inside the pt and was retrieved. The harmonic curved shears instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | A portion of the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm of the harmonic curved shears insert was missing and not returned. Engineering also observed that the outer tube had a minor deformation at the horse collar. The inner tube does not appear to be deformed or damaged. It was indicated that the missing component was retrieved from the pt and no pt injury occurred. | HARMONIC CURVED SHEARS INSERT 5PACK | ENDOSCOPIC INSTRUMENT | 685386 | 696276 | 662620 | 2955842200500096 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-09-13 | 2005-10-13 | No | Yes | Health Professional | 40016905 | Device Returned To Manufacturer | No | 2005-09-13 | Yes | Yes | Initial | 2005
40017204 | 2005-09-13 | Malfunction |  | It was reported that during the surgical procedure, the bipolar maryland tip forceps instrument was smoking at the wirst inside the pt. Thinking that it was the pt's blood that was causing the smoke, the customer removed and cleaned the instrument, then re-inserted the instrument into the pt. When the instrument was inside the pt for the second time, it was noticed that the tips clugged [sic]. At that point, the instrument was removed and replaced with the same type of instrument and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the base of both yaw pulley grips are burnt and melted to the extent that they can no longer open and close. Engineering also discovered an arc path, which extends from one grip to the other. The evidence of an arc path has led engineering to conclude that the instrument arced during the surgical procedure, thus causing the "smoking" experienced by the customer. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 685394 | 696284 | 662628 | 2955842200500097 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-09-13 | 2005-10-13 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | No | 2005-09-13 | Yes | No | Initial | 2005
IS1200 A43P5 | 2005-09-12 | Other | Other | It was reported that during the surgical procedure the section of the master tool manupulator (mtm) grip that holds the velcro loop fell off. The mtm is not a pt contact component. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Investigation was conducted by isi field engineering. The isi field engineer repaired the hosp's system by repairing the master tool manipulator grip assembly. As of 2005, no recurrences of this failure mode have reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 630768 | 641261 | 609857 | 2955842200500095 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-09-12 | 2005-10-12 | Yes | Yes | Health Professional | IS1200 A43P5 | No | No | 2005-09-12 | Device Not Retu | No | Reuse | 2005
IS1200 A4 3P4 | 2005-09-08 | Other | Other | It was reported that during the surgical procedure the system arm was slow in response and was having trouble opening and closing the instrument tips. No error messages were seen. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | Investigation was conducted by isis field engineering. Inspection and testing was performed on the pt side manipulator (which the customer referred to as the system "arm") by an isi field engineer. The isi field engineer was not able to reproduce the faults experienced by the customer, therefore, no conclusion can be drawn regarding the cause of the customer reported complaint. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 629858 | 640347 | 608960 | 2955842200500093 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-09-08 | 2005-10-07 | Yes | Yes | Health Professional | IS1200 A4 3P4 | No | Yes | 2005-09-08 | Device Not Retu | No | Reuse | 2005
40011007 | 2005-08-24 | Malfunction | Other | During the surgical procedure, it was noticed that the joint on the end of the precise bipolar pyramid tip forceps instrument was broken. The instrument would not come out of the cannula, therefore, the cannula was removed from the patient with the instrument still inserted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found that the connection from the proximal clevis to the main tube was broken. Engineering also observed that the main tube was frayed and buckled. Engineering also discovered that a section of the proximal clevis collar was missing and not returned with the instrument. This complaint is being reported due to the fact that it cannot be determined if any components separated from the instrument during the surgical procedure or during handling after the procedure. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 685101 | 695992 | 662347 | 2955842200500094 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2005-08-24 | 2005-10-07 | No | Yes | Invalid Data | 40011007 | Device Returned To Manufacturer | No | 2005-08-24 | Yes | No | Reuse | 2005
IS1200 | 2005-07-11 | Malfunction |  | A prostatectomy procedure performed using the da vinci surgical system was interrupted midway through the surgery due to a system malfunction. The surgeon was unable to complete the procedure with the robotic surgical system. The malfunction was reported as difficulty experienced when removing an instrument from one arm of the system. It has been alleged that the procedure was converted to an "open incision prostatectomy" due to the malfunction of the robot. It was also alleged that the patient sustained injuries. Hospital personnel have indicated that the patient did not sustain any injuries during and from the procedure. The surgery occurred in 2005. Isi was informed of the malfunction, but not of any alleged injury to the patient in 2005. Isi was informed of conversion to open on in 2005. | The investigation conducted by isi field engineering on july 15, 2005 concluded that the patient side manipulator (psm) insertion counterweight bearing malfunctioned. The cause of the malfunction is as of yet uncertain. The isi field engineer repaired the hospital's system by replacing the patient side manipulator (psm) assembly. A review of the surgical record indicates that the conversion to open surgery was not required due to the system malfunction. The surgeon stated that the patient had other treatment options. Moreover, hospital personnel have indicated that the patient did not sustain any injuries during and from the procedure. Isi was first notified of an alleged injury in 2005 when it received a patient letter dated in 2005. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 688379 | 699291 | 665526 | 2955842200500092 | 1 | NAY | Manufacturer | Other | Invalid Data | Initial | 2005-09-27 | 2005-09-30 | No | Yes | Health Professional | IS1200 | No | No | 2005-09-02 | Yes | No | Reuse | 2005
40017204 | 2005-08-22 | Malfunction |  | During the surgical procedure, it was noticed that the bipolar maryland tip forceps instrument was smoking at the side. The instrument was removed and the procedure was continued with another bipolar maryland tip forceps instrument. It was noticed that the cable was broken on the second instrument. The second instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering eval observed that the instrument plastic jaw pulley was burnt. Engineering also discovered an arc track at the base of the instrument metal grip, which extends from one grip to the other. The evidence of arc tracking had led engineering to concluded that the instrument arced during the surgical procedure, thus causing the "smoking" experienced by the customer. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 636340 | 646851 | 615265 | 2955842200500089 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-08-26 | 2005-09-23 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | No | 2005-08-26 | Yes | No | Reuse | 2005
40011007 | 2005-08-03 | Malfunction | Other | The hosp reported that during the surgical procedure, while trying to remove the precise bipolar pyramid tip forceps instrument for cleaning, the rim above the working mechanism fractured. The instrument and trocar were removed from the patient's abdomen connected together and the fragments were retrieved. Additiional details provided by the hospital indicate that all components were retrieved from the patient. The istrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument proximal clevis and main tube were damaged. It appears that the proximal clevis fractured radially first, followed by the main tube when the clevis was no longer equally supported on all sides. The hospital indicated that all components were retrieved from the patient and no patient injury occurred. | PRECISE BIOPLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640391 | 650917 | 619180 | 2955842200500090 | 1 | GEI | Manufacturer | Other | Other | Initial | 2005-08-26 | 2005-09-23 | No | Yes | Invalid Data | 40011007 | Device Returned To Manufacturer | No | 2005-08-26 | Yes | No | Reuse | 2005
40017204 | 2005-08-22 | Malfunction |  | During the surgical procedure it was noticed that the cable was broken on the bipolar maryland tip forceps instrument. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation observed that the conductor wire was broken at the instrument yaw pulley. Engineering also found a small bit of localized burning at the break. The discovery of the localized burning has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640404 | 650930 | 619192 | 2955842200500091 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-08-26 | 2005-09-23 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | No | 2005-08-26 | Yes | No | Reuse | 2005
40018301 | 2005-08-24 | Malfunction |  | Almost immediately after the start of the surgical procedure, during the surgeon's first attempt to divide the pelvic fascia, there was smoke coming out of the wrist joint of the permanent cautery hook instrument. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the conductor wire was broken at the pulley interface and that the insulation was damaged underneath the conductor cap. It apperars that the conductor wire kinked prior to breaking. Signs of arc tracking were also found near the location of the exposed bare wire. The discovery of arc tracking has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640420 | 650946 | 619208 | 2955842200500086 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2005-08-24 | 2005-09-23 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2005-08-24 | Yes | No | Reuse | 2005
40012607 | 2005-08-26 | Malfunction |  | There was a crack near the wrist joint of the permanent cautery hook instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering discovered signs of arc tracking on the instrument yaw pulley and clevis, which begins at the instrument gooseneck. No "crack" was observed. The "crack" identified by the customer was most likely the visual surface appearance of the arc tracks that make it appear to be cracked. This complaint is being reported due to the evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640431 | 650957 | 619219 | 2955842200500088 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-26 | 2005-09-23 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2005-08-26 | Yes | No | Reuse | 2005
40017904 | 2005-08-25 | Malfunction |  | During the surgical procedure it was noticed that the monopolar curved scissors instrument had a hole in its tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument main tube was burned through. Engineering also found an arc path through a longitudinal crack in the main tube. The discorvery of an arc path has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640441 | 650967 | 619229 | 2955842200500087 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-25 | 2005-09-23 | No | Yes | Health Professional | 40017904 | Device Returned To Manufacturer | Yes | 2005-08-25 | Yes | No | Reuse | 2005
40018301 | 2005-08-22 | Malfunction |  | During the first part of the surgical procedure the permanent cautery hook instrument was arcing. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of arc tracking beginning at the conductor wire, underneath the conductor cap. It appears that the conductor wire kinked, thus damaging the conductor wire insulation. The discovery of arc tracking has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 640016 | 650547 | 618820 | 2955842200500085 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-08-22 | 2005-09-21 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2005-08-22 | Yes | No | Reuse | 2005
40017203 | 2005-08-12 | Malfunction |  | It was reported that during the surgical procedure, exposed wires were noticed at the "wrist" of the bipolar maryland tip forceps instrument. When the bipolar pedal was pressed, the instrument began arcing and a spray of sparks was seen. The instrument was removed and the planned surgical procedure was completed with a different instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evalution observed that the instrument yaw pulleys were charred and deformed. Engineering also discovered an arc track path at the base of the instrument grips, which extends from one grip to the other. The evidence of arc tracking has led engineering to conclude that the instrument arced during the surgical procedure, thus causing the "spray of sparks" experienced by the customer. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 659515 | 670201 | 637547 | 2955842200500084 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-08-12 | 2005-09-09 | No | Yes | Health Professional | 40017203 | Device Returned To Manufacturer | No | 2005-08-12 | Yes | No | Reuse | 2005
40012608 | 2005-08-09 | Malfunction |  | It was reported that the plug broke off of the permanent cautery hook instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument plug riser was broken off. Engineering also found signs of arc tracking evident at the instrument gooseneck. This complaint is being reported due to the evidence of arc tracking found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 660519 | 671199 | 638507 | 2955842200500083 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-09 | 2005-09-08 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-08-09 | Yes | No | Reuse | 2005
40017903 | 2005-08-09 | Malfunction |  | During the beginning of the surgical procedure the monopolar curved scissors instrument was smoking. The instrument was immediately removed and replaced with a different type of instrument to complete the planned surgicl procedure. It was also reported that there was a visible indication of melting of the outer shaft near the tip of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation determined that the seal between the main tube and the tube extension failed, which allowed arcing within the main tube. This discovery has led engineering to conclude that the instrument arced during the surgical procedure, thus creating the smoke and melted appearance noticed by the customer. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 660528 | 671208 | 638516 | 2955842200500082 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-08-09 | 2005-09-08 | No | Yes | Health Professional | 40017903 | Device Returned To Manufacturer | No | 2005-08-09 | Yes | No | Reuse | 2005
40012608 | 2005-08-08 | Malfunction |  | It was reported that the permanent cautery hook instrument does not work. No further details were provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of arc tracking through the plastic at the instrument gooseneck. It was also observed that the instrument teflon shrink tubing was missing and not returned with the instrument. This complaint is being reported due to the evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 659120 | 669794 | 637165 | 2955842200500080 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-08 | 2005-09-07 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-08-08 | Yes | No | Reuse | 2005
40018301 | 2005-08-02 | Malfunction |  | It was reported that the hook of the permanent cautery hook instrument dislodged and fell into the pt cavity. The surgeon was able to locate the hook and retrieve it prior to closure. It was also reported that it was difficult to read the writing of the lot number of the instrument. No pt harm, adverse outcome or injury was reported. This incident was also reported by the hospital via medwatch report. | The instrument was returned and evaluated. The hook and ceramic sleeve were not returned with the instrument. It was noticed that the lot number was smeared, however, the correct lot number could be confirmed. Engineering evaluation observied that the instrument hook and ceramic sleeve had pulled completly out of the plastic yaw pulley, and there were remnants of silicone adhesive on the base of the yaw pulley. There was no indication that the plastic yaw pulley had broken. Engineering also observed that there was incomplete fill of the ultem yaw pulley material around the base of the hook. Charring was also noticed at the base of the yaw pulley. The charring appears to have happened after the hook broke free from the conductor cable and pulley. This incident was also reported by the hospital. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 940432 | 971255 | 931967 | 2955842200500081 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2005-08-08 | 2005-09-07 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-08-08 | Yes | No | Reuse | 2005
IS1200 A43P4 | 2005-08-04 | Other | Other | It was reported that during the surgical procedure the customer experienced "temporary data sensor mismatch" errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field service engineering concluded that the system errors experienced by the customer were associated with the extended pt side manipulator (psm). The system was repaired by replacing the affected extended psm. The extended psm was returned to isi for failure analysis investigation. It was discovered that the pinion gear on axis 6 of the extended psm was assembled with the wrong glue which allowed the pinion system errors experienced by the customer. As of september 2005, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 623022 | 633453 | 602268 | 2955842200500079 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-04 | 2005-09-02 | Yes | Yes | Health Professional | IS1200 A43P4 | No | Yes | 2005-08-04 | Device Not Retu | No | Reuse | 2005
40012608 | 2005-08-02 | Malfunction |  | During the surgical procedure the permanent cautery hook instrument was arcing. The instrument was removed and replaced with a different type of instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of arc tracking, which begins at the instrument gooseneck and goes to the drive cable. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. Engineering also observed that the instrument distal clevis ear was broken off. It cannot be determined if any components separated from the instrument during the surgical procedure or during handling after the procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 675231 | 686024 | 652786 | 2955842200500078 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-08-02 | 2005-09-01 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-08-02 | Yes | No | Reuse | 2005
40018301 | 2005-07-26 | Malfunction |  | Event Description: --/-- It was reported that during the surgical procedure the permanent cautery hook instrument become covered with debris and started smoking from the sides. The surgeon began to clean the permanent cautery hook instrument with another instrument and during cleaning, the surgeon pulled the hook off of the instrument. The hook was retrieved from the patient and the planned surgical procedure was completed. No pt harm, adverse outcome of injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated by engineering. It appears that the instrument hook began to arc due to a break in the conductor wire insulation. The insulation break was caused by the conductor wire kinking underneath the cap, thus damaging the conductor wire insulation and causing the "smoking" experienced by the customer. The detached hook was not returned with the instrument, however, the customer reported complaint indicates that the hook was pulled out of the yaw pulley as a result of the customer trying to clean the charred material off of the instrument with another instrument. The permanent cautery instrument instructions for use (ifu) specifically states under general precautions and warnings, "do not use an instrument to clean debris off another instrument intraoperatively. This may result in damage to the instruments. --/-- Additional Manufacturer Narrative: --/-- The instrument was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 628068 | 638520 | 607214 | 2955842200500077 | 1 | GEI | Manufacturer | Company Representative | Other | InitialFollowup | 2005-07-26 | 2005-08-25 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2005-07-26 | Yes | No |  | 2005
40000419 | 2005-07-13 | Malfunction |  | It was reported that during the surgical procedure, the scalpel cautery instrument was arcing. The instrument was removed and replaced with another scalpel cautery instrument, however, the replacement instrument also began to arc. The planned surgical procedure was completed with a permanent cautery hook instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering concluded that the instrument arcing may have been due to poor insulation around the hook. This instrument was manufactured in june 2002. The affected instrument model is no longer distributed by intuitive surgical, inc. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 690386 | 701278 | 667453 | 2955842200500076 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-07-13 | 2005-08-12 | No | Yes | Health Professional | 40000419 | Device Returned To Manufacturer | Yes | 2005-07-13 | Yes | No | Reuse | 2005
40000419 | 2005-07-13 | Malfunction |  | It was reported that during the surgical procedure, the scalpel cautery instrument was arcing. The instrument was removed and replaced with another scalpel cautery instrument, however, the replacement instrument also began to arc. The planned surgical procedure was completed with a permanent cautery hook instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering concluded that the instrument arcing may have been due to poor insulation around the hook. This instrument was manufactured in june 2002. The affected instrument model is no longer distributed by intuitive surgical, inc. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 690389 | 701281 | 667456 | 2955842200500075 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-07-13 | 2005-08-12 | No | Yes | Health Professional | 40000419 | Device Returned To Manufacturer | Yes | 2005-07-13 | Yes | No | Reuse | 2005
IS1200 A43P4 | 2005-07-11 | Other | Other | It was reported that the emergency stop was activated during the procedure and could not be overridden. In an effort to resolve this issue, the system was restarted. After the system was restarted and self-test had been completed, the following error message was received, "emergency stop activated restart system". The system was then restarted several more times, however, the same issue continued to recur. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field engineering concluded that emergency stop cabling was causing the system to fault. The isi field engineer repaired the hospital's system by replacing the emergency stop cables. The hospital's da vinci system passed testing after repair was performed. As of august 2005, no recurrency of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 616853 | 627226 | 596170 | 2955842200500072 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-07-11 | 2005-08-10 | Yes | Yes | Health Professional | IS1200 A43P4 | No | Yes | 2005-07-11 | Device Not Retu | No | Reuse | 2005
40018401 | 2005-07-11 | Malfunction |  | During the surgical procedure the permanent cautery spatula instrument wrist was not working properly. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the ceramic was cracked near the top of the spatula. Engineering also found signs of arc tracking on the instrument yaw pulley and clevis, possibly beginning from the cable crimp. This discovery has led engineering to concluded that the instrument. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 660030 | 670707 | 638033 | 2955842200500073 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2005-07-11 | 2005-08-10 | No | Yes | Health Professional | 40018401 | Device Returned To Manufacturer | Yes | 2005-07-11 | Yes | No | Reuse | 2005
40017204 | 2005-07-11 | Malfunction | Other | During the surgical procedure a bipolar maryland tip forceps instrument was inserted through a cannula into the patient. Immediately upon endoscopic visualization, it was noticed that the instrument was bent over to one side. The instrument had not yet been used to perform any surgical tasks. The surgical staff immediately chose to remove the instrument from the patient. While attempting to remove the instrument, resistance was felt through the cannula and the proximal clevis of the instrument fragmented. Two instrument fragments were removed from the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and elevated. Per the customer reporter complaint, engineering observed that the proximal clevis broke at the main tube of the instrument. The side of the proximal clevis broke out and the broken pieces were returned with the instrument. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 660043 | 670720 | 638046 | 2955842200500074 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-07-11 | 2005-08-10 | No | Yes | Invalid Data | 40017204 | Device Returned To Manufacturer | No | 2005-07-11 | Yes | No | Reuse | 2005
40018301 | 2005-06-27 | Malfunction |  | Event Description: --/-- It was reported that during the surgical procedure the permanent cautery hook instrument snagged on the endo catch. A small plastic piece from the tip of the instrument was noted missing. The instrument was removed and the planned surgical procedure was completed without further incident. The small plastic piece was not found. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed that the ear of the instrument distal clevis was broken off. The broken piece and the conductor wire cap were not returned with the instrument. --/-- Additional Manufacturer Narrative: --/-- H10 the instrument was not returned for evaluation, therefore, the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 623117 | 633552 | 602359 | 2955842200500071 | 1 | GEI | Manufacturer | User facility | Nurse | InitialFollowup | 2005-06-30 | 2005-07-29 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-06-30 | Yes | No |  | 2005
40018301 | 2005-06-23 | Malfunction |  | The permanent cautery hook instrument was charred. No further details were provided. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Engineering found signs of arc tracking on the instrument distal clevis. It was determined that the conductor wire insulation became damaged at the conductor cap due to kinking and then getting pinched. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure, thus causing the "charred" appearance of the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 708970 | 719902 | 685413 | 2955842200500070 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-06-23 | 2005-07-22 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-06-23 | Yes | No | Reuse | 2005
40017204 | 2005-06-21 | Malfunction |  | During the middle of the surgical procedure, the bipolar maryland tip forceps instrument was not opening / closing. The instrument was removed and replaced and the planned surgical procedure was completed. Upon closer inspection of the removed instrument, it was noticed that the insulation of the instrument was melted and charred. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Engineering discovered signs of arc tracking between the base of the grips of the melted plastic yaw pulleys. The arc tracking has led engineering to conclude that the instrument arced during the surgical procedure, which caused the plastic to burn at the instrument grip base and resulted in the instrument not functioning as intended. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 708987 | 719919 | 685430 | 2955842200500068 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-06-21 | 2005-07-22 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | No | 2005-06-21 | Yes | No | Reuse | 2005
AESOP 3000 | 2005-05-11 | Malfunction |  | It was reported that the aesop system was "skipping" without voice command. The arm was not responding well to voice commands and the arm was moving when no voice command was issued. No pt harm, adverse outcome or injury was reported. | An isi field engineer went to the customer site and reprogrammed the surgeon's voicecard. After the voicecard was reprogrammed, the system correctly recognized the surgeon's commands. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 731709 | 743912 | 708915 | 2955842200500069 | 1 | GCJ | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-11 | 2005-07-22 | No | Yes | Health Professional | AESOP 3000 | No | Yes | 2005-05-11 | Yes | No | Reuse | 2005
40012608 | 2005-06-16 | Malfunction |  | It was reported that the tip of the permanent cautery hook instrument was not functioning properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of arc tracking on several parts of the instrument wrist, which appears to begin at the instrument gooseneck. This discovery has led engineering to conclude hat the instrument may have arced during a previous surgical procedure, thus contributing to the tip not functioning as intended. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 611688 | 622043 | 591110 | 2955842200500067 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-06-16 | 2005-07-15 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-06-16 | Yes | No | Reuse | 2005
IS1200 A43P4 | 2005-06-10 | Other | Other | It was reported that during the surgical procedure the customer experienced a recurring "20008" system error. No pt harm was reported. The procedure was converted to tradition open surgical techniques to finish the planned surgery. | H10: the investigation conducted by field service engineering concluded that the recurring system error experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. | DA VINCI SURGICAL SYSTEMS | ENDOSCOPIC INSTRUMENT CONTROL | 610420 | 620742 | 589901 | 2955842200500066 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-06-10 | 2005-07-11 | Yes | Yes | Health Professional | IS1200 A43P4 | No | No | 2005-06-10 | Device Not Retu | No | Reuse | 2005
40012608 | 2005-06-07 | Malfunction |  | It was reported that during the surgical procedure, sparks were seen coming from the permanent cautery hook instrument. The instrument was removed and the customer observed burnt pieces and cracking on the housing. The planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. The customer reported that the instrument had sparked and that the housing was burned and cracked. The failure analysis investigation determined that the conductor wire insulation was damage. No components were found to be missing from the instrument. It was concluded that the instrument arced from the damaged conductor wire. The "cracked" idenified by the customer were actually arc tracks on the instrument wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 610424 | 620746 | 589905 | 2955842200500065 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-06-07 | 2005-07-07 | No | Yes | Health Professional | 40012608 | No | No | 2005-06-07 | Yes | No | Reuse | 2005
40017204 | 2005-05-26 | Malfunction |  | Prior to beginning the surgical procedure, the customer noticed that the bipolar maryland tip forceps instrument had a broken tip. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument grip was bent at the yaw pulley, which prevented the tips from meeting. The instrument was returned intact, with no pieces missing. Engineering also discovered signs of arc tracking, which spans from grip to grip at the instrument yaw pulley. This complaint is being reported due to evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during previous surgical procedure. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 723253 | 735362 | 700613 | 2955842200500063 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-26 | 2005-06-23 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | Yes | 2005-05-26 | Yes | No | Reuse | 2005
40017204 | 2005-05-26 | Malfunction |  | It was reported that prior to beginning the surgical procedure, the customer noticed that the bipolar maryland tip forceps instrument had a broken tip. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering discovered signs of arc tracking, which spans from grip to grip at the instrument yaw pulley. This complaint is being reported due to the evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during previous surgical procedure. The instrument was returned intact, with no pieces missing. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 723270 | 735379 | 700630 | 2955842200500064 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-26 | 2005-06-23 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | Yes | 2005-05-26 | Yes | No | Reuse | 2005
40017204 | 2005-05-26 | Malfunction |  | It was reported that prior to beginning the surgical procedure, the customer noticed that the bipolar maryland tip forceps instrument had a broken tip. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument grip was bent at the yaw pulley. Which prevented the tips from meeting. Engineering also discovered signs of arc tracking between the grips at the base of the instrument yaw pulley. This complaint is being reported due to the evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 723315 | 735425 | 700676 | 2955842200500062 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-26 | 2005-06-23 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | Yes | 2005-05-26 | Yes | No | Reuse | 2005
IS1200 | 2005-06-02 | Malfunction | Other | It was reported that during a surgical procedure, the jaws of instruments controlled by one arm of the da vinci surgical system would not close fully. The surgery was converted to an open surgical techniques in order to complete the planned surgical procedure. No pt harm, injury or other adverse event were reported. | An isi field engineer evaluated the system and did not find any mechanical defects. In addition, no errors were found in the system error log. As a precaution, the field engineer replaced the corresponding pt side manipulator (psm) arm. The returned psm was also evaluated by isi engineering. Engineering was also unable to duplicate the customer reported symptom when testing the psm with multiple instruments. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 629624 | 640103 | 608727 | 2955842200500060 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-06-02 | 2005-06-16 | No | Yes | Health Professional | IS1200 | No | No |  | Yes | No | Reuse | 2005
40012608 | 2005-05-20 | Malfunction |  | During sterilization it was noticed that the insulation at the tip of the permanent cautery hook instrument was very poor. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of initial arc tracking on the instrument hook tip, which begins at the shrink tubing and continues down towards the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 713026 | 723945 | 689368 | 2955842200500058 | 1 | GEI | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2005-05-20 | 2005-06-16 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-05-20 | Yes | No | Reuse | 2005
IS1200 A4 3P4 | 2005-05-24 | Other | Other | During the surgical procedure the customer experienced several "temporary motion sensor mismatch" errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Engineering discovered a potentiometer out of specification and that the lubrication used on the mtml was contaminated, causing the system to fault. As of june 15, 2005, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 713032 | 723951 | 689374 | 2955842200500061 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-24 | 2005-06-16 | Yes | Yes | Health Professional | IS1200 A4 3P4 | No | Yes | 2005-05-24 | Device Not Retu | No | Reuse | 2005
40018301 | 2005-05-20 | Malfunction |  | During the beginning of the surgical procedure, the permanent cautery hook instrument began smoking at the joint, while the instrument was inside of the pt. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found evidence of arc tracking around and under the silicone potting of the instrument wire and yaw pulley interface. The discovery of arc tracking has led engineering to conclude that the instrument arced during the surgical procedure, which generated the smoke noticed by the customer. The electrical arcing may have also compromised the adhesion of the potting to the yaw pulley. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 713049 | 723968 | 689390 | 2955842200500059 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-20 | 2005-06-16 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-05-20 | Yes | No | Reuse | 2005
40015403 | 2005-05-17 | Malfunction |  | After the surgical procedure was completed, the physician noticed a burn on the pt's skin around the area where the right arm instrument trocar was placed. Additional information was provided by an isi account manager in 5/2005. The isi account manager stated that no instrument arcing was observed. The burn was approximately 1 millimeter in depth and on the dermis level only. No permanent pt injury was reported. | The instrument was not returned for evaluation, therefore, the cause of the pt burn cannot be determined. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 713057 | 723976 | 689397 | 2955842200500057 | 1 | GEI | Manufacturer | Company Representative | Other | InitialFollowup | 2005-05-17 | 2005-06-16 | No | Yes | Health Professional | 40015403 | No | No | 2005-05-17 | Device Not Retu | No | Unkown | 2005
IS1000 A33 | 2005-05-16 | Other | Other | It was reported that during the surgical procedure the right eye of the system went black. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by isi field engineering concluded that the right camera control unit (ccu) had lost power and could not be restarted. In may 2005, the isi field engineer repaired the hospital's system by replacing the right ccu and right camera cable. As of june 2005, no recurrences of this failure mode have been reported at this site since the repair of the ccu. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 604744 | 614975 | 584515 | 2955842200500056 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-05-16 | 2005-06-15 | Yes | Yes | Health Professional | IS1000 A33 | No | No | 2005-05-16 | Device Not Retu | No | Reuse | 2005
40018301 | 2005-05-12 | Malfunction |  | During the middle of the surgical procedure the permanent cautery hook instrument was smoking inside the pt. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. During failure analysis investigation, engineering found evidence of arc tracking. It was determined that the conductor wire kinked and became pinched between the cap and the yaw pulley, thus damaging the insulation on the conductor wire. This discovery has led engineering to conclude that the instrument arced during the surgical procedure, which generated the smoke noticed by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 624179 | 634610 | 603402 | 2955842200500054 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-05-12 | 2005-06-10 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | Yes | 2005-05-12 | Yes | No | Reuse | 2005
40012608 | 2005-05-13 | Malfunction |  | During the surgical procedure the customer noticed burn marks on the permanent cautery hook instrument tip. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found signs of arc tracking on the instrument yaw pulley and distal clevis, which may have begun at the instrument gooseneck. This discovery has led engineering to conclude that the instrument arced during the surgical procedure, thus causing the burn marks observed by the customer. The instrument was returned intact, with no pieces missing. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 624188 | 634619 | 603411 | 2955842200500055 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2005-05-13 | 2005-06-10 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-05-13 | Yes | No | Reuse | 2005
40012608 | 2005-05-10 | Malfunction |  | It was reported that during the surgical procedure with permanent cautery hook instrument the amplitude of hood movement was limited in one axis. The planned surgical procedure was completed with the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the yaw pulley crimp had pulled out of the yaw pulley, thus limiting the instrument movement as experienced by the customer. Engineering also discovered signs of arc tracking on the instrument gooseneck and wrist. Material degradation from the electrical arcing may have allowed the yaw pulley crimp to pull out of the yaw pulley. This complaint is being reported due to the evidence of arc tracking found. The arc tracking has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 624488 | 634915 | 603692 | 2955842200500053 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2005-05-10 | 2005-06-09 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-05-10 | Yes | No | Reuse | 2005
IS1000 A33 | 2005-05-24 | Other | Other | During a procedure, the surgeon experienced an intermittent vision problem in the right eye of the stereo viewer. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation by isi field engineering concluded that the right monitor in the high resolution stereo viewer (hrsv) was having intermittent failures. The isi field engineer repaired the hospital's system by replacing the hrsv. The hospital's da vinci vision system passed testing after the repair was performed. As of june 2005, no recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 605780 | 616042 | 585502 | 2955842200500052 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-05-24 | 2005-06-03 | Yes | Yes | Health Professional | IS1000 A33 | No | No | 2005-05-24 | Device Not Retu | No | Reuse | 2005
40017204 | 2005-05-06 | Malfunction |  | It was reported that during a surgical procedure the customer observed firing off of the sides of the wires of the bipolar maryland tip instrument. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Engineering discovered an arc track at the base of the instrument grips, which extends from one grip to the other. The material at the base of the instrument grips is also charred and slightly melted away. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 607711 | 618025 | 587387 | 2955842200500051 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2005-05-06 | 2005-06-02 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | Yes | 2005-05-06 | Yes | No | Reuse | 2005
40012608 | 2005-04-26 | Malfunction |  | It was reported that the permanent cautery hook instrument was dead on arrival. Prior to beginning the surgical procedure, the customer opened the instrument, and sparks shot out of the instrument when it was plugged in. The instrument was not inside the pt when the sparks shot out. No pt harm, adverse outcome or injury was reported. | H6 conclusion: the instrument was returned and evaluated. Engineering found that the permanent cautery hook instrument had been used previously and had 4 lives remaining. Signs of arc tracking beginning at the instrument gooseneck were also discovered. Engineering observed that the fep insulation was damaged, which allowed arcing to occur. No other damage to the instrument was noted. The instrument malfunction occurred during the setup. It is possible that the electrosurgical pedal was depressed while the instrument was being setup. As of may, 2005, there habe been no reported recurrences of this failure mode at this hosp or any other hosp. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 604231 | 614463 | 584037 | 2955842200500050 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-04-26 | 2005-05-26 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-26 | Yes | No | Reuse | 2005
40012608 | 2005-04-20 | Malfunction |  | It was reported that during the beginning of the surgical procedure the permanent cautery hook instrument was smoking inside the patient. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | H10 the instrument was returned and evaluated per the customer reported complaint, engineering found signs of arc tracking which begins at the instrument gooseneck. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 598163 | 608366 | 578210 | 2955842200500044 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2005-04-20 | 2005-05-20 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-20 | Yes | No | Reuse | 2005
40015402 | 2005-04-22 | Malfunction |  | Event Description: --/-- It was reported that the permanent cautery spatula instrument was arching during the surgical procedure. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- H10 the permanent cautery spatula instrument was not returned for failure analysis investigation, therefore, the customer reported complaint cannot be confirmed or denied. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found evidence of arc tracking on the base of the instrument spatula, at the bottom edge of the shrink tubing. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY SPATULA INSTRUMENT | Invalid Data | 598165 | 608368 | 578212 | 2955842200500049 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | InitialFollowup | 2005-04-22 | 2005-05-20 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | Yes | 2005-04-22 | Yes | No |  | 2005
40012608 | 2005-04-21 | Malfunction |  | It was reported that during the surgical procedure the permanent cautery hook instrument cables were broken at the instrument wrist. No patient harm, adverse outcome or injury was reported. | H10 the instrument was returned and evaluated. Engineering found that all of the instrument cables are intact. No cables were broken. Engineering also discovered signs of arc tracking at the instrument gooseneck. This complaint is being reported due to arc tracking found, which has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 598170 | 608373 | 578217 | 2955842200500046 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-04-21 | 2005-05-20 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-04-21 | Yes | No | Reuse | 2005
4001006 | 2005-04-22 | Malfunction |  | It was reported that during a surgical procedure the tip of the precise bipolar pyramid tip instrument melted during prostate dissection. The instrument tip was in contact with the prostate when the instrument tip meleted. The instrument was removed and replaced and the planned surgical procedure was completed. | H10: the instrument was returned and evaluated investigation observed the the plastic was burnt at the instrument grip base and the plastic was burnt at the instrument grip base and the plastic pulleys were melted and charred. Engineering also discovered arc tracking at the base of the instrument grips. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 598281 | 608484 | 578317 | 2955842200500047 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2005-04-22 | 2005-05-20 | No | Yes | Health Professional | 4001006 | Device Returned To Manufacturer | No | 2005-04-22 | Yes | No | Reuse | 2005
40018401 | 2005-04-22 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery spatula instrument was arcing. The instrument was removed and replaced and replaced adn the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, engineering found no signs of arc tracking on the permanent. Cautery spatula, however, engineering observed that the tip of the metal spatula had melted and resolidified. The melting and resolidification of the instrument tip may have led the customer to believe that the instrument arced during the surgical procedure. No definitive conclusion has been determined regarding the cause of the melted instrument tip. Additional details provided by an isi account manager revealed that the voltage setting on the electrosurgical unit (esu) was set to 65 and the esu was set to coagulate. Isi has not received reports of any similar incidents haviang occurred with instruments produced in the same manufacturing lot. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 598284 | 608487 | 578320 | 2955842200500048 | 1 | GEI | Manufacturer | User facility | Physician | Initial | 2005-04-22 | 2005-05-20 | No | Yes | Health Professional | 40018401 | Device Returned To Manufacturer | Yes | 2005-04-22 | Yes | No | Reuse | 2005
IS1200 A43P4 | 2005-04-20 | Other | Other | It was reported that during a surgical procedure the customer experienced no control of the left arm immediately after an instrument change. The customer installed a clip applier instrument, an icon appeared, the instrument spun up and the customer pushed the instrument into use. When the surgeon attempted to take control of the instrument, the instrument would not respond. The customer performed another instrument change with a different clip applier instrument, with the same result. At this point, due to patient bleeding, the surgeon decided to convert the case and performed a sternotomy. Technical support was not called until after the case ws converted and finished. | H10: the investigation conducted by field service engineering found that the system performed as intended. No conclusion can be drawn regarding the cause of the customer reported complaint. As of may 2005, no additional recurrences of this failure mode have been reported at this site. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 598297 | 608497 | 578329 | 2955842200500045 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2005-04-20 | 2005-05-20 | Yes | Yes | Health Professional | IS1200 A43P4 | No | Yes | 2005-04-20 | Device Not Retu | No | Reuse | 2005
40012607 | 2005-04-18 | Malfunction |  | Prior to sterilization it was noticed that the insulation of the permanent cautery hook instrument was cracked. No pt harm, adverse outcomeor injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found signs of arc tracking with begins at the instrument gooseneck. The customer's reference to "cracking" was due to the visual surface appearance of the plastic after arcing. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. There were no missing pieces from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 596233 | 606451 | 576330 | 2955842200500042 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-04-18 | 2005-05-18 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2005-04-18 | Yes | No | Reuse | 2005
40012607 | 2005-04-18 | Malfunction |  | Prior to sterilization it was noticed that the insulation of the permanent cautery hook instrument was cracked. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found signs of arc tracking which begins at the instrument gooseneck. The customer's reference to "cracking" was due to the visual surface appearance of the plastic after arcing. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. There were no missing pieces from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 596241 | 606459 | 576338 | 2955842200500043 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-04-18 | 2005-05-18 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2005-04-18 | Yes | No | Reuse | 2005
40012608 | 2005-04-13 | Malfunction |  | It was reported that during the surgical procedure the permanent cautery hook instrument cracked at the wrist and was arcing while inside the pt. The permanent cautery hook instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported that the instrument was "cracked at the wrist" and was "arcing". Inspection of the instrument found that the "cracks" were actually arc tracks on the surface of the distal pulley and distal clevis. No components of the instrument were missing. It was concluded that the conductor wire kinked and became pinched between the cap and the yaw pulley, thus damaging the insulation on the conductor wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 593267 | 603431 | 573441 | 2955842200500041 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-04-13 | 2005-05-13 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-13 | Yes | No | Reuse | 2005
IS1200 A42 | 2004-05-12 | Malfunction | Other | During a procedure, the da vinci surgical would not home. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finished the surgery. | The isi field engineer determined that the right endoscopic camera manipulator (ecm) had malfunctioned. The hospital's system was repaired by replacing the ecm and the system became fully operational. The ecm was returned to isi for failure analysis. Engineering concluded that a motor in the ecm had malfunctional. The rotator was making contact with the motor stator. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 586824 | 596996 | 567322 | 2955842200400040 | 1 | KOG | Manufacturer | Health Professional | Other | Initial | 2004-05-12 | 2005-05-11 | Yes | Yes | Health Professional | IS1200 A42 | Yes | Yes | 2004-05-12 | Yes | No | Reuse | 2005
40017801 | 2005-03-08 | Malfunction |  | It was reported that when the surgeon went to use the scissors instrument it would not open or close. Upon inspection of the instrument by the intuitive surgical rep on site, it was determined that one scissors blade had fragmented and was missing. Upon exploring the abdomen, the instrument fragment was successfully retrieved and removed from the pt. The damaged instrument was removed and replaced and the planned surgical procedure was completed. The instrument was replaced with an intuitive surgical round tip scissors to continue the case to closing. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Visual examination confirmed that one curved scissors blade had separated from the curved scissors instrument. The fragment was returned with the instrument. Based upon analysis by an independent lab, the material failure experienced is believed to be induced by exposure of the curved scissors instrument (version 400178-01) to non-recommended cleaning methods and chemicals that may accelerate corrosion cracking. Uf report # 0502620000-2005-8006. | CURVED SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 573493 | 603953 | 554720 | 2955842200500039 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-03-08 | 2005-05-09 | No | Yes | Health Professional | 40017801 | Device Returned To Manufacturer | No | 2005-03-08 | Yes | No | Reuse | 2005
40011006 | 2005-04-01 | Malfunction | Other | During a surgical procedure the precise bipolar pyramid tip forceps instrument would not respond. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found a beak in the proximal clevis which caused high friction on the instrument wrist cables, thereby causing the instrument not to respond. Engineering also discovered that a section of the proximal clevis was broken off and not returned with the instrument. This complaint is being reported ude to the fact that it cannot be determined if any components separated from the instrument during the surgical procedure or during handling after the procedure. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 590836 | 601002 | 571183 | 2955842200500033 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-04-01 | 2005-05-06 | No | Yes | Invalid Data | 40011006 | Device Returned To Manufacturer | No | 2005-04-01 | Yes | No | Reuse | 2005
40014206 | 2005-04-06 | Malfunction |  | During a surgical procedure the damaged distal end of the 5mm cannula rubbed against the 5mm monopolar cautery instrument. Due to the rubbing, some "rubber" shaft of the instrument fell into the patient. The surgeon was able to grab most of the "rubber" that fell into the patient and the 5mm cannula and 5mm monopolar cautery instrument were removed and replaced and the procedure was completed. At the end of the procedure it was noticed that the replacement 5mm cannula was damaged/disfigured at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found that a small section of the instrument insulation (which the customer referred to as "rubber") on the main tube was scratched and peeled away. Engineering concluded that the damage to the instrument was clearly due to the use of the damaged cannula. The 5mm cannula malfunction is also being reported via mdr report #2955842-2005-00037. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 590867 | 601034 | 571212 | 2955842200500036 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-04-06 | 2005-05-06 | No | Yes | Health Professional | 40014206 | Device Returned To Manufacturer | No | 2005-04-06 | Yes | No | Reuse | 2005
40012608 | 2005-04-07 | Malfunction |  | It was reported that the customer was told by an isi representative to return the permanent cautery hook instrument. No details regarding the reason for return were provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated and failure analysis investigation found evidence of arc tracking. It was determined that the conductor wire insulation became damaged at the conductor cap due to kinking and then getting pinched. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 591550 | 601730 | 571848 | 2955842200500038 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-04-07 | 2005-05-06 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-07 | Yes | No | Reuse | 2005
40012608 | 2005-04-06 | Malfunction |  | It was reported that prior to beginning the surgical procedure the customer noticed that the permanent cautery hook instrument had a broken tip. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evlauated. Per the customer reported complaint, engineering found signs of arc tracking on the instrument wrist, whcih begins at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUEMNT | 591562 | 601742 | 571858 | 2955842200500035 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-04-06 | 2005-05-06 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-06 | Yes | No | Reuse | 2005
40012608 | 2005-04-06 | Malfunction |  | It was reported that during the beginning of the surgical procedure the surgeon was cleaning tissue off of the permanent cautery hook instrument with a prograsp instrument. During the cleaning the tip of the permanent cautery hook instrument snapped and fell inside the pt, somewhere on top of the bladder. The surgeon was able to remove the tip by using a large needle driver instrument. The permanent cautery hook instrment was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument and detached tip were returned and evaluated. Per the customer reported complaint, it was determined that the user failed to follow the permanent cautery instrument instructions for use (ifu). The ifu specifically states under general precautions and warnings, "do not use an instrument to clean debris off another instrument intraoperatively. This may result in damage to the instruments. An isi account manager was requested to review the ifu with hospital staff in 2005. The following month, this issue has not recurred at the hospital. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 591567 | 601747 | 571862 | 2955842200500034 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-04-06 | 2005-05-06 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-04-06 | Yes | No | Reuse | 2005
31314701 | 2005-04-06 | Malfunction |  | It was reported that during a surgical procedure the damaged distal end of the 5mm cannula rubbed against the 5mm monopolar cautery instrument. Due to the rubbing, some "rubber" shaft of the instrument fell into the pt. The surgeon was able to grab most of the "rubber" that fell into the pt and the 5mm cannula and 5mm monopolar cautery instrument were removed and replaced and the procedure was completed. At the end of the procedure it was noticed that the replacement 5mm cannula was damaged/disfigured at the distal end. No pt harm, adverse outcome or injury was reported. | The 5mm cannula was returned and evaluated. Per the customer reported complaint, failure analysis investigation found that the cannula has a noticeably bent tip. The metal is bent inward creating a smaller than normal opening with a sharp edge that could damage the insulation on 5mm instruments. Engineering concluded that the damage to the cannula was due to customer mishandling. Engineering also concluded that the damage to the 5mm instrument was clearly due to the use of the damage cannula. The 5mm trocar system product information data sheet provides the following warning, "prior to use, inspect the distal tip of the cannula for any dents or tip damage that may have occurred during handling and cleaning. Dents, nicks, or cracks in the distal tip of the cabnnula may increase the potential for cannula fracture, and may render the cannula compatabile with certain instruments. Cannulas with damage tips should to be used. " the 5mm monopolar cautery instrument malfunction is also being reported via mdr report #2955842-2005-00036. | 5MM CANNULA | INSTRUMENT ACCESSORY | 591579 | 601759 | 571874 | 2955842200500037 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2005-04-06 | 2005-05-06 | No | Yes | Health Professional | 31314701 | Device Returned To Manufacturer | No | 2005-04-06 | Yes | No | Reuse | 2005
IS1000 A33 | 2005-04-04 | Other | Other | During a surgical procedure the customer experienced a "temporary motion sensor mismatch" error during tool change and was able to override the error. The customer rebooted the system, however, the error reappeared when a sterile adaptor was installed. The second recurrence of the error could not be overridden by the customer. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The system errors experienced by the customer were associated with the pt side manipulator (psm). The system was repaired by replacing the affected psm. The system alarm functioned as designed. The da vinci system stopped functioning, due to a system malfunction, but gave an appropriate alarm (i. E. Error code visible on the monitor and an audible alarm). As of april 2005, this issue has not recurred at the hospital since the system repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 587809 | 597978 | 568276 | 2955842200500031 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2005-04-04 | 2005-04-29 | Yes | Yes | Health Professional | IS1000 A33 | No | Yes | 2005-04-04 | Device Not Retu | No | Reuse | 2005
40017204 | 2005-04-01 | Malfunction |  | During a surgical procedure the bipolar maryland tip instrument was arcing. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation discovered an arc path between the grips at the base of the instrument. Engineering also observed that the plastic was burned around the base of each grip. It was determined that the instrument arced during the surgical procedure. | BIPOLAR MARYLAND TIP FROCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 588755 | 598927 | 569182 | 2955842200500032 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-04-01 | 2005-04-29 | No | Yes | Health Professional | 40017204 | Device Returned To Manufacturer | No | 2005-04-01 | Yes | No | Reuse | 2005
40012608 | 2005-03-23 | Malfunction |  | It was reported that during the end of the surgical procedure the permanent cautery hook instrument was smoking inside the pt. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, evidence of arc tracking was found, which appears to begin at the instrument gooseneck and continues on to the grip cable. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 585414 | 595588 | 565979 | 2955842200500029 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-03-23 | 2005-04-20 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-03-23 | Yes | No | Reuse | 2005
40012608 | 2005-03-23 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument cannot work in or with some angle movements. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found that the instrument move as intended. Engineering also observed evidence of arc tracking on the instrument wrist, which begins at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 585423 | 595597 | 565986 | 2955842200500028 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2005-03-23 | 2005-04-20 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-03-23 | Yes | No | Reuse | 2005
40012608 | 2005-03-11 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument malfunctioned and became stuck in a right angle. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Failure analysis investigation found no evidence of the customer reported complaint and that the instrument moves as intended. Engineering also observed evidence of arc tracking which begins at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 585427 | 595601 | 565990 | 2955842200500030 | 1 | NAY | Manufacturer | User facilityDistributor | Other Health Care Professional | Initial | 2005-03-11 | 2005-04-20 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-03-11 | Yes | No | Reuse | 2005
40015402 | 2005-03-21 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery spatula instrument wouldn't activate. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, engineering found signs of are tracking all over the instrument wrist, which begins at a break in the cautery conductor wire. This discovery had led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 585707 | 595884 | 566257 | 2955842200500027 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-21 | 2005-04-20 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | Yes | 2005-03-21 | Yes | No | Reuse | 2005
40015402 | 2005-03-04 | Malfunction |  | Event Description: --/-- It was reported that during a surgical precedure the permanent cautery spatula instrument tip was smoking. The instrumnet was replaced and the procedure was continued. No pt harm, odverse outcome or injury was reported. --/-- Additional Event Description: --/-- During a surgical procedure the permanent cautery spatula instrument tip was smoking. The instrument was replaced and the procedure was continued. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was not returned to isi for failure analysis investigation; however, a video of the instrument malfunction was provided by the hospital. A review of the video was conducted by engineering and instrument arcing was observed during the surgical procedure. Because the instrument was not returned, the cause of th arcing cannot be determined. --/-- Additional Manufacturer Narrative: --/-- H10: 400, 67 - the instrument was returned and evaluated. A video of the instrument malfunction was also provided by the hospital. A review of the video was conducted by engineering and instrument arching appeared to have occurred during the surgical procedure, however, during failure analysis investigation of the instrument, engineering found no signs of arching and there was no indication that the insultation was damaged. Engineering also observed that the instrument was returned with considarable burnt blood and tissue attached to the wrist, which the customer may have mistaken for arc tracking. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 583406 | 593561 | 564091 | 2955842200500020 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2005-03-04 | 2005-04-15 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | No | 2005-03-04 | Yes | No |  | 2005
IS1000 A33 | 2005-03-17 | Other | Other | It was reported that during the middle of a surgical procedure the customer lost focus control. Nl o pt harm was reported. The procedure was coverted to traditional open surgical techniques to finish the planned surgery. | H10: the investigation conducted by field service engineering concluded that the loss of focus control experienced by the customer was associated with the system camera. The system was repaired by replacing the affected camera. The camera was returned to isi for failure analysis investigation. Engineering discovered that the camera focus motor was rough and may have resulted in poor focus control. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 583603 | 593763 | 564279 | 2955842200500026 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-17 | 2005-04-15 | Yes | Yes | Health Professional | IS1000 A33 | No | Yes | 2005-03-17 | Device Not Retu | No | Reuse | 2005
40012608 | 2005-03-16 | Malfunction |  | It was reported taht during the start of the surgical procedure the permanent cautery hook instrument was smoking inside the pt. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire was damaged at the observed that the conductor wire was damaged at the idler pulley, causing a break in the insulation. Evidence of arc tracking was found, which begins at the break in the insulation. This discovery has lead engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 583613 | 593773 | 564288 | 2955842200500025 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-16 | 2005-04-15 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-03-16 | Yes | No | Reuse | 2005
40018301 | 2005-03-04 | Malfunction |  | Event Description: --/-- It was reported that during a surgical procedure the permanent cautery hook instrument tip was smoking. The instrument was replced and the procedure was continued. No pt harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- It was reported that during a surgical procedure the permanent cautery hook instrument tip was smoking. The instrument was replaced and the procedure was continued. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- H10: the instrument was not returned to isi for failure analysis investigation; however , a video of the instrument malfunction was provided by the hosp. A review of the video has led engineering to conclude that a high amount of heat was generated in the proximal clevis area of the instrument, which created the smoke noticed by the customer. Because the instrument was not returned, the cause of the heat generated cannot be determined. --/-- Additional Manufacturer Narrative: --/-- H10: 47 - the instrument was returned and evaluated. A video of the instrument malfunction was also provided by the hospital. During a review of the video, engineering observed that a high amount of heat was generated in the proximal clevis area of the insrtument. Which created the smoke noticed by the customer. During failure analysis investigation of the insrtument, engineering discovered an arc track beginning at a break in the conductor cap. Engineering has determined that the conductor wire may have kinked prior to the insulation becomming damaged. This discovery has led engineering to conclude that the instrument arced during the surgical procedure, which generated the high amount of heat in the proximal clevis area. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 583627 | 593787 | 564302 | 2955842200500019 | 1 | GEI | Manufacturer | Company Representative | Other | InitialFollowup | 2005-03-04 | 2005-04-15 | No | Yes | Health Professional | 40018301 | Device Returned To Manufacturer | No | 2005-03-04 | Yes | No |  | 2005
ISI000 A33 | 2005-03-07 | Other | Other | During a surgical procedure the customer experienced a constant "209" system error code. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the "209" system error experienced by the customer was associated with the multiple servo driver (msd) pca. The system was repaired by replacing the affected msd pca. The msd pca was returned to isi for failure analysis investigation. Engineering discovered that the msd pca had a capacitor that malfunctioned, which caused the system to fault. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 579816 | 589967 | 560714 | 2955842200500022 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2005-03-07 | 2005-04-08 | Yes | Yes | Health Professional | ISI000 A33 | No | Yes | 2005-03-07 | Device Not Retu | No | Reuse | 2005
40012608 | 2005-03-11 | Malfunction |  | During a surgical procedure the permanent cautery hook instrument was giving off smoke, making it difficult to see. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, engineering found evidence of arc tracking which began at the instrument gooseneck. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 581700 | 591838 | 562519 | 2955842200500024 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-11 | 2005-04-08 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-03-11 | Yes | No | Reuse | 2005
40012608 | 2005-03-10 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument was arcing and the instrument was cauterizing above where it should be. A hospital representative stated that they did not notice any fracture on the instrument clevis nor anything else unusual about the instrument prior to returning it to isi. No patient harm, adverse outocme or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found evidence of an arc track beginning at the instrument gooseneck. It was determined that the instrument arced during the surgical procedure. Also during the failure analysis investigation, it was noticed that a section of the proximal clevis was missing and not returned with the instrument. It cannot be determined if any components separated from the instrument during the surgical procedure or during handling after the procedure. A hospital representative stated that they did not notice any fracture on the instrument clevis nor anything else unusual about the instrument prior to returning it to isi. The relationship between the electrical arcing and the proximal clevis fragment cannot be conclusively determined. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 581939 | 592092 | 562752 | 2955842200500023 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-10 | 2005-04-08 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-03-10 | Yes | No | Reuse | 2005
40015402 | 2005-03-06 | Malfunction |  | Event Description: --/-- During a surgical procedure the wrist of the permanent cautery spatula instrument broke. The instrument was inserted into the cannula and the wrist broke during its initial movement. Upon breakage, two components detached from the instrument and fell into the pt. Both components were retrieved from the pt and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was not returned to isi for failure analysis investigation, therefore, the cause of the instrument malfunction cannot be determined. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis broke at the base of the instrument ear. All pieces of the instrument were returned. | PERMANENT CAUTERY SPATULA INSTRUMENT | Invalid Data | 579865 | 590016 | 560761 | 2955842200500021 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Physician | InitialFollowup | 2005-03-06 | 2005-04-05 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | Yes | 2005-03-06 | Yes | No |  | 2005
40017204 | 2005-03-02 | Malfunction |  | It was reported that during a surgical procedure the bipolar maryland tip instrument became stuck in the trocar and the customer broke the instrument shaft when trying to remove it. Additional details provided to the isi account manager from the customer revealed that the customer was using the instrument with no problems when it seized up on them. The customer had to remove the trocar from the pt's abdomen to then retrieve the instrument from the trocar. During this process, a piece of the instrument fell into the pt's abdomen and was retrieved. The customer felt that no additional pieces had fallen into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was not returned to isi for failure analysis investigation. A description of the instrument malfunction provided by the hospital has led engineering to conclude that the proximal clevis of the instrument had fractured and a section of the proximal clevis detached from the instrument. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 579086 | 589238 | 559999 | 2955842200500018 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-03-04 | 2005-04-01 | No | Yes | Health Professional | 40017204 | No | No | 2005-03-04 | Device Not Retu | No | Unkown | 2005
40012608 | 2005-03-02 | Malfunction |  | It was reported that during a surgical procedure a cloud of smoke was coming from the plastic portion of the permanent cautery hook instrument, located before the hook area of the instrument. The instrument was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found evidence of arc tracking. It was determined that the conductor wire insulation became damaged at the conductor cap due to kinking and then getting pinched. This discovery has led engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 579127 | 589278 | 560040 | 2955842200500017 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2005-03-02 | 2005-04-01 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-03-02 | Yes | No | Reuse | 2005
40012607 | 2005-03-02 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument was arcing at the wrist. After the procedure was completed, it was noticed that the instrument was charred at the housing. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found evidence of arc tracking at the instrument gooseneck, causing the instrument proximal clevis (which the customer referred to as the "housing") to appear charred. This discovery leads engineering to conclude that the instrument arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 579133 | 589284 | 560046 | 2955842200500013 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-03-02 | 2005-04-01 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2005-03-02 | Yes | No | Reuse | 2005
40012608 | 2005-02-22 | Malfunction |  | Distributor reported after a surgical procedure that two abdominal incisions made for instrument cannula appeared cauterized. No patient harm, adverse outcome or permanent injury was reported. | The instrument was not returned for failure analysis investigation, therefore, no difinitive conclusion has been determined regarding the cause of the incision cauterization at the cannula ports. Additionally, no information has been provided regarding the voltage setting used and whether the esu was set to dessicate, fulgurate, or spray. Isi has not received report of any similar incidents having occurred with instruments produced in the same manufacturing lot. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 576799 | 586951 | 557806 | 2955842200500016 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2005-02-22 | 2005-03-25 | No | Yes | Health Professional | 40012608 | No | No | 2005-02-22 | Device Not Retu | No | Unkown | 2005
40012608 | 2005-02-25 | Malfunction |  | It was reported that during a surgical procedure the tip of the permanent cautery hook instrument burned. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation observed that the plastic at the instrument gooseneck showed a little bit of melting, however, no evidence of the instrument having arced was found. Engineering also noticed that the shrink sleeving on the instrument hook was missing and the conductor wire was kinked at the conductor cap. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 576808 | 586960 | 557815 | 2955842200500014 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-02-25 | 2005-03-25 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-02-25 | Yes | No | Reuse | 2005
40012608 | 2005-02-25 | Malfunction |  | It was reported that during a surgical procedure the tip of the permanent cautery hook instrument burned. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found evidence of an arc track at the instrument gooseneck. It was also noticed that the instrument shrink tubing was completely gone and the arc tracking began at the interface where the shrink tubing should have been. Engineering determined that the instrument arced during the surgical procedure, which caused the instrument tip to appear burned. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 577441 | 587594 | 558431 | 2955842200500015 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-02-25 | 2005-03-25 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-02-25 | Yes | No | Reuse | 2005
40012608 | 2005-02-11 | Malfunction |  | After a surgical procedure it was noticed that the tip of the permanent cautery hook instrument was burnt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation found evidence of an arc track beginning at the instrument gooseneck. It was also noticed that the conductor wire was kinked at the cap. Engineering determined that the instrument arced during the surgical procedure, which caused the instrument tip to appear burnt. No info could be provided regarding whether the instrument hook was being used to apply heat to another instrument. No deficiencies in instrument performance were alleged. The instrument damaged was recognized by the hosp after the planned surgery was completed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 575880 | 586032 | 556925 | 2955842200500012 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-02-15 | 2005-03-17 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-02-15 | Yes | No | Reuse | 2005
| 2005-03-08 | Malfunction |  | A patient was in the or for a robotic prostatectomy. The physician went to use the scissor instrument and noted that one of the blades was missing. Upon exploring the abdomen, the broken blade of the scissor was retrieved successfully. The broken piece matched the rest of the instrument. The scissor (both pieces) was removed from the sterile field and terminalized in the autoclave. The event was reported to intuitive surgical since the curved scissor is a new instrument in their inventory. The scissor that broke today was used only once before and that was yesterday. No problems were noticed that day. Surgery progressed as scheduled. |  | DA VINCI | CURVED SCISSORS ROBOTIC SURGIC | 573493 | 583666 | 554720 | 583666 | 1 | NAY | User Facility |  |  | Initial | 2005-03-14 | 2005-03-14 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2005
40012608 | 2005-02-11 | Malfunction |  | It was reported that during a surgical procedure the tip of the permanent cautery hook instrument began smoking. The instrument was removed from the pt and the planned surgery was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Per the customer reported complaint, faiure analysis investigation found evidence of an arc track beginning at the instrument gooseneck. It was determined that the instrument arced during the surgical procedure, which generated the smoke noticed by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 573443 | 583616 | 554673 | 2955842200500011 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-02-11 | 2005-03-11 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-02-11 | Yes | No | Reuse | 2005
40017801 | 2005-03-01 | Malfunction |  | It was reported that during a surgical procedure the tip of the 8mm curved scissors instrument broke off inside the pt. The surgeon was able to retrieve and remove the instrument tip and complete the planned surgical procedure with a replacement instrument. No pt harm, adverse outcome or injury was reported. | Visual examination confirmed that one curved scissors blade had separated from the curved scissors instrument. Investigation of the root cause of the component separation is ongoing. A market withdrawal of all distributed curved scissors instruments was initiated on 3/2005. An initial analysis conducted by an independent lab identified trace elements in the cross section of the parts separated, inconsistent to the stainless steel composition, leading to the conclusion that the instruments were exposed to chemicals other than the recommended enzymatic cleaners during the cleaning and sterilization processing. The external lab identified stress corrosion cracking as the primary root cause for the part separation. | 8MM CURVED SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 573502 | 583675 | 554729 | 2955842200500010 | 1 | KOG | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-03-01 | 2005-03-11 | No | Yes | Health Professional | 40017801 | Device Returned To Manufacturer | Yes | 2005-03-01 | Yes | No | Reuse | 2005
40009304 | 2005-01-31 | Malfunction |  | It was reported that during a surgical procedure the wrist of the endowrist prograsp instrument broke. No details were provided regarding how the instrument was being use. The hospital reported that the prograsp instrument was not in the field of vision when the wrist broke. No patient harm, adverse outcome or injury was reported. The procedure was converted to non-robotic laparoscopic techniques to finish the planned surgery. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation idscovered that the instrument had broken at the main tube to clevis interface. A section of the proximal clevis was broken off and not returned with the instrument. It cannot be determined if any components separated from the instrument during the surgical procedure or duringhandling after the procedure. Standard surgical techniques require that caution be taken when any surgical instrument is not under direct or endoscopic visualization. The hospital acknowledged that the instrument damage occurred while the instrument was outside the field of view. It is believed that user handling may have contributed to the event. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 571349 | 581497 | 552702 | 2955842200500009 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2005-01-31 | 2005-03-02 | No | Yes | Health Professional | 40009304 | Device Returned To Manufacturer | Yes | 2005-01-31 | Yes | No | Initial | 2005
400007115 | 2005-01-31 | Malfunction |  | It was reported that during a surgical procedure the wrist of the rount tip scissors instrument broke. No patient harm, adverse outcome or injury was reported. The procedure was converted to non-robotic laparoscopic techniques to finish the planned surgery. | The instrument was returned abd evaluated. Failure analysis investigation discovered that a section of the proximal clevis was broken off and not returned with the instrument. It cannot be determined if any components separated from the instrument during the surgical procedure or during handling after the procedure. | ROUND TIP SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 571354 | 581502 | 552706 | 2955842200500008 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2005-01-31 | 2005-03-02 | No | Yes | Health Professional | 400007115 | Device Returned To Manufacturer | Yes | 2005-01-31 | Yes | No | Initial | 2005
IS1200 A43P4 | 2005-02-04 | Other | Other | During a surgical procedure the customer experienced recurring "temporary motion sensor mismatch" errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered that the lubrication used on the mtmr was contaminated, causing the system to fault. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 568692 | 578845 | 550205 | 2955842200500007 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-02-04 | 2005-02-23 | Yes | Yes | Health Professional | IS1200 A43P4 | No | No | 2005-02-04 | Device Not Retu | No | Reuse | 2005
IS1000 A33 | 2005-01-28 | Other | Other | During a surgical procedure the customer experienced a problem with image depth perception. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering found a damaged endoscope at the customer site. The lack of 3-d depth perception was likely due to a focus problem with the endoscope. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 568706 | 578859 | 550218 | 2955842200500006 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-01-28 | 2005-02-23 | Yes | Yes | Health Professional | IS1000 A33 | No | No | 2005-01-28 | Device Not Retu | No | Reuse | 2005
40012608 | 2005-01-25 | Malfunction |  | Housing for the permanent cautery hook instrument tip was cracked. The instrument was inside the pt throughout the surgical procedure, however, the crack was not noticed until the procedure was completed. It was also reported that no components were missing from the housing, nor did any pieces fall inside the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation revealed that the instrument proximal clevis (which the customer referred to as the "housing") is not cracked, however, engineering discovered arc track lines on the clevis that make it appear to be cracked. This discovery had led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 568763 | 578916 | 550275 | 2955842200500005 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2005-01-25 | 2005-02-23 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2005-01-25 | Yes | No | Reuse | 2005
40012608 | 2005-01-20 | Malfunction |  | It was reported that the instrument began to smoke during a procedure. The instrument was removed from the patient and the surgery was completed with another permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. The surgical procedure was not significantly lengthened due to the instrument exchange. | The instrument was returned and evaluated. Engineering found evidence of an arc track at the gooseneck of the instrument. It was determined that the instrument may have arced and thus generated the smoke noticed by the customer. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 567599 | 577760 | 549119 | 2955842200500002 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-01-20 | 2005-02-18 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2005-01-20 | Yes | No | Reuse | 2005
40017203 | 2005-01-18 | Malfunction |  | It was reported that the instrument was not articulating properly. No other details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering determined that damage to a cable caused the poor instrument articulation. In addition, engineering found evidence of insulation degradation at the base of the yaw pulleys. One of the yaw pulleys had fragmented and was not returned. | BIPOLAR MARYLAND TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 567605 | 577766 | 549125 | 2955842200500003 | 1 | GEI | Manufacturer | User facility | Other Health Care Professional | Initial | 2005-01-18 | 2005-02-18 | No | Yes | Health Professional | 40017203 | Device Returned To Manufacturer | Yes | 2005-01-20 | Yes | No | Reuse | 2005
40009304 | 2005-01-20 | Malfunction |  | It was reported that the tip of the instrument was broken. No details were provided regarding whether the instrument was damaged during surgery, instrument handling, or during cleaning and sterilization. No patient harm was reported. No adverse outcome or injury was reported. | Failure analysis of the returned instrument determined that the instrument's proximal clevis had fractured at the distal end of the instrument. The clevis fragment was not returned for analysis. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 567619 | 577780 | 549139 | 2955842200500004 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2005-01-20 | 2005-02-18 | No | Yes | Health Professional | 40009304 | Yes | Yes |  | Yes | No | Reuse | 2005
40012608 | 2004-12-30 | Malfunction |  | Wires of the permanent cautery hook instrument were exposed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found evidence of an arc track on the instrument wrist. It was determined that the conductor wire insulation became damaged due to a kink in the conductor wire. The damage to the insulation exposed the conductor wire and allowed the instrument to arc. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 563703 | 573834 | 545397 | 2955842200400140 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-12-30 | 2005-02-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-12-30 | Yes | No | Reuse | 2005
40012608 | 2004-12-17 | Malfunction |  | Permanent cautery hook instrument was "bad. " no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated and evidence of an arc track was found under the conductor wire cap. Engineering determined that the conductor insulation was damaged due to the wire kinking under the conductor wire cap. The damage to the insulation exposed the conductor wire and allowed the instrument to arc. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 564173 | 574307 | 545854 | 2955842200400139 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2004-12-17 | 2005-02-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-12-17 | Yes | No | Reuse | 2005
IS1200 A43P4 | 2005-01-13 | Other | Other | During a surgical procedure the customer experienced "temporary motion sensor mismatch" errors. The customer restarted the system, however experienced the same errors when trying to home the system. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer were associated with the right master tool manipulator (mtmr). The system was repaired by replacing the affected mtmr. The mtmr was returned to isi for failure analysis investigation. Engineering discovered that the gear clamp on axis 5 was loose and the gears were disengaged. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 558981 | 569133 | 540807 | 2955842200500001 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2005-01-13 | 2005-02-01 | Yes | Yes | Health Professional | IS1200 A43P4 | No | No | 2005-01-13 | Device Not Retu | No | Reuse | 2005
40012608 | 2004-12-29 | Malfunction |  | It was reported that the permanent cautery hook instrument began smoking. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated and evidence of an arc track was found. Engineering determined that the conductor wire became pinched at the conductor wire cap, which damaged the instrument insulation and allowed the instrument to arc. This discory was led engineering to conclude that the instrument may have arced. The customer observation of the instrument "began smoking" was likely the customer's interpretation of the arching event. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 562964 | 573096 | 544689 | 2955842200400138 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-12-29 | 2005-02-01 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-12-29 | Yes | No | Reuse | 2005
40011005 | 2004-12-13 | Malfunction |  | During a surgical procedure the jaw of the precise bipolar pyramid tip instrument broke. Upon breakage, one of the graspers detached from the instrument and was retrieved. It was also reported that the instrument insulation was missing. No pt harm, adverse outcome or injury was reported. | The instrument was not returned and no info could be provided on how the instrument was being used, therefore the customer reported complaint cannot be confirmed or denied. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 557943 | 568115 | 539808 | 2955842200400135 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-12-13 | 2005-01-14 | No | Yes | Health Professional | 40011005 | No | No | 2004-12-13 | Device Not Retu | No | Unkown | 2005
40015402 | 2004-11-18 | Malfunction |  | It was reported that during a surgical procedure the side yoke holding the tip of the permanent cautery spatula instrument broke off. Additional information provided by clinical specialist indicated that no components separated from the instrument during the surgical procedure, however, components did separate from the instrument during instrument handling. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation discovered that the post on the yaw pulley was sheared off. The conductor wire cap and broken ear on the distal clevis were not returned with the instrument. It cannot be determined if any components separated from the instrument during a surgical procedure. This complaint is being reported due to the fact that if this failure mode were to recur, patient safety may be compromised. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 555426 | 565637 | 537412 | 2955842200400134 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-12-08 | 2005-01-07 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | No | 2004-12-08 | Yes | No | Reuse | 2005
40012608 | 2004-12-01 | Malfunction |  | The permanent cautery hook instrument was leaking. No further problems were alleged. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation found signs of arc tracking at the base of the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 553760 | 564026 | 535839 | 2955842200400132 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-12-01 | 2004-12-30 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-12-01 | Yes | No | Reuse | 2004
40012608 | 2004-12-02 | Malfunction |  | During a surgical procedure the permanent cautery hook instrument was arcing. The instrument was replaced and the procedure was completed. No pt harm,adverse outcome or injury was reported. | The instrument was not returned, therefore the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 553770 | 564036 | 535849 | 2955842200400133 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-12-02 | 2004-12-30 | No | Yes | Health Professional | 40012608 | No | No | 2004-12-02 | Device Not Retu | No | Unkown | 2004
40012608 | 2004-12-01 | Malfunction |  | The permanent cautery hook instrument was broken. No further problems were alleged. No pt harm,adverse outcome or injury was reported. | Failure analysis investigation found signs of minor arcing at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 553774 | 564040 | 535853 | 2955842200400131 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-12-01 | 2004-12-30 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-12-01 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-12-01 | Other | Other | It was reported that during a surgical procedure the customer experienced a system mismatch error. The customer restarted the system, but experienced another error when trying to home the system. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the system errors experienced by the customer was associated with the left master tool manipulator (mtml). The system was repaired by replacing the affected mtml. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 553478 | 563748 | 535573 | 2955842200400130 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-12-01 | 2004-12-29 | Yes | Yes | Health Professional | IS1000 A33 | No | No | 2004-12-01 | Device Not Retu | No | Reuse | 2004
| 2004-12-09 | Malfunction |  | Instrument broke while in use in pt's pelvis. Area thoroughly investigated and all pieces apparently retrieved. Upon later inspection of equipment a very small piece of insulation not seen and possible, though unlikely, retained in pt. |  | DE VINCI ENDOWRIST PRECISE BIPOLAR |  | 554253 | 564502 | 536308 | MW1034279 | 1 | GEI | Voluntary |  | Physician | Initial | 2004-12-13 | 2004-12-22 | No | Yes | Invalid Data |  | Yes | Yes |  |  |  |  | 2004
40012607 | 2004-11-22 | Malfunction |  | The permanent cautery hook instrument was not working. No further problems were alleged. No patient harm, adverse outcome or injruy was reported. | The instrument was returned and evaluated. Failure analysis investigation found evidence of arce tracking beginning at the instrument cable crimp. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 731016 | 743221 | 708233 | 2955842200400129 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-11-22 | 2004-12-21 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-11-22 | Yes | No | Reuse | 2004
40011006 | 2004-11-15 | Malfunction | Other | It was reported that the precise bipolar pyramid tip forceps instrument moved in a non-intuitive manner. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found no defects. The instrument was recognized and functioned normally, when installed on a da vinci system. All motions were intuitive and as expected. The allegation that the instrument moved non-intuitively could not be confirmed. No conclusion can be drawn regarding whether user error, system malfunction or instrument malfunction may have contributed to the event. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 776971 | 789293 | 753174 | 2955842200400128 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-11-15 | 2004-12-15 | No | Yes | Invalid Data | 40011006 | Device Returned To Manufacturer | No | 2004-11-15 | Yes | No | Reuse | 2004
40015403 | 2004-11-10 | Malfunction |  | It was reported that during a surgical procedure the permanet cautery spatula instrument began to arc. The instrument was removed and replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation discovered that the conductor insulation was damaged due to kinking under the conductor wire cap. Engineering determined that the damage to the insulation exposed the conductor wire and allowed the instrument to arc. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 782808 | 795208 | 758966 | 2955842200400126 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-11-10 | 2004-12-10 | No | Yes | Health Professional | 40015403 | Device Returned To Manufacturer | Yes | 2004-11-10 | Yes | No | Reuse | 2004
40015402 | 2004-11-12 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery spatula instrument was arcing. The instrument was removed and replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation discovered arc tracking at the cable crimp in the yaw pulley. Engineering determined that the cable crimp came into contact with the body of the spatula within the molded section and this contact allowed the instrument to arce. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 785511 | 797978 | 761664 | 2955842200400127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-11-12 | 2004-12-10 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | Yes | 2004-11-12 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-11-08 | Other | Other | It was reported that during a surgical procedure the system sustained multiple recoverable faults. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the recoverable system errors experienced by the customer were associated with the endoscopic camera manipulator (ecm). The system was repaired by replacing the affected ecm. The system was tested and passed the applicable dance and functional tests. As of 2004, no additional recurrences of this failure mode have been reported at this site. | DE VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 549066 | 559358 | 531431 | 2955842200400125 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2004-11-08 | 2004-12-07 | Yes | Yes | Health Professional | IS1000 A33 | No | Yes | 2004-11-08 | Device Not Retu | No | Reuse | 2004
IS1000 A33 | 2004-10-26 | Other | Other | It was reported that during a surgical procedure the customer experienced a green and grainy image problem. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The investigation conducted by field service engineering concluded that the green and grainy image experienced by the customer was due to low light transmission from the light guides. The system was repaired by replacing the affected light guides. As of 2004, no additional recurrences of this failure mode have been reported. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 548318 | 558623 | 530729 | 2955842200400120 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2004-10-26 | 2004-11-24 | Yes | Yes | Health Professional | IS1000 A33 | No | Yes | 2004-10-26 | Device Not Retu | No | Reuse | 2004
40012608 | 2004-10-27 | Malfunction |  | It was reported that the permanent cautery hook instrument was broken. No further problems were alleged. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis discovered that a piece of the proximal clevis circumference was broken off and not returned with the instrument. There is no indication from the customer that the missing fragment separated from the instrument during a procedure. Failure analysis investigation also found evidence of arc tracking on the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. This case is being reported due to the evidence that the instrument may have arced during previous surgical use. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 793480 | 805989 | 769541 | 2955842200400121 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-10-27 | 2004-11-24 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-10-27 | Yes | No | Reuse | 2004
40012608 | 2004-10-29 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument was arcing. The instrument was removed and replaced and the procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusion: the instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation discovered that the conductor wire was kinked inside the conductor cap and that the instrument insulation had been damaged. The exposure of the conductor wire allowed the instrument to arc. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 793491 | 805999 | 769551 | 2955842200400122 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-10-29 | 2004-11-24 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-10-29 | Yes | No | Reuse | 2004
40012608 | 2004-11-01 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument was arcing. The instrument was removed and replaced and the procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, failure analysis investigation discovered that the conductor wire was kinked inside the conductor cap and that the instrument insulation had been damaged. The exposure of the conductor wire allowed the instrument to arc. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 793503 | 806011 | 769563 | 2955842200400123 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-11-01 | 2004-11-24 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-11-01 | Yes | No | Reuse | 2004
40009304 | 2004-10-27 | Malfunction |  | It was reported that the pin came out of a grasping instrument. The pin was pushed back in by hand by a hosp employee. Customer cannot provide info regarding when the event occurred, with which surgical procedure or with which surgeon. No adverse event or injury were reported. | Conclusion: the returned instrument was evaluated by isi engineering. It was determined that an mfg error during instrument assembly contributed to the device malfunction. All instrument components were returned to isi. | PROGRASP ENDOWRIST INSTRUMENT | ENDOSCOPIC INSTRUMENT | 793595 | 806103 | 769655 | 2955842200400124 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-10-27 | 2004-11-24 | No | Yes | Invalid Data | 40009304 | Device Returned To Manufacturer | Yes | 2004-10-27 | Yes | No | Reuse | 2004
40012607 | 2004-10-19 | Malfunction |  | It was reported that the permanent cautery hook instrument had a broken cable. No further problems were alleged. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis confirmed that the pitch cable was broken at a location where it enters the proximal clevis. Failure analysis investigation also found evidence of an arc track on the distal and proximal clevis. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure, and the arcing may have caused the damage to the pitch cable. This case is being reported due to the evidence that the instrument may have arced during previous surgical use. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 795728 | 808243 | 771721 | 2955842200400118 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-10-19 | 2004-11-18 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-10-19 | Yes | No | Reuse | 2004
40012608 | 2004-10-15 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument was smoking. The instrument was removed and replaced, and the procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation found evidence of an arc track at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. The customer observation of the instrument '"was smoking" was likely the customer's interpretation of the arcing at the gooseneck of the monopolar hook. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 795740 | 808255 | 771732 | 2955842200400119 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-10-19 | 2004-11-18 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-10-19 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-10-14 | Injury | Required Intervention | During a minimally invasive surgical procedure, a small tear occurred in the anterior portion of the superior vena cava (svc) when dissecting the svc pericardial reflection. It appeared that a portion of the svc was pinched between the instrument jaws when the surgeon was attempting to blunt dissect through the pericardial reflection. A small right thoracotomy was made to control the bleeding from the svc. No system or instrument malfunction were reported. It was reported that the pt recovered and was released from the hosp. | The investigation conducted by isi clinical development engineering determined that the pt had pericardial adhesions, which is a contraindication for minimally invasive surgery. Isi clinical development engineering concluded that no product problems (defects or malfunctions) contributed to the adverse event. The system and instrument operated as expected. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 544622 | 554986 | 527346 | 2955842200400116 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-10-14 | 2004-11-12 | Yes | No | Health Professional | IS1000 A33 | No | No | 2004-10-14 | Device Not Retu | No | Reuse | 2004
IS10000 A33 | 2004-10-18 | Other | Other | During a surgical procedure the customer experienced critical fault errors. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. | The investigation conducted by field service engineering concluded that the critical fault errors experienced by the customer was due to the master tool manipulator (mtm), a non-patient contact system component, being damaged. The system alarm (i. E. Critical fault) performed as designed to indicate that the system had been damaged. The system was repaired by replacing the affected mtm. As of november, 2004, no additional recurrences of this failure mode have been reported. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 544668 | 555030 | 527389 | 2955842200400117 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2004-10-18 | 2004-11-12 | Yes | Yes | Health Professional | IS10000 A33 | No | Yes | 2004-10-18 | Device Not Retu | No | Reuse | 2004
40011005 | 2004-10-06 | Malfunction |  | A precise bipolar pyramid tip instrument was returned to isi. No instrument problems were alleged and no details regarding the instrument return were provided. No pt harm, adverse outcome or injury was reported. | The instrument was evaluated. Failure analysis investigation found evidence of arc tracking on one grip. This discovery has led engineering to believe that the instrument may have arced during a previous surgical procedure or may have been used in conjunction with another electrosurgical instrument which may have arced to the bipolar forceps. Since no details regarding how the instrument was being used could be provided, no definite conclusion can be drawn. | PRECISE BIPOLAR PYRAMID TIP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 806295 | 818811 | 782119 | 2955842200400115 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2004-10-06 | 2004-11-12 | No | Yes | Health Professional | 40011005 | Device Returned To Manufacturer | Yes | 2004-10-06 | Yes | No | Reuse | 2004
IS1200 A4 3P4 | 2004-10-21 | Other | Other | It was reported that during a surgical procedure the customer experienced a critical error which forced them to restart the de vinci surgical system. After restarting the system, the system failed to home. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The investigation conducted by field service engineering concluded that the critical error experienced by the customer was due to the patient side manipulator (psm), a system component, being damaged. Field service engineering determined that the psm axis 1 cable was broken, thus preventing the system from homing. The system alarm (i. E. Critical fault) performed as designed to indicate that the system has been damaged. The system was repaired by replacing the affected psm. The psm was returned io isi for failure analysis investigation. Engineering confirmed that the axis 1 2nd cables were broken and frayed. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 544262 | 554634 | 527002 | 2955842200400114 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2004-10-21 | 2004-11-10 | Yes | Yes | Health Professional | IS1200 A4 3P4 | No | Yes | 2004-10-21 | Device Not Retu | No | Reuse | 2004
40012608 | 2004-10-04 | Malfunction | Other | It was reported that the permanent cautery hook instrument moved in a non-intuitive manner. It has been alleged that the non-intuitive instrument movement caused patient bleeding. The surgeon chose to continue the case with the same permanent cautery hook instrument. It is not believed that the instrument was energized when the non-intuitive movement occurred. The alleged non-intuitive movement of the instruments was experienced 2-3 minutes into the procedure. No permanent injury was reported. The patient was reported as being "o. K. " and the surgeon chose to continue using the same instrument during the surgery. | The allegation that the instrument moved non-intuitively could not be confirmed. No conclusion can be drawn regarding whether user error, system malfunction or instrument malfunction may have contributed to the event. The hospital did not return the instrument until the instrument and system associated with this complaint had been used in 3 additional surgeries. Inspection of the returned instrument revealed damage at the proximal clevis/instrument shaft interface. The clevis damage interferes with instrument function and precludes further evaluation of whether a user error, system malfunction or instrument malfunction contributed to the event. Isi has concluded that the instrument damage could not have been present and its current presence is unrelated to the alleged event. Isi cannot determine if the instrument damage occurred during handling at the hospital or during shipping. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 811039 | 823634 | 786821 | 2955842200400113 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-10-04 | 2004-11-08 | Yes | Yes | Invalid Data | 40012608 | Device Returned To Manufacturer | No | 2004-10-04 | Yes | No | Reuse | 2004
40012608 | 2004-10-06 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument sparked every time the surgeon put his foot on the pedal to initiate use of the instrument. The instrument was removed and replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation confirmed that the instrument had previously arced due to evidence of an arc track at the gooseneck. The customer observation of the instrument "sparked" was likely the customer's interpretation of the arcing at the gooseneck of the monopolar hook. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 809356 | 821947 | 785155 | 2955842200400109 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2004-10-06 | 2004-11-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-10-06 | Yes | No | Reuse | 2004
40012608 | 2004-10-07 | Malfunction |  | It was reported that the permanent cautery hook instrument was faulty, in that it displayed 6 uses left instead of 7. No further problems were alleged. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis confirmed that the instrument displayed 6 uses remaining. This failure mode is not associated with a fda reportable event. Failure analysis investigation also found evidence of an arc track starting at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 809373 | 821963 | 785171 | 2955842200400110 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-10-07 | 2004-11-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-10-07 | Yes | No | Reuse | 2004
40011006 | 2004-09-13 | Malfunction |  | It was reported that the permanent bipolar forceps instrument was broken. No patient harm, adverse outcome or injury was reported. | The bipolar forceps instrument was returned and evaluated. Engineering determined that the plastic insulation showed evidence of burning at the base of the two metal forcep grips. From additional investigation by the isi clinical specialist, it was concluded that arcing likely occurred when a monopolar cautery instrument was energized in close proximity to the affected bipolar forceps instrument. The energized monopolar cautery instrument likely arced to the nearby bipolar forceps instrument and damaged the plastic insulation at the base of the metal grips. Medical personnel at the hospital have been instructed to keep instruments (e. G. The bipolar forceps instrument) at least 1 centimeter away from an energized monopolar cautery instrument to prevent arcing. | PRECISE BIPOLAR PYRAMID TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 809388 | 821978 | 785186 | 2955842200400112 | 1 | GEI | Manufacturer | Company Representative | Other Health Care Professional | Initial | 2004-10-05 | 2004-11-04 | No | Yes | Health Professional | 40011006 | Device Returned To Manufacturer | Yes | 2004-09-13 | Yes | No | Reuse | 2004
40012608 | 2004-10-08 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument began to arc. The instrument was removed and replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation discovered that the wire was partially kinked in the wrist and came into contact with the drive cable. Consequently, the instrument insulation was damaged due to the contact with the cable. Engineering concluded that the insulation damage was evidence that the instrument had arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 809441 | 822032 | 785239 | 2955842200400111 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-10-08 | 2004-11-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-10-08 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-09-30 | Other | Other | During a procedure the surgeon experienced a vision synchronization issue with the right eye of the camera. It was noted by the isi sales rep present that both the right and left cameras were not properly set at the front panel before the case had started. The sales rep properly adjusted the cameras before the case started. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The investigation conducted by field service engineering concluded that the loss of synchronization was due to a miswired cable. Since no cable wiring or camera adjustment issues were a problem when the system was used to perform procedure two days prior, it was determined that the system had been tampered with at the hosp. This hosp has a known sabotage issue. The hosp typically stores the system in its own locked room. The hosp's system was repaired by correcting the cable that had been miswired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 542696 | 553118 | 525544 | 2955842200400107 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-10-29 | 2004-10-29 | Yes | Yes | Health Professional | IS1000 A33 | No | No | 2004-09-30 | Device Not Retu | No | Reuse | 2004
IS1200 A43 | 2004-09-30 | Other | Other | While the system was being prepared for surgery, the power cord was inadvertently unplugged. The unplugged power cord was not noticed until the system backup battery ran down while the system was being docked to the pt. The hosp was unable to power back on the system until the battery had been recharged. The surgeon chose to convert to traditional open surgical techniques to perform the planned surgery. No pt harm was reported. | Inadvertant error by hosp staff during system setup caused the system backup to run down. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 542733 | 553155 | 525579 | 2955842200400106 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-10-29 | 2004-10-29 | Yes | Yes | Health Professional | IS1200 A43 | No | No | 2004-09-30 | Device Not Retu | No | Reuse | 2004
40012602 | 2004-10-01 | Malfunction |  | It was reported that the permanent cautery hook instrument was broken. No other details were provided. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Engineering determined that during prior use, the electrical insulation at the instrument wrist had begun to degrade. Engineering concluded that the insulation degradation was evidenced that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 868739 | 891702 | 854056 | 2955842200400108 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-10-29 | 2004-10-29 | No | Yes | Health Professional | 40012602 | Device Returned To Manufacturer | Yes | 2004-10-01 | Yes | No | Reuse | 2004
40012608 | 2004-09-23 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument was arcing. The instrument was replaced and the procedure was completed. No pt harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Engineering determined that the insulation on the conductor wire was damaged due to kinking of the wire in and around the conductor wire cap. This observation leads engineering to conclude that the insulation damage was evidence that the instrument had arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 877526 | 902198 | 864491 | 2955842200400104 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-23 | 2004-10-22 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-23 | Yes | No | Reuse | 2004
40012608 | 2004-09-23 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument was arcing. The instrument was replaced and the procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was not returned therefore, the customer reported complaint cannot be confirmed or denied. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 877536 | 902264 | 864556 | 2955842200400105 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-24 | 2004-10-22 | No | Yes | Health Professional | 40012608 | No | No | 2004-09-24 | Device Not Retu | No | Unkown | 2004
40012608 | 2004-09-14 | Malfunction |  | It was reported that the permanent cautery hook instrument arced when the suction irrigator was placed close to it. No patient harm, adverse outcome or injury was reported. The hospital has reported 5 separate incidents of arcing that occurred in 2004. Ref. : mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | The instrument was returned and evaluated. Engineering cannot confirm whether or not the instrument had arced during the surgical procedure. Additional investigation has found that the arcing occurred only when the cautery instrument was energized and the suction irrigator had been placed less than 1 cm from the energized permanent cautery hook instrument. Medical personnel at the hospital have been instructed to keep the suction irrigator at least 1 centimeter away from the permanent cautery hook instrument to prevent arcing. Ref: mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 865712 | 888302 | 850673 | 2955842200400102 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-14 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-14 | Yes | No | Reuse | 2004
40012608 | 2004-09-14 | Malfunction |  | It was reported that the permanent cautery hook instrument arced when the suction irrigator was placed close to it. No patient harm, adverse outcome or injury was reported. The hospital has reported 5 separate incidents of arcing that occurred on the following dates: fives times in 2004. Ref. : mdr 2955845-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | The instrument was returned and evaluated. Engineering cannot confirm whether or not the instrument had arced during the surgical procedure. Additional investigation has found that the arcing occurred only when the cautery instrument was energized, and the suction irrigator had been placed less than 1 cm from the energized permanent cautery hook instrument. Medical personnel at the hospital have been instructed to keep the suction irrigator at least 1 centimeter away from the permanent cautery hook instrument to prevent arcing. Ref: mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102 and mdr 2955842-2004-00103. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 865717 | 888305 | 850676 | 2955842200400099 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-14 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-14 | Yes | No | Reuse | 2004
40012608 | 2004-09-14 | Malfunction |  | It was reported that the permanent cautery hook instrument arced when the suction irrigator was placed close to it. No patient harm, adverse outcome or injury was reported. The hospital has reported 5 separate incidents of arcing that occurred with 5 different instruments. The incidents occurred on the following dates: five times in 2004. Ref. : mdr 2955845-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | The instrument was returned and evaluated. During failure analysis investigation, engineering found evidence of an arc track beginning at the gooseneck. This observation confirms the customer reported complaint that the instrument had previously arced during a surgical procedure. Additional investigation has found that the arcing occurred only when the cautery instrument was energized, and the suction irrigator had been placed less than 1 cm from the energized permanent cautery hook instrument. Medical personnel at the hospital have been instructed to keep the suction irrigator at least 1 centimeter away from the permanent cautery hook instrument to prevent arcing. Ref: mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102 and mdr 2955842-2004-00103. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 866216 | 888783 | 851161 | 2955842200400100 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-14 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-14 | Yes | No | Reuse | 2004
40012607 | 2004-09-14 | Malfunction |  | It was reported that the permanent cautery hook instrument arced when the suction irrigator was placed close to it. No patient harm, adverse outcome or injury was reported. The hospital has reported 5 separate incidents of arcing that occurred with 5 different instruments. The incidents occurred on the following dates: five times in 2004. Ref. : mdr 2955845-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | The instrument was returned and evaluated. During failure analysis investigation, engineering found evidence of an arc track. It was also noticed that the shrink sleeve and most of the plastic gooseneck of the instrument were gone. These obvservations confirm the customer reported complaint that the instrument had previously arced during a surgical procedure. Additional investigation has found that the arcing occurred only when the cautery instrument was energized and the suction irrigator had been placed less than 1 cm from the energized permanent cautery hook instrument. Medical personnel at the hospital have been instructed to keep the suction irrigator at least 1 centimeter away from the permanent cautery hook instrument to prevent arcing. Ref: mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102 and mdr 2955842-2004-00103. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 866225 | 888788 | 851166 | 2955842200400103 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-14 | No | Yes | Health Professional | 40012607 | Yes | No | 2004-09-14 | Yes | No | Reuse | 2004
40012608 | 2004-09-14 | Malfunction |  | It was reproted that the permanent cautery hook instrument arced when the suction irrigator was placed close to it. No pt harm, adverse outcome or injury was reported. The hosp has reported 5 separate incidents of arcing that occurred with 5 different instruments. The incidents occurred on the following five days in 2004. Ref. : mdr 2955842-2004-00099, mdr 2955842-2004-00100, mdr 2955842-2004-00101, mdr 2955842-2004-00102, and mdr 2955842-2004-00103. | Conclusions - the instrument was returned and evaluated. During failure analysis investigation, engineering found evidence of an arc track beginning at the idler pulley. This observation confirms the customer reported complaint that the instrument had previously arced during a surgical procedure. Additional investigation has found that the arcing occurred only when the cautery instrument was energized and the suction irrigator had been placed less than 1 cm from the energized permanent cautery hook instrument. Medical personnel at the hosp have been instructed to keep the suction irrigator at least 1 centimeter away from the permanent cautery hook instrument to prevent arcing. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 867357 | 889740 | 852110 | 2955842200400101 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-14 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-14 | Yes | No | Reuse | 2004
IS1000 | 2004-09-17 | Malfunction | Other | It was reported that during a procedure, the surgeon experienced an intermittent signal loss on the assistant monitor and in the left eye of the surgeon's console. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The investigation conducted by field service engineering concluded that the vision system synchronizer had failed. The hospital's system was repaired by replacing the synchronizer. The hospital's da vinci vision system passed testing after the repair was performed. As of october 13, 2004, this issue has not occurred at the hospital since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 879575 | 904783 | 867036 | 2955842200400098 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-09-17 | 2004-10-14 | Yes | Yes | Health Professional | IS1000 | Yes | No | 2004-09-17 | Yes | No | Reuse | 2004
IS1000 | 2004-09-14 | Malfunction | Other | It was reported that during a procedure the surgeon experienced no vision in the left eye of the camera. No patient harm was reported. The procedure was converted to traditional open surgical technique to finish the surgery. | The investigation conducted by field service engineering concluded that the camera head and camera controller. The hospital's system was repaired by replacing the camera head and camera controller failed. The hospital's da vinci vision system passed testing after the repairs were performed. The camera head and camera controller were returned and evaluated. Failure analysis found an intermittent short in the camera head cable, which caused the fuse to blow in the camera controller. As of october 7, 2004, this issue has not recurred at the hospital since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 875216 | 899177 | 861483 | 2955842200400097 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-09-14 | 2004-10-08 | Yes | Yes | Health Professional | IS1000 | Yes | No | 2004-09-14 | Yes | No | Reuse | 2004
40012608 | 2004-09-09 | Malfunction |  | It was reported that the permanent cautery hook instrument housing is cracked. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering determined that the conductor wire was kinked at the wrist and became pinched between the conductor cap and distal clevis, causing a break in the insulation. Engineering also concluded that the break in the insulation was evidence that the instrument had arced during a previous surgical procedure. The customer observation of "housing is cracked" was likely the customer's interpretation of the housing at the tip of the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 868907 | 892043 | 854397 | 2955842200400096 | 1 | GEI | Manufacturer | Health Professional | Nurse | Initial | 2004-09-09 | 2004-10-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-09-09 | Yes | No | Reuse | 2004
40012608 | 2004-09-14 | Malfunction |  | It was reported that the permanent cautery hook instrument was broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering determined that during prior use, the electrical insulation at the instrument wrist had begun to degrade. Degradation of the insulation reduced the mechanical stability of the wrist assembly components and allowed the cable crimp to come off the yaw pulley. This failure code caused the instrument to become non-functional and substantiates the customer complaint that the instrument was broken. Engineering also concluded that the insulation degradation was evidence that the instrument had arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 868911 | 892045 | 854399 | 2955842200400093 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2004-09-14 | 2004-10-04 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-09-14 | Yes | No | Reuse | 2004
40012607 | 2004-09-14 | Malfunction |  | It was reported that during a surgical procedure the permanent cautery hook instrument broke. The procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation found evidence of an arc track at the instrument gooseneck. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 868914 | 892046 | 854400 | 2955842200400092 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2004-09-14 | 2004-10-04 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-09-14 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-08-30 | Malfunction | Other | It was reported that a system fault occurred while exchanging instruments on the system's left arm. The system was powered down and would not restart. No patient harm was reported. It was reported that the surgeon had just entered the patient's abdomen when the system fault occurred. The surgeon chose to stop the procedure and close the patient's incisions. | The investigation by isi field engineering concluded that a multiple servo driver (msd) printed circuit assembly had failed. The hospital's system was repaired by replacing the faulty msd. The hospital's da vinci system passed testing after the repair was performed. The msd was evaluated by engineering. It was determined that an internal short within the pca fabrication was causing an over current fault. As of october 5, 2004, this issue has not returned at the hospital since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 869297 | 892290 | 854639 | 2955842200400095 | 1 | KOG | Manufacturer | User facility | Other | Initial | 2004-09-07 | 2004-10-04 | Yes | Yes | Health Professional | IS1000 A33 | Device Returned To Manufacturer | No | 2004-09-07 | Yes | No | Reuse | 2004
40012608 | 2004-08-12 | Malfunction |  | It was reported that the "plastic part at end of hook has come off". Follow-up correspondence with the hospital has revealed that the instrument damage occurred when the surgeon attempted to use the hook instrument to clean charred tissue off another surgical instrument. No information has been provided regarding whether the plastic component was retrieved. No patient harm, adverse outcome or injury was reported. | As of september 30, 2004 the instrument has not been returned to isi for evaluation. Discussion with the site revealed that misuse of the instrument occurred. During the surgical procedure, the permanent cautery hook instrument was used to clean tissue from another surgical instrument. Misuse of the instrument likely contributed to the alleged instrument damage. The instrument ifu specifically states: "do not use an instrument to clean debris off another instrument intraoperatively. This may result in damage to the instruments. ". | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 889439 | 916535 | 878769 | 2955842200400094 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2004-08-12 | 2004-10-04 | No | Yes | Health Professional | 40012608 | No | Yes | 2004-08-12 | Device Not Retu | No | Unkown | 2004
40011005 | 2004-08-25 | Malfunction | Other | It was reported that the electrical connector on the back of the instrument broke. A backup bipolar forceps was used until the electrical connector on the back of the second instrument broke. Neither instrument malfunction constituted an fda reportable event. The surgery was converted to an open procedure since no other appropriate backup instrument was available. No additional surgical procedure was required. No patient injury or adverse event was reported. | Failure analysis of the returned instrument confirmed that the electrical connector on the back of the instrument was changed. Damage to the electrical connector likely occurred while the user was disconnecting the instrument from the electrosurgical unit (esu). It is unlikely the damage occurred while the instrument was in use. The damage to the electrical connector caused the instrument's cautery function to become non-operational. The da vinci user manual explicity states that "environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. ". | BIPOLAR FORCEPS PYRAMID TIP | ELECTROSURGICAL INSTRUMENT | 890434 | 917620 | 879850 | 2955842200400090 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-08-25 | 2004-10-04 | Yes | Yes | Health Professional | 40011005 | No | Yes | 2004-08-25 | Yes | No | Reuse | 2004
40011005 | 2004-08-25 | Malfunction |  |  |  | BIPOLAR FOECEPS PYRAMID TIP | ELECTROSUTGICAL INSTRUMENT | 901993 | 929627 | 891760 | 2955842200400091 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-08-25 | 2004-10-04 | No | Yes | Health Professional | 40011005 | No | Yes | 2004-08-25 | Yes | No | Initial | 2004
40008304 | 2004-07-21 | Malfunction |  | A distributor reported that the instrument tip broke during a surgical procedure. No additional details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned without the ultrasonic probe. Failure analysis confirmed the ultrasonic shears instrument to have been damaged. An inspection of the instrument revealed that the polymer grip section had melted where the grip and ultrasonic probe meet. A thin 0. 05" x 0. 5" fragment of the grip was missing. The evaluation concluded that the instrument was likely misused. The observed instrument damage is known to occur when the instrument is energized and there is no tissue between the grip and the probe. | ULTRASONIC SHEARS | ULTRASONIC SHEARS | 897320 | 925368 | 887542 | 2955842200400089 | 1 | NAY | Manufacturer |  | Other | Initial | 2004-07-21 | 2004-09-29 | No | Yes | Invalid Data | 40008304 | Device Returned To Manufacturer | No |  | Yes | No | Reuse | 2004
40012608 | 2004-08-25 | Malfunction |  | It was reported that during a surgical procedure, the permanent cautery hook instrument sparked. The instrument was replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation determined that the instrument had previously arced. Evidence of an arc track was observed at the gooseneck. The customer observation of "the instrument sparked" was likely the customer's interpretation of the arcing at the gooseneck of the monopolar hook. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 897328 | 925372 | 887546 | 2955842200400088 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-09-01 | 2004-09-29 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-09-01 | Yes | No | Reuse | 2004
40012607 | 2004-07-29 | Malfunction |  |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 922460 | 951053 | 913076 | 2955842200400081 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-08-12 | 2004-09-03 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-08-12 | Yes | No | Reuse | 2004
40012607 | 2004-08-05 | Malfunction |  |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 922646 | 951404 | 913427 | 2955842200400080 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-08-11 | 2004-09-03 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-08-11 | Yes | No | Initial | 2004
40012608 | 2004-08-20 | Malfunction |  |  |  | PERMANENT CAUTERY BOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 927448 | 956883 | 918718 | 2955842200400082 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-08-20 | 2004-09-03 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-08-20 | Yes | No | Reuse | 2004
IS1200 A43P2 | 2004-08-06 | Malfunction |  |  |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 928649 | 958411 | 920225 | 2955842200400083 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2004-08-06 | 2004-09-03 | No | Yes | Health Professional | IS1200 A43P2 | Yes | Yes | 2004-08-06 | Yes | No | Reuse | 2004
IS1000 | 2004-07-23 | Malfunction | Other |  |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 864657 | 887046 | 849421 | 2955842200400074 | 1 | ACJ | Manufacturer | Health Professional | Physician | Initial | 2004-07-23 | 2004-08-20 | No | Yes | Health Professional | IS1000 | No | Yes | 2004-07-23 | Yes | No | Reuse | 2004
AESOP 3000 | 2003-07-15 | Malfunction |  |  |  | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 864664 | 887060 | 849435 | 2955842200400078 | 1 | GCJ | Manufacturer | Distributor | Other | Initial | 2003-07-15 | 2004-08-20 | No | Yes | Health Professional | AESOP 3000 | No | No | 2003-07-15 | Yes | No An | Reuse | 2004
40012607 | 2004-06-25 | Malfunction | Other |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 864672 | 887078 | 849452 | 2955842200400075 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2004-07-06 | 2004-08-20 | Yes | Yes | Health Professional | 40012607 | No | Yes | 2004-07-06 | Device Not Retu | No | Unkown | 2004
40011005 | 2004-06-25 | Malfunction | Other |  |  | PRECISE BIPOLAR PYRAMID TIP FORCEPS | ELECTROSURGICAL INSTRUMENT | 864683 | 887105 | 849477 | 2955842200400076 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2004-07-06 | 2004-08-20 | Yes | Yes | Health Professional | 40011005 | No | Yes | 2004-07-06 | Device Not Retu | No | Unkown | 2004
40000402 | 2004-07-09 | Malfunction |  |  |  | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 864688 | 887164 | 849536 | 2955842200400077 | 1 | GEI | Manufacturer | Distributor | Other | Initial | 2004-07-09 | 2004-08-20 | No | Yes | Health Professional | 40000402 | No | No | 2004-07-09 | Device Not Retu | No | Unkown | 2004
40012607 | 2004-07-27 | Malfunction |  | It was reported that during a surgical procedure the instrument broke. No pt harm. Adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation found evidence of arcing on the gooseneck. This discovery has led engineering to conclude the instrument may have arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 615913 | 626290 | 595236 | 2955842200400071 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-07-27 | 2004-08-16 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-07-27 | Yes | No | Reuse | 2004
40012607 | 2004-07-27 | Malfunction |  | It was reported that during a surgical procedure the instrument cautery flow was not working. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis investigation found evidence of arc tracking on the distal clevis and yaw pulley. This discovery has led engineering to conclude that the instrument may have arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELCTROSURGICAL INSTRUMENT | 615926 | 626303 | 595249 | 2955842200400070 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-07-27 | 2004-08-16 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-07-27 | Yes | No | Reuse | 2004
40012607 | 2004-08-05 | Malfunction |  | It was reported that during a surgical procedure the instrument would not respond. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis confirmed that the instrument had zero live remaining. This failure mode is not associated with a fda reportable event. Additional investigation found evidence of arc tracking on the yaw pulley. This discovery has led engineering to conclude that the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELCTROSURGICAL INSTRUMENT | 615935 | 626312 | 595258 | 2955842200400069 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-08-05 | 2004-08-16 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-08-05 | Yes | No | Reuse | 2004
IS1200 | 2004-07-21 | Malfunction | Other |  |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 878202 | 903251 | 865544 | 2955842200400073 | 1 | GCJ | Manufacturer | Health Professional | Other | Initial | 2004-07-21 | 2004-08-16 | Yes | Yes | Health Professional | IS1200 | Yes | No | 2004-07-21 | Device Not Retu | No | Reuse | 2004
40012607 | 2004-07-22 | Malfunction |  |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 878210 | 903284 | 865577 | 2955842200400066 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-07-22 | 2004-08-16 | No | Yes | Health Professional | 40012607 | No | No |  | Device Not Retu | No | Unkown | 2004
40012607 | 2004-07-12 | Malfunction |  |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 878217 | 903286 | 865579 | 2955842200400067 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-07-12 | 2004-08-16 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-07-12 | Yes | No | Initial | 2004
40012608 | 2004-07-22 | Malfunction |  |  |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 878225 | 903289 | 865582 | 2955842200400068 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-07-22 | 2004-08-16 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | No | 2004-07-22 | Yes | No | Reuse | 2004
40012608 | 2004-07-26 | Malfunction |  | During a surgical procedure, the instrument tip burnt and melted. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis discovered that the conductor wire insulation was compromised in the service loop area and evidence of arc tracking was found. This discovery had led engineering to conclude that the instrument may have arced during the surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 634592 | 645104 | 613559 | 2955842200400065 | 1 | NAY | Manufacturer | User facility | Physician | Initial | 2004-07-26 | 2004-08-13 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-08-02 | Yes | No | Reuse | 2004
40012608 | 2004-07-16 | Malfunction |  | During a procedure, the cables at the end of the instrument were bulging out and about to break. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis confirmed that the pitch cable is off of the back end idler pulleys, causing slack that is noticed at the wrist. This failure mode is not associated with an fda reportable event. Also during failure analysis, it was discovered that a fragment of the proximal clevis was missing and was not returned with the instrument. There is no indication from the customer that the missing fragment separated from the instrument during a procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 634601 | 645113 | 613568 | 2955842200400061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-07-16 | 2004-08-13 | No | Yes | Health Professional | 40012608 | Device Returned To Manufacturer | Yes | 2004-07-22 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-07-21 | Malfunction | Other | During a procedure, the surgeon experienced an intermittent vision problem in the right eye of the stereo viewer. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The investigation by field engineering concluded that the right monitor in the hrsv was having intermittent failures. The hospital's system was repaired by replacing the hrsv. The hospital's da vinci vision system passed testing after the repair was performed. As of aug 2004, this issue has not recurred at the hospital since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 693068 | 703973 | 670001 | 2955842200400063 | 1 | KOG | Manufacturer | Health Professional | Other | Initial | 2004-07-21 | 2004-08-13 | Yes | Yes | Health Professional | IS1000 A33 | Yes | No | 2004-07-21 | Yes | No | Reuse | 2004
40011006 | 2004-07-12 | Malfunction |  | An instrument broke. The customer does not know how it was broken, but believes that it was discovered prior to the procedure and that the procedure was completed with another instrument. No further details were provided. No patient harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. The investigation found evidence of an arc track at the base of the grip to the distal clevis. This arc track has led engineering to conclude the instrument may have arced during a previous surgical procedure. | PRECISE BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 693111 | 704015 | 670042 | 2955842200400062 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-07-12 | 2004-08-13 | No | Yes | Health Professional | 40011006 | Device Returned To Manufacturer | Yes | 2004-07-12 | Yes | No | Initial | 2004
400004 | 2004-07-17 | Malfunction |  | During a procedure, a cautery spatula accessory separated from the scalpel cautery instrument and fell into the patient. The cautery spatula accessory was retrieved with no patient injury. No patient harm, adverse outcome or injury was reported. | The instrument was not returned to isi for evaluation. However, upon investigation by a company representative, it was discovered that during the surgical procedure, the physician pulled on the tip of the cautery spatula accessory to clean it, and the tip came off. The company's instructions for use manual for endowrist instruments specifically states: "do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. ". | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 693152 | 704057 | 670082 | 2955842200400064 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2004-07-17 | 2004-08-13 | No | Yes | Health Professional | 400004 | No | Yes |  | Device Not Retu | No | Unkown | 2004
40012607 | 2004-03-24 | Malfunction |  | During a surgical procedure it was noticed that the distal end of the instrument would not rotate through it's full range of motion and that the instrument tip was broken. The instrument was returned to the manufacturer for analysis. No patient harm was reported. No adverse outcome or injury was reported. | Failure analysis of the returned instrument confirmed that the instrument would not rotate through it's full range of motion and that the instrument tip was broken. Engineering determined that the proximal clevis and instrument shaft were no longer aligned thus restricting the instrument range of motion. In addition, it was determined that the proximal clevis had fractured at the distal end of the instrument. A small fragment of the proximal clevis was missing and was not returned with the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 607530 | 617842 | 587216 | 2955842200400060 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-03-24 | 2004-08-05 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-03-24 | Yes | No | Initial | 2004
IS1200 | 2004-06-21 | Malfunction | Other | The da vinci surgical system experienced a fault condition when the operation room personnel exchanged instruments to be used with the system. No patient harm was reported. The procedure was converted to traditional open surgical techniques to complete the planned surgery. | The isi field engineer determined that the system fault was due to a system component having been damaged. The system alarm (i. E. System fault) performed as designed to indicate the system had been damaged. The isi field engineer determined that aggressive handling of the system by operating room personnel damaged the patient side manipulator (psm). The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis. Engineering confirmed that the cables in the psm had come off their corresponding pulleys. The cables had no visual defects. Engineering concluded that the user likely dislodged the cables accidentally. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 607520 | 617832 | 587207 | 2955842200400059 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-06-21 | 2004-08-03 | Yes | Yes | Health Professional | IS1200 | No | Yes | 2004-06-21 | Yes | No | Reuse | 2004
40012607 | 2004-06-10 | Malfunction |  | An instrument broke during a procedure and that there was no harm to the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis determined that the instrument had previously arced. No components were reported missing from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 606183 | 616452 | 585886 | 2955842200400057 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-10 | 2004-08-02 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-06-10 | Yes | No | Reuse | 2004
40012607 | 2004-06-10 | Malfunction |  | An instrument broke during a procedure and that there was no harm to the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis determined that the instrument had previously arced. No compoments were reported missing from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 606190 | 616459 | 585893 | 2955842200400056 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-10 | 2004-08-02 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-06-10 | Yes | No | Reuse | 2004
40012607 | 2004-07-14 | Malfunction |  | An instrument broke during a procedure and that there was no harm to the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. As per the customer reported complaint, failure analysis confirmed that the instrument housing, a non-patient contact portion of the instrument, had fractured. This failure mode is not associated with a fda reportable event. Additional investigation found evidence of an arc track at the instrument tip that has led engineering to conclude the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 606195 | 616464 | 585898 | 2955842200400058 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-07-14 | 2004-08-02 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-07-14 | Yes | No | Reuse | 2004
40012607 | 2004-06-25 | Malfunction |  | During a surgical procedure it was noticed that the distal end of the instrument had fractured. The surgeon located and removed an instruement fragment from the patient. The instrument and fragment were returned to the manufacturer for analysis. No patient harm was reported. No adverse outcome or injury was reported. | Failure analysis of the returned instrument and fragment confirmed that the proximal clevis had fractured at the distal end ot the instrument. It was also determined that a second small fragment, 3mmx4mm, was not returned for analysis. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 605499 | 615743 | 585246 | 2955842200400053 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-25 | 2004-07-29 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-06-25 | Yes | No | Reuse | 2004
40012607 | 2004-06-25 | Malfunction |  | It was reported that an instrument arced during a procedure. The instrument was replaced with another permanent cautery hook instrument and the procedure was completed without harm to the patient. No patient harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Evidence of an arc track at the cable crimp was observed. Failure analysis confirmed that the instrument had previously arced. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 605503 | 615747 | 585250 | 2955842200400052 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-06-25 | 2004-07-29 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-06-25 | Yes | No | Reuse | 2004
40012607 | 2004-05-27 | Malfunction |  | It was reported that the instrument was not cutting or coagulating well. The surgeon completed the surgical procedure with the instrument although its performance was alleged to be suboptimal. No pt harm, adverse outcome or injury was reported. | H10: the instrument was returned and evaluated. Evidence of an arc track on the gooseneck portion of the yaw pulley was observed. Failure analysis concluded that the instrument had likely previously arced. The energy shunted through this alternate path would explain why the instrument was not cutting or coagulating tissue well. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 605946 | 616214 | 585659 | 2955842200400055 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-27 | 2004-07-29 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-05-27 | Yes | No | Reuse | 2004
IS1000 | 2004-07-16 | Malfunction | Other | During surgery, the surgeon experienced a system fault while an instrument was being inserted into the da vinci system. The surgeon chose to convert the procedure to an open surgical procedure. No pt injury or adverse outcome was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | H10: upon review of the system error logs, isi field engineering determined that an intermittent issue with the mtm and the msd was likely initiating the system fault. The hosp's system was repaired by replacing both the mtm and the msd. The potentially defective components were returned to intuitive surgical for further eval. Failure analysis investigation at intuitive surgical determined that a broken motor wire in the mtm was only making intermittent contact. Engineering confirmed that intermittent electrical contact was the likely source of the customer reported system fault. As of july 2004, this issue has not recurred at the hosp since the system was repaired. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 605953 | 616221 | 585665 | 2955842200400054 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-07-16 | 2004-07-29 | Yes | Yes | Health Professional | IS1000 | No | No | 2004-07-16 | Yes | No | Reuse | 2004
40012607 | 2004-06-11 | Malfunction |  | It was reported that an instrument broke during a procedure. No further details were provided. No pt harm, adverse outcome or injury was reported. | H10: 47 the instrument was retuned and evaluated. No components were reported missing from the instrument. Evidence of an arc track on the instrument wrist was observed. Failure analysis concluded that the instrument had previously arced. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 602005 | 612195 | 581925 | 2955842200400051 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2004-07-12 | 2004-07-16 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-06-11 | Yes | No | Reuse | 2004
IS1000 A33 | 2004-06-08 | Other | Other | It was reported that the system performed "sluggish". The site was advised by an isi rep via phone to re-insert the instrument and sterile adapter, but it was unclear if this operation was performed by the site. The customer decided to convert the case. The customer reported that the system was set-up away from the pt after the conversion, and the system performed properly. | One pt side manipulator (psm) arm of the da vinci system was returned for eval. Although the psm was determined to have low cable tension, the customer reported symptom is not related to the low cable tension. Engineering has concluded that the customer reported symptom of "sluggish" movement was likely due to interference (between arms, cables, drapes) or improper engagement of the instrument or sterile adapter. As of 2004, no add'l complaints of "sluggish" movement have been reported by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 524405 | 535155 | 508071 | 2955842200400049 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-06-08 | 2004-07-09 | Yes | Yes | Health Professional | IS1000 A33 | Yes | Yes | 2004-06-08 | Yes | No | Reuse | 2004
40012607 | 2004-06-08 | Malfunction |  | It was reported that during a surgical procedure the instrument froze at the "neck" and the joint where the wrist attaches to the shaft was broken. After cleaning of the instrument at the hosp, it was noted that a piece was missing and never found. | H10: 47- the instrument was returned for evaluation. Failure analysis of the returned instrument had fractured and separated from the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROCAUTERY INSTRUMENT | 602013 | 612203 | 581932 | 2955842200400050 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2004-06-08 | 2004-07-09 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-06-08 | Yes | No | Reuse | 2004
40012607 | 2004-06-02 | Malfunction |  | It was reported that an instrument broke during a procedure and that there was no harm to the patient. No further details were provided. | H10: 47 - the instrument was returned and evaluated. Failure analysis determined that the instrument had arced. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 594375 | 604588 | 574495 | 2955842200400048 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2004-06-02 | 2004-07-01 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-06-02 | Yes | No | Reuse | 2004
40012607 | 2004-06-01 | Malfunction |  | It was noticed by the surgeon that the instrument arced during a procedure. The instrument was removed and replaced to complete the procedure robotically. No patient harm, adverse outcome or injury was reported. | H10: 47 - the instrument was returned and evaluated. Failure analysis confirmed evidence that the instrument had arced. An arc track was visible at the yaw pulley interface. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 594380 | 604593 | 574500 | 2955842200400047 | 1 | NAY | Manufacturer | Health Professional | BIOMEDICAL ENGINEER | Initial | 2004-06-01 | 2004-07-01 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-06-01 | Yes | No | Reuse | 2004
40000417 | 2004-06-01 | Malfunction |  | The customer noticed a crack in the instrument tip after opening the product prior to a procedure. | Failure analysis of the returned instrument confirmed evidence that the yaw pulley was cracked. In addition engineering detected evidence that the instrument had likely arced during a previous use. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 594384 | 604597 | 574504 | 2955842200400046 | 1 | NAY | Manufacturer | Health Professional | Other Health Care Professional | Initial | 2004-06-01 | 2004-07-01 | No | Yes | Health Professional | 40000417 | No | Yes | 2004-06-01 | Yes | No | Reuse | 2004
40012607 | 2004-05-26 | Malfunction |  | It was alleged that the instrument wrist arced during a procedure. The instrument was removed and replaced to complete the procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis confirmed evidence that the instrument had arced. An arc track was visible at the yaw pulley interface. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 591264 | 601443 | 571574 | 2955842200400044 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-26 | 2004-06-25 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-05-26 | Yes | No | Reuse | 2004
40012608 | 2004-05-26 | Malfunction |  | The instrument wrist arced during a procedure. Although the instrument was alleged to have arced, the surgeons chose to continue to use the instrument to complete the case with no patient injury. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for analysis. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 591371 | 601550 | 571678 | 2955842200400045 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-26 | 2004-06-25 | No | Yes | Health Professional | 40012608 | No | No | 2004-05-26 | Device Not Retu | No | Unkown | 2004
40015402 |  | Malfunction |  | Prior to surgery it was observed that a loop of wire protruded from the tip of the permanent spatula instrument. In spite of the presence of the protruding wire loop, the site choose to use the instrument in a surgery since: 1) the instrument was on it's 7th of 8 lives and 2) the hosp believed the loop was present during previous surgeries and had caused no problems. It was reported that part way through the ensuing surgery, the permanent spatula instrument began to "disintegrate. " the instrument was removed and replaced to complete the procedure robotically. It was reported that there were no injuries to the pt. | The instrument was returned and evaluated. Failure analysis determined that no pieces were missing from the instrument and that all components of the instrument remained intact. Engineering discovered evidence of an arc track on the wrist and tip of the instrument. It has been concluded that a loop of the conductor wire came in contact with the device cable. Movement of the drive cable wore through the insulation on the conductor cable allowing the instrument to arc. Root cause analysis has determined that the device malfunction was caused by a new failure mode. There have been no known previous occurrences of this failure mode. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 591912 | 602084 | 572161 | 2955842200400042 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2004-05-04 | 2004-06-22 | No | Yes | Health Professional | 40015402 | Device Returned To Manufacturer | No | 2004-05-04 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | The instrument was returned with the complaint that the insulation was "rubbing off". The instrument was not being used in a surgical procedure when the damaged insulation was noticed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis confirmed that a portion of the insulation was missing. Evidence of an arc track has led engineering to believe the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 588295 | 598458 | 568741 | 2955842200400041 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-20 | 2004-06-17 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | No | 2004-05-05 | Yes | No | Reuse | 2004
IS1000A33 | 2004-05-12 | Malfunction | Other | During a procedure , the surgeon experienced diffculty while manipulating a surgical instrument with da vinci surgical system. The surgeon perceived force feedback when he attempted to roll the instrument. Furthermore, the instrument did not always respond to the movement of the da vinci surgical system masters. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The isi field engineer determined that the right master tool manipulator (mtm) had malfunctioned. The hospital's system was repaired by replacing the right mtm. The mtm was returned to isi for failure analysis. Engineering concluded that during initial assembly of the mtm that clamping screw in the assembly may not have been seated properly. Wear marks on the screw indicate that the clamping screw may have gripped the shaft with in-line contact versus the larger intended surfact contact as designed. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 587023 | 597189 | 567507 | 2955842200400039 | 1 |  | Manufacturer | Health Professional | Physician | Initial | 2004-05-12 | 2004-06-15 | Yes | Yes | Health Professional | IS1000A33 | Yes | Yes | 2004-05-21 | Yes | No | Reuse | 2004
40012607 | 2004-04-20 | Injury | Required Intervention | Three days after a laparoscopic sigmoidectomy was performed at a hosp, an injury of the pt's ureter was discovered. At the time of the surgery: 1) no malfunctions of the instrument or surgical system were observed or reported and 2) no arcing of the instrument was observed or reported. After discovery of the injury, it was alleged that the permanent cautery hook instrument may have arced during the surgical procedure. The alleged defective instrument was discarded by the hosp and cannot be provided to intuitive surgical for testing. | No failure analysis evaluation of the permanent cautery hook instrument can be performed. The alleged defective instrument was discarded by the hosp. As of june 2004, no conclusion can be made regarding whether surgeon error or instrument malfunction contributed to the pt's injury. Intuitive surgical is continuing to investigate this incident. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 518931 | 529703 | 502850 | 2955842200400035 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2004-05-05 | 2004-06-09 | Yes | Yes | Health Professional | 40012607 | No | Yes | 2004-05-05 | Device Not Retu | No | Unkown | 2004
40012306 |  | Malfunction |  | It was reported that the instrument was not working. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Evidence of an arc track has led engineering to believe the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 584132 | 594305 | 564780 | 2955842200400036 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2004-02-09 | 2004-06-09 | No | Yes | Health Professional | 40012306 | Device Returned To Manufacturer | Yes | 2004-02-09 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | It was reported that the instrument was borken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Evidence of an arc track has led engineering to believe the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 584136 | 594310 | 564784 | 2955842200400037 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-02-23 | 2004-06-09 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-02-23 | Yes | No | Initial | 2004
40012607 |  | Malfunction |  | The instrument was returned for credit. No specific complain was alleged. No patient harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Evidence of an arc track has led engineering to believe the instrument may have arced during a previous surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 584141 | 594315 | 564788 | 2955842200400038 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-03-04 | 2004-06-09 | No | Yes | Health Professional | 40012607 | Device Returned To Manufacturer | Yes | 2004-03-04 | Yes | No | Reuse | 2004
40012607 | 2004-05-07 | Malfunction |  | It was reported that the instrument began to acr during the case. There were no precipitating factors identified. The instrument was removed and replaced with a backup permanent hook cautery instrument to continue the case to completion. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 581365 | 591507 | 562201 | 2955842200400034 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-06-04 | 2004-06-04 | No |  | Health Professional | 40012607 | No | Yes | 2004-05-07 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | It was reported that the coating on the instrument hook had worn off. There was no record whether the observation was noticed at : a) instrument cleaning and sterilization; b) instrument inspection immediately prior to surgery or c) during surgery. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis confirmed that the distal insulation was missing from the instrument. Evidence of an arc track on the yaw pulley had led engineering to believe the instrument may have arced during a previous surgical procedure. No detrmination could be made regarding whether the distal insulation separated from the instrument during: a) instrument cleaning and setrilization or b) during surgery. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 581367 | 591509 | 562203 | 2955842200400033 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-04 | 2004-06-04 | No | Yes | Health Professional | 40012607 | No | No | 2004-04-06 | Yes | No | Reuse | 2004
40012607 | 2004-05-03 | Malfunction |  | It was alleged that the insulation on the instrument began to smoke at the beginning of the case. The instrument was removed and replaced with a backup permanent hook cautery instrument to continue the case to completion. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 580922 | 591065 | 561776 | 2955842200400032 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-02 | 2004-06-03 | No | Yes | Health Professional | 40012607 | No | No | 2004-05-03 | Yes | No | Reuse | 2004
40015402 | 2004-05-03 | Malfunction |  | It was alleged that the instrument arced 15 minutes into use. The instrument was removed and the case was completed with a backup permanent cautery spatula instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The evaluation showed no visual evidence of arcing. There was no evidence of insulation damage and no discoloration indicating that arcing had occurred. The instrument was tested with respect to the operational environment by submersed the instrument's electrode and wrist into saline and energizing the instrument. The instrument passed the test without any signs of arcing. The device was evaluated and the alleged failure could not be duplicated. The reported symptom could not be reproduced. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 580926 | 591069 | 561780 | 2955842200400031 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-06-02 | 2004-06-03 | No | Yes | Health Professional | 40015402 | No | Yes | 2004-05-03 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | It was alleged that a cable on the instrument had broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis concluded that no instrument cables were severed or broken. Evidence of an arc track between the gooseneck to the grip cable has led engineering to believe the instrument may have arced during a previous surgical procedure. Engineering has theorized that the discoloration of the instrument wrist (the aforementioned arc track) was interpreted by hospital as a broken segment of cable. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 580963 | 591106 | 561817 | 2955842200400030 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-06-02 | 2004-06-03 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-12 | Yes | No | Initial | 2004
40012607 | 2004-04-01 | Malfunction |  | During a surgical procedure, the surgeon used a grasping forceps type instrument to remove tissue from the permanent cautery hook instrument and continued the case. During an instrument exchange, it was noticed that a plastic piece was missing from the permanent cautery hook instrument. The doctor retrieved and removed the plastic fragment from the pt. The case was completed with no pt injury or adverse event. No additional surgical procedure was required. | Both the instrument and the component fragment were returned for eval. Failure analysis of the returned instrument confirmed that the distal clevis of the instrument had broken at the base. Failure to follow the instructions for use which were shipped with the instrument likely contributed to the instrument malfunction. The instructions for use includes a general precaution and warning which states: "do not use an instrument to clean debris off another instrument intraoperatively. This may result in damage to the instruments. ". | PERMANENT CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT | 552398 | 562668 | 534513 | 2955842200400025 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-15 | 2004-05-17 | No | Yes | Health Professional | 40012607 | No | No | 2004-04-01 | Yes | No | Initial | 2004
40012607 | 2004-04-22 | Malfunction |  | An instrument would be returned due its having a broken tip. No other details relating to the malfunction of the instrument were provided. | Failure analysis of the returned instrument confirmed that the distal tip of the instrument had fragmented. Not all components were returned. There was no indication from the customer that the instrument malfunctioned inside a pt. | PERMANENT CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT | 552401 | 562671 | 534516 | 2955842200400026 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-15 | 2004-05-17 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-22 | Yes | No | Reuse | 2004
40012607 | 2004-04-26 | Malfunction |  | It was alleged that the instrument arced. The instrument was replaced with a backup instrument and the procedure was completed without further incident. No pt harm, adverse outcome or injury was reported. | The instrument has not yet been returned for eval. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 552404 | 562674 | 534519 | 2955842200400027 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2004-05-15 | 2004-05-17 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-26 | Device Not Retu | No | Unkown | 2004
40012607 | 2004-04-27 | Malfunction |  | It was alleged that the instrument shorted. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 552406 | 562676 | 534521 | 2955842200400028 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-05-15 | 2004-05-17 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-27 | Yes | No | Reuse | 2004
40012607 | 2004-04-28 | Malfunction |  | It was reported that the instrument tip bent. No other details relating to the malfunction of the instrument were provided. No pt harm was reported. No adverse outcome or injury was reported. | Failure analysis of the instrument confirmed that the instrument had malfunctioned and that the proximal clevis had fractured. The wrist remained attached to the instrument shaft. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 552408 | 562678 | 534523 | 2955842200400029 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-05-15 | 2004-05-17 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-28 | Yes | No | Reuse | 2004
IS1200 A42P5 | 2004-03-30 | Malfunction | Other | During a procedure, the surgeon experienced an intermittent vision problem in the right eye of the stereo viewer. No pt harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | The intermittent vision problem was reported as having occurred on the right eye of the stereo viewer. Investigation by field engineering concluded that the right monitor in the hrsv was having intermittent failures. The hosp's system was repaired by replacing the hrsv and a cable that had been damaged. The hosp's da vinci vision system passed testing after the repair was performed. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 552412 | 562682 | 534527 | 2955842200400024 | 1 | GEI | Manufacturer | Health Professional | Nurse | Initial | 2004-05-15 | 2004-05-17 | Yes | Yes | Health Professional | IS1200 A42P5 | No | Yes | 2004-03-30 | Yes | No | Reuse | 2004
40012607 | 2004-04-05 | Malfunction |  | The instrument arced. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 551116 | 561407 | 533329 | 2955842200400023 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-05-12 | 2004-05-12 | No | Yes | Health Professional | 40012607 | No | No | 2004-04-05 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | It was reported that the tip of the instrument was broken. No pt injury or adverse event was reported. No additional surgical procedure was required. | Failure analysis of the returned instrument confirmed that the distal tip of the instrument had fragmented. | PERMANENT CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT | 548988 | 559274 | 531355 | 2955842200400022 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-05-06 | 2004-05-06 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-04-02 | Yes | No | Reuse | 2004
40012607 |  | Malfunction |  | It was alleged that the instrument arced. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 549271 | 559570 | 531621 | 2955842200400021 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-05-05 | 2004-05-05 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-02-05 | Yes | No | Reuse | 2004
40011707 | 2004-03-29 | Malfunction |  | While the surgeon was using the needle driver instrument to suture soft tissue, one grasping jaw of the instrument fractured. It was reported that the jaw fractured where the teeth transition to a smooth surface. The surgeon retrieved and removed the separated component from the pt. A back-up needle driver was used to finish the procedure. The procedure was successfully completed with no harm to the pt. | Failure analysis of the instrument confirmed the customer reported complaint that a component had separated from the instrument. Evaluation of the instrument determined that the carbide insert component had separated from the grip jaws. Based upon visual inspection of the returned carbide insert and instrument, engineering has concluded that the separation was due to suboptimal braze quality. | NEEDLE DRIVER INSTRUMENT 5MM | ENDOSCOPIC INSTRUMENT | 552009 | 562292 | 534144 | 2955842200400020 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-04-29 | 2004-04-29 | No | Yes | Health Professional | 40011707 | No | No | 2004-03-29 | Yes | No | Initial | 2004
40000316 | 2003-09-15 | Malfunction |  | During a procedure, deployed clips would not disengage easily from three clip applier instruments. No pt injury or adverse outcome was reported. | Visual inspection revealed that the grips of the returned instrument were damaged. The damaged grips likely prevented clips from easily releasing from the instrument. It is believed that the instrument was damaged at the customer site due to the following reasons: 1) the intrument passed functional testing prior to being shipped; and 2) this instrument had been previously used at the customer site to apply over 49 staples without report of difficulty. This incident was originally reported to the fda in 10/2003 as 2955842-2203-00063. At the request of the fda, the mdr has been split into three separate mdr reports. Ref. : 2955842-2003-00063, 2955842-2003-00077. | CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 546163 | 556492 | 528737 | 2955842200300078 | 1 | GDO | Manufacturer | Company Representative | Other | Initial |  | 2004-04-28 | No | Yes | Health Professional | 40000316 | No | No | 2003-09-15 | Yes | No | Reuse | 2004
40012607 | 2003-10-02 | Malfunction |  | During a surgical procedure, a plastic fragment of the permanent cautery instrument separated from the instrument. The surgeon was unable to locate the fragment. The case was completed with the da vinci system. No pt harm was reported. No adverse outcome or injury was reported. | The hosp returned the instrument to isi for eval. Failure analysis confirmed that the distal clevis component separated from the instrument. This mdr is being submitted as both: 1) an update for the device eval and 2) a correction to the mdr number. Device eval: the original mdr for this incident was submitted prior to receipt and eval of the returned instrument. The report has been updated to include the subsequent eval. Correction due to a typographical error, the mdr was originally submitted under the wrong number. Please update your files with the corrected mdr. 2955842-2003-00063 (incorrect). | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 546167 | 556496 | 528741 | 2955842200300064 | 1 | NAY | Manufacturer | Company Representative | Other | Initial |  | 2004-04-28 | No | Yes | Health Professional | 40012607 | No | No | 2003-10-02 | Yes | No | Reuse | 2004
40000421 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel instruments in bulk with the general complaint that a spatula accessory had a separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom may have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 7 lbs. The returned instrument no longer passed the tensile specification of 2 lbs. A tolerance analysis of the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009, mdr 2955842-2004-00014, mdr 2955842-2004-00010, mdr 2955842-2004-00015,mdr 2955842-2004-00011, mdr 2955842-2004-00016, mdr 2955842-2004-00012, mdr mdr 2955842-2004-00013, mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544086 | 554458 | 526849 | 2955842200400017 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40000423 | 2003-10-01 | Malfunction |  | It was reported that during surgery, cautery spatula accesory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. He hospital returned 10 scalpel cautery instruments in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure anaysis concluded that the customer reported symptom may have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 2 lbs. The returned instrument no longer passed the tensile specifications of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref mdr 2955842-2004-00009, mdr 2955842-2004-00014, mdr 2955842-2004-00010 , mdr 2955842-2004-00015,mdr 2955842-2004-00011, mdr 2955842-2004-00016, mdr 2955842-2004-00012, mdr 2955842-2004-00017. Mdr 2955842-2004-00013. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544096 | 554468 | 526858 | 2955842200400018 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000423 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40012607 | 2004-03-25 | Malfunction |  | It was alleged that insulation may have degraded and the instrument arced during a case. The surgeons removed the instrument and completed the case with a replacement instrument without further incident. No pt harm, adverse outcome or injury was reported. | Instrument has not been returned for isi for evaluation. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544103 | 554475 | 526865 | 2955842200400019 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-27 | 2004-04-27 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-03-25 | Device Not Retu | No | Unkown | 2004
40000423 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reportedthe spatula blade was discarded by the customer. The hospital retured 10 scalpel cautery instruments in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom could have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 6 lbs. The returned instrument no longer passed the tensile specifications of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref mdr 2955842-2004-00009 mdr 2955842-2004-00014 mdr 2955842-2004-00015 mdr 2955842-2004-00011 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544105 | 554477 | 526849 | 2955842200400010 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000423 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40000421 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instrument in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom could have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument 1. 6lbs. The returned instrument no longer passed the tensile specification of 2lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref mdr 2955842-2004-00009, mdr 2955842-2004-00014, mdr 2955842-2004-00010, mdr 2955842-2004-00015, mdr 2955842-2004-00011, mdr 2955842-2004-00016, mdr 2955842-2004-00017, mdr 2955842-2004-00013, mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544128 | 554500 | 526875 | 2955842200400012 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-19 | Yes | No | Initial | 2004
40000423 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separted from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instruments in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury was reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom could have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 5 lbs. The returned instrument no longer passed the tensile specifications of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009 mdr 2955842-2004-00014 mdr 2955842-2004-00010 mdr 2955842-2004-00015 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544134 | 554506 | 526881 | 2955842200400011 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000423 | No | No | 2004-03-19 | Yes | No | Initial | 2004
40000421 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instruments in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom may have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 2 lbs. The returned instrument no longer passed the tensile specificaton of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009 mdr 2955842-2004-00010 mdr 2955842-2004-00015 mdr 2955842-2004-00011 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544140 | 554512 | 526887 | 2955842200400014 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40000421 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instrument in bulk with general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis could not conclude that the customer reported symptom occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 3. 5 lbs. The returned instrument passed the tensile specifications of 2 lbs. A tolerance analysis on the actual instument and accessory used during the surgical procedures cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009 mdr 2955842-2004-00014 mdr 2955842-2004-00010 mdr 2955842-2004-00015 mdr 2955842-2004-00011 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544145 | 554517 | 526892 | 2955842200400013 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-19 | Yes | No | Initial | 2004
40000423 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instrument in bulk with the general complaint that a spatula accessory had separtated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom could have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 9 lbs. The returned instrument no longer passed the tensile specification of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00014 mdr 2955842-2004-00010 mdr 2955842-2004-00015 mdr 2955842-2004-00011 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544152 | 554525 | 526899 | 2955842200400009 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000423 | No | No | 2004-03-19 | Yes | No | Initial | 2004
40000417 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery insturment. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instruments in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom may have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 5 lbs. The returned instrument no longer passed the tensile specifications of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be perfomed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009 mdr 2955842-2004-00014 mdr 2955842-2004-00010 mdr 2955842-2004-00011 mdr 2955842-2004-00016 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544159 | 554533 | 526906 | 2955842200400015 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000417 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40000421 | 2003-10-01 | Malfunction |  | It was reported that during surgery, a cautery spatula accessory separated from the scalpel cautery instrument. The cautery spatula was retrieved and removed from the patient. No patient harm or injury was reported. The spatula blade was discarded by the customer. The hospital returned 10 scalpel cautery instrument in bulk with the general complaint that a spatula accessory had separated from each instrument. The hospital was not able to provide the date of occurrence or case details associated with the alleged incidents. Nonetheless, the site was able to confirm that all cautery spatula accessories were retrieved and removed from all patients. No patient harm or injury were reported. | The hospital returned the instrument for evaluation but scrapped the associated cautery spatula accessory. Failure analysis concluded that the customer reported symptom may have occurred. The scalpel cautery instrument was tensile tested with a new cautery spatula accessory. The maximum retention force measured for the instrument was 1. 3 lbs. The returned instrument no longer passed the tensile specification of 2 lbs. A tolerance analysis on the actual instrument and accessory used during the surgical procedure cannot be performed since the cautery spatula accessory was not returned to isi. Ref. Mdr 2955842-2004-00009 mdr 2955842-2004-00014 mdr 2955842-2004-00010 mdr 2955842-2004-00015 mdr 2955842-2004-00011 mdr 2955842-2004-00012 mdr 2955842-2004-00017 mdr 2955842-2004-00013 mdr 2955842-2004-00018. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 544164 | 554537 | 526910 | 2955842200400016 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-23 | 2004-04-27 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-19 | Yes | No | Reuse | 2004
40004815 | 2004-03-22 | Malfunction |  | It was reported that the instrument wrist broke and a fragment separated from the instrument during surgery. The surgeons removed the instrument and retrieved the fragment from the patient. The case was completed without further incident with the da vinci system. No patient harm was reported. No adverse outcome or injury was reported. | Failure analysis of the instrument confirmed the customer reported complaint that a component fragment had separated from the instrument. Eval of the instrument determined that a fragment from the proximal clevis was missing and the instrument shaft had begun to fray. Engineering was unable to distinguish which component (shaft or proximal clevis) failed first to cause the fragmentation. | LONG TIP FORCEPS | ENDOSCOPIC INSTRUMENT | 542753 | 553174 | 525598 | 2955842200400008 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-04-21 | 2004-04-26 | No | Yes | Health Professional | 40004815 | No | No | 2004-03-22 | Yes | No | Reuse | 2004
40000421 | 2004-03-03 | Malfunction |  | During surgery with a scalpel cautery instrument it was noticed that after 30 minutes of use that the instrument tip began to heat up and melt. The cautery instrument was removed from the pt and replaced with another isi scalpel cautery instrument to continue to case to completion. No pt injury or adverse outcome was reported. The pt was not injured and the procedure was not significantly lengthened due to the incident. | Failure analysis of the returned instrument detected evidence of arcing on the top and side of the yaw pulley. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 536879 | 547469 | 520024 | 2955842200400006 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-04-01 | 2004-04-02 | No | Yes | Health Professional | 40000421 | No | No | 2004-03-03 | Yes | No | Reuse | 2004
40012607 | 2004-02-03 | Malfunction |  | It was alleged that an instrument arced during a case. The surgeons removed the instrument and completed the case with a replacement instrument without further incident. No pt harm, adverse outcome or injury was reported. | Evidence of arcing was confirmed at the distal end of the instrument. In addition, mechanicl damage was noted at the insulation surrounding the cautery hook. Engineering has concluded that instrument mishandling at the customer site caused the mechanical damage to the insulation and resultant arcing. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 537997 | 548529 | 521044 | 2955842200400007 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2004-04-02 | 2004-04-02 | No | Yes | Health Professional | 40012607 | No | Yes | 2004-02-13 | Yes | No | Reuse | 2004
40000423 | 2004-02-18 | Malfunction |  | Upon removal of the scalpel cautery instrument from a patient, the cautery spatula accessory was missing. The surgeons proceeded with the procedure. The case was continued with no further incident. At the end of the procedure the surgeon retrieved and removed the cautery spatula from the patient. The spatula blade was discarded by the customer. | The hospital returned two instruments and their remaining stock of unused spatula accessories. The site was unable to recall which instrument was associated with the complaint. The affected spatula accessory was not returned. Neither of the returned instruments nor the unused spatula accessories had any visual damage. Unfortunately, since the affected cautery spatula accessory was not returned isi cannot conclusively determine whether user error or other issue contributed to the alleged device malfunction. Please note that intuitive surgical no longer distributes the scalpel cautery instrument or the corresponding cautery spatula accessory. Add'l lot# 1006031 10. Add'l device mfr date: 06/2003. ). | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 534194 | 544786 | 517460 | 2955842200400004 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-03-29 | 2004-03-29 | No | Yes | Health Professional | 40000423 | No | No | 2004-02-23 | Yes | No | Reuse | 2004
40000423 | 2004-02-23 | Malfunction |  | The scrub nurse inserted a 15 deg. Scalpel blade accessory into a scalpel cautery instrument and handed the instrument to the patient side assistant. The patient side assistant introduced the instrument with the blade into the patient. When the instrument appeared on the video monitor, no blade was found in the instrument. The blade had separated from the instrument and fallen into the patient. The blade was retrieved using other intuitive surgical endoscopic instruments and removed from the patient. Examination of the blade at the site determined that the blade had fractured and a 1mm section at the base of the blade was missing. The sharp portion of the blade was completely retrieved. The missing piece could not be found by surgeons. A new blade was inserted into the same scalpel cautery instrument and the case was continued to completion with no further incident. | Although the instrument and disposable blade were not returned, an intuitive surgical clinical engineer was on site and inspected and photographed the 15 degree blade accessory. Evaluation of the photo confimed the blade had fractured. Since the blade and instrument were not returned for evaluation, no conclusion can be made regarding whether a user error or other issue contributed to the device malfunction. Please note that intuitive surgical no longer distributes the scalpel cautery instrument or the corresponding scalpel accessory. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 534214 | 544806 | 517480 | 2955842200400005 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2004-03-26 | 2004-03-29 | No | Yes | Health Professional | 40000423 | No | No | 2004-02-23 | Device Not Retu | No | Unkown | 2004
40011706 | 2004-03-11 | Malfunction |  | It was reported that during a surgical procedure, the jaws of a 5 mm needle driver instrument would not close. No pt harm was reported. No adverse outcome or injury was reported. | Failure analysis concluded that the cable stop component had failed and was responsible for preventing the instrument jaws from closing (i. E. Loss of grip function). The component failure caused the instrument to become non-functional. Inspection of the instrument under magnification confirmed that the cable stop component remained attached to the device. A small crack was discovered emanating from the corner of a slot that had been machined in the component. Engineering has theorized that a localized stress concentration developed at the site due to stresses imparted by a dimensional interference fit. The cable stop component on all intuitive surgical 5mm instruments has been redesigned to eliminate the potential interference fit. | NEEDLE DRIVER INSTRUMENT 5MM | ENDOSCOPIC INSTRUMENT | 531162 | 541834 | 514547 | 2955842200400003 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2004-03-26 | 2004-03-26 | No | Yes | Health Professional | 40011706 | No | No | 2004-03-11 | Yes | No | Initial | 2004
| 2000-02-29 | Malfunction |  | It was reported that the positioner was moving on its own with no command. The behavior was attributed to a defective positioner cable. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | In 2000 the positioner cable was replaced and no further difficulty was reported at the site. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516032 | 526868 | 500117 | 2029233200400004 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2000-02-29 | No Answer Provi | No | Reuse | 2004
| 2000-01-21 | Malfunction |  | It was reported that the positioner moved "quickly" on its own with no command. The behavior was attributed to a defective positioner. It was not confirmed if the malfunction occurred during a clinical procedure. No adverse event was reported. | Troubleshooting was done over the phone and alleged failure could not be duplicated. Cause of the failure could not be determined. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516033 | 526869 | 500118 | 2029233200400002 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2000-01-21 | Device Not Retu | No | Reuse | 2004
| 2000-02-08 | Malfunction |  | It was reported that the positioner moved "quickly" on its own with no command. The behavior was attributed to a defective positioner. It was not confirmed if the malfunction occurred during a clinical procedure. No adverse event was reported. | During the troubleshooting the alleged failure could not be duplicated. Cause of the failure could not be determined. In 2000 the positioner was replaced and no further difficulty was reported at the site. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516034 | 526870 | 500119 | 2029233200400003 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2000-02-08 | Yes | No | Reuse | 2004
| 2002-04-18 | Malfunction |  | Event Description: --/-- It was reported that a microjoint scissors used in a procedure development study had a broken head failure. It was not confirmed if the malfunction occurred during a clinical procedure. No adverse event was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No further evaluation of the system is expected. | MICRO JOINT POTTS SCISSORS | Reuse | 516036 | 526872 | 500121 | 2029233200400028 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 1999-08-19 | Malfunction |  | It was reported that the positioner failed to stop during the procedure. The behavior was attributed to a defective positioner. It was confirmed that the malfunction occurred during a surgery. No adverse event was reported. | During troubleshooting it was discovered that the speaker had open voice coil. Speaker was replaced. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516079 | 526915 | 500164 | 2029233200400001 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 1999-08-19 | Yes | No | Reuse | 2004
| 2001-07-02 | Malfunction |  | It was reported that the positioner moved back without any command. The behavior was attributed to a defective positioner. It was not confirmed whether the malfunction occurred during a surgery. No adverse event was reported. | No further eval of the system is expected. | AESOP SYSTEM | AUTOMATED ENDOSCOPIC POSITIONI | 516080 | 526916 | 500165 | 2029233200400009 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2001-07-02 | Yes | No | Reuse | 2004
| 2004-02-02 | Malfunction |  | Event Description: --/-- It was reported that a microjoint instrument had a broken head failure. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | MICRO JOINT POTTS SCISSORS | Reuse | 516084 | 526920 | 500169 | 2029233200400026 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2003-06-26 | Malfunction |  | Event Description: --/-- It was reported that the micro joint heavy needle driver's jaw tip broken and the worn jaws would not grasp the tissue. The articulink screw also is broken, but its threads remained on the articulink. This instrument was used in a clinical case; however, there is no info available to confirm if the failure occurred during the case. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | Reuse | 516086 | 526922 | 500171 | 2029233200400037 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-06-26 | No | No |  | 2004
| 2002-11-12 | Malfunction |  | It was reported that the articulink was broken and the screw was missing. This insturment was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome occurred. | Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516090 | 526926 | 500175 | 2029233200400039 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2002-11-12 | Yes | No | Reuse | 2004
| 2003-03-07 | Malfunction |  | It was reported that the 12mm fenestrated grasper articulink was broken and the screw broke by the threads that remained on the articulink. This instrument was used in a clinical case; however, there is no info available to confirm if the failure occurred during the case. No pt injury or adverse outcome occurred. | Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 516092 | 526928 | 500177 | 2029233200400040 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-03-07 | Yes | No | Reuse | 2004
| 2003-03-07 | Malfunction |  | It was reported that the heavy needle driver articulink was broken and the screw was missing. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome occurred. | Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516095 | 526931 | 500180 | 2029233200400041 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-03-07 | Yes | No | Reuse | 2004
| 2003-07-07 | Malfunction |  | It was reported that the heavy needle driver articulink screw was missing and jaw gear was broken. This is an instrument used in a clinical case, but there is no info available about whether the failure occurred during the case. No pt injury or adverse otucome occurred. | Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516103 | 526939 | 500188 | 2029233200400042 | 1 | GCJ | Manufacturer | Distributor | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-07-07 | Yes | No | Reuse | 2004
| 2000-06-28 | Malfunction |  | It was reported that the positioner made sudden movements without any command. The behavior was attributed to a defective positioner. It was confirmed that the movement occurred during a surgery. No adverse event was reported. | No further eval of the sysem is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516108 | 526944 | 500193 | 2029233200400005 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2000-06-28 | Yes | No | Reuse | 2004
| 2003-03-05 | Malfunction |  | It was reported that the positioner "erratically. " the behavior was attributed to a defective positioner. It was not confirmed if the malfunction occurred during a clinical procedure. No adverse event was reported. | No further eval of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516112 | 526948 | 500197 | 2029233200400013 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-03-05 | Yes | No | Reuse | 2004
| 2001-12-07 | Malfunction |  | It was reported that the mj articulink failed and separated in the lung of a pt. The procedure was converted to an endoscopic surgery. The procedure was completed successfully. No pt injury or adverse outcome occurred. | Engineering confirmed the alleged malfunction. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516113 | 526949 | 500198 | 2029233200400038 | 1 | GCJ | Manufacturer |  | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2001-12-07 | Yes | No | Reuse | 2004
| 2003-05-05 | Malfunction |  | Event Description: --/-- It was reported that the microjoint fenestrated grasper had a broken articulink screw, but the threads remained in the articulink. This instrument was used in a clinical case; however, there is no info available to confirm if the failure occurred during the case. No pt injury or adverse outcome occurred. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. According to the reported event details, the instrument failed in a safe mode. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT FENESTRATED GRASPER | Reuse | 516116 | 526952 | 500201 | 2029233200400036 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-05 | No | No |  | 2004
| 2003-04-28 | Malfunction |  | It was reported that a microjoint metzenbaum scissors had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink and the articulink screw were broken. Part of the screw remained in hole on the main jaw. No definitive root cause of the malfunction was determined. No further evaluation of the system is expected. | ZEUS MICRO JOINT METZENBAUM SCISSORS | ENDOSCOPIC INSTRUMENT | 516119 | 526955 | 500204 | 2029233200400043 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2003-04-28 | Malfunction |  | It was reported that a microjoint micro needle driver had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink screw was broken and half of the screw was still inside the articulink. No definitive root cause of the malfunction was determined. No further evalauton of the instrument is expected. | ZEUS MICRO JOINT MICRO NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516123 | 526959 | 500208 | 2029233200400044 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2003-04-28 | Malfunction |  | It was reported that a microjoint heavy needle driver had a broken articulink. This instrument was not used in a clinical case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink was broken, both sides remained intact. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516124 | 526960 | 500209 | 2029233200400047 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2002-09-26 | Malfunction |  | Event Description: --/-- It was reported that during a zeus clinical case the disposable and reusable motor packs started to fail and give error messages. No pt injury or adverse outcome occurred. This is a safe failure because the motorpack is not inserted in the pt's body and any failure of the motorpack will only cause the instrument not to function. However, due to dr's annoyance to the malfunction, the case was proceeded without zeus per dr's request. Because of the limited info provided, whether the procedure was converted to laparoscopic or open surgery cannot be confirmed. | Manufacturer Narrative: --/-- Although an rma number was issued by computer motion, the return of the alleged defective unit cannot be confirmed. No definitive root cause of the alleged malfunction was determined. According to the reported event details, the zeus system provided the appropriate system error messages to the user as the reusable motor packs began to fail. No further evaluation of the system is expected. --/-- Additional Manufacturer Narrative: --/-- Unk. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516125 | 526961 | 500210 | 2029233200400017 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2003-04-28 | Malfunction |  | It was reported that a microjoint fenestrated grasper had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink and the articulink screw were broken and part of the screw remained in one side. No definitive root cause of the malfunction was determined. No further evalaution of the instrument is expected. | ZEUS MICRO JOINT FENESTRATED GRASPER 18MM | ENDOSCOPIC INSTRUMENT | 516127 | 526963 | 500212 | 2029233200400048 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2003-05-02 | Malfunction |  | It was reported that a microjoint debakey grasper had a bent articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed the alleged malfunction. No definitive root cause of the malfunction was determined. No further evalaution of the instrument is expected. | ZEUS MICRO JOINT DEBAKEY GRASPER | ENDOSCOPIC INSTRUMENT | 516137 | 526973 | 500222 | 2029233200400049 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-02 | Yes | No | Reuse | 2004
| 2003-04-14 | Malfunction |  | It was reported that the 12mm fenestrated grasper was inserted into a cannula with the grasping jaws closed. During the surgeon's first attempt to move the instrument tip, the instrument would not articulate. While the surgeon was attempting to articulate the instrument with the zeus instrument control system, the jaws of the grasper opened and could not be closed. Fragments were noted to have separated from the distal end of the instrument. The instrument fragments were removed from the pt via the assistant port. The surgeon is confident all pieces of retained metal were removed from the pt. In order to remove the fenestrated grasper, the instrument and cannula were removed from the pt simultaneously while the pneumoperitoneum was maintained manually. The fenestrated grasper was removed and replaced with a large needle driver. The case was completed using the zeus instrument control system without further incident. There were no other reported complications due to the breakage or the resultant instrument removal. No pt harm, pt injury or other adverse event was reported. | Per the case details, it was confirmed that an articulink separated from its distal hinge. One of the instrument fragments was concluded as being the corresponding hinge pin. | ZEUS MJ FENESTRATED GRASPER 12MM | ENDOSCOPIC INSTRUMENT | 516151 | 526987 | 500236 | 2029233200400035 | 1 | GCJ | Manufacturer | Health Professional | Physician | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | Yes | 2003-04-29 | Yes | No | Unkown | 2004
| 2003-06-12 | Malfunction |  | Event Description: --/-- It was reported that a zeus instrument driver will jump when the jaws were closed completely. It was not confirmed whether the malfunction occurred during a procedure. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- According to the reported event details, the zeus system provided the appropriate system error messages to the user as the instrument driver began to fail. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516156 | 526992 | 500241 | 2029233200400019 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2002-07-12 | Malfunction |  | Event Description: --/-- It was reported that during a zeus clinical case the instrument driver started to fail and give error messages. No pt injury or adverse outcome occurred. This is a safe failure because the instrument driver is not inserted in the pt's body and any failure of the instrument driver will only cause the instrument not to function. However, the case was proceeded without zeus per dr's request. Because of the limited info provided, whether the procedure was converted to laparoscopic or open surgery cannot be confirmed. | Manufacturer Narrative: --/-- Engineering confirmed the failure. According to the reported event details, the zeus system provided the appropriate system error messages to the user as the instrument driver began to fail. No further evaluation of the system is expected. --/-- Additional Manufacturer Narrative: --/-- Unk. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516159 | 526995 | 500244 | 2029233200400016 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2002-01-30 | Malfunction |  | It was reported that a microjoint instrument had a broken head failure. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No further evaluation of the system is expected. | MICRO JOINT FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 516165 | 527001 | 500250 | 2029233200400025 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No | Reuse | 2004
| 2002-01-28 | Malfunction |  | Event Description: --/-- It was reported that three different types of micro joint instruments had broken heads. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | MICRO JOINT FENESTRATED GRASPER | Reuse | 516166 | 527002 | 500251 | 2029233200400023 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2003-04-15 | Malfunction |  | Event Description: --/-- It was reported that the microjoint heavy needle driver has a broken jaw and was missing piece at the very tip. This instrument was used in a clinical case; however, there is no info available to confirm if the failure occurred during the case. No pt injury or adverse outcome occurred. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | Reuse | 516169 | 527005 | 500254 | 2029233200400034 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-15 | No | No |  | 2004
| 2002-05-28 | Malfunction |  | Event Description: --/-- It was reported that the fenestrated grasper tip separated from the instrument due to torsioned jaws resulting from hard tissue removal. No pt injury or adverse outcome occurred. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No further evaluation of the system is expected. | MICRO JOINT FENESTRATED GRASPER | Reuse | 516173 | 527009 | 500258 | 2029233200400030 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2002-05-28 | No | No |  | 2004
| 2002-06-18 | Malfunction |  | Event Description: --/-- It was reported that the microjoint instrument jaws were locked on tissue; the head had a cavity at the tip. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT FENESTRATED GRASPER | Reuse | 516177 | 527013 | 500262 | 2029233200400031 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2002-07-29 | Malfunction |  | Event Description: --/-- It was reported that a microjoint instrument had broken head and a missing jaw. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the reported malfunction. Failure analysis of other instruments exhibiting similar failure mode concluded that the material used for the head may have caused the problem. However, no definitive root cause of the alleged jaw malfunction was determined. No further evaluation of the system is expected. | MICRO JOINT MICRO NEEDLE DRIVER | Reuse | 516184 | 527020 | 500269 | 2029233200400032 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2002-07-29 | No | No |  | 2004
| 2002-02-22 | Malfunction |  | Event Description: --/-- It was reported that microjoint instrument jaws were locked on a needle. It was not confirmed whether the malfunction occurred during a procedure. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516188 | 527024 | 500273 | 2029233200400027 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
ZEUS MICROWRIST | 2002-01-04 | Malfunction |  | Event Description: --/-- It was reported that four different types of micro joint instruments had broken parts including broken head failure. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516202 | 527038 | 500287 | 2029233200400022 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional | ZEUS MICROWRIST | No | No |  | No | No |  | 2004
| 2001-02-27 | Malfunction |  | Event Description: --/-- It was reported that the microjoint scissors would not cut and could not be removed through the trocar. The scissors and the trocar were removed from the pt simultaneously, and jaws of the scissors were not sheathed as the instrument was removed. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Computer motion, sent a micro joint scissors to the hosp to replace the alleged defective unit. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516209 | 527045 | 500294 | 2029233200400020 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2003-05-02 | Malfunction |  | It was reported that a microjoint heavy debakey grasper had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink and articulink screw were broken. Threaded part to the articulink screw remained inside the articulink hole. Instrument back was also broken around the metal insert. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT DEBAKEY GRASPER | ENDOSCOPIC INSTRUMENT | 516215 | 527051 | 500300 | 2029233200400053 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-02 | Yes | No | Reuse | 2004
| 2002-12-02 | Malfunction |  | Event Description: --/-- It was reported that the microassist debakey grasper has a broken jaw. This instrument was used in a lab case; however, there is no info available to confirm if the failure occurred during a clinical use. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516220 | 527056 | 500305 | 2029233200400033 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2002-12-02 | No | No |  | 2004
| 2001-11-19 | Malfunction |  | Event Description: --/-- 5mm cautery probe shows signs of arcing around 3/4 of the insulation. It was not clear whether the arcing occurred during the procedure and under what circumstances. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed failure. No definitive root cause of the alleged malfunction was determined. No further evaluation of the system is expected. | ZEUS MICROWRIST SURGICAL SYSTEM | Reuse | 516223 | 527059 | 500308 | 2029233200400021 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2001-11-19 | No | No |  | 2004
| 2003-05-02 | Malfunction |  | A microjoint heavy needle driver had a broken articulink. This is a lab instrument and was not used in a clinical case. No adverse outcome reported. | Engineering confirmed that the articulink and the articulink screw were broken. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT HEAVY NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516254 | 527086 | 500341 | 2029233200400051 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-02 | Yes | No | Reuse | 2004
| 2003-05-02 | Malfunction |  | A microjoint micro needle driver had a bent articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink was bent and cracked. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT MICRO NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516256 | 527088 | 500343 | 2029233200400050 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-02 | Yes | No | Reuse | 2004
| 2001-09-13 | Malfunction |  | The positioner moved in the wrong direction or failed to stop. The behavior was attributed to a defective headset. It was not confirmed if the problem occurred during a clinical procedure. No adverse event was reported. | It was discovered that headset/microphone assembly was cause of the problem. Computer motion sent a new headset/microphone assembly to replace the alleged defective unit. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516259 | 527091 | 500346 | 2029233200400011 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No Answer Provided | 2001-09-13 | No Answer Provi | No | Reuse | 2004
| 2003-04-18 | Malfunction |  | The positioner moved without command during surgery. The behavior was attributed to the user error use of incorrect headset. No adverse event was reported. | During the conversation with the hospital staff, it was found out that the doctor was using a wireless microphone, from a vendor not approved by computer motion for use with aesop. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516261 | 527093 | 500348 | 2029233200400014 | 1 | GDC | Manufacturer | Company Representative | UNKNOWN | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | Yes | 2003-04-18 | Device Not Retu | No | Reuse | 2004
| 2001-01-12 | Malfunction |  | The positioner moved without command. The behavior was attributed to the user error. It was not confirmed if the problem occurred during a clinical procedure. No adverse event was reported. | During the evaluation of the system by a computer motion clinical specialist, it was discovered that the doctor was not using the aesop hermes ready system correctly. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516263 | 527095 | 500350 | 2029233200400007 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No Answer Provided | 2001-01-12 | Device Not Retu | No | Reuse | 2004
| 2000-10-30 | Malfunction |  | The positioner moved without command during training. The behavior was attributed to a defective controller. No adverse event was reported. | Aesop and hermes controllers were returned to computer motion for functional testing. Engineering was unable to reproduce the alleged failure. No further evaluation of the system is expected. | AESOP SYSTEM | AUTOMATED ENDOSCOPIC POSITIONI | 516268 | 527100 | 500355 | 2029233200400006 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2000-10-30 | Yes | No | Reuse | 2004
| 2003-05-30 | Malfunction |  | The positioner moved without command, during a procedure. The behavior was attributed to a defective controller. The alleged unit was then disengaged and the procedure was completed using manually positioned endoscope. No adverse event was reported. | Both aesop and hermes controllers, as well as the flex cable connector were returned to computer motion for further analysis. Engineering was unable to replicate the alleged failure. Reference was made to possible inadvertent contact of the hermes pendant during the procedure. Consequently, the engineering analysis concludes that such contact could clearly result in a similar unintended motion. Controller was replaced, no further difficulty reported. | AESOP HERMESEADY SYSTEM | AUTOMATED ENDOSCOPIC POSITIONI | 516270 | 527102 | 500357 | 2029233200400015 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | Device Returned To Manufacturer | No | 2003-05-30 | Yes | No | Reuse | 2004
| 2001-08-16 | Malfunction |  | The positioner moved without command. The behavior was attributed to a defective controller. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | In 08/2001 the controller was replaced, and no further difficulty was reported at the site. No further evaluation of the system is expected. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516272 | 527104 | 500359 | 2029233200400010 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2001-08-16 | Device Not Retu | No | Reuse | 2004
| 2001-01-16 | Malfunction |  | The positioner moved without command. The behavior was attributed to a defective controller. It was not confirmed if the malfunction occurred during a clinical procedure. No adverse event was reported. | During the troubleshooting over the phone it was discovered that the controller was cause of the problem. Controller was replaced, no further difficulty reported. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516274 | 527106 | 500361 | 2029233200400008 | 1 | GDC | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2001-01-16 | Yes | No | Reuse | 2004
| 2002-01-31 | Malfunction |  | The positioner moved without command. The behavior was attributed to a defective controller. It was not confirmed whether the malfunction occurred during a clinical procedure. No adverse event was reported. | During the troubleshooting over the phone it was discovered that the controller was cause of the problem. Controller was replaced, no further difficulty reported. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 516276 | 527108 | 500363 | 2029233200400012 | 1 | GDC | Manufacturer | Company Representative | Physician | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | Yes | 2002-01-31 | Yes | No | Reuse | 2004
| 2003-04-28 | Malfunction |  | A microjoint debakey grasper had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injjury or adverse outcome reported. | Engineering confirmed that the articulink was fractured and the articulink screw was broken. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT DEBAKEY GRASPER | ENDOSCOPIC INSTRUMENT | 516278 | 527110 | 500365 | 2029233200400046 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2003-04-28 | Malfunction |  | A microjoint micro needle driver had a broken articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink and the articulink screw were broken. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT MICRO NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 516290 | 527122 | 500377 | 2029233200400045 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-04-28 | Yes | No | Reuse | 2004
| 2002-05-28 | Malfunction |  | Micro joint fenestrated grasper broke during the surgery. The instrument head bent and broke upon removal from the pt. The surgeon was able to retrieve the instrument fragments with laparoscopic instruments. No adverse outcome reported. | Engineering confirmed the alleged malfunction. Failure analysis of other zeus instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | MICRO JOINT FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 516292 | 527124 | 500379 | 2029233200400029 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No Answer Provi | No | Reuse | 2004
| 2002-02-08 | Malfunction |  | Event Description: --/-- A micro joint heavy needle driver used for a procedure development lab study has a broken and missing jaw. It was not confirmed whether the malfunction occurred during a clinical procedure or in vivo study. No adverse event was reported. | Manufacturer Narrative: --/-- Unk. --/-- Additional Manufacturer Narrative: --/-- Engineering confirmed the alleged malfunction. Failure analysis of other zeus instruments exhibiting similar failure concluded that the material used for the head may have caused the problem. No further evaluation of the system is expected. | MICRO JOINT HEAVY NEEDLE DRIVER | Reuse | 516295 | 527127 | 500382 | 2029233200400024 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No |  | No | No |  | 2004
| 2003-05-02 | Malfunction |  | A microjoint heavy metzenbaum scissors had a bent articulink. This instrument was used in a clinical case; however, there is no info available to confirm whether the failure occurred during the case. No pt injury or adverse outcome reported. | Engineering confirmed that the articulink was slightly bent, causing the instrument to articulate unevenly in two directions. No definitive root cause of the malfunction was determined. No further evaluation of the instrument is expected. | ZEUS MICRO JOINT METZENBAUM SCISSORS | ENDOSCOPIC INSTRUMENT | 516296 | 527128 | 500383 | 2029233200400052 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2004-02-20 | 2004-02-23 | No | Yes | Health Professional |  | No | No | 2003-05-02 | Yes | No | Reuse | 2004
40015402 | 2004-01-15 | Malfunction |  | A permanent cautery spatula instrument was alleged to have arced during a procedure. The instrument was removed and replaced to continue the case. No pt injury or adverse outcome occurred. | No conclusion can be drawn. Failure analysis of the instrument has not been completed since the instrument has not yet been returned. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 513549 | 524415 | 497761 | 2955842200400002 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-02-13 | 2004-02-13 | No | Yes | Health Professional | 40015402 | No | No | 2004-01-15 | Device Not Retu | No | Reuse | 2004
40015402 | 2004-01-06 | Malfunction |  | A permanent cautery spatula instrument was alleged to have arced during a procedure. No pt injury or adverse outcome occurred. | The permanent cautery spatula instrument was returned and evaluated. Damaged to the plastic sleeving around the cautery spatula was observed. The damaged plastic may have been related to the arcing reported by the customer. It cannot be definitively determined whether the sleeving damaged cause the arcing or resulted from the arcing. There are no indications that the instrument did not meet the manufacturing specifications. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 511873 | 522745 | 496117 | 2955842200400001 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2004-02-06 | 2004-02-10 | No | Yes | Health Professional | 40015402 | No | Yes | 2004-01-09 | Yes | No | Reuse | 2004
40000317 | 2003-11-01 | Malfunction |  | Event Description: --/-- During a procedure, deployed clips would not disengage easily from three clip applier instruments. No pt injury or adverse outcome was reported. --/-- Additional Event Description: --/-- It was alleged that during a procedure a clip applier dropped a clip unexpectedly. The clip was retrieved and removed without harm to the pt. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- No conclusion can be drawn. The hosp has not returned the alleged discrepant instrument. Failure analysis of the instrument could not be performed. --/-- Additional Manufacturer Narrative: --/-- Visual inspection revealed that the grips of the returned instrument were damaged. The damaged grips likely prevented clips from easily releasing from the instrument. It is believed that the instrument was damaged at the customer site due to following reasons: 1) the instrument passed functional testing prior to being shipped; and 2) this instrument had been previously used at the customer site to apply over 24 staples without report of difficulty. Note: this incident was originally reported to the fda on 10/03 as 2955842-2203-00063. At the request of the fda, the mdr has been split into three separate mdr reports. Ref: 2995842-2003-00063, 2955842-2003-00077, 2955842-2003-00078. | CLIP APPLIER INSTRUMENT | Invalid Data | 510489 | 521379 | 494794 | 2955842200300077 | 1 | GDO | Manufacturer | Health ProfessionalCompany Representativ | Other | InitialFollowup | 2004-02-04 | 2004-02-05 | No | Yes | Health Professional | 40000317 | No | No | 2004-01-05 | Yes | No |  | 2004
40015402 | 2003-12-11 | Malfunction |  | Event Description: --/-- A permenent cautery spatula instrument was alleged to have arced during a procedure. No patient injury or adverse outcome occurred. --/-- Additional Event Description: --/-- A permanent cautery spatula instrument was alleged to have arced during a procedure. No pt injury or adverse outcome occurred. | Manufacturer Narrative: --/-- No conclusion can be drawn. Failure analysis of the instrument has not been completed. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Failure analysis of the instrument confirmed evidence of the instrument having arced. No other visual defects were noted. No conclusion of the cause of the arcing has been determined. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 498710 | 509743 | 483582 | 2955842200300075 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2004-01-09 | 2004-01-13 | No | Yes | Health Professional | 40015402 | No | No | 2003-12-11 | Yes | No |  | 2004
40000421 | 2003-12-04 | Malfunction |  | It was reported that a scalpel cautery instrument began to arc during a case. During removal of the instrument from the pt, the insulated electrocautery spatula separated from the instrument. The spatula tip was retrieved by the surgeon and removed from the pt. No pt injury or adverse outcome occurred. The case was continued to completion with no further incident. | Although the scalpel cautery instrument was returned, the insulated cautery spatula accessory was not returned for eval. Failure analysis of the returned scalpel cautery instrument confirmed evidence of arcing. Tensile testing of the scalpel cautery instrument with a new cautery spatula demonstrated that retention was within specification. Since the spatula was not returned, no further eval of retention is expected. The relationship between the arcing and the spatula separation cannot be conclusively determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 498961 | 509991 | 483816 | 2955842200300076 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2004-01-09 | 2004-01-13 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-12-11 | Yes | No | Reuse | 2004
40012607 | 2003-12-05 | Malfunction |  | It was alleged that an instrument arced during a case. The surgeons removed the instrument and completed the case with a replacement instrument without further incident. No pt harm, adverse outcome or injury was reported. | The hospital has not returned the alleged discrepant instrument. Failure analysis of the scalpel cautery instrument could not be performed. Instrument arcing cannot be confirmed. No definitive root cause of the alleged instrument arcing can be determined at this time. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 498376 | 509411 | 483245 | 2955842200300074 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2004-01-07 | 2004-01-09 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-12-08 | Device Not Retu | No | Unkown | 2004
DAVINCI STANDARD INTUITIVE 7 MM HOOK CAUTERY | 2003-08-20 | Malfunction |  | During a davinci assisted laparoscopic cholecystectomy, the tip of the intuitive davinci 7 mm hook cautery came apart. Three separate plastic pieces were identified and withdrawn from the surgical field. X-rays were taken to confirm that all pieces were retrieved. Procedure continued to amputate the gallbladder from the hepatic bed and remove the gallbladder through the epigastric port. The patient tolerated the procedure well. |  | ENDOWRIST PERMANENT CAUTERY HOOK | ROBOTIC SURGICAL CAUTERY | 545520 | 555872 | 528145 | 555872 | 1 | NAY | User Facility |  |  | Initial | 2004-01-05 | 2004-01-05 | No | Yes | Physician | DAVINCI STANDARD INTUITIVE 7 MM HOOK CAUTERY | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2004
IS1000 | 2003-11-24 | No Answer Provided |  | It was reported that during a procedure, a system critical fault occurred. The surgeon could not override the fault. In order to remove a needle driver instrument from a pt, the emergency grip release on a needle driver instrument was used. The procedure was converted to a traditional laparoscopic surgery without any pt injury. | An error (fault) was reported as having occurred. Preliminary investigation concluded that an iod board likely caused the fault. Nevertheless, the customer reported complaint could not be replicated. In 2003, the hospital's system was repaired replacing the iod board. Two weeks later no further difficulty had been reported at the site. There was no alleged malfunction of the needle driver instrument and device eval was not necessary. The emergency grip release on the needle driver instrument worked as designed. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 495474 | 506598 | 480498 | 2955842200300072 | 1 | NAY | Manufacturer | User facility | Physician | Initial | 2003-12-23 | 2003-12-24 | No | Yes | Health Professional | IS1000 | No | Yes | 2003-11-24 | Device Not Retu | No | Reuse | 2003
40012607 | 2003-11-28 | Malfunction |  | It was reported that a permanent cautery hook instrument broke while the surgeon was dissecting in an extreme postition. All instrument components were removed from the pt. The instrument and trocar were removed from the pt simultaneously. No pt injury or adverse event was reported. No additional surgical procedure was required. | Failure analsyis of the returned instrument confirmed that the distal tip of the instrument had fragmented. | PERMANENT CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT | 495476 | 506600 | 480500 | 2955842200300073 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-12-23 | 2003-12-24 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-11-28 | Yes | No | Reuse | 2003
AESOP 3000 | 2003-11-10 | Malfunction |  | Arm of an aesop system moved forward and rotated to the right without command. The surgeon grabbed the arm and prevented it from moving further. The surgeon chose to continue the procedure manually without voice control. No pt injury or adverse event was reported. | During eval of the system by an intuitive surgical field engineer, the alleged failure mode could not be reproduced. In 2003, the controller was replaced. As of one month later, no further difficulty had been reported at the site. | AESOP 3000 | AUTOMATED ENDOSCOPIC POSITIONI | 493543 | 504722 | 478686 | 2955842200300071 | 1 | GCJ | Manufacturer | Health Professional | UNKNOWN | Initial | 2003-12-10 | 2003-12-10 | No | Yes | Health Professional | AESOP 3000 | No | Yes | 2003-11-10 | Yes | No | Reuse | 2003
40000417 | 2003-10-23 | Malfunction | Other | An instrument fragment separated from the instrument but was retrieved. No pt injury or adverse event was reported. The procedure was converted to an open procedure since an appropriate backup instrument was not available. No additional surgical procedure was required. | Failure analysis of the returned instrument confirmed that the distal tip of the instrument had fragmented. No pt damage was reported. It was reported that there was no adverse outcome or injury. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 489598 | 500830 | 474948 | 2955842200300070 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-12-02 | 2003-12-04 | Yes | Yes | Health Professional | 40000417 | No | No | 2003-10-23 | Yes | No | Reuse | 2003
40000315 | 2003-10-17 | Malfunction |  | During a procedure, two clip appliers were dropping clips unexpectedly. The dropped clips were retrieved and removed without harm to the pt. A third clip applier instrument (same model) was used with the da vinci system to complete the case. No pt harm, adverse outcome or injury was reported. | Failure analysis confirmed that the instruments intermittently dropped clips. D6: lot # 1207011 31. | CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 489604 | 500836 | 474954 | 2955842200300068 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-12-02 | 2003-12-04 | No | Yes | Health Professional | 40000315 | No | Yes | 2003-10-17 | Yes | No | Reuse | 2003
40012607 |  | Malfunction |  | The instrument arced during a case. No pt harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument confirmed evidence of arcing. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 489609 | 500841 | 474959 | 2955842200300069 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-12-01 | 2003-12-04 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-10-28 | Yes | No | Reuse | 2003
IS1000 | 2003-10-20 | Malfunction | Other | During a procedure, the da vinci system faulted and could not be restarted. No patient harm was reported. The procedure was converted to traditional open surgical techniques to finish the surgery. | Errors (faults) were reported as having occurred on the patient side manipulator (psm). Preliminary investigation concluded that the faults were likely the cause of the unrecoverable system fault. In 2003, the hospital's system was repaired by replacing the psm. The hospital's system was then tested and passed all homing and system dance tests. As of three weeks later, no further difficulty reported at the site. The psm was returned to isi for evaluation. During failure analysis evaluation of the psm at isi no trouble was found. The psm was tested and inspected using software, electrical, mechanical and performance test methods. No unusual observations were noted. The customer reported complaint could not be replicated. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 488872 | 500105 | 474237 | 2955842200300065 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-11-26 | 2003-11-26 | Yes | Yes | Health Professional | IS1000 | No | No | 2003-10-20 | Yes | No | Reuse | 2003
40012607 | 2003-10-22 | Malfunction |  | The instrument leaked current and arced during a case. The surgeons removed the instrument and completed the case with a replacement instrument without further incident. No patient harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure of the permanent hook instrument confirmed evidence of arcing at goose-neck. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 488947 | 500180 | 474311 | 2955842200300066 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-11-26 | 2003-11-26 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-10-22 | Yes | No | Reuse | 2003
40012607 | 2003-10-24 | Malfunction |  | The instrument tip broke during a surgery and the tip was retrieved. The surgeons removed the instrument and completed the case with a replacement instrument without further incident. The case was completed with the da vinci system. No patient harm was reported. No adverse outcome or injury was reported. | Failure analysis of the instrument confirmed the customer reported complaint that the hook was no longer attached to the instrument. In addition, the base of the instrument showed evidence of having melted. The relationship between the melted base and the instrument fragmentation cannot be conclusively determined. The hook was not returned to isi for evaluation. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 489596 | 500828 | 474946 | 2955842200300067 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2003-11-26 | 2003-11-26 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-10-24 | Yes | No | Unkown | 2003
400126XX | 2003-10-02 | No Answer Provided |  | During a surgical procedure, a plastic fragment of the permanent cautery instrument separated from the instrument. The surgeon was unable to locate the fragment. The case was completed with the da vinci system. No pt harm was reported. No adverse outcome or injury was reported. |  | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 482831 | 494125 | 468486 | 2955842200300063 | 1 | GEI | Manufacturer |  | Other | Initial | 2003-10-31 | 2003-10-31 | No | Yes | Health Professional | 400126XX | No | No |  | No Answer Provi | No An | Invalid Data | 2003
| 2002-10-10 | Injury | DisabilityLife ThreateningRequired Intervention | Event Description: --/-- Pt underwent minimal invasion, robotics assisted, coronary bypass surgery -lima to lad-, with very bad results, and pt wants to be sure that their situation has been reported properly. First, the lima artery, grafted to the lad, reclogged approx. 90% -in the area where they stitched-, and the leg vein they used as a backuup, reclogged 100% all within 4 months. Secondly, pt has been, and continues to be in chronic pain, since the operation, throughout the left side of their chest. Pt has seen two pain drs and is under the care of a pain dr now, and waiting for a thoracic epidural. --/-- Additional Event Description: --/-- The da vinci surgical system is currently being used for mobilization of the internal mammary artery (ima) during coronary artery bypass procedures under fda clearance k002489 (issued 2001), and for totally endoscopic coronary artery bypass procedures under the auspices of ide # g010011. Based on the info provided, the event does not match any complaint recorded for the da vinci surgical system or its instruments, nor does it match any procedure performed under the ide. The medwatch report indicates that the user facility was notified and intuitive surgical has not been contacted by the facility, leading tothe conclusion that a relationship between the pt condition and the da vinci system used during the procedure could not be established by the hospital. Company's conclusion is that this is not an mdr reportable event based on the following: no death or serious injury attributed to the device occurred. No malfunction of the device was reported to have occurred. | Manufacturer Narrative: --/-- | DA VINCI ROBOTICS SYSTEM |  | 481477 | 492776 | 467168 | MW1029863 | 1 | NAY | Voluntary |  | Patient | InitialFollowup | 2003-10-21 | 2003-10-21 | Yes | Yes | Health Professional |  | Yes | No |  |  |  |  | 2003
40011005 | 2003-09-15 | Malfunction |  | During a procedure, the instrument shaft broke proximal to the wrist. No fragments or components were reported to have fallen in the pt. The procedure was completed using another forceps instrument. No pt harm or pt injury was reported. | Failure analysis confirmed the main tube was damaged at the distal end. The affected section of tube was immediately adjacent to the proximal "clevis". There was no damaged to the "clevis". No conclusion as to the cause of the tube damage has been determined. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 479492 | 490765 | 465257 | 2955842200300060 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-10-15 | 2003-10-17 | No | Yes | Health Professional | 40011005 | No | No | 2003-09-15 | Yes | No | Reuse | 2003
40000423 | 2003-09-17 | Malfunction |  | During a procedure, the cautery spatula separated form the cautery instrument. The cautery spatula fell into the pt and was retrieved. The procedure was completed using the same instrument but with a new cautery spatual. No pt harm or pt injury was reported. The surgery was completed without further incident. | The instrument has not been returned. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 479932 | 491219 | 465686 | 2955842200300061 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-10-17 | 2003-10-17 | No | Yes | Health Professional | 40000423 | No | No | 2003-09-17 | Device Not Retu | No | Unkown | 2003
40009304 | 2003-09-10 | Malfunction |  | During a procedure, the distal end of the instrument shaft broke. No fragments or components were reported to have fallen in the patient. The instrument and trocar cannula were removed from the patient stimultaneously. Once outside the patient, the surgeon cut and removed the forceps's distal grasping tip in order to remove the damaged forceps through the trocar. No patient harm, adverse outcome or injury was reported. | To remove the damaged forceps through the trocar, the surgeon was forced to cut and remove the forcep's distal tip. No definitive conclusion as to the cause of the forceps breakage could be drawn since the surgeon had cut the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 478378 | 489629 | 464189 | 2955842200300058 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2003-10-10 | 2003-10-10 | No | Yes | Health Professional | 40009304 | No | Yes | 2003-09-12 | Yes | No | Reuse | 2003
40012606 | 2003-09-12 | Malfunction |  | During a surgical procedure, the surgeon caught the hook with a forceps instrument causing the hook to "snap". The fragments were removed with the forceps instrument and handed to a lap grasper for removal. A second permanment cautery hook instrument was used to complete the procedure. The case was completed with the da vinci system. No pt harm, adverse outcome or injury was reported. | Failure analysis of the returned instrument confirmed the distal clevis had fractured. No evaluation of the clevis fragment could be performed since the fragment was not returned to isi. Evaluation of the customer reported scenario has led isi to conclude that the product problem was caused by user error, i. E. Mishandling of the surgical instrumrnts. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURICAL INSTRUMENT | 478543 | 489800 | 464351 | 2955842200300057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2003-10-10 | 2003-10-10 | No | Yes | Health Professional | 40012606 | No | Yes | 2003-09-12 | Yes | No | Reuse | 2003
40015402 | 2003-09-12 | Malfunction |  | In 2003, a permanent cautery spatula instrument was alleged to have arced during a procedure. No patient injury or adverse outcome occurred. The incident was reported to the intuitive surgical (isi) account manager and they instructed them to remove the instrument, tag it and prepare to return the instrument to isi. Six days later the same surgeons used the same instrument for another procedure, where it arced again. No patient injury or adverse event occurred. The alleged discrepant instrument is no longer at the hospital. No patient injury or adverse outcome was reported for either patient. | No conclusion can be drawn. Failure analysis of the instrument has not been completed. The instrument is currently in transit to isi for evaluation. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 478800 | 490058 | 464590 | 2955842200300059 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2003-10-10 | 2003-10-10 | No | Yes | Health Professional | 40015402 | No | Yes | 2003-09-12 | Device Not Retu | No | Reuse | 2003
40012607 | 2003-09-09 | Malfunction |  | During a surgical procedure, the lower tip of insulation broke on a permanent cautery hook instrument. A second permanent cautery hook instrument was used to complete the procedure. The case was completed with the da vinci system. No patient harm, adverse outcome or injury was reported. | Failure analysis confirmed that the distal clevis was fractured. The component that had been reported as the "lower tip of insulation" was determined by engineering to be the distal clevis. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 478314 | 489565 | 464125 | 2955842200300056 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2003-10-09 | 2003-10-09 | No | Yes | Health Professional | 40012607 | No | Yes | 2003-09-09 | Yes | No | Reuse | 2003
IS1000 | 2003-09-07 | Malfunction |  | During use of the da vinci system, it was alleged that the camera arm went limp and "fell on the patient". The surgeon was able to recover by clutching the arm and moving it back to where it needed it be. No patient injury or adverse outcome was reported. | No trouble was found with the da vinci surgical system during evaluation. The system passed homing and system dance tests. The system error log contained only one entry recording that the emergency stop button had been depressed. There were no other entries in the log. The field engineer was unable to duplicate the customer alleged complaint. Further research regarding the hospital's instrument setup for the procedure, has led engineering to conclude that an instrument arm may have collided with the camera arm thereby not allowing the camera to move and causing the arm to look "limp". | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 478740 | 489999 | 464531 | 2955842200300055 | 1 | KOG | Manufacturer | Health Professional | Other | Initial | 2003-10-08 | 2003-10-08 | No | Yes | Health Professional | IS1000 | No | Yes | 2003-09-08 | Yes | No | Initial | 2003
40000317 | 2003-09-08 | Malfunction |  | During use of a clip applier, it was noticed that the deployed clip "would skip to one side of jaw". The clip applier was removed and patient bleeding was managed by another method. No patient injury or adverse outcome was reported. | Visual inspection revealed that the return instrument was damaged. Failure analysis concluded that a dented groove at the distal end of the clip applier prevented clips from releasing from the instrument. | CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 478774 | 490032 | 464564 | 2955842200300054 | 1 | GDO | Manufacturer | Company Representative | Other | Initial | 2003-10-08 | 2003-10-08 | No | Yes | Health Professional | 40000317 | No | No | 2003-09-08 | Yes | No | Initial | 2003
40011005 | 2003-09-04 | Malfunction |  | The distal end of the instrument shaft broke. No fragments or components were reported to have fallen in the pt. The procedure was completed with another bipolar forceps instrument. No pt harm or pt injury was reported. | Failure analysis confirmed the main tube was damaged at the proximal clevis. No conclusion as to the cause of the fracture has been determined. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 476829 | 487969 | 462644 | 2955842200300053 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-09-30 | 2003-09-30 | No | Yes | Health Professional | 40011005 | No | No | 2003-09-04 | Yes | No | Reuse | 2003
40000421 | 2003-08-27 | Malfunction |  | The cautery spatula separated from the cautery instrument. The cautery spatula fell into the patient and was retrieved 5 minutes after the incident. The spatula and instrument were removed and the procedure was completed with another instrument. No patient harm or patient injury was reported. The surgery was completed without further incident. | No discrepancies of the instrument were observed during failure analysis. The cautery spatula used during the surgery was not returned and therefore could not be evaluated. When tested with a new cautery spatula, the insertion and retention forces for the spatula and instrument were within specification. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 475987 | 487099 | 461810 | 2955842200300052 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2003-09-26 | 2003-09-26 | No | Yes | Health Professional | 40000421 | No | No |  | Yes | No | Reuse | 2003
40012602 | 2003-08-26 | Malfunction |  | During a surgical procedure a permanent cautery hook instrument cracked and several pieces fell into the pt. They were able to retrieve all but 1 of the fragments. It was reported that the remaining fragment was very small and could not be located using x-ray. The case was continued and completed with the da vinci system. It was reported that there was no pt harm. | No conclusion can be drawn. The hosp has not returned the alleged discrepant instrument. Failure analysis of the instrument could not be performed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 475738 | 486851 | 461571 | 2955842200300051 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-09-25 | 2003-09-25 | No | Yes | Health Professional | 40012602 | No | No | 2003-08-26 | Device Not Retu | No | Unkown | 2003
40000421 | 2003-08-11 | Malfunction |  | During a procedure, the cautery spatula separated from the cautery instrument. The cautery spatula was retrieved and removed from the pt. No pt harm or pt injury was reported. The surgery was completed without further incident. | No discrepancies of the instrument were observed during failure analysis. The cautery spatula used during the surgery was not returned and therefore could not be evaluated. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 473172 | 484282 | 459092 | 2955842200300050 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2003-09-11 | 2003-09-12 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-08-14 | Yes | No | Reuse | 2003
40000316 | 2003-07-17 | Malfunction |  | During use of a clip applier, it was noticed that the deployed clips remained affixed to the clip applier at one side of the instrument tip. The clip applier instrument was removed from the patient and replaced with another to continue the case to completion. No patient injury or adverse outcome was reported. | Failure analysis of the returned instrument confirmed that deployed clips did not easily fall out of the clip applier groove when the instrument was opened. | CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 468155 | 479315 | 454301 | 2955842200300048 | 1 | GDO | Manufacturer | Health Professional | Physician | Initial | 2003-08-15 | 2003-08-15 | No | Yes | Health Professional | 40000316 | No | Yes |  | Yes | No | Initial | 2003
40000417 | 2003-07-15 | Malfunction |  | During use of an scalpel cautery instrument, it was noticed that the instrument tip was melting. The cautery instrument was removed from the pt and replaced with another to continue the case to completion. No pt injury or adverse outcome was reported. | Failure analysis of the returned instrument confirmed that a distal clevis had melted. Evidence of arcing from a crack in the yaw pulley was detected. The arcing likely caused the clevis to melt. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 467832 | 479019 | 454033 | 2955842200300047 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-08-14 | 2003-08-14 | No | Yes | Health Professional | 40000417 | No | No | 2003-07-15 | Yes | No | Reuse | 2003
400012607 | 2003-07-09 | Malfunction |  | After a permanent cautery hook was removed from a patient and placed on the instrument tray, it was noticed that the tip of the instrument was missing. According to the surgeon and a company representative present at the case, the tip was attached to the instrument as it was being removed from the patient. Both confirmed the tip was intact while still in view on the monitor. The surgeon reported that he does not believe the tip was left in the pateint. No patient harm or patient injury was reported. A post-procedure flat x-ray was performed. The x-ray was negative for radio-opaque foreign bodies, including any artifact resembling a fragment from a permanent cautery hook. The patient was released without incident. | The instrument was not returned for failure analysis. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 466653 | 477845 | 452884 | 2955842200300046 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-08-08 | 2003-08-08 | No | Yes | Health Professional | 400012607 | No | Yes | 2003-07-09 | No | No | Unkown | 2003
400004XX | 2003-07-01 | Malfunction |  | During a procedure, the 15 degree scalpel blade separated from the scalpel cautery instrument. The blade accessory was retrieved and the case was completed. No patient harm or patient injury was reported. | The hospital did not return the instrument in question or the 15 degree blade. Failure analysis of the instrument and accessory could not be performed. The customer complaint cannot be confirmed. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 465327 | 476499 | 451623 | 2955842200300045 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2003-07-31 | 2003-07-31 | No | Yes | Health Professional | 400004XX | No | Yes | 2003-07-01 | Device Not Retu | No | Unkown | 2003
40000417 | 2003-05-20 | Malfunction |  | During a procedure, the tip of an instrument broke. It was alleged that "part of the instrument [was] missing during the procedure but the piece was immediately removed. " it is suspected that the breakage occurred during a collision with another instrument. No pt harm or pt injury was reported. | Failure analysis confirmed the proximal clevis was broken at the side of the main tube. However, the identity and source of the fragment retrieved by the surgeon cannot be established. It was reported that the instrument damage might have been caused by a collision between instruments. Nevertheless, no definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 465097 | 476270 | 451399 | 2955842200300044 | 1 | NAY | Manufacturer | Distributor | UNKNOWN | Initial | 2003-07-30 | 2003-07-30 | No | Yes | Health Professional | 40000417 | No | No | 2003-06-30 | Yes | No | Reuse | 2003
40000421 | 2003-06-12 | Malfunction |  | While the surgeon was mobilizing the cystic duct for transection, it was reported that a scalpel cautery instrument arced from the instrument wrist and transected the cystic duct. The ends of the duct were then tied off and there were no complications for the pt. The cautery instrument was removed from the pt and replaced with another scalpel cautery instrument to continue the case to completion. No pt injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument confirmed evidence of arcing at the yaw pulley located at the instrument wrist. Evidence of arcing emanates from a crack in the corner of the accessory slot in the yaw pulley. The crack in the yaw pulley was likely due to repeated insertion of accessories. Arcing was likely due to the integrity of the insulation at the yaw pulley becoming compromised. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 463957 | 475087 | 450317 | 2955842200300043 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2003-07-24 | 2003-07-24 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-06-19 | Yes | No | Reuse | 2003
IS1200 | 2003-06-17 | Malfunction |  | While performing an instrument tool exchange during a procedure, the da vinci system faulted without providing a system alarm and could not be restarted. The surgery was completed laparoscopically. No patient harm, injury or other adverse events were reported. | Root cause analysis determined that the system fault and resulting absence of an alarm was due to a software issue on the is1200 system. Internal investigation has determined this to be a rare event. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 462474 | 473625 | 448921 | 2955842200300042 | 1 | NAY | Manufacturer | Company Representative | Other | Initial |  | 2003-07-16 | No | Yes | Health Professional | IS1200 | No | No |  | Yes | No | Reuse | 2003
40012604 | 2003-06-16 | Malfunction |  | During a surgical procedure, the surgeon had a sling around the esophagus and was handling it with a cadiere forceps controlled by his left hand, and a permanent cautery hook instrument controlled by his right hand. The surgeon moved the position of the endoscopic camera so as to center the esophagus in the field of view and did not maintain visual contact with the instruments. In this position, the surgeon attempted to move the permanent cautery hook back into view, and a noise was heard and movement of the instrument arm was noticed. The permanent hook cautery instrument appeared to be damaged after this point and was replaced with another instrument to complete the case. No pt harm or injury was reported. | Failure analysis confirmed that the distal end of the permanent hook instument had been damaged. The damaged to the instrument was likely caused by the surgeon moving the permanent cautery hook when it was outside his field of view, resulting in the hook getting caught in the cadiere forceps. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 462603 | 473753 | 449046 | 2955842200300041 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-07-16 | 2003-07-16 | No | Yes | Health Professional | 40012604 | No | No | 2003-06-16 | Yes | No | Reuse | 2003
IS1200 | 2003-06-09 | Malfunction |  | During a surgical procedure with the da vinci surgical system, unintended motion with a 5mm monopolar electrosurgical instrument was reported to have caused a hematoma. The surgeon removed the 5mm instrument from the da vinci system and replaced it with an 8mm instrument to complete the case. The surgeon stated that the aforementioned incident was a "clinically insignificant event. " the incident did not cause serious injury to the pt, did not alter the operation nor compromise the clinical outcome of the case. | Upon investigation, root cause analysis revealed that two data values in the software parameters for the 5mm monopolar cautery instrument were transported. These two data files control how the system enforces internal limits when the instrument is used at the end of it's range of motion. The result is that if the tip is articulated to the end of it's range of motion, the instrument may move in the insertion axis when surgeon is attempting to roll the instrument tip. The da vinci system at facility was the only site where this data file had been activated. The instrument data file has since been corrected on the facility system. This file is not available for use on any other systems in the field. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 460753 | 471903 | 447293 | 2955842200300037 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Other | Initial | 2003-07-08 | 2003-07-14 | No | Yes | Health Professional | IS1200 | No | Yes | 2003-06-09 | Yes | No | Reuse | 2003
40000423 | 2003-06-12 | Malfunction |  | A scalpel cautery instrument was alleged to have an electrical problem in the shaft of the instrument. After 5 minutes of usage, it was reported that there was smoke coming out of the instrument shaft. The cautery instrument was removed from the patient and replaced with another scalpel cautery instrument to continue the case to completion. No patient injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument confirmed evidence of arcing at the yaw pulley, which is located at the distal end of the instrument. There were no signs of arcing elsewhere on the instrument (i. E. No signs of arcing within the shaft of the instrument). Smoke can occur during operation of an electrosurgical device. Insufflation pressure may have forced the smoke to exit via the instrument shaft. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 461474 | 472621 | 447988 | 2955842200300040 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-07-11 | 2003-07-11 | No | Yes | Health Professional | 40000423 | No | No | 2003-06-12 | Yes | No | Reuse | 2003
IS1000 | 2003-06-11 | Malfunction |  | During a procedure, the da vinci system faulted and could not be restarted. The surgery was converted to a traditional laparoscopic procedure. No patient harm, injury or other adverse event were reported. | Upon investigation, the root cause analysis revealed that a potentiometer on a patient side manipulator (psm) arm failed. The da vinci system was repaired by replacing the psm. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 461484 | 472631 | 447998 | 2955842200300039 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-07-11 | 2003-07-11 | No | Yes | Health Professional | IS1000 | No | No | 2003-06-11 | Yes | No | Reuse | 2003
40012606 | 2003-06-11 | Malfunction |  | During a procedure, insulation from a permanent cautery hook instrument separated from the instrument. It was reported that the insulation was immediately retrieved from the patient. No portion of the permanent cautery hook instrument is believed to have been left in the patient. No patient harm, adverse outcome or injury was reported. A second permanent cautery hook instrument was used to complete the procedure. The case was completed with the da vinci system. | Analysis of the instrument confirmed that insulation was missing from the distal tip. Visual inspection of the insulation remaining on the instrument showed signs of mechanical abrasion. The cause of the abrasion could not be determined. The insulation fragment was not returned for evaluation. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 461493 | 472640 | 448007 | 2955842200300038 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2003-07-11 | 2003-07-11 | No | Yes | Health Professional | 40012606 | No | No | 2003-06-11 | Yes | No | Reuse | 2003
40011002 | 2003-06-03 | Malfunction |  | During a procedure, one grip (i. E. Jaw) of the bipolar forceps separated from the instrument. The grip was immediately retrieved from the patient by the patient side surgeon. No portion of the bipolar forceps instrument was left in the patient. No patient harm was reported. The case was completed with the da vinci system. | Analysis of the instrument determined that the grip failure occurred where the grip was crimped to the conductor wire. | BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 462441 | 473592 | 448888 | 2955842200300036 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-07-03 | 2003-07-03 | No | Yes | Health Professional | 40011002 | No | No | 2003-06-03 | Yes | No | Reuse | 2003
IS1000 | 2003-06-02 | Malfunction |  | During a procedure, the da vinci system faulted and could not be restarted. No pt harm was reported. The procedure was converted to traditional laparoscopic surgical techniques to finish the surgery. | H6- a printed circuit board was determined to be the cause of the system fault. The board was replaced. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 459583 | 470743 | 446158 | 2955842200300035 | 1 | KOG | Manufacturer | Company Representative | Other | Initial | 2003-07-02 | 2003-07-02 | No | Yes | Health Professional | IS1000 | No | No | 2003-06-03 | Yes | No | Reuse | 2003
40011002 | 2003-05-30 | Malfunction |  | A bipolar forceps instrument was reported as working intermittently. After the case was completed, blackened tissue was found on the surface at two abdominal incisions for the surgical ports. The blackened tissue was removed and the ports were closed. A mono-polar instrument supplied by another co was also used on the pt prior to the bipolar forceps instrument. There was no reported pt harm or adverse event. | Failure analysis determined the root cause of the intermittent functioning of the instrument was due to poor conductor wire continuity. During evaluation it was determined that the conductor wire had become detached at the distal end of the instrument. No definitive conclusion has been determined regarding the cause of the blackened tissue at the ports. A mono-polar instrument supplied by another co was also used on the prior to the bipolar forceps instrument through these same ports. | BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 458455 | 469587 | 445062 | 2955842200300034 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-06-27 | 2003-06-27 | No | Yes | Health Professional | 40011002 | No | No | 2003-05-30 | Yes | No | Reuse | 2003
40004816 | 2003-05-29 | Malfunction |  | After successful placement of an electrical lead using a long tip forceps instrument, the surgical staff was not able to remove the instrument through the cannula as expected. There were no indications that the instrument fragmented. There were no reported complications due to the instrument removal. No pt harm or pt injury was reported. | No conclusion can be drawn. The hosp has not returned the alleged discrepant instrument. Failure analysis of the instrument could not be performed. | LONG TIP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 458456 | 469588 | 445063 | 2955842200300033 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2003-06-27 | 2003-06-27 | No | Yes | Health Professional | 40004816 | No | No | 2003-05-29 | Device Not Retu | No | Unkown | 2003
IS1200 A42P4 | 2003-05-22 | Malfunction |  | In 2003, one of the is1200 systems shut-down due to an error code fault, and there were subsequent persistent faults with the same error code. This malfunction occurred prior to starting a surgical procedure and there was no adverse event. | Upon investigation, the root cause analysis revealed that a printed circuit board assembly, unique to the system configuration, was the cause of the malfunction. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 457018 | 468092 | 443628 | 2955842200300032 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-06-20 | 2003-06-20 | No | Yes | Health Professional | IS1200 A42P4 | No | No |  | Yes | No | Reuse | 2003
40000419 |  | Malfunction |  | A scalpel cautery instrument was alleged to have arced. The cautery instrument was removed from the pt and replaced with another scalpel cautery instrument to continue the case to completion. No pt injury or adverse outcome was reported. There was no impact on the pt or the procedure. | No conclusion can be drawn. Although failure analysis of the scalpel cautery instrument confirmed evidence of arcing on the yaw pulley, user mishandling may have contributed to the incident. Discussion with the site revealed that misuse of the instrument likely occurred. During the surgical procedure, another surgical instrument was used to clean tissue from the scalpel cautery instrument. The site stated that such cleaning occasionally requires that the cautery spatula accessory tip be reinserted into the instrument. The instrument ifu specifically states: "do not use an instrument to clean debris off another instrument intraoperatively. This may result in damgae to the instruments. ". | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 456910 | 467984 | 443522 | 2955842200300031 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2003-06-19 | 2003-06-19 | No | Yes | Health Professional | 40000419 | No | Yes | 2003-05-20 | Yes | No | Reuse | 2003
40000417 | 2003-04-09 | Malfunction |  | The tip of a scalpel cautery instrument with a cautery spatula attached, was alleged to have "burned" during a procedure. No pt injury or adverse outcome was reported. | The hosp has not returned the alleged discrepant instrument. Failure analysis of the scalpel cautery instrument could not be performed. Instrument "burning" cannot be confirmed. No definitive root cause of the alleged instrument "burning" can be determined at this time. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 455104 | 466162 | 441762 | 2955842200300030 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2003-06-11 | 2003-06-11 | No | Yes | Health Professional | 40000417 | No | Yes | 2003-05-12 | Device Not Retu | No | Reuse | 2003
40000421 | 2003-05-09 | Malfunction |  | A scalpel cautery instrument was alleged to have arced. The cautery instrument was removed from the pt and replace with another scalpel cautery instrument to continue the case to completion. No pt injury or adverse outcome was reported. | The hospital has not returned the alleged discrepant instrument. Failure analysis of the scalpel cautery instrument could not be performed. Instrument arcing cannot be confirmed. No definitive root cause of the alleged instrument arcing can be determined at this time. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 454504 | 465561 | 441194 | 2955842200300029 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-06-05 | 2003-06-05 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-05-09 | Device Not Retu | No | Unkown | 2003
40000421 | 2003-03-26 | Malfunction |  | During a procedure, the cautery spatula separated from the cautery instrument. The spatula was retrieved and the case was completed. No patient harm or patient injury was reported. The patient was released without further incident. | No discrepancies were observed during failure analysis of the instrument. The cautery spatula from the incident was not returned and therefore could not be evaluated. When a new cautery spatula was inserted into the instrument, the pullout force met specification. No definitive conclusion as to the cause of the event can be determined. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 453922 | 464997 | 440631 | 2955842200300028 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2003-06-04 | 2003-06-04 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-04-30 | Yes | No | Reuse | 2003
40009303 | 2003-04-02 | Malfunction |  | On may 1, 2003 the site reported that they were unable to remove the prograsp grasper instrument from the trocar and it broke. The instrument was removed and replaced with another instrument and the procedure was completed. There were no reported complications due to the resultant instrument removal. No patient harm or patient injury was reported. In 05/2003 the following information regarding the same incident was received from the site. "patient in hospital for myomectomy. Davinci prograsp endowrist broke while in the patient. Unable to remove from trocar - insulation from programmer flaked off and fell into abdomen. ". | Failure analysis concluded that the proximal clevis fractured at the tube tip. A chip was missing from the proximal clevis. | PROGRASP GRASPER | ENDOSCOPIC INSTRUMENT | 453305 | 464367 | 440042 | 2955842200300027 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2003-05-30 | 2003-05-30 | No | Yes | Health Professional | 40009303 | No | Yes | 2003-05-01 | Yes | No | Reuse | 2003
PROGRASP ENDOWRIST | 2003-04-02 | Malfunction |  | Pt in hospital for myomectomy. Davinci prograsp endowrist broke while in the pt. Unable to remove from trocar - insulation from programmer flaked off and fell into abdomen. |  | PROGRASP ENDOWRIST | PROGRASP ENDOWRIST | 453448 | 464526 | 440183 | MW1028619 | 1 | GEI | Voluntary |  | Other | Initial | 2003-05-07 | 2003-05-30 | No | Yes | Health Professional | PROGRASP ENDOWRIST | Yes | Yes |  |  |  |  | 2003
PROGRASP ENDOWRIST | 2003-04-02 | No Answer Provided |  | Pt in hospital for myomectomy. Davinci prograsp endowrist broke while in the pt. Unable to remove from trocar - insulation from programmer flaked off and fell into abdomen. |  | PROGRASP ENDOWRIST | PROGRASP ENDOWRIST | 453305 | 462554 | 440042 | MW1028512 | 1 | GEI | Voluntary |  | Other | Initial | 2003-05-07 | 2003-05-27 | No |  | Invalid Data | PROGRASP ENDOWRIST | No Answer Provided | Yes |  |  |  |  | 2003
40012606 | 2003-04-22 | Malfunction |  | The instrument leaked energy during a case. The surgeons removed the instrument and complete the case with a replacement instrument without patient harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument confirmed evidence of arcing at goose-neck and proximal clevis. The insulation sleeve on the goose-neck had been mechanically damaged and is the likely cause of the arcing. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 451447 | 462498 | 438205 | 2955842200300026 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2003-05-15 | 2003-05-21 | No | Yes | Health Professional | 40012606 | No | Yes | 2003-04-22 | Yes | No | Reuse | 2003
400110 | 2003-04-21 | Injury | Required Intervention | Biopollar electrosurgery malfunction requiring change in surgical procedure approach from laparoscopic to open laparotomy procedure. |  | DAVINCI | BIPOLAR ELECTROSURGERY | 454671 | 465728 | 441356 | 465728 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2003-04-25 | 2003-05-21 | Yes | Yes | Invalid Data | 400110 | Device Returned To Manufacturer | Yes |  |  |  |  | 2003
40011004 | 2003-04-21 | Injury |  | During a procedure, it was alleged that the bipolar forceps instrument was no longer becoming energized. The surgeon removed the forceps and replaced it with a like instrument. The same problem occurred with the second instrument. Due to the instruments only being intermittently energized, the procedure was converted to an open laparotomy procedure. | Failure analysis of both instruments determined that the conductor wires were detached at the distal end of the instrument. Add'l lot #1803031 49. | BIOPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 454677 | 465734 | 441356 | 2955842200300025 | 1 | GEI | Manufacturer | Health ProfessionalCompany Representativ | Other | Initial | 2003-05-15 | 2003-05-21 | No |  | Health Professional | 40011004 | No | No | 2003-04-21 | Yes | No | Reuse | 2003
40000417 | 2003-04-16 | Malfunction |  | It was alleged that while the instrument was being rotated to reach tissue that the distal shaft broke and the distal end was abnormally "twisted". There were no reported complications due to the breakage. No patient harm or patient injury was reported. | No patient damage was reported. It was reported that there was no adverse outcome or injury. | SCALPELELECTROCAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 450410 | 461468 | 437206 | 2955842200300024 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-15 | 2003-05-16 | No | Yes | Health Professional | 40000417 | No | No | 2003-04-16 | Yes | No | Reuse | 2003
40000417 | 2003-04-15 | Malfunction |  | It was alleged that while the instrument was loaded and under axial-rotation, that an audible "crack" was heard and the instrument shaft broke at the distal end. The instrument wrist was reported as being abnormally bent. There were no reported complications due to the breakage. No patient harm or patient injury was reported. | No patient damage was reported. It was reported that there was no adverse outcome or injury. | SCALPELELECTROCAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 450415 | 461473 | 437211 | 2955842200300023 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-15 | 2003-05-16 | No | Yes | Health Professional | 40000417 | No | No | 2003-04-16 | Yes | No | Reuse | 2003
IS1200 A42 | 2003-04-16 | Malfunction |  | During installation of an upgrade to add an additional arm to the pt side cart, a malfunction was observed which resulted in uncontrolled motion of a master took manipulator (mtm). Intuitive surgical field personnel observed the malfunction during a power-on "homing" test during installation and not during surgery. | Upon investigation, the root cause analysis revealed that a printed circuit board assembly, unique to the is1200 system configuration, was the cause of the malfunction. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 450289 | 461347 | 437085 | 2955842200300022 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-15 | 2003-05-15 | No | Yes | Health Professional | IS1200 A42 | No | No | 2003-04-16 | Yes | No | Reuse | 2003
40000421 | 2003-04-08 | Malfunction |  | A scalpel cautery instrument was alleged to have arced at the wrist of the instrument. The cautery instrument was removed from the pt and replaced with another scalpel cautery instrument to continue the case to completion. No pt injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument confirmed evidence of arcing on the distal pulley. A crack in the pulley indicates a compromise in the insulation integrity. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 448328 | 459386 | 435186 | 2955842200300021 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-08 | 2003-05-08 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-04-08 | Yes | No | Reuse | 2003
40000417 | 2003-04-04 | Malfunction |  | It was alleged that a scalpel cautery instrument began to arc at the "elbow of the instrument". The affected instrument was removed and replaced to complete the procedure. During the alleged incident the insertable cautery hook accessory was used with the scalpel cautery instrument. No patient injury or adverse outcome was reported. | The alleged discrepant instrument was not returned. Therefore, failure analysis of the instrument could not be performed. Failure analysis of instruments returned for similar complaints have showed evidence of arcing at a knitline crack. The insertable hook is no longer commercially available from intuitive surgical. The permanent hook cautery instrument has replaced the insertable cautery hook. No further evaluation is expected. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 447595 | 458632 | 434468 | 2955842200300020 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-02 | 2003-05-02 | No | Yes | Health Professional | 40000417 | No | No | 2003-04-04 | Device Not Retu | No | Reuse | 2003
40000421 | 2003-04-03 | Malfunction |  | Event Description: --/-- Two scalpel cautery instruments were alleged to have arced. The following was alleged: "screws burning tissue in area where it should not be cauterizing. " the cautery instruments were removed from the patient and replaced with a third scalpel cautery instrument to continue the case to completion. No patient injury or adverse outcome was reported. --/-- Additional Event Description: --/-- Two scalpel cautery instruments were alleged to have arced. The following was alleged: "screws buring tissue in areas where it should not be cauterizing. " the cautery instruments were removed from the pt and replaced with a third scalpel cautery instrument to continue the case to completion. No pt injury or adverse outcome was reported. | Manufacturer Narrative: --/-- No screws are present on the distal end of the instrument. The customer may have noticed arcing from a a pin the distal end of the device. Failure analysis of the scalpel cautery instrument confirmed evidence of arcing on the distal clevis. No other visual flaws were noticed. Root cause of arching is indeterminate. --/-- Additional Manufacturer Narrative: --/-- No screws are present on the distal end of the instrument. The customer may have noticed arcing from a pin at the distal end of the instrument. Failure analysis of the scalpel cautery instrument confirmed evidence of arcing on the pulley. Cracks in the pulley indicate a compromise in the insulation integrity. | SCALPEL CAUTERY INSTRUMENT | Reuse | 447599 | 458636 | 434472 | 2955842200300018 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2003-05-02 | 2003-05-02 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-04-03 | Yes | No |  | 2003
40012604 | 2003-04-01 | Malfunction |  | It was alleged that the instrument arced during the case. The surgeons chose to continue the case with the arcing instrument. No patient harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument confirmed evidence of arcing at the pulley and distal clevis. No definitive root cause of the instrument arcing can be determined at this time. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 447591 | 458628 | 434464 | 2955842200300017 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-05-01 | 2003-05-01 | No | Yes | Health Professional | 40012604 | No | No | 2003-04-01 | Yes | No | Reuse | 2003
IS100004 | 2003-03-26 | Injury | Required Intervention | System failure during a da vinci surgical procedure. Surgeon was unable to restart the system. System error message functioned appropriately. The case was converted and completed. No other patient injury or adverse outcome was reported. | Failure analysis of the system concluded that the system failure and error message were due to a malfunctioning psm. The da vinci surgical system was repaired by replacing the psm. Failure analysis of the faulty psm determined that the motor was defective. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 445669 | 456676 | 432580 | 2955842200300016 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-04-25 | 2003-04-25 | Yes | No | Health Professional | IS100004 | No | No | 2003-03-26 | Yes | No | Reuse | 2003
40000421 | 2003-02-20 | Malfunction |  | A scalpel cautery instrument was alleged to have arced during a procedure. The instrument was removed from the patient and replaced with another to continue the case to completion. No patient injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument noted a crack in the yaw pulley. The crack indicates a compromise in the insulatioin integrity. Arcing may have occured at the yaw pully crack, cannot be confirmed. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 443869 | 454873 | 430848 | 2955842200300014 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-04-11 | 2003-04-15 | No | Yes | Health Professional | 40000421 | No | Yes | 2003-03-13 | Yes | No | Reuse | 2003
40000417 | 2003-03-19 | Malfunction |  | A scalpel cautery instrument was alleged to have smoked during procedure. The cautery instrument was removed from the patient and replaced with a permanent cautery hook instrument to continue the case to completion. No patient injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument confirmed evidence of surface arcing on the yaw pully and distal clevis. Cracks in the yaw pulley indicates a compromise in the insulation integrity. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURIGAL ACCESSORY | 443876 | 454879 | 430854 | 2955842200300015 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-04-11 | 2003-04-15 | No | Yes | Health Professional | 40000417 | No | No | 2003-03-19 | Yes | No | Reuse | 2003
40012604 | 2003-02-11 | Malfunction |  | It was alleged there was a high frequency isolation problem at the "heel" of the cautery hook. No pt harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument showed evidence of arcing at the base of the hook along the mold parting line. No defintive root cause of the instrument arcing can be determined at this time. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 438127 | 449143 | 425290 | 2955842200300013 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-03-12 | 2003-03-12 | No | Yes | Health Professional | 40012604 | No | No | 2003-02-11 | Yes | No | Reuse | 2003
40000419 | 2003-02-10 | Malfunction |  | A scalpel cautery instrument was alleged to have arcing during a procedure. The instrument was removed from the patient and replaced with another to continue the case to completion. No patient injury or adverse outcome was reported. | - failure analysis of the scalpel cautery instrument confirmed the failure. The top of the yaw pulley was also noted as having cracked. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 437449 | 448467 | 424637 | 2955842200300012 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-03-06 | 2003-03-11 | No | Yes | Health Professional | 40000419 | No | No | 2003-02-10 | Yes | No | Reuse | 2003
400004XX | 2003-01-30 | Injury |  | In 2003, it was reported that a pt complained of abdominal pain and an x-ray revealed an artifact that resembled a round cautery hook in the right lower quadrant. Laparoscopic surgery was performed to remove the object in february, 2003. It was also reported that this pt had undergone abdominal surgery with the intuitive surgical system utilizing a scalpel cautery instrument with a round cautery hook in 2002. Prior to the receipt of this report, intuitive received no notice that there was any issue related to the system or instruments used during the 2002 surgery. | The hosp has chosen to retain the object removed from the pt, and it was not returned to intuitive surgical for eval. No add'l info is anticipated. | SCALPELELECTROCAUTERY | ENDOSCOPIC INSTRUMENT | 434636 | 445660 | 421946 | 2955842200300011 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-02-28 | 2003-02-28 | Yes | No | Health Professional | 400004XX | No | Yes | 2003-01-30 | Device Not Retu | No | Unkown | 2003
IS1000 | 2003-01-21 | Other | Required Intervention | During a thoracoscopic procedure it was reported that an epicardial/tissue injury (small hole) occurred, resulting in conversion from a minimally invasive to an open procedure. The procedure was completed and the pt recovered without incident. The injury was noticed as the surgeon was transferring an ablation probe (a non-isi medical product) from a long tip forceps instrument controlled with the right hand to another long tip forceps controlled with the left hand. There was no reported malfunction of the da vinci surgical system or associated instruments. Based on available info, the specific cause of this injury cannot be determined. No add'l info regarding this event is anticipated. |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 433122 | 444144 | 420451 | 2955842200300006 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-02-13 | 2003-02-14 | Yes | No | Health Professional | IS1000 | No | No | 2003-01-22 | Device Not Retu | No | Reuse | 2003
40000421 | 2003-01-21 | Malfunction |  | A scalpel cautery instrument was alleged to have leaked energy during a procedure and was sticking to tissue. No pt injury or adverse outcome was reported. | Failure analysis of the scalpel cautery instrument confirmed that there were signs of arcing on the cautery pulley and evidence of arcing at a knitline crack. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 433178 | 444201 | 420507 | 2955842200300009 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40000421 | No | No | 2003-01-21 | Yes | No | Reuse | 2003
40012603 | 2002-10-28 | Malfunction |  | It was alleged that the insulation surrounding the instrument hook came off. The sterilization dept removed the instrument from service after noticing the insulation was missing. The 2002 incident date corresponds to the date when burn marks were first alleged to have been noticed on the instrument. Since the scalpel cautery instrument still functioned, the site continued to use the instrument until 2 mos after the event date when the insulation was noticed as missing. No pt harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument confirmed that the distal insulation sleeve was missing from the hook. In addition, the pulley showed signs of having melted and an arc track was present between both the distal and proximal clevis. No definitive root cause of why the sleeve insulation was missing from the instrument can be determined at this time. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 433243 | 444266 | 420572 | 2955842200300004 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40012603 | No | Yes | 2003-01-15 | Yes | No | Reuse | 2003
40000421 | 2003-01-16 | Malfunction |  | It was alleged that the wrist of the scalpel cautery instrument did not have full range of motion. During removal of the instrument to exchange it with a backup, the tip (cautery spatula accessory) of the original instrument fell into the pt. After the procedure was converted to an open surgery, the cautery spatula accessory was retrieved. The conversion of the surgical procedure was planned and was not due to the alleged incident with the scalpel cautery instrument. No harm to the pt or adverse events were reported. | The instrument associated with the alleged incident was not returned. Failure analysis of the discrepant instrument could not be performed. No pt damage was reported. It was reported that there was no adverse outcome or injury. No additional info regarding the alleged incident is expected. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 433296 | 444319 | 420624 | 2955842200300007 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40000421 | No | No | 2003-01-16 | Device Not Retu | No | Unkown | 2003
40012604 | 2003-01-20 | Malfunction |  | It was alleged that the wrist of a permanent hook cautery instrument would no longer articulate. Forceps were used to straighten the instrument wrist in order to remove the permanent hook cautery instrument through the cannula. The surgeon identified two instrument fragments; one fragment was recovered from the pt. The case was completed successfully. No pt harm, adverse outcome or injury was reported. | The permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument confirmed that the conductor cap and one ear of the distal clevis were missing. Lack of these components prevented the instrument wrist from articulating properly. None of the missing components were returned for failure analysis. In addition, the insulation sleeve was found to be missing from the hook. Evidence of arcing was also present on the pulley and the remaining distal clevis. Arcing was likely caused by the lack of sleeve insulation. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 433304 | 444327 | 420632 | 2955842200300008 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40012604 | No | No | 2003-01-20 | Yes | No | Reuse | 2003
40012602C | 2003-01-16 | Malfunction |  | It was alleged that while cauterizing, the wrist of a permanent hook cautery instrument arced during a procedure. The instrument was removed and replaced with another to continue the case to completion. No pt harm, adverse outcome or injury was reported. | H6: the permanent hook cautery instrument was returned and evaluated. Failure analysis of the permanent hook instrument showed evidence of the pulley having melted. The melted pulley was likely the result of arcing. Nonetheless, no definitive root cause of the instrument arcing can be determined at this time. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 433327 | 444350 | 420653 | 2955842200300003 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40012602C | No | No | 2003-01-16 | Yes | No | Reuse | 2003
40009303 | 2003-01-22 | Malfunction |  | It was alleged that a prograsp grasper broke at the "cuff" (1 cm proximal to the wrist) where the wrist and shaft are connected. The instrument could not be removed from the pt without also removing the trocar cannula. The case was completed using another isi instrument. No other difficulties were reported. There was no reported complications due to breakage or the resultant instrument removal. No pt harm or pt injury was reported. | The alleged deficient instrument was not returned to isi for failure analysis. No additional info is anticipated. | PROGRASP GRASPER | ENDOSCOPIC INSTRUMENT | 433336 | 444359 | 420662 | 2955842200300010 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2003-02-14 | 2003-02-14 | No | Yes | Health Professional | 40009303 | No | No | 2003-01-22 | Device Not Retu | No | Reuse | 2003
IS1000 | 2003-01-07 | Other | Required Intervention | During a thoracoscopic procedure, it was reported that an epicardial/tissue injury (samll hole) occurred, resulting in conversion from a minimally invasive to an open procedure. The procedure was completed and the patient recovered without incident. The injury was noticed during the initial surgical procedure and may have been exacerbated as a result of probing with an intuitive surgical instrument (cautery spatula). There were no reported malfunctions of the da vinci surgical system or associated instruments. Based on available information, the specific cause of this injury cannot be determined. No addtional information regarding this event is anticipated. |  | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 431478 | 442496 | 418853 | 2955842200300005 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-02-06 | 2003-02-07 | Yes | No | Health Professional | IS1000 | No | No | 2003-01-22 | Device Not Retu | No | Reuse | 2003
40012604 | 2003-01-06 | Malfunction |  | It was alleged that while cauterizing, the wrists of two permanent hook cautery instruments became "leaky". The instruments were removed and replaced with another to continue the case to completion. No patient harm, adverse outcome or injury was reported. | The permanent hook cautery instruments were returned and evaluated. Failure analysis of one permanent hook instrument confirmed signs of arcing at the instrument wrist and the goose-neck of the hook. No definitive root cause of the instrument arcing can be determined at this time. Arcing is believed to be associated with the customer reported symptom. Failure analysis of the second permanent hook instrument showed no visible signs of arcing, burning or other evidence of instrument failure. No trouble was found during failure analysis. Occurrence of the alleged arcing incident could not be confirmed. No definitive root cause of the alleged instrument arcing can be determined at this time. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 431106 | 442128 | 418487 | 2955842200300001 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-01-31 | 2003-02-03 | No | Yes | Health Professional | 40012604 | No | No | 2003-01-06 | Yes | No | Reuse | 2003
40000419 | 2003-01-06 | Malfunction |  | A co rep attending a surgery at the site reported that a cautery spatula instrument exhibited leaking during a surgical procedure. Upon removing the instrument, the tip was noticed as missing. The spatula and two instrument fragments were recovered from the pt. It was reported that there was no adverse outcome or injury. | Failure analysis of the returned instrument confirmed signs of arcing over the top of the yaw pulley were visible. Arcing was most likely due to the integrity of the insulation and seal at the yaw pulley becoming compromised. In addition, failure analysis confirmed that the distal clevis, conductor cap and spatula were separated from the instrument. All instrument component fragments are believed to be accounted for during examination. No components fragments were known to have been left in situ. The relationship between the arcing and the instrument fragmentation (i. E. Distal clevis and conductor cap failure and resultant spatula retention issue) cannot be conclusively determined. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective aciton. No pt damage was reported. It was reported that there was no adverse outcome or injury. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 431285 | 442307 | 418664 | 2955842200300002 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2003-01-31 | 2003-02-03 | No | Yes | Health Professional | 40000419 | No | No | 2003-01-06 | Yes | No | Reuse | 2003
40003615 | 2002-11-06 | Malfunction |  | It was alleged that the debakey forceps fractured during surgery allowing two fragments to separate from the instrument. The two fragments were removed from the pt. The debakey forceps could not be removed through the cannula. The instrument was retrieved by removing the trocar together with the instrument through the thorax. It was alleged that the surgeon experienced difficulty maintaining insufflation prior to removal of the trocar. The case was converted due to poor insufflation. There were no reported complications due to the breakage, resultant retrieval or poor insufflation. No pt harm, adverse outcommee or injury was reported. | Failure analysis of the returned product determined that the instrument had broken at the proximal clevis and tube. The reported malfunction is consistent with a failure of the proximal clevis to main tube interface that was fixed with correct action to debakey forceps mfg after 9/2002. The instrument involved in this incident was mfg in 1/2001, and did not have the improved design. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. No definitive conclusion could be determined to identify the root cause of the insufflation difficulty. No pt damage was reported. It was reported that there was no adverse outcome or injury. | DEBAKEY FORCEPS | ENDOSCOPIC INSTRUMENT | 422666 | 433727 | 410404 | 2955842200200089 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2002-12-06 | 2002-12-11 | No | Yes | Health Professional | 40003615 | No | No | 2002-11-11 | Yes | No | Reuse | 2002
40000615 | 2002-11-04 | Malfunction |  | A component of the large needle driver broke off the instrument during a procedure. The fragment was retrieved. Since the tip of the instrument could not be straightened fully, the instrument and trocar cannula were removed from the patient simultaneously. There were no reported complications due to the breakage or the resultant retrieval. No patient harm or patient injury was reported. It was reported that there was no adverse outcome or injury. | All remnants of the large needle driver were retrieved from the patient. Although the instrument was returned to intuitive surgical, the fragment was not returned. Failure analysis determined that the proximal clevis fractured at the tube tip. A design improvement for the proximal clevis/tube interface was implemented in june 2002. This is the first reported failure received on the improved design. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. No patient damage was reported. It was reported that there was no adverse outcome or injury. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 420484 | 431504 | 408253 | 2955842200200088 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-11-26 | 2002-12-02 | No | Yes | Health Professional | 40000615 | No | No |  | Yes | No | Reuse | 2002
40000417 | 2002-10-24 | Malfunction |  | The cautery spatula tip was noticed to have fallen off the cautery instrument immediately before cauterizing was to begin. The blade was retrieved and reinserted into the instrument to continue the case. The blade was lost a second time. The blade was located on the floor of the operating room. A new tip was inserted into the instrument to complete the case. No patient harm or patient injury was reported. The patient was released without further incident. | Failure analysis determined that the yaw pulley was cracked at the knitline. No signs of arcing were observed. When a new spatula was inserted into the instrument, the average pullout force was below specification. The failure may have potentially occurred during device handling prior to the procedure (i. E. At sterilzation or spatula insertion). However, no definitive conclusion as to the cuase of the mechanical failure can be determined. All scalpel blades were retrieved from the patient. No patient damage was reported. It was reported that there was no adverse outcome or injury. | CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 418982 | 429993 | 406791 | 2955842200200087 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-11-13 | 2002-11-21 | No | Yes | Health Professional | 40000417 | No | No | 2002-10-24 | Yes | No | Reuse | 2002
40000615 | 2002-10-15 | Malfunction |  | It was alleged that the tip of the large needle driver partially separated from the instrument. During retraction into the cannula, it was reported that the tip completely separated from the instrument. The tip was visualized and retrieved using the right instrument. There were no complications due to the tip breakage or the resultant retrieval. No patient harm or patient injury was reported as related to the product malfunction. | All remnants of the needle driver were retrieved from the patient and returned for failure analysis. Failure analysis confirmed that one grip had been broken off. The clinical specialist in attendance at the case stated that the breakage of the large needle driver tip was precipitated when the physician attempted to throw a stitch inadvertently in the posterior aspect of the pubic bone. No patient damage was reported. It was reported that there was no adverse outcome or injury. | LARGE NEEDLE DRIVER INSTRUMENT | ENDDOSCOPIC INSTRUMENT | 417661 | 428629 | 405508 | 2955842200200085 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-11-13 | 2002-11-14 | No | Yes | Health Professional | 40000615 | No | No | 2002-10-15 | Yes | No | Initial | 2002
40012602 | 2002-10-21 | Malfunction |  | The permanent hook cautery instrument was damaged while attempting to remove the instrument through the cannula. The damaged instrument tip prevented removal through the cannula. As a result, the cannula and instrument were removed from the patient simultaneously. The instrument tip was then manually straightened and the damaged instrument was removed. The cannula was reinserted in the chest of the patient without any further incidents to finish a successful case. No patient harm or patient injury was reported. The patient was released without further incident. | Damage to the permanent hook cautery instrument was not noticed during the surgical procedure. Failure analysis of the returned instrument confirmed that the distal clevis had broken off. One photographic image was taken immediately after the cannula and instrument were removed from the patient. No video footage of the procedure was available. Analysis determined that the distal clevis likely fractured within the patient while the or staff was attempting to remove the instrument, but this conclusion cannot be confirmed. It was theorized that the instrument tip was not straightened before removal was attempted. No patient damage was reported. It was reported that there was no adverse outcome or injury. The post-procedure chest x-ray was unremarkable. The patient was discharged without further incident. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 417665 | 428633 | 405512 | 2955842200200086 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-11-13 | 2002-11-14 | No | Yes | Health Professional | 40012602 | No | No | 2002-10-21 | Yes | No | Reuse | 2002
40000419 | 2002-10-07 | Malfunction |  | It was reported that a scalpel cautery instrument began to leak at the wrist during a prostatectomy. No patient injury or adverse outcome was reported. | The scalpel cautery instrument was returned and evaluated. Failure analysis confirmed that there were signs of arcing over the top of the yaw pulley on the scalpel cautery instrument. The average retention force on the scalpel cautery instrument was measured with a new spatula and found to be slightly below specification. Arcing was most likely due to the integrity of the insulation and seal between these components becoming compromised, most likely due to low retention force of this instrument. No pt injury was reported. It was reported that there was no adverse outcome or injury. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 415090 | 426038 | 403029 | 2955842200200083 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-30 | 2002-10-31 | No | Yes | Health Professional | 40000419 | No | No | 2002-10-07 | Yes | No | Reuse | 2002
40000419 | 2002-10-08 | Malfunction |  | It was alleged that a scalpel cautery instrument began to arc. Upon inspection at the site it was noticed that the cautery spatula was possibly not properly inserted into the cautery instrument. A second scalpel cautery instrument was used to complete the procedure. No patient injury or adverse outcome was reported. | The scalpel cautery instrument was not returned with the original blade cited in the complaint. Consequently, failure analysis was unable to confirm whether the cautery spatula had been incorrectly inserted as suggested by the clinical specialist. No user error can be confirmed or denied. Failure analysis of the scalpel cautery instrument confirmed that there were signs of arcing over the top of the yaw pulley on the scalpel cautery instrument. The average retention force on the scalpel cautery instrument was measured with a new spatula and found to be slightly below specification. Arcing was most likely due to the integrity of the insulation and seal between these components becoming compromised, most likely due to low retention force of this instrument. | SCALPEL CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 415091 | 426039 | 403030 | 2955842200200084 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-30 | 2002-10-31 | No | Yes | Health Professional | 40000419 | No | No | 2002-10-08 | Yes | No | Reuse | 2002
40012602 | 2002-10-01 | Malfunction |  | It was alleged that the housing (i. E. Plastic cover) of the permanent cautery hook instrument partially separated from the instrument during use. No patient harm or patient injury was reported. The procedure was converted to an open procedure since no back up instrument was available. No components completely separated from the instrument. The instrument was retrieved by removing the cannula together with the instrument. It was reported that there was no adverse outcome or injury. | Evaluation pending device return. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL DEVICE | 415104 | 426054 | 403043 | 2955842200200082 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-30 | 2002-10-31 | No | Yes | Health Professional | 40012602 | No | No | 2002-10-01 | No | No | Unkown | 2002
40012602C | 2002-09-30 | Malfunction |  | It was alleged that during a surgical procedure the permanent hook cautery instrument began to arc. No patient harm was reported. It was reported that there was no adverse outcome or injury. | H6; evaluation pending device return. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 415121 | 426071 | 403060 | 2955842200200079 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-28 | 2002-10-30 | No | Yes | Health Professional | 40012602C | No | No | 2002-09-30 | No | No | Reuse | 2002
40012603 | 2002-10-01 | Malfunction |  | It was reported that during use the articulating wrist of the instrument split in half. Half of the instrument wrist and pin fell into the patient. Both components were retrieved with no injury to the patient. The procedure was continued with another permanent cautery hook without incident. It was reported that there was no adverse outcome or injury. | H6: all remnants of the permanent cautery hook were retrieved from the patient. Nonetheless, the instrument wrist fragment was not returned for failure analysis. Review of the video footage and failure analysis performed on the remainder of the instrument concluded that the distal clevis broke, causing a piece of the clevis and pin to break away from the instrument. The failure occurred toward the beginning of the procedure and there was no indication of rough or abusive use of the instrument intra-operatively. The clevis was likely weakened during previous handling, but this conclusion cannot be confirmed. No patient damage was reported. It was reported that there was no adverse outcome or injury. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL DEVICE | 415125 | 426075 | 403064 | 2955842200200081 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-29 | 2002-10-30 | No | Yes | Health Professional | 40012603 | No | No | 2002-10-01 | Yes | No | Reuse | 2002
40012603 | 2002-09-27 | Malfunction |  | It was alleged that while cauterizing, the entire hook became active and the instrument cauterized more than it was intended. The instrument was removed and replaced with another to continue the case to completion. No patient harm was reported. It was reported that there was no adverse outcome or injury. | The permanent hook cautery instrument was returned and evaluated. Failure analysis confirmed that there were signs of arcing around the potting of the wire to the hook. The arcing likely caused the wire to disconnect from hook. Based upon analysis it appears the potting was compromised as to create a conductive path to the wire and hook. No definitive cause of the compromised potting can be determined at this time. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL ACCESSORY | 415130 | 426080 | 403069 | 2955842200200080 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-24 | 2002-10-30 | No | Yes | Health Professional | 40012603 | No | No | 2002-09-30 | Yes | No | Reuse | 2002
40012603 | 2002-09-23 | Malfunction |  | The tip of a permanent hook cautery device broke inside the pt. No pt damage was reported. It was reported that there was no adverse outcome or injury. | To remove the damaged hook through the trocar, the surgeon was forced to cut and remove the hook's distal tip. No definitive conclusion as to the cause of the hook breakage could be drawn since the surgeon had cut the instrument tip. All remnants of the permanent cautery hook were retrieved from the pt and returned for failure analysis. No pt damage was reported. It was reported that there was no adverse outcome or injury. | PERMANENT CAUTERY HOOK | ELECTROSURGICAL DEVICE | 413674 | 424625 | 401650 | 2955842200200076 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2002-10-21 | 2002-10-23 | No | Yes | Health Professional | 40012603 | Device Returned To Manufacturer | No | 2002-10-03 | No | No | Initial | 2002
40003615 | 2002-09-25 | Malfunction |  | It was reported that the tip of a debakey forceps instrument broke inside the pt. No remnants were left in the pt. No pt damage was reported. It was reported that there were no adverse outcome or injury. | To remove the damaged forceps through the trocar, the surgeon was forced to cut and remove the forcep's distal tip. No definitive conclusion as to the cause of the forceps breakage could be drawn since the surgeon had cut the instrument tip. The reported malfunction is consistent with a failure of the proximal clevis to main tube interface that was fixed with corrective action on debakey forceps manufactured after 9/2002. The instrument involved in this incident was manufactured in 6/2001, and did not have the improved design. Intuitive surgical will continue to investigate similar incidents for trends and potential corrective action. All remnants of the debakey forceps were retrieved from the pt and returned for failure analysis. No pt damage was reported. It was reported that there was no adverse outcome or injury. | DEBAKEY FORCEPS | ENDOSCOPIC INSTRUMENT | 413787 | 424738 | 401756 | 2955842200200077 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2002-10-23 | 2002-10-23 | No | Yes | Health Professional | 40003615 | Device Returned To Manufacturer | No | 2002-10-03 | No | No | Reuse | 2002
40000419 | 2002-09-27 | Malfunction |  | It was reported that the scalpel cautery instrument began to smoke during a radical prostatectomy. The instrument was removed and replaced with another to continue the case to completion. No pt injury was reported. It was reported that there was no adverse outcome or injury. | The scalpel cautery instrument was returned and evaluated. Failure analysis confirmed that there were signs of arcing over the top of the distal yaw pulley on the scalpel cautery instrument. The average retention force on the scalpel cautery instrument was measured with a new spatula and found to be slightly below specification. Arcing was most likely due to the integrity of the insulation and seal between these components becoming compromised, most likely due to low retention force of this instrument. | SCALPEL CAUTERY INSTRUMENT | ELCTROSURGICAL ACCESSORY | 413633 | 424578 | 401613 | 2955842200200078 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-22 | 2002-10-22 | Yes | No | Health Professional | 40000419 | No | No | 2002-09-27 | Yes | No | Reuse | 2002
IS1000 | 2002-09-18 | Other | Required Intervention | A superficial epicardial injury occurred during a totally endoscopic coronary artery bypass (tecab) procedure. It was reported that bleeding was stopped by applying pressure and the case was successfully completed without further incident. It was reported that the paitent was discharged without any complications. | Review of a video tape of the procedure and discussion with the surgeon confirmed that the surgeon was controlling forceps in his right hand and was intending to swap control to another instrument (scalpel cautery instrument) in a different robotic arm. The video tape review confirmed that control to the scalpel cautery did not occur and the surgeon moved the camera so that the forceps were no longer in view. The surgeon most likely moved his right hand, intending to move the scalpel cautery but still actively controlling the forceps, and moved the forceps causing the epicardial injury. The surgeon was able to control bleeding by applying pressure on the site, and the procedure was successfully completed without further incident. There was no alleged malfunction of the device and device evaluation was not necessary. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 412134 | 423080 | 400168 | 2955842200200074 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2002-10-16 | 2002-10-17 | Yes | No | Health Professional | IS1000 | No | No | 2002-09-18 | Device Not Retu | No | Reuse | 2002
40000417 | 2002-09-21 | Malfunction |  | During a case one 15 degree scalpel blade fell into the patient. A post-op ct scan confirmed the presence of the blade in the patient. The surgeons and patient decided together to not retrieve the blade. It was reported that the patient was discharged without any complications. | Device not returned for evaluation. Review of a video tape of the procedure confirmed that the scalpel cautery instrument with the 15 degree blade was in the surgeon's right hand, and that the surgeon had re-adjusted the camera so that the scalpel cautery instrument was not visible at the time the blade fell out. Potential causes include improper insertion of the blade into the instrument or inadvertently bumping another instrument with the blade while the instrument was out of view. The scalpel cautery instrument involved in the incident was not returned for evaluation and no conclusions could be drawn with respect to the condition of the instrument. It was reported that the customer continued to use the instrument without any additional issues. | SCALPEL CAUTERY INSTRUMENT DEVICE 1 | ELECTROSURGICAL DEVICE | 412692 | 423639 | 400700 | 2955842200200075 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2002-10-16 | 2002-10-17 | No | Yes | Health Professional | 40000417 | No | No | 2002-09-24 | Device Not Retu | No | Unkown | 2002
40012602C |  | Malfunction |  | It was reported that the tip of a permanent hook cautery device broke. The surgeon did not believe the breakage occurred within the patient. It was reported that there was no adverse outcome or injury. | Failure analysis confirmed that the distal clevis was broken. It is suspected that the failure was caused by user handling but no definitive conclusion as to the cause of the fracture could be determined. | PERMANENT CAUTERY HOOK | ELECTROSURGICAL DEVICE | 412696 | 423643 | 400704 | 2955842200200073 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-10-16 | 2002-10-17 | No | Yes | Health Professional | 40012602C | Device Returned To Manufacturer | Yes | 2002-09-17 | Yes | No | Reuse | 2002
40000419 | 2002-09-11 | Malfunction |  | During a surgical procedure, arcing was observed from the scalpel cautery instrument with the cautery spatula attached. It was reported that there was no adverse event or injury to the pt. | Failure analysis confirmed that there were signs of arcing between the distal yaw pulley on the scalpel cautery instrument and where the insulation on the single use cautery spatula mates with the pulley. The retention force on the scalpel cautery instrument was measured and found to be slightly below specs. Arcing was most likely due to the integrity of the insulation and seal between these components becoming compromised, most likely resulting from the low retention force on the instrument. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 411175 | 422120 | 399246 | 2955842200200072 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-10 | 2002-10-10 | No | Yes | Health Professional | 40000419 | No | No | 2002-09-11 | Yes | No | Reuse | 2002
40000419 | 2002-09-03 | Malfunction |  | During a surgical procedure, it is alleged that the cautery spatula fell out of the scalpel cautery instrument and into the pt. It was reported that the cautery spatula was retrieved and there was no adverse outcome or pt injury. | The instrument was returned and evaluated, and failure analysis revealed no visible defects. The retention force on the instrument was measured and found to be on the low end of specification. The cause of the cautery spatula falling out of the scalpel cautery instrument could not be conclusively determined, but it is likely that the retention force of this instrument was not sufficient to hold the spatula for the forces experienced during this procedure. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 409771 | 420724 | 397906 | 2955842200200071 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-10-03 | 2002-10-03 | No | Yes | Health Professional | 40000419 | Yes | No | 2002-09-03 | Yes | No | Reuse | 2002
40000419 | 2002-08-30 | Malfunction |  | During a surgical procedure, it is alleged that the cautery spatula fell out of the scalpel cautery instrument and into the patient. It was reported that the cautery spatula was retrieved and there was no adverse outcome or patient injury. | The device was returned and evaluated, and failure analysis revealed that the distal yaw pulley on the scalpel cautery instrument was cracked and that the instrument did not meet specification for retention of the cautery spatula. The cause of the cracked yaw pulley could not be conclusively determined, but it is likely that excessive loading during use of the instrument caused the receptacle to loosen and the yaw pulley to crack, leading to reduced ability to hold the spatula. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 409120 | 420064 | 397270 | 2955842200200070 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2002-09-30 | 2002-10-01 | No | Yes | Health Professional | 40000419 | Yes | Yes | 2002-08-30 | Yes | No | Reuse | 2002
40000419 | 2002-08-27 | Malfunction |  | During a prostatectomy case, arcing was observed from the scalpel cautery instrument with a cautery spatula attached. | Failure analysis confirmed that there were signs of arcing between the distal yaw pulley on the scalpel cautery instrument and where the insulation on the single use cautery spatula mates with the pulley. The retention force of the instrument was measured and found to be slightly below specification. Arcing was most likely due to the integrity of the insulation and seal between these components becoming compromised. This may have been caused by the low retention force of this instrument, excessive force and the wet environment of this particular procedure, or the user not inserting the spatula properly. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 408841 | 419797 | 397013 | 2955842200200069 | 1 | NAY | Manufacturer | Company Representative | UNKNOWN | Initial | 2002-09-25 | 2002-09-27 | No | Yes | Health Professional | 40000419 | No | No Answer Provided | 2002-08-27 | Yes | No | Reuse | 2002
40000417 | 2002-08-23 | Malfunction |  | During a surgical procedure, the cautery spatula fell out of the scalpel cautery instrument and into the pt. It was reported that this occurred twice during the same procedure with the same instrument. In the first occurrence, it was reported that the cautery spatula fell out during dissection. In the second occurrence, it was not reported when the cautery spatula fell out during dissection. In the second occurrence, it was not reported when cautery spatula fell out but it was found while retrieving a needle. | The instrument was returned and evaluated, and failure analysis revealed no defects in the instrument, and the retention force was measured and found to be within specification. No visible signs of damage were evident on the returned instrument. The cautery spatula that fell out of the instrument is a single use disposable accessory to the instrument and was not returned for evaluation. No conclusion can be made as to the specific cause of this incident. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 408807 | 419763 | 396986 | 2955842200200068 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2002-09-24 | 2002-09-24 | No | Yes | Health Professional | 40000417 | No | Yes | 2002-08-23 | Yes | No | Reuse | 2002
40012602 | 2002-08-22 | Malfunction |  | During a case the permanent hook cautery instrument began "leaking" at the wrist during a prostatectomy case. It was reported that there was no injury or harm to the pt. | Device was returned and evaluated. Visual inspection confirmed that there was a very small pinhole with some cracking in the insulation at the base of the hook. There was also evidence of arcing in this area. | PERMANENT HOOK CAUTERY INSTRUMENT | ELECTROSURGICAL DEVICE | 408382 | 419335 | 396589 | 2955842200200067 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-09-20 | 2002-09-23 | No | Yes | Health Professional | 40012602 | No | Yes | 2002-08-22 | Yes | No | Reuse | 2002
40000417 | 2002-08-15 | Malfunction |  | It was reported that during a prostatectomy case, arcing was observed from the scalpel cautery instrument with a cautery spatula attached. | Failure analysis confirmed that there were signs of arcing between the distal yaw pulley on the scalpel cautery instrument and the insulation on the single use spatula. Arcing was most likely due to the insulation integrity of the seal between these components becoming compromised. This could have been caused by excessive forces and the wet environment of this particular surgical procedure (radical prostatectomy), or from the user not inserting the spatula properly. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 407021 | 417986 | 395300 | 2955842200200066 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2002-09-16 | 2002-09-17 | No | Yes | Health Professional | 40000417 | Device Returned To Manufacturer | No Answer Provided | 2002-08-15 | Yes | No | Reuse | 2002
40000419 | 2002-08-13 | Malfunction |  | During an ima takedown procedure, it is alleged that the cautery spatula fell out of the scalpel cautery instrument and into the patient. It was reported that spatula was retrieved and there was no adverse effects to the patient. | The instrument was returned to the manufacturer and failure analysis revealed no defects in the instrument. The spatula that fell out of the instrument is a single use disposable accessory, and was not returned or available for evaluation. It was reported that the spatula was inserted, removed, and then re-inserted into the instrument before the incident. Multiple insertions and removals of a single use spatula into the instrument could potentially deform the spatula so the inserted portion is slightly out of tolerance, resulting in lower retention force when inserted into the instrument after the first time. | SCALPEL CAUTERY | ELECTROSURGICAL DEVICE | 406259 | 417211 | 394562 | 2955842200200064 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2002-09-13 | 2002-09-13 | No | Yes | Health Professional | 40000419 | Device Returned To Manufacturer | No Answer Provided | 2002-08-13 | Yes | No | Reuse | 2002
40000419 | 2002-08-02 | Malfunction |  | It is alleged that during a case the scalpel electrocautery instrument began to arc at the wrist while being used with the cautery spatula. The cautery spatula was switched to the cautery hook and the case was completed with no pt harm reported. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 405160 | 416117 | 393513 | 2955842200200063 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2002-08-30 | 2002-09-03 | No | Yes | Health Professional | 40000419 | Yes | No | 2002-08-02 | Yes | No | Unkown | 2002
40000419 | 2002-07-16 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument began leaking at the insertion point. It was reported that the instrument was removed and the case was completed with a second scalpel/electrocautery instrument. No pt harm was reported. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 397591 | 408599 | 386263 | 2955842200200062 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-07-25 | 2002-07-29 | No | Yes | Health Professional | 40000419 | Yes | No Answer Provided | 2002-07-16 | Yes | No | Unkown | 2002
40000419 | 2002-07-19 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that the procedure was completed with no injury to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 397595 | 408603 | 386267 | 2955842200200061 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-07-25 | 2002-07-29 | No | Yes | Health Professional | 40000419 | Yes | No | 2002-07-19 | Yes | No | Unkown | 2002
40000417 | 2002-07-23 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument began arcing at the wrist. It was reported that there was no patient harm. | Device has not yet been evaluated so no conclusion can be drawn at this time. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 397693 | 408701 | 386362 | 2955842200200060 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-07-25 | 2002-07-29 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-07-23 | No | No | Unkown | 2002
40000418 | 2002-06-26 | Malfunction |  | It is alleged that during a case the scalpel cautery instrument began to smoke at the wrist assembly. It was reported that the case was completed with no harm to the patient. | Device evaluation pending product return. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 394303 | 405301 | 383077 | 2955842200200059 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-07-11 | 2002-07-12 | No | Yes | Health Professional | 40000418 | Yes | No | 2002-06-26 | No | No | Unkown | 2002
40000417 | 2002-06-12 | Malfunction |  | It is alleged that the distal clevis was discovered to be broken after the instrument was inserted into the pt. It was reported that the surgeon searched for the part within the pt, but was not able to locate it, and it was not confirmed that the piece, in fact, broke off after insertion into the pt. No pt harm was reported. | Device was visually examined at the site by the isi sales rep. It could not be confirmed if the clevis broke before or after insertion into the cannula. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 394406 | 405407 | 383175 | 2955842200200058 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-07-11 | 2002-07-12 | No | Yes | Health Professional | 40000417 | Yes | No | 2002-06-12 | Device Not Retu | No | Unkown | 2002
40000418 | 2002-07-02 | Malfunction |  | It is alleged that the hook fell out of the instrument into the pt. It was reported that the instrument wrist was not properly straightened out before trying to remove it from the cannula requiring excessive force and causing the hook to fall into the pt. It was reported that the hook was retrieved from the pt with no injury. |  | SCALPELELECTROCAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 394418 | 405419 | 383187 | 2955842200200057 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-07-11 | 2002-07-12 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-07-02 | No | No | Unkown | 2002
40000418 | 2002-06-04 | Malfunction |  | It is alleged that the instrument arced during a case. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | INTUITIVE SURGICAL SCALPELELEC | 392743 | 403750 | 381588 | 2955842200200053 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-06-26 | 2002-06-27 | No | Yes | Health Professional | 40000418 | Yes | No Answer Provided |  | No Answer Provi | No | Unkown | 2002
40000419 | 2002-06-20 | Malfunction |  | It is alleged that the instrument began arcing at the wrist during a case. It was reported that the instrument was replaced and the case was completed with no other arcing incidences. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | INTUITIVE SURGICAL SCALPELELEC | 392744 | 403751 | 381589 | 2955842200200054 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-06-26 | 2002-06-27 | No | Yes | Health Professional | 40000419 | Yes | No Answer Provided | 2002-06-20 | No | No | Unkown | 2002
40009303 | 2002-06-19 | Malfunction |  | It is alleged that during a case a screw on the prograsp tip popped out into the pt. It was reported that the screw was retrieved with no pt injury. | Device evaluation pending instrument return. | PROGRASP | ENDOSCOPIC INSTRUMENT | 392745 | 403752 | 381590 | 2955842200200055 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-06-26 | 2002-06-27 | No | Yes | Health Professional | 40009303 | Yes | No Answer Provided |  | No | No | Unkown | 2002
40000419 | 2002-06-19 | Malfunction |  | Event Description: --/-- It is alleged that during a case the scalpel/electrocautery began smoking at the wrist. It was reported that the case was continued with no injury to the pt. | Manufacturer Narrative: --/-- Failure analysis confirmed that there were signs of arcing between the distal yaw pulley on the scalpel cautery instrument and the insulation on the single use spatula. Arcing was most likely due to the insulation integrity of the seal between these components becoming compromised. This could have been caused by excessive forces and the wet environment of the particular surgical procedure (prostatectomy), or from the user not inserting the spatula properly. --/-- Additional Manufacturer Narrative: --/-- Root cause evaluation pending. | SCALPELELECTROCAUTERY | Unkown | 392746 | 403753 | 381591 | 2955842200200056 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | InitialFollowup | 2002-06-26 | 2002-06-27 | No | Yes | Health Professional | 40000419 | Yes | No Answer Provided | 2002-06-19 | No | No |  | 2002
40001102 | 2002-06-10 | Malfunction |  | It is alleged that during a case the 15 degree scalpel/electrocautery blade was lost inside the pt and never found. It was reported that there was no other pt harm and the case was aborted for reasons not associated with the loss of the blade. | Device not returned to mfr. Incident is currently under investigation. | 15 DEGREE SCALPEL BLADE | ELECTROSURGICAL ACCESSORY | 391853 | 402843 | 380718 | 2955842200200052 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2002-06-20 | 2002-06-25 | No | Yes | Health Professional | 40001102 | No | Yes | 2002-06-10 | Device Not Retu | Yes | Initial | 2002
40005006 | 2002-05-15 | Malfunction |  | It is alleged that during a case the round cautery hook became lost inside the patient and the patient had to be opened up for the hook to be retrieved. It was reported that the hook was retrieved with no harm to the patient. | Device not returned to manufacturer. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 392522 | 403525 | 381375 | 2955842200200049 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2002-06-19 | 2002-06-25 | No | Yes | Health Professional | 40005006 | No | Yes | 2002-05-29 | Device Not Retu | Yes | Initial | 2002
40000418 | 2002-05-31 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was leaking and the round cautery hook fell off. The hook was retrieved and no adverse events to the patient were reported. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 392529 | 403532 | 381381 | 2955842200200048 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-06-19 | 2002-06-25 | No | Yes | Health Professional | 40000418 | Yes | No | 2002-05-31 | No | Yes | Initial | 2002
40000418 | 2002-06-04 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument began to smoke. The instrument was continued to be used and the case was completed with no patient harm reported. | Device not returned to manufacturer. Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 392539 | 403542 | 381390 | 2955842200200051 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-06-19 | 2002-06-25 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-06-04 | Device Not Retu | Yes | Initial | 2002
40000418 | 2002-06-04 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument began to arc and the instrument was continued to be used until the wrist was burnt. The case was completed with another scalpel/electrocuatery instrument and no harm to the patient was reported. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 392544 | 403547 | 381393 | 2955842200200050 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-06-19 | 2002-06-25 | No | Yes | Health Professional | 40000418 | Yes | No Answer Provided | 2002-06-04 | No | Yes | Initial | 2002
40000417 | 2002-05-08 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 388027 | 398983 | 377016 | 2955842200200036 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-05-13 | 2002-06-05 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-05-08 | No | No | Unkown | 2002
40000418 |  | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386126 | 397075 | 375187 | 2955842200200047 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40000418 | 2002-03-29 | Malfunction | Congenital Anomaly | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386129 | 397078 | 375190 | 2955842200200046 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Initial | 2002
40000418 | 2002-04-02 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386134 | 397083 | 375195 | 2955842200200045 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40000418 | 2002-04-01 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386136 | 397085 | 375197 | 2955842200200044 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40000418 | 2002-04-02 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the patient. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386138 | 397087 | 375199 | 2955842200200043 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40000418 | 2002-04-01 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386140 | 397089 | 375197 | 2955842200200042 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40000418 | 2002-04-04 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced. It was reported that the case was completed without harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 386164 | 397113 | 375224 | 2955842200200041 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40000418 | No | No Answer Provided | 2002-05-22 | No | No | Reuse | 2002
40003615 | 2002-05-14 | Malfunction |  | It is alleged that during a case the black plastic piece of the tip of the debakey forceps broke off. The piece was retrieved and no adverse events to the patient were reported. | Evaluation pending device return. | DEBAKEY FORCEPS | INSTRUMENT FORCEPS | 386212 | 397161 | 375280 | 2955842200200040 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2002-05-22 | 2002-05-24 | No | Yes | Health Professional | 40003615 | Yes | No Answer Provided | 2002-05-14 | No | No | Invalid Data | 2002
40000417 | 2002-05-17 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that there was no harm to the pt. | Root cause eval pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 385811 | 396765 | 374874 | 2955842200200039 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-20 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-05-17 | No | No | Unkown | 2002
40000417 | 2002-05-11 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that there was no harm to the pt. | Root cause eval pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 385813 | 396767 | 374876 | 2955842200200038 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2002-05-13 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No |  | No | No | Unkown | 2002
40000417 | 2002-05-07 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that the procedure was completed without harm to the pt. | Root cause eval pendiing completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 385815 | 396769 | 374878 | 2955842200200035 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-07 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-05-07 | No | No | Unkown | 2002
40000417 | 2002-05-07 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that the procedure was completed without harm to the pt. | Root cause eval pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 385818 | 396772 | 374880 | 2955842200200034 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-13 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-05-07 | No | No | Unkown | 2002
40000417 | 2002-05-07 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist, causing the wrist to completely melt. It was reported that there was no injury to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 385822 | 396776 | 374884 | 2955842200200033 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-07 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-05-07 | No | No | Unkown | 2002
40000417 | 2002-05-10 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that there was no harm to the pt. | Root cause evaluation pending completion. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 402725 | 413692 | 391174 | 2955842200200065 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2002-05-13 | 2002-05-23 | No | Yes | Health Professional | 40000417 | Yes | No | 2002-05-10 | No | No | Unkown | 2002
38001307 | 2002-04-16 | Malfunction |  | During routine manufacturing, a small fire occurred on one of the printed circuit boards inside the surgeon console electronic chassis. The system was in the initial steps of system level testing, and was in operation for approximately an hour and a half when the operator noticed a board on fire in the electronic chassis. The system was in an operating mode that is not representative of a finished device, and the chassis cover and enclosure was not yet installed. | Root cause investigation determined that the fire was due to a failure of a defective tantalum capacitor on the printed circuit assembly. | SURGEON SIDE SYSTEM | SURGEON SIDE SYSTEM | 384762 | 395755 | 373909 | 2955842200200037 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2002-05-09 | 2002-05-15 | No | Yes | Other | 38001307 | Yes | No | 2002-04-16 | Yes | No | Invalid Data | 2002
37026903 | 2002-04-24 | Malfunction |  | It is alleged that during a case the camera cannula mount snapped open and a screw fell into the pt. It was reported that the screw was retrieved with no injury to the pt. | Corrective action is in process. | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 384130 | 395122 | 373295 | 2955842200200032 | 1 | GCJ | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-08 | 2002-05-14 | No | Yes | Health Professional | 37026903 | Yes | No Answer Provided | 2002-04-24 | Yes | No | Unkown | 2002
40000417 | 2002-04-30 | Malfunction |  | It is alleged that during a case the tip of the scalpel/electrocautery instrument broke. It was reported that there was no harm to the pt. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 382980 | 393969 | 372173 | 2955842200200029 | 1 | NAY | Manufacturer | Health Professional | Invalid Data | Initial | 2002-05-01 | 2002-05-07 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-04-30 | No | No | Unkown | 2002
40004202 | 2002-04-19 | Malfunction |  | It is alleged that during a case, the tip of the scalpel/electrocautery spatula fell into the pt. The tip was retrieved and it was reported that there were no injuries to the pt. | Cause of event has been determined to be user error. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 382983 | 393972 | 372175 | 2955842200200028 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2002-05-01 | 2002-05-07 | No | Yes | Health Professional | 40004202 | No | No | 2002-04-23 | No | No | Unkown | 2002
40004202 | 2002-04-17 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery spatula fell into the pt. The spatula was retrieved and no adverse events to the pt were reported. | Cause of event has been determined to be user error. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 382991 | 393980 | 372183 | 2955842200200027 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-29 | 2002-05-07 | No | Yes | Health Professional | 40004202 | No | No Answer Provided | 2002-04-17 | No | No | Unkown | 2002
40000417 | 2002-05-02 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. It was reported that there was no harm to the pt. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 382994 | 393983 | 372186 | 2955842200200031 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-02 | 2002-05-07 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-05-02 | No | No | Unkown | 2002
40000417 | 2002-04-19 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced at the wrist. The instrument was replaced and the case was completed. It was reported that there was no harm to the pt. | Corrective action is in process. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 382996 | 393985 | 372188 | 2955842200200026 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-29 | 2002-05-07 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-04-25 | Yes | No | Unkown | 2002
40000417 | 2002-04-30 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was arcing at the wrist. The instrument was replaced and the case continued. No injuries to the pt were reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 382999 | 393988 | 372191 | 2955842200200030 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-05-01 | 2002-05-07 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-04-30 | No | No | Unkown | 2002
40000417 | 2002-04-23 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument, which was used with the round cautery hook, was smoking/arcing at the wrist. No adverse events to the pt were reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 380506 | 391455 | 369751 | 2955842200200025 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-04-24 | 2002-04-25 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-04-23 | No | No | Unkown | 2002
40000416 | 2002-02-27 | Malfunction |  | It is alleged that during a case the scapel/electrocautery instrument broke at wrist angle during a tool change. No adverse events to the pt were reported. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 380148 | 391099 | 369402 | 2955842200200023 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2002-04-22 | 2002-04-23 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided | 2002-02-27 | No | No | Unkown | 2002
40000417 | 2002-04-22 | Malfunction |  | Event Description: --/-- It is alleged that during a case the pt's liver was slightly burned due to arcing at the wrist of the scalpel/electrocautery instrument. It was reported that there was no serious injury to the pt. | Manufacturer Narrative: --/-- Product was redesigned to reduce the likelihood of future recurrences. --/-- Additional Manufacturer Narrative: --/-- The letter from fda stated that intuitive's medwatch report number 2955842-2002-00024 indicated that a remedial action was initiated; however, according to co's records, the report in question does not indicate that remedial action was taken, nor did the co initiate a remedial action for the product and incident reported. | SCALPELELECTROCAUTERY | Unkown | 380156 | 391107 | 369410 | 2955842200200024 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | InitialFollowup | 2002-04-22 | 2002-04-23 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-04-22 | No | No |  | 2002
40008303 | 2002-03-25 | Malfunction |  | It is alleged that the ultrasonic shears instrument broke off. No adverse events were reported. | Device evaluation pending product return. | ULTRASONIC SHEARS | ENDOSCOPIC INSTRUMENT | 380201 | 391152 | 369453 | 2955842200200022 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-22 | 2002-04-23 | No | Yes | Health Professional | 40008303 | Yes | No Answer Provided | 2002-03-25 | No | No | Unkown | 2002
40000417 | 2002-03-19 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument was smoking at the wrist. It was reported that the case was completed with no injury to the patient. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 379001 | 389954 | 368279 | 2955842200200019 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-16 | 2002-04-17 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-03-19 | No | No | Unkown | 2002
40000417 | 2002-03-26 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced at the wrist. No adverse events to the patient were reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 379021 | 389974 | 368298 | 2955842200200020 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2002-04-16 | 2002-04-17 | No | Yes | Health Professional | 40000417 | Yes | No | 2002-03-26 | No | No | Unkown | 2002
40000417 | 2002-04-05 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument arced at the wrist. It was reported that there was no injury to the patient. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 379035 | 389988 | 368312 | 2955842200200021 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-16 | 2002-04-17 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-04-05 | No | No | Unkown | 2002
40009301C | 2002-03-25 | Malfunction |  | It is alleged that during a case the wrist of the prograsp instrument was broken at an angle inside the pt and was impossible to get out of the cannula. In this particular case they detached the cannula from the instrument cannula mount, and took the cannula out with the instrument still inside it. It was reported that there were no more sterile case due to lack of cannulas, they broke the wrist off the end of the instrument in order to get the wrist off the end of the instrument in order to get the instrument out of the cannula. No adverse events to the pt were reported. | Evaluation pending device return. | ATRAUMATIC FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 377768 | 388719 | 367088 | 2955842200200018 | 1 | GEI | Manufacturer | Foreign | UNKNOWN | Initial | 2002-04-11 | 2002-04-15 | No | Yes | Health Professional | 40009301C | Yes | No Answer Provided | 2002-04-02 | No | No | Unkown | 2002
40000417 | 2002-03-22 | Malfunction | Required Intervention | It is alleged that during a case, the left internal mammary vein was punctured when switching from the clip applier to the scalpel/electrocautery instrument. It was reported that the case was converted to a sternotomy and there was no injury to the patient. It was further reported that the puncture was due to surgeon error while switching the instruments. | Device evaluation is not necessary, the cause of the incident has been found to be surgeon error. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 377896 | 388848 | 367212 | 2955842200200017 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-11 | 2002-04-15 | Yes | No | Health Professional | 40000417 | No | No Answer Provided | 2002-03-25 | No | No | Unkown | 2002
40000417 | 2002-03-21 | Malfunction |  | It is alleged that during a case there was arcing at the wrist with two scalpel/electrocautery instruments. It was reported that the case was completed with a third scalpel/electrocautery and there was no injury to the pt. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 377439 | 388399 | 366769 | 2955842200200016 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-04-11 | 2002-04-11 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-03-21 | No | No | Unkown | 2002
40000417 | 2002-03-21 | Malfunction |  | It is alleged that during a case there was arcing at the wrist with two scalpel/electrocautery instruments. It was reported that the case was completed with a third scalpel/electrocautery and there was no injury to the pt. | Product was redesigned to reduce the likelihood of future recurrences. This is a correction to the original report. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 377445 | 388405 | 366775 | 2955842200200014 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-03-27 | 2002-04-11 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2002-03-21 | No | No | Unkown | 2002
40004812 | 2002-02-04 | Malfunction |  | It is alleged that during a case the long tip forceps would not straighten out causing the surgeon to make the port incision larger in order to remove the instrument. No adverse events to the pt were reported. | Device is pending evaluation. | LONG TIP FORCEPS | ENDOSCOPIC INSTRUMENT FORCEPS | 376260 | 387242 | 365674 | 2955842200200013 | 1 | GCJ | Manufacturer | Foreign | UNKNOWN | Initial | 2002-03-27 | 2002-04-01 | No | Yes | Health Professional | 40004812 | Yes | No Answer Provided | 2002-02-06 | No | No | Unkown | 2002
40004915 | 2002-03-21 | Malfunction |  | It is alleged that during a case, a black plastic piece of the cadiere forceps fell off into the pt. Once the piece fell off, the instrument got stuck at a 45 degree angle and could not be removed from the trocar. In order to get the instrument out, the cannula had to be removed with the instrument still attached to the sterile adapter. It was reported that the case was completed with no harm to the pt. | Device is pending eval. | CADIERE FORCEPS | ENDOSCOPIC INSTRUMENT FORCEPS | 376266 | 387248 | 365680 | 2955842200200015 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-03-27 | 2002-04-01 | No | Yes | Health Professional | 40004915 | Yes | No Answer Provided | 2002-03-21 | No | No | Unkown | 2002
40005004 | 2002-03-01 | Malfunction |  | It is alleged that during a case, the cautery hook came off of the scalpel/electrocautery in the patient's abdomen when the instrument was removed from the arm. It was reported that the tip of the cautery hook hit against the cannula and no adverse events to the patient were reported. | Product was redesigned to reduce the likelihood of future recurrences. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 373020 | 383954 | 362501 | 2955842200200010 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2002-03-14 | 2002-03-15 | No | Yes | Health Professional | 40005004 | No | Yes | 2002-03-01 | Device Not Retu | Yes | Initial | 2002
40000615 | 2002-03-13 | Malfunction |  | It is alleged that during a case the pulley came off of the large needle driver and fell into the patient. It was reported that the pulley was retrieved and there was no injury to the pt. | Pending device evaluation. | LARGE NEEDLE DRIVER | ENDOSCOPIC NEEDLE DRIVER | 373375 | 384323 | 362857 | 2955842200200011 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-03-14 | 2002-03-15 | No | Yes | Health Professional | 40000615 | No Answer Provided | Yes | 2002-03-13 | No | No | Unkown | 2002
40005005 | 2002-03-05 | Malfunction |  | It is alleged that during a case the cautery hook was unwillingly inserted slantwise, and as a result, the cautery hook fell into the patient's abdomen. It was reported that there was no injury to the patient. | Visual examination via photographs send by hospital. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 373424 | 384372 | 362899 | 2955842200200012 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2002-03-14 | 2002-03-15 | No | Yes | Health Professional | 40005005 | No | Yes | 2002-03-14 | Device Not Retu | Yes | Initial | 2002
40000417 | 2002-02-07 | Malfunction |  | It is alleged that during a case, when the scalpel/electrocautery was at an angle, it would fry at the housing. When it was straight, there were no problems. It was reported that a different scalpel/electrocautery and hook were used to complete the case and there was reportedly no injury to the patient. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 366540 | 377477 | 356231 | 2955842200200009 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-02-07 | 2002-02-11 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-02-07 | No | No | Unkown | 2002
40000417 | 2002-02-06 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery, which was used with the round cautery hook, shorted out and started to arc into the instrument head. It was reported that a different scalpel/electrocautery and hook were used to complete the case. No injury to the pt was reported. | Device is pending eval. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 366679 | 377613 | 356367 | 2955842200200008 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-02-07 | 2002-02-11 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-02-06 | No | No | Unkown | 2002
40000414 | 2002-01-11 | Malfunction |  | It is alleged that the tip of the scapel/electrocautery is completely gone except for the black plastic part. | Device was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 364524 | 375417 | 354251 | 2955842200200006 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2002-01-21 | 2002-01-31 | No | Yes | Health Professional | 40000414 | Yes | No Answer Provided | 2002-01-15 | No | No | Unkown | 2002
40000417 | 2002-01-18 | Malfunction |  | It is alleged that during a case a piece of the tip of the scalpel/electrocautery broke off inside the patient. It was reported that the piece was retrieved with no injury to the patient. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 364624 | 375519 | 354345 | 2955842200200007 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial |  | 2002-01-31 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-01-18 | No | No | Unkown | 2002
40000416 | 2001-12-20 | Malfunction |  | It is alleged that during a case there was arcing at the wrist of the scalpel/electrocautery. No adverse events to the patient were reported. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 362129 | 373025 | 351917 | 2955842200200003 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-01-15 | 2002-01-17 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided | 2001-12-20 | No | No | Unkown | 2002
40000417 | 2002-01-08 | Malfunction |  | It is alleged that during a case there was arcing at the distal clevis where the cautery hook inserts into the mouth of the scalpel/electrocautery instrument. No adverse events to the patient were reported. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 362131 | 373027 | 351919 | 2955842200200005 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-01-15 | 2002-01-17 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-01-08 | No | No | Unkown | 2002
40000417 | 2002-01-03 | Malfunction |  | It is alleged that during a case there is smoke coming out of the end of the scalpel/electrocautery where the hook is inserted and burnt through the seal. The case was completed with another scalpel/electrocautery and hook and it was reported that there was no injury to the patient. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 362451 | 373344 | 352238 | 2955842200200004 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2002-01-15 | 2002-01-17 | No | Yes | Health Professional | 40000417 | Yes | Yes | 2002-01-03 | No | No | Unkown | 2002
40000417 | 2001-12-18 | Malfunction |  | It is alleged that during a case the movement of the instrument arm was reportedly "jerky" and there was difficulty in removing the instrument. It was reported that the patient's liver was punctured due to the difficulty in moving the instrument. | A complete investigation report from the console surgeon found that at no time during the case was the liver damaged. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 362074 | 372969 | 351863 | 2955842200200002 | 1 | NAY | Manufacturer | Company Representative | UNKNOWN | Initial | 2002-01-14 | 2002-01-15 | No | Yes | Health Professional | 40000417 | Yes | No Answer Provided | 2001-12-18 | Yes | No | Unkown | 2002
40008302 | 2001-12-28 | Malfunction |  | It is alleged that during a case the tip of the rod of the ultrasonic shears broke off into the patient. It was reported that the tip was retrieved from the patient with no adverse events. |  | ULTRASONIC SHEARS | ENDOSCOPIC INSTRUMENT | 359982 | 370851 | 349840 | 2955842200200001 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2002-01-04 | 2002-01-07 | No | Yes | Health Professional | 40008302 | Yes | No Answer Provided | 2001-12-28 | Yes | No | Unkown | 2002
40000315 | 2001-12-03 | Malfunction |  | It is alleged that during a case the clip applier was not holding the clips. No adverse events to the patient were reported. |  | CLIP APPLIER | CLIP APPLIER | 359073 | 369951 | 348944 | 2955842200100068 | 1 | GDO | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-12-27 | 2001-12-27 | No | Yes | Health Professional | 40000315 | Yes | No Answer Provided | 2001-12-03 | No | No | Initial | 2001
38001306 | 2001-11-28 | Malfunction |  | It is alleged that when a customer was transporting the surgeon's console from the lab to the operating room, the customer rec'd a strong static shock when the customer touched the doorknob. It was reported that the shock left a red mark on their skin inside the customer's wrist. |  | SURGEONS SIDE SYSTEM | SURGEONS SIDE SYSTEM | 358548 | 369421 | 348421 | 2955842200100066 | 1 | NAY | Manufacturer | Company Representative | UNKNOWN | Initial | 2001-12-26 | 2001-12-26 | No | Yes | Other | 38001306 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40000416 | 2001-12-19 | Malfunction |  | It is alleged that during a case the distal idler pulley fell off the scalpel cautery and fell into the pt. The pulley was retrieved and the case was completed with no adverse events reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPEL ELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 358550 | 369423 | 348423 | 2955842200100067 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-12-26 | 2001-12-26 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40004812 | 2001-08-31 | Injury | Required Intervention | It is alleged that during a case a hole in the patient's stomach was discovered. It was reported that the surgeon had to convert to an open case in order to fix the hole. | Method: device not returned for evaluation. It was determined that user error/technique was the cause of this incident. | LONG TIP FORCEPS | LONG TIP FORCEPS | 356942 | 367823 | 346872 | 2955842200100065 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2001-12-18 | 2001-12-18 | Yes | No | Health Professional | 40004812 | No | Yes | 2001-08-31 | Device Not Retu | No | Unkown | 2001
40003102 | 2001-10-11 | Malfunction |  | While reviewing a case video tape, it was noticed by an employee that a paddle blade separated from a scalpel/electrocautery instrument. | Device not available for eval. No conclusion can be drawn. | PADDLE SCALPEL BLADE | ELECTROSURGICAL ACCESSORY | 357725 | 368610 | 347622 | 2955842200100064 | 1 | GEI | Manufacturer | ForeignCompany Representative | Other | Initial | 2001-12-18 | 2001-12-18 | No | Yes | Health Professional | 40003102 | No | No | 2001-12-04 | Device Not Retu | Yes | Unkown | 2001
4000416 | 2001-11-06 | Malfunction |  | It is alleged that the hook would not engage in the scalpel/electrocautery instrument and when a new hook was inserted, it broke off. No injury to the patient was reported and the case was delayed by 45 minutes. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 352875 | 363767 | 342974 | 2955842200100063 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-21 | 2001-11-26 | No | Yes | Health Professional | 4000416 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40000416 | 2001-10-30 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery instrument is arcing at the distal wrist. It was reported that there was no injury to the patient and the case was completed using a different scalpel/electrocautery instrument. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 352880 | 363772 | 342979 | 2955842200100062 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-13 | 2001-11-26 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40000416 | 2001-11-05 | Malfunction |  | It is alleged that during a case there was arcing at the proximal end of the scalpel cautery wrist. No adverse events to the patient were reported. | Device not being returned for evaluation. No conclusion can be drawn at this time. Pending evaluation of other device experiencing similar incidence. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 352884 | 363776 | 342983 | 2955842200100061 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-13 | 2001-11-26 | No | Yes | Health Professional | 40000416 | No | No Answer Provided |  | Device Not Retu | No | Unkown | 2001
40000416 | 2001-09-10 | Malfunction |  | It is alleged that during a case, the hook broke off at the tip of the scalpel electrocautery and the remaining spring piece could not be removed from the tip. No adverse events to the patient were reported. | Pending device evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 352888 | 363780 | 342986 | 2955842200100059 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-13 | 2001-11-26 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40000416 | 2001-11-06 | Malfunction |  | It is alleged that during a case the scalpel cautery began to arc at the pivot points in the wrist. Surgery was completed and no harm to the patient was reported. | Device is pending evaluation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 352893 | 363785 | 342991 | 2955842200100058 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-13 | 2001-11-26 | No | Yes | Health Professional | 40000416 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40003614 | 2001-11-06 | Malfunction |  | It is alleged that during a case the yaw broke off into the patient when the the distal clevis of the stabilizer was gripped by the debakey forceps. No adverse events to the patient were reported. | Pending device evaluation. | DEBAKEY FORCEPS | INSTRUMENT FORCEPS | 352896 | 363788 | 342993 | 2955842200100060 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-11-13 | 2001-11-26 | No | Yes | Health Professional | 40003614 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40004202 | 2001-10-19 | Malfunction |  | It is alleged that during a case the scalpel/electrocautery spatula fell into the pt. The spatula was retrieved and no adverse events to the pt were reported. | Product was redesigned to reduce the likelihood of future recurrences. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 349899 | 360727 | 340001 | 2955842200100056 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-10-25 | 2001-10-30 | Yes | No | Health Professional | 40004202 | Yes | No Answer Provided |  | Device Not Retu | No | Unkown | 2001
40000414 | 2001-10-24 | Malfunction |  | It is alleged that during a case the distal clevis on the scalpel/electrocautery broke off into the pt's abdomen and was retrieved. No adverse events to the pt were reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 349904 | 360732 | 340006 | 2955842200100057 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2001-10-26 | 2001-10-30 | Yes | No | Invalid Data | 40000414 | No | Yes |  | Device Not Retu | No | Unkown | 2001
40003615 | 2001-10-09 | Malfunction |  | It is alleged that during a case the distal clevis became detached from the instrument and could not be removed easily via the instrument trocar. It was reported that the left instrument arm was detached from the cannula and was carefully removed out of the cannula with no harm to the patient. | Device not being returned to manufacturer for physical investigation. Pending root cause analysis investigation. | DE BAKEY FORCEPS | INSTRUMENT FORCEPS | 347219 | 358005 | 337374 | 2955842200100053 | 1 | NAY | Manufacturer | ForeignHealth Professional | Physician | Initial | 2001-10-17 | 2001-10-19 | No | Yes | Health Professional | 40003615 | No | Yes |  | Device Not Retu | No | Unkown | 2001
40000414 | 2001-07-13 | Malfunction |  | It is alleged that the hook broke at the tip of the scalpel/electrocautery. No adverse events to the patient were reported. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 347220 | 358006 | 337375 | 2955842200100055 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-10-17 | 2001-10-19 | No | Yes | Health Professional | 40000414 | Yes | No Answer Provided |  | No | No | Unkown | 2001
40000416 | 2001-10-12 | Malfunction |  | It is alleged that during a case the scapel/electrocautery blade tip fell into the patient and was reportedly not located. | Product was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 347223 | 358009 | 337378 | 2955842200100054 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-10-17 | 2001-10-19 | No | Yes | Health Professional | 40000416 | No | No Answer Provided |  | Device Not Retu | No | Unkown | 2001
40000414 | 2001-09-28 | Malfunction |  | It is alleged that the black plastic part on the side of the tip broke off. It was reported that it was removed from the pt with no adverse effects. | Device was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 344726 | 355507 | 334972 | 2955842200100051 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-10-02 | 2001-10-02 | No | Yes | Health Professional | 40000414 | Yes | No | 2001-09-28 | No | No | Invalid Data | 2001
40000414 | 2001-09-28 | Malfunction |  | It is alleged that the black plastic part on the side of the tip broke off into the pt. It was reported that the part was retrieved with no adverse effects to the pt. | Device was redesigned to reduce the likelihood of future recurrences. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 344739 | 355520 | 334985 | 2955842200100052 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-10-02 | 2001-10-02 | No | Yes | Health Professional | 40000414 | Yes | No | 2001-09-28 | No | No | Invalid Data | 2001
40000315 | 2001-09-20 | Malfunction |  | It is alleged that after the case it was noticed that a piece of plastic was missing on the black shaft of the clip applier. It was reported that there is a possibility that the plastic may have fallen into the pt. The surgeon reported that there were no adverse effects to the pt. | Device investigation pending product return. | CLIP APPLIER | CLIP APPLIER | 344746 | 355527 | 334992 | 2955842200100049 | 1 | GDO | Manufacturer | Foreign | UNKNOWN | Initial | 2001-10-02 | 2001-10-02 | No | Yes | Health Professional | 40000315 | Yes | No Answer Provided | 2001-09-22 | No | No | Unkown | 2001
40000314 | 2001-09-26 | Malfunction |  | It is alleged that the clip appliers are not holding clips. | Device mfg process was modified to reduce the likelihood of future recurrences. | CLIP APPLIER | CLIP APPLIER | 344753 | 355534 | 334999 | 2955842200100050 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-10-02 | 2001-10-02 | No | Yes | Health Professional | 40000314 | Yes | No Answer Provided | 2001-09-26 | No | No | Invalid Data | 2001
40000414 | 2001-08-30 | Malfunction |  | It is alleged that two spatula blades fell out of the scalpel/electrocautery and into a pt during a case. It was reported that both blades were retrieved with no adverse effects to the pt. | Device has since been redesigned to reduce the likelihood of recurrences for this event. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 344759 | 355540 | 335005 | 2955842200100048 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-09-27 | 2001-10-02 | No | Yes | Health Professional | 40000414 | No | No Answer Provided | 2001-08-30 | No | No | Unkown | 2001
40005003 | 2001-08-08 | Malfunction |  | It is alleged that during the procedure, the round cautery hook fell into the patient. It was reported that they were able to retrieve the hook with injury or complications. |  | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 338418 | 349101 | 328807 | 2955842200100046 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-08-20 | 2001-08-21 | No | Yes | Health Professional | 40005003 | No | No Answer Provided | 2001-08-08 | Device Not Retu | Yes | Initial | 2001
40000416 | 2001-08-13 | Malfunction |  | It is alleged that the joint that is not supposed to move on the shaft of the instrument broke. No adverse event was reported. | Incident is currently under investigation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 338420 | 349103 | 328809 | 2955842200100045 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-08-20 | 2001-08-21 | No | Yes | Health Professional | 40000416 | Yes | No | 2001-08-13 | Device Not Retu | No | Unkown | 2001
40000414 | 2001-08-02 | Malfunction |  | It is alleged that the instruments arced during a case. It was reported taht the case was completed with no injury to the patient. | Device is pending evaluation. No conclusion can be drawn. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 336612 | 347304 | 327079 | 2955842200100044 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-08-08 | 2001-08-08 | No | Yes | Health Professional | 40000414 | Yes | No | 2001-08-07 | No | No | Unkown | 2001
40000415 | 2001-06-27 | Malfunction |  | It is alleged that during a case the black plastic part on the side of the scalpel/cautery instrument broke, and a 13mm cautery blade broke at the tip. It was reported that both pieces were removed from the pt with no injury. | Incident is currently under investigation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 334480 | 345148 | 325035 | 2955842200100040 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-07-24 | 2001-07-26 | No | Yes | Health Professional | 40000415 | Yes | No Answer Provided | 2001-07-02 | No | No | Unkown | 2001
40005003 | 2001-07-09 | Malfunction |  | It is alleged that during a case the scalpel hook came completely off the instrument two times. It was reported that the hook was retrieved both times and there were no adverse effects to the pt. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 334571 | 345239 | 325122 | 2955842200100041 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2001-07-24 | 2001-07-26 | No | Yes | Health Professional | 40005003 | No | No | 2001-07-10 | Device Not Retu | Yes | Initial | 2001
40005003 | 2001-07-10 | Malfunction |  | It is alleged that during a case a metal piece snapped off the round cautery hook. It was reported that the product was retrieved from the pt with no adverse effects. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 334574 | 345242 | 325125 | 2955842200100043 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2001-07-24 | 2001-07-26 | No | Yes | Health Professional | 40005003 | No | No | 2001-07-10 | No | Yes | Initial | 2001
IS1000 | 2001-07-03 | Malfunction |  | It is alleged that during the use of the instrument control system, the arms had delays in movement and erratic movement. It was reported that there was no adverse event. | Incident is currently under investigation. | INTUITIVE SURGICAL INSTRUMENT CONTROL SYSTEM | INSTRUMENT CONTROL SYSTEM | 334577 | 345245 | 325128 | 2955842200100042 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2001-07-24 | 2001-07-26 | No | Yes | Health Professional | IS1000 | Yes | Yes | 2001-07-03 | No | No | Reuse | 2001
IS1000 | 2001-06-12 | Malfunction |  | It is alleged that during a case the system received a critical fault which was not overridable. It was reported that the set up joints wound not raise up, so they had to be manually forced up and removed from the patient. |  | INSTRUMENT CONTROL SYSTEM | INSTRUMENT CONTROL SYSTEM | 329582 | 340261 | 320373 | 2955842200100035 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2001-06-25 | 2001-06-28 | No | Yes | Health Professional | IS1000 | Yes | Yes | 2001-06-12 | Yes | No | Reuse | 2001
400004 | 2001-06-19 | Malfunction |  | It is alleged that during a case the instrument was arcing and reportedly burned the patient's bladder in several places. It was reported that the cautery hook was switched out but continued to arc. It was also reported that the case was completed with no injury to the patient. | Device is pending evaluation. No conclusion can be drawn. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 329585 | 340264 | 320376 | 2955842200100037 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2001-06-25 | 2001-06-28 | No | Yes | Health Professional | 400004 | Yes | No | 2001-06-19 | No Answer Provi | No | Unkown | 2001
400042 | 2001-06-20 | Malfunction |  | It is alleged that during a case the cautery tip came off. | Device is in the process of being redesigned. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 329679 | 340358 | 320469 | 2955842200100038 | 1 | GEI | Manufacturer | Health Professional | Other | Initial | 2001-06-27 | 2001-06-28 | No | Yes | Health Professional | 400042 | No Answer Provided | No | 2001-06-20 | Device Not Retu | No An | Initial | 2001
40004202 | 2001-06-20 | Malfunction |  | It is alleged that during a case the 13mm cautery blade fell out of the scalpel/electrocautery instrument and into the pt. It was reported that the blade was found and removed from the pt with no injury. | Product is in the process of being redesigned. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 329683 | 340362 | 320473 | 2955842200100039 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2001-06-27 | 2001-06-28 | No | Yes | Health Professional | 40004202 | Yes | Yes | 2001-06-20 | Device Not Retu | Yes | Initial | 2001
40004202 | 2001-06-14 | Malfunction |  | It is alleged that during a case two 13mm cautery blades fell into the pt. It was reported that one was removed from the pt and the other was not found. | Product is in the process of being redesigned. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 329802 | 340483 | 320590 | 2955842200100036 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-06-25 | 2001-06-28 | No | Yes | Health Professional | 40004202 | No | No Answer Provided | 2001-06-14 | Device Not Retu | Yes | Initial | 2001
IS1000 | 2001-05-31 | Malfunction |  | During a case the customer reported that the system had motion sensor mismatches and the camera arm moved laterally, which caused the laparoscope to lean against the abdominal wall. It was reported that the camera arm wouldn't move and the pt had to be lowered off the cannulas. |  | INSTRUMENT CONTROL SYSTEM | INSTRUMENT CONTROL SYSTEM | 329815 | 340496 | 320603 | 2955842200100034 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2001-06-25 | 2001-06-28 | No | Yes | Health Professional | IS1000 | Yes | Yes | 2001-05-31 | Yes | No | Reuse | 2001
40005003 | 2001-06-06 | Malfunction |  | It is alleged that at the end of a lapcholy the surgeon tried to move the cannula while the instrument arm was still in the pt. It was reported that the cautery hook got caught and the tip fell into the pt. It was reported that the tip was not able to be removed from the pt. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROCAUTERY HOOK | 327713 | 338409 | 318563 | 2955842200100030 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40005003 | No | No | 2001-06-06 | Device Not Retu | Yes | Initial | 2001
40005003 | 2001-05-30 | Malfunction |  | It is alleged that a round cautery hook was bending during a case and after it was removed from the patient, broke off into two pieces. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT | 327731 | 338427 | 318581 | 2955842200100033 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40005003 | Yes | Yes | 2001-06-08 | No | Yes | Initial | 2001
40000314 | 2001-05-21 | Malfunction |  | It is alleged that before the surgical procedure the clip applier was dropping clips. It was reported that the case was completed successfully. | Manufacturing process has been changed to reduce the likelihood of a recurrence. | CLIP APPLIER | CLIP APPLIER | 327740 | 338436 | 318588 | 2955842200100025 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40000314 | Yes | No Answer Provided | 2001-05-21 | No | No | Unkown | 2001
40005003 | 2001-05-30 | Malfunction |  | It is alleged that during the case the surgeon went through ten round cautery hooks due to bending, snapping and arcing. It was reported that the case was completed using the conventional laparoscopic methods. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROCAUTERY INSTRUMENT | 327768 | 338464 | 318616 | 2955842200100029 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40005003 | Yes | No | 2001-05-31 | No | Yes | Initial | 2001
40000314 | 2001-05-24 | Malfunction |  | It is alleged that prior to starting the surgical procedure the clip applier was dropping clips. It was reported that the case was completed successfully. | Manufacturing process has been changed to reduce the likelihood of a recurrence. | CLIP APPLIER | CLIP APPLIER | 327776 | 338472 | 318624 | 2955842200100026 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40000314 | No | No Answer Provided | 2001-05-24 | No | No | Unkown | 2001
40000414 | 2001-05-31 | Malfunction | Required Intervention | It is alleged that the black plastic part on the tip of the scalpel/electrocautery broke off and fell into the patient during a surgical procedure. It was reported that the tip was retrieved with no injury to the pt. | Incident is currently under investigation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 327780 | 338476 | 318628 | 2955842200100027 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-06-12 | 2001-06-14 | Yes | No | Health Professional | 40000414 | Yes | No Answer Provided | 2001-05-31 | No | No | Unkown | 2001
40000414 | 2001-06-01 | Malfunction |  | It is alleged that during a right adrenalectomy the scalpel/electrocautery was arcing at the wrist. It was reported that the pt suffered minor burning of the gallbladder, but the case continued and was completed successfully. | Incident is currently under investigation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 327784 | 338480 | 318631 | 2955842200100028 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40000414 | Yes | No | 2001-06-01 | No | No | Unkown | 2001
40004202 | 2001-06-07 | Malfunction |  | It is alleged that the cautery spatula tip fell out of the scalpel/electrocautery instrument during a case. It was reported that the tip could not be located. | Product is in the process of being redesigned. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT | 327785 | 338481 | 318632 | 2955842200100031 | 1 | NAY | Manufacturer | Other | UNKNOWN | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40004202 | No | No Answer Provided | 2001-06-07 | No | Yes | Initial | 2001
40000414 | 2001-05-21 | Malfunction |  | It is alleged that during a surgical procedure the tip of the scalpel/electrocautery instrument broke and fell into the pt. It was reported that the tip was removed from the pt. | Incident is currently under investigation. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 327788 | 338484 | 318635 | 2955842200100032 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2001-06-12 | 2001-06-14 | No | Yes | Health Professional | 40000414 | Yes | Yes | 2001-06-08 | No | No | Unkown | 2001
40000414 | 2001-05-18 | Malfunction |  | It is alleged that a small piece broke off the side of the electrocautery/scalpel tip. It was reported that the tip was removed from the patient without any incident. No other info is available at this time. | Product is in the process of being redesigned. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 324807 | 335460 | 315739 | 2955842200100023 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-05-23 | 2001-05-24 | No | Yes | Health Professional | 40000414 | No | No Answer Provided | 2001-05-18 | Device Not Retu | No | Unkown | 2001
40000413 | 2001-05-15 | Malfunction |  | It is alleged that during a case two cautery hooks broke while in the abdomen. It was reported that the broken part of the hook remained in the instruments so a new hook could not be inserted. No additional information is available at this time. | This product is currently being redesigned. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 324810 | 335463 | 315742 | 2955842200100024 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | Initial | 2001-05-23 | 2001-05-24 | No | Yes | Health Professional | 40000413 | No | No Answer Provided | 2001-05-22 | No | No | Unkown | 2001
40005003 | 2001-05-16 | Malfunction |  | It is alleged that the electrocautery hook fell off into the patient and was not retrieved. There is no additional information at this time. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROCAUTERY HOOK | 324578 | 335233 | 315515 | 2955842200100022 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-05-22 | 2001-05-22 | No | Yes | Health Professional | 40005003 | No | No Answer Provided | 2001-05-16 | Device Not Retu | Yes | Initial | 2001
400050 | 2001-04-26 | Malfunction |  | It is alleged that two cautery hooks were arcing during a case. No other information is available at this time. | Device has not yet been evaluated so no conclusion can be drawn. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 319896 | 330532 | 311068 | 2955842200100020 | 1 | GEI | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 400050 | No | No Answer Provided | 2001-04-26 | No | Yes | Initial | 2001
400004 | 2001-03-26 | Malfunction |  | It is alleged that during a nephrectomy the cautery instrument began to arc. It was reported that the surgeon continued with the procedure and there were no adverse effects to the pt. | Device has not yet been evaluated so no conclusion can be drawn. | SCALPELELECTROCAUTERY | ELECTROSURGICAL DEVICE | 319908 | 330544 | 311080 | 2955842200100016 | 1 | GEI | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 400004 | Yes | No Answer Provided | 2001-03-26 | No | No | Unkown | 2001
40000300 | 2001-03-22 | Malfunction |  | Event Description: --/-- It is alleged that the clips dropped from the clip applier. There was no adverse pt incident reported. --/-- Additional Event Description: --/-- It is alleged that the clips dropped from the clip applier. There was no adverse pt incident reported. | Manufacturer Narrative: --/-- Incident is currently under investigation. | CLIP APPLIER | Unkown | 319915 | 330551 | 311087 | 2955842200100015 | 1 | NAY | Manufacturer | Foreign | UNKNOWN | InitialFollowup | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 40000300 | Yes | No Answer Provided | 2001-03-22 | No | No |  | 2001
400004 | 2001-04-16 | Malfunction |  | It is alleged that a piece of the electrocautery instrument fell off and was found inside the pt during the examination after the case. It was reported that there was no injury to the pt. |  | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 319916 | 330552 | 311088 | 2955842200100019 | 1 | GEI | Manufacturer | Health Professional | Physician | Initial | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 400004 | Yes | Yes | 2001-04-16 | No | No An | Unkown | 2001
400050 | 2001-03-29 | Malfunction |  | It is alleged that during a case the cautery hook broke. There was no instance of an adverse incident with the pt reported. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROSURGICAL ACCESSORY | 320101 | 330739 | 311258 | 2955842200100017 | 1 | GEI | Manufacturer | Company Representative | UNKNOWN | Initial | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 400050 | No | No | 2001-03-29 | Device Not Retu | No An | Unkown | 2001
370269 | 2001-04-10 | Malfunction |  | It is alleged that during a case the screw under the left arm of the cannula amount fell off. It was reported that the screw was retrieved with no adverse effects to the pt and the case was completed successfully. |  | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 320112 | 330748 | 311267 | 2955842200100018 | 1 | NAY | Manufacturer | Consumer | UNKNOWN | Initial | 2001-04-27 | 2001-04-30 | No | Yes | Health Professional | 370269 | Yes | No Answer Provided | 2001-04-10 | Yes | No | Unkown | 2001
IS1000 | 2001-03-05 | Malfunction |  | It is alleged that the counterweight cable on the set up joint is disconnected, which could cause the set up joint to fall on a pt or user. It was reported that there was no adverse event involved with this incident. | Device eval anticipated, but has not yet begun. | INTUITIVE SURGICAL INSTRUMENT CONTROL SYSTEM | INSTRUMENT CONTROL SYSTEM SET | 311053 | 321513 | 302438 | 2955842200100009 | 1 | NAY | Manufacturer | Company Representative | UNKNOWN | Initial | 2001-03-14 | 2001-03-14 | No | Yes | Health Professional | IS1000 | Yes | No Answer Provided | 2001-03-05 | No | No | Unkown | 2001
400004 | 2001-03-07 | Malfunction |  | During a cadaver lab it was alleged that the cautery spatula fell out of the instrument. | Product is in the process of being redesigned. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 311061 | 321521 | 302445 | 2955842200100014 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2001-03-14 | 2001-03-14 | No | Yes | Health Professional | 400004 | No | No Answer Provided | 2001-03-08 | No | Yes | Initial | 2001
40004202 | 2001-03-07 | Malfunction |  | It is alleged that during a case the tip of the 13mm cautery blade popped off into the pt. It was reported that the blade was retrieved and the case was completed with no adverse effects to the pt. | Product is in the process of being redesigned. | 13MM CAUTERY BLADE | ELECTROSURGICAL INSTRUMENT ACC | 311065 | 321525 | 302449 | 2955842200100011 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2001-03-14 | 2001-03-14 | No | Yes | Health Professional | 40004202 | Yes | No Answer Provided | 2001-03-07 | No | Yes | Initial | 2001
40000300 | 2001-03-06 | Malfunction |  | It is alleged that during a lab the clip was loaded into the clip applier, then the clip applier was put into the cannula and loaded into the sterile adapter. As the sterile adapter engages the clip applier, it reportedly spins the shaft and jaws around and drops the clip out. | Incident is currently under investigation. | CLIP APPLIER | CLIP APPLIER | 311074 | 321535 | 302457 | 2955842200100013 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-03-14 | 2001-03-14 | No | Yes | Health Professional | 40000300 | Yes | No | 2001-03-07 | No | No | Unkown | 2001
40000314 | 2001-03-07 | Malfunction |  | It is alleged that before the procedure, the clip was loaded into the clip applier, then the clip applier was put into the cannula and loaded into the sterile adapter. As the sterile adapter engages the clip applier, it reportedly spins the shaft and jaws and drops the clip out. | Incident is currently under investigation. | CLIP APPLIER | CLIP APPLIER | 311080 | 321541 | 302461 | 2955842200100012 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2001-03-14 | 2001-03-14 | No | Yes | Health Professional | 40000314 | Yes | No | 2001-03-07 | No | No | Unkown | 2001
400050 | 2001-03-06 | Malfunction | Required Intervention | It is alleged that when the scalpel/electrocautery instrument was pulled out the round cautery hook was missing. It was reported that an x-ray was taken and the round cautery hook was then seen inside the pt. The pt was reportedly opened with a small incision and the round cautery hook was removed. It was reported that there was no injury to the pt. | Round cautery hook is in the process of being redesigned. | ROUND CAUTERY HOOK | ELECTROSURGICAL INSTRUMENT ACC | 311089 | 321550 | 302469 | 2955842200100010 | 1 | GEI | Manufacturer |  | UNKNOWN | Initial | 2001-03-14 | 2001-03-14 | Yes | No | Health Professional | 400050 | No | No Answer Provided | 2001-03-06 | Device Not Retu | Yes | Initial | 2001
400004 | 2001-02-20 | Malfunction |  | It is alleged that the electrocautery/scalpel broke during a prostatectomy. It was reported that there were no adverse events relating incident. | Device eval has not yet begun, so no conclusion can be drawn at this time. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 313988 | 324529 | 305323 | 2955842200100007 | 1 | NAY | Manufacturer | Health Professional | UNKNOWN | Initial | 2001-03-13 | 2001-03-13 | No | Yes | Health Professional | 400004 | Yes | No Answer Provided | 2001-02-20 | No | No | Unkown | 2001
37026902 | 2001-02-21 | Malfunction |  | It is alleged that during a tecab the camera cannula mount broke. It was reported that there was no injury to the pt. | The design of the camera cannula mount has been modified to increase the load capacity, reducing the likelihood of breakage due to forces encountered during normal use. | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 313991 | 324532 | 305326 | 2955842200100006 | 1 | GCJ | Manufacturer | Foreign | UNKNOWN | Initial | 2001-03-13 | 2001-03-13 | No | Yes | Health Professional | 37026902 | Yes | No Answer Provided | 2001-02-23 | No | No | Unkown | 2001
400004 | 2001-02-21 | Malfunction |  | It is alleged that the electrocautery instrument grounds near the hook over surrounding tissues if they come in contact. It was also reported that the electrocautery head also heats up quickly as seen by the production of steam and smoke from the head, which is not in contact with the tissues. There was no reported info that there was an adverse event relating to the incident. | Device has not yet been evaluated so no conclusion can be drawn. | SCALPELELECTROCAUTERY | ELECTROSURGICAL INSTRUMENT | 314010 | 324551 | 305345 | 2955842200100008 | 1 | NAY | Manufacturer | Company Representative | UNKNOWN | Initial | 2001-03-13 | 2001-03-13 | No | Yes | Health Professional | 400004 | No | No Answer Provided | 2001-02-21 | No | No | Unkown | 2001
IS100004 | 2001-01-18 | Malfunction |  | It is alleged that during a lap nissen case, the cadiere forceps began moving by themselves. It was reported that one jaw would open and the other would quiver. It was also reported that the surgeon moved his head out of the stereo viewer and the instrument continued to move. Reportedly, the staff was able to replace the forceps and the case proceeded for no further incidences. There were no reported pt injuries. | This incident is currently under investigation. A follow up report will be sent that includes the completed investigation findings and corrective actions. | IS1000 INSTRUMENT CONTROL SYSTEM | INSTRUMENT CONTROL SYSTEM | 306556 | 317019 | 298029 | 2955842200100005 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Physician | Initial | 2001-02-12 | 2001-02-14 | No | Yes | Health Professional | IS100004 | Yes | Yes | 2001-01-18 | Yes | No | Invalid Data | 2001
40005003 | 2001-01-11 | Malfunction |  | It is alleged that during a case the round cautery hook broke off below the plastic covering and a piece fell off inside the pt. It was reported that the piece was immediately retrieved and removed from the pt with no injury observed. |  | ROUND CAUTERY HOOK | ELECTROCAUTERY HOOK | 305260 | 315670 | 296765 | 2955842200100002 | 1 | GEI | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2001-02-07 | 2001-02-07 | No | Yes | Health Professional | 40005003 | Yes | No Answer Provided | 2001-01-11 | Yes | No | Unkown | 2001
370269 | 2001-01-18 | Malfunction |  | It is alleged that during a case the pins on the camera cannula mount broke off and fell into the sterile field. It was reported by the surgeon that there were no long term sequelae or adverse events. | A more robust product design has been completed. | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 305262 | 315672 | 296767 | 2955842200100003 | 1 | GCJ | Manufacturer | Health ProfessionalCompany Representativ | Physician | Initial | 2001-02-07 | 2001-02-07 | No | Yes | Health Professional | 370269 | Yes | Yes | 2001-01-18 | Yes | No | Unkown | 2001
370269 | 2001-01-10 | Malfunction |  | It is alleged that the camera cannula mount broke and a small piece of black plastic reportedly fell inside the pt. It was reported by the surgeon that the piece was immediately removed and that no pt injury was observed. | A more robust product design has been completed. | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 305266 | 315676 | 296771 | 2955842200100001 | 1 | GCJ | Manufacturer | Health ProfessionalCompany Representativ | Physician | Initial | 2001-02-07 | 2001-02-07 | No | Yes | Health Professional | 370269 | Yes | Yes | 2001-01-10 | Yes | No | Unkown | 2001
40005003 | 2001-01-23 | Malfunction |  | It is alleged that during a case a round cautery hook broke and a piece fell into the pt. It was reported that the broken piece was removed with no injury to the pt. |  | ROUND CAUTERY HOOK | ELECTROCAUTERY HOOK | 305365 | 315782 | 296870 | 2955842200100004 | 1 | GEI | Manufacturer | Health ProfessionalCompany Representativ | UNKNOWN | Initial | 2001-02-07 | 2001-02-07 | No | Yes | Health Professional | 40005003 | Yes | No Answer Provided | 2001-01-23 | Yes | No | Unkown | 2001
37026902 | 2000-11-07 | Malfunction | Required Intervention | It is alleged that while connecting the ethicon cannula into the camera cannula mount that the two hinges on the side of the camera cannula mount reportedly broke off when snapped down. It was reported that four pieces total fell off. Two pieces were reportedly retrieved from the pericardium, one piece was found in the fold of the instrument arm drape, and the remaining piece was reportedly not found. | The design of the camera cannula mount will be modified to increase the load capacity, reducing the likelihood of breakage due to forces encountered during normal use. | CAMERA CANNULA MOUNT | CAMERA CANNULA MOUNT | 300635 | 310775 | 292104 | 2955842200000001 | 1 | GCJ | Manufacturer | Health Professional | Physician | Initial | 2000-12-27 | 2000-12-27 | No | Yes | Health Professional | 37026902 | Yes | Yes | 2000-11-07 | Yes | No | Unkown | 2000
40004201 | 2000-07-27 | Malfunction | Required Intervention | It is alleged that after a successful mobilization of the left internal mammary artery with the da vinci surgical system, the 13mm cautery blade dropped off the instrument into the patient's thoracic cavity. It was reported that the pt was reoperated due to bleeding from one of the side branches of the left internal mammary artery. It was reported that this was most likely a result of the anastomosis (non-robotically performed). It was also reported that the 13mm blade was retrieved during the reoperation and the pt did not suffer any complications as a result of the da vinci surgical system, nor its instruments. | Method - device was not evaluated: it was concluded that the device was not the cause of the incident; however potential causes for the blade dropping off the instrument are being addressed through training and revision of the instructions for use. | 13MM CAUTERY BLADE | ELECTROCAUTERYSCALPEL | 300656 | 310796 | 292125 | 2955842200000002 | 1 | NAY | Manufacturer | Health Professional | Physician | Initial | 2000-12-27 | 2000-12-27 | No | Yes | Health Professional | 40004201 | No | Yes |  | Device Not Retu | Yes | Initial | 2000