## FDA MAUDE Reports
Model Number | Event Date | Event Type | Patient Outcome | Event Description | Manufacturer Narrative | Brand Name | Type of Device | Device Event Key | MDR Report Key | Event Key | Report Number | Device Sequence Number | Product Code | Report Source | Source Type | Reporter Occupation | Type of Report | Report Date | Date FDA Received | Is This An Adverse Event Report? | Is This A Product Problem Report? | Device Operator | Device MODEL Number | Was Device Available For Evaluation? | Is The Reporter A Health Professional? | Date Manufacturer Received | Was Device Evaluated By Manufacturer? | Is The Device Single Use? | Type of Device Usage | Year Reported
IS3000 A60P6 | 2012-10-11 | Death | Death | It was reported 2 days post a successful da vinci si myomectomy procedure, the patient expired. The intuitive surgical clinical sales representative (csr) that reported this incident indicated that on (b)(6) 2012, the surgeon who performed the surgical procedure reported to him that a tenaculum forceps instrument being used during the surgical procedure was not moving as it should. The csr indicated that he instructed the surgeon to have the instrument removed, cleaned and re-installed. The csr indicated that the site proceeded using the instrument and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported to have occurred at the time of the surgical procedure. On (b)(6) 2012, the intuitive surgical csr was notified by the surgeon that the patient was re-admitted to the hospital on (b)(6) 2012 due to a perforated bowel and the patient expired on (b)(6) 2012. The csr indicated that the surgeon indicated the surgical procedure went fine. The surgeon indicated to the csr that a video recording of the surgical procedure is not available and that the cause of the patient's injury is unknown. Late submission of the medwatch report is due to an reporting oversight by the responsible complaint handler. | On (b)(6) 2012 an inspection of the site's system was performed by an intuitive surgical field service engineer(fse). Inspection of the site's da vinci si surgical system by the fse found no issues and that the site's system functioned within specification. The site has indicated to the intuitive surgical clinical sales representative that the site is currently performing an internal investigation into the reported incident and is unable to provide any other details concerning the incident. Several attempts have been made to verify additional details concerning the reported event; however, no additional information was provided. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2916912 | 2890369 | 2786905 | 2955842201201482 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-10-11 | 2012-12-29 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No | 2012-10-11 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 |  | Injury | Required Intervention | It was reported that after a da vinci s hysterectomy procedure, a bowel perforation was discovered post surgery. The surgeon reportedly stated that the injury may have occurred while inserting the camera cannula, or when cutting through the fascia. No malfunction of the da vinci system, instrument(s) and/or accessories were reported during the course of this procedure. | According to follow up information obtained on (b)(4) 2012 from the clinical sales representative who was present during the surgery, there were no patient complications observed during the case. The bowel perforation was discovered after the patient presented with pain at the emergency room, after which a secondary, non-robotic procedure was conducted to treat the injury. The surgeon believes the perforation may have occurred when the camera cannula was inserted, or when cutting through the fascia. The patient has since recovered and is doing well. No further information was available. There were no reports of any malfunction of the da vinci s surgical system, instruments and/or accessories, and the hospital has continued to use the system to perform surgery on patients. The instruments and accessories user manual specifically states: general precautions and warnings to minimize the risks associated with port placement, ensure the following: appropriate patient positioning to shift organs away from the port placement site. An adequate level of insufflation. Obturator tip is pointing away from major vessels, organs, and other anatomic structures. When possible, visualization of the entire insertion of the cannula using the endoscope is preferred. Moderate, controlled pressure is employed when placing the cannula and obturator. As of (b)(4) 2012, there were no reports of the same recurrence at the this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2916308 | 2889862 | 2786378 | 2955842201201468 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-11-28 | 2012-12-28 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2012-11-28 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-11-26 | Injury | Other | It was reported that during a da vinci si transoral robotic surgery (tors), the patient suffered a lip burn after the cannula came into contact with the patient's lip at the beginning of the tongue base resection procedure. At the time of this incident, the surgeon reportedly utilized the 5 mm maryland dissector and 5 mm monopolar cautery instruments connected to the secondary grounding pad, in conjunction with the recommended flared cannula. To prevent further harm to the patient, intraoperative repositioning of the da vinci system was conducted, such that the cannula was pulled well above the level of the patient's lip. The intended robotic procedure was completed successfully without conversion. | Based on the information provided by the physician assistant (pa) on (b)(6) 2012, the patient has not returned to the hospital due to any post-surgical complications related to this event. The surgery was staged by performing the transoral robotic surgery (tors) for tonsil cancer first, then undergoing a separate (non-robotic) neck dissection a few weeks thereafter. It is her professional belief that the lip burn was caused due to the patient's jaw being too narrow. Per the pa, as soon as the lip burn was noticed, a retractor tool was used to widen the jaw area to properly position the instruments without any further harm to the patient. The pa was unsure whether the energy from the instrument caused the lip burn or whether it was caused by contact with the cannula. However, no arcing from any instrument was observed during the procedure. An attempt was made to retrieve the instruments back for analysis, however, it was stated that they have since been used in other cases and no malfunctions were observed. The electrosurgical unit (esu) mode and setting (coagulation/cut levels) were unknown. Reportedly, the patient is doing very well at this time. No further information was available. The tors procedure reference guide recommends that the 5mm flared cannula be used in conjunction with the endowrist 5mm cautery instruments during da vinci si surgical procedures. Investigation has shown that it is possible for the esu current to pass through the patient side manipulator (psm) arms to ground, and through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2914569 | 2888137 | 2784673 | 2955842201201440 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | PHYSICIAN ASSISTANT | Initial | 2012-11-27 | 2012-12-27 | Yes | Yes | Health Professional | IS3000 A60P7 | No | Yes | 2012-11-27 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-11-27 | Injury | Required Intervention | It was reported that during a davinci si hysterectomy procedure, while the surgical staff was undocking the system from the patient, it was noted that the patient suffered a bowel tear during the procedure. According to the information provided by initial reported clinical sales representative, the tips of the prograsp forceps instrument were not used to grab the bowel, but the shaft of the instrument was used to retract the bowel. | According to the information provided by the clinical sales representative, the gyn surgeon (he had performed 4-5 cases) was performing a lengthy dvh. There was a bowel perforation and tears in the serosa of the sigmoid colon, which were not noticed until the robotic portion was complete and the surgeon was ready to undock the davinci robot system. The csr was in the operating room for most of the case and did not see a specific cause for the injury. The surgeon was using a prograsp forceps instrument to push/sweep the bowel back with the shaft/wrist, but did not grab bowel. The assistant was blindly inserting a lap suction/irrigator throughout the case. Once the surgeon noticed some unusual fluid in the abdomen at the end of the case, she called a colorectal surgeon who repaired the bowel laparoscopically. The patient recovered well and had no further complications. A review of the system logs showed no system malfunction occurred during the procedure performed on (b)(6) 2012. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2910295 | 2883607 | 2780143 | 2955842201201401 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-11-27 | 2012-12-21 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-11-27 | No | No | Reuse | 2012
40018010 |  | Injury | Required Intervention | On (b)(6) 2012, dr (b)(6) provided the article titled electrocautery-associated vascular injury during robotic-assisted surgery published by dr. Beatrice cormier, farr nezhat, jason sternchos, yukio sonoda, and mario leitao. The article alleges that the mcs tip cover accessory, installed on the monopolar curved scissors instruments failed intra-operatively during da vinci robotic procedures. Three of the tip cover accessory failures were alleged to have resulted in patient injuries. The following details are related to the second patient injury as alleged in the article: during a da vinci hysterectomy procedure, a (b)(6) woman with a body mass index (b)(6) presented with stabe ib1 vervical cancer. Robotic-assisted radical hysterectomy, bilateral salpingo-oophorectomy, sentinel lymph node mapping followed by complete pelvic lymph node dissection and para-aortic lymph node sampling were performed using the robotic surgical platform system. When proceeding to right-sided pelvic lymph node dissection, the surgeon accidently activated the scissors monopolar coagulation-type current, set at 35w, when intending to activate the bipolar forceps. A spark traveled through the protective sheath, creating a hole in the right external iliac artery. The artery was grasped with a cadiere forceps to tamponade bleeding, and the defect was sutured robotically with a 5. 0 prolene interrupted stitch. Estimated blood loss was 250 ml, and laparotomy was not required. The patient had an uncomplicated post operative course and was discharged home 2 days after surgery without any subsequent injury. Medwatch reports 2955842-2012-01314 and 2955842-2012-01405 were submitted for the first and third reports of patient injury. | Isi has contacted the initial reporter concerning the reported event. A follow-up mdr will be submitted if additional information is provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2910926 | 2884067 | 2780603 | 2955842201201406 | 1 | NAY | Manufacturer | LiteratureHealth Professional | Physician | Initial | 2012-08-02 | 2012-12-21 | Yes | Yes | Health Professional | 40018010 | No | Yes |  | No | Yes | Unkown | 2012
IS3000 A60P7 | 2012-09-24 | Other |  | It was reported that during a da vinci si cystocele repair procedure, the site experienced system error code 25003. The site contacted isi for technical support assistance. The site indicated to the technical support engineer (tse) that they pressed the fault override button multiple times; however, the issue persisted. The tse asked to review the site's system log; however, there was no connection to the site's system available. While the tse was attempting to troubleshoot the issue with the surgical staff, the tse overheard the surgeon state that the procedure would be converted to open surgical techniques. At this time the site indicated that they would call back. Approximately 30 minutes later, the site called isi technical support back for additional troubleshooting assistance. The site confirmed that the surgeon made the decision to complete the planned surgical procedure using traditional open surgical techniques. Review of the site's system log by the tse found no communication errors associated with the system error code 25003 as stated by the site. The tse explained to the site that system error code 25003 is a recoverable fault and could have been resolved by performing a power cycle of the system. No patient harm was reported. | The investigation conducted by an isi field service engineer (fse) was unable to replicate the system error code 25003 experienced by the site. Review of the site's system log by the fse found that system error code 25003 occurred; however, functional testing of the system found no issues with the site's da vinci si surgical system and that the system functioned within specifications. Additional engineering review of the site's system log found that system error code 25100 also occurred during the surgical procedure. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25003 is a sympathetic error that may occur in conjunction with system error code 25100, when the system detects a disruption in the communication between the different components. The fault may be resolved by power cycling the system. On (b)(4) 2012, isi contacted rn, (b)(6). (b)(6) indicated that no patient harm, adverse outcome or injury to the patient occurred and the patient tolerated the open procedure well. (b)(4). | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2910942 | 2884091 | 2780627 | 2955842201201411 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-09-24 | 2012-12-21 | No | Yes | Health Professional | IS3000 A60P7 | Yes | No | 2012-09-24 | Device Not Retu | No | Reuse | 2012
40018010 |  | Injury | Required Intervention | On (b)(6) 2012, dr. (b)(6) provided the article titled electrocautery-associated vascular injury during robotic-assisted surgery published by dr. (b)(6). The article alleges that the mcs tip cover accessory, installed on the monopolar curved scissors instruments failed intra-operatively during da vinci robotic procedures. Three of the tip cover accessory failures were alleged to have resulted in patient injuries. The following details are related to the third patient injury as alleged in the article: during a da vinci hysterectomy procedure, a (b)(6) woman with stage ib1 cervical cancer was undergoing robotic-assisted radical hysterectomy, bilateral salpingectomy, pelvic lymph node dissection, and para-aortic lymph node sampling using the robotic surgical platform system. During lymphadenectomy, activation of monopolar coagulation-type current, set at 35 w, caused a spark to extrude laterally through the scissors insulating plastic sheath and onto the right external iliac artery. The surface of the vessel was visibly charred, and a defect was noted in the tunica media of the vessel. Because of concern for delayed necrosis, the area was oversewn with a 4. 0 prolene stitch; estimated blood loss was 100 ml. The patient was discharged from the hospital on the third postoperative day because of unrelated respiratory issues requiring nebulizer treatments and chest physical therapy. She did not encounter further complications or sequelae. Medwatch report 2955842-2012-01314 and 2955842-2012-01406 were submitted for the first and second reports of patient injury. | Isi has contacted the initial reporter concerning the reported event. A follow-up mdr will be submitted if additional information is provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2910946 | 2884094 | 2780630 | 2955842201201405 | 1 | NAY | Manufacturer | LiteratureHealth Professional | Physician | Initial | 2012-08-02 | 2012-12-21 | Yes | Yes | Health Professional | 40018010 | No | Yes |  | No | Yes | Unkown | 2012
400180 |  | Injury | Required Intervention | On (b)(6) 2012, dr. (b)(6) at (b)(6) hospital in (b)(6) forwarded to intuitive surgical an article, titled robotic instrument insulation failure: initial report of a potential source of patient injury published by dr. (b)(6). The article alleges that from (b)(6) 2008 - (b)(6) 2009, 12 mcs tip cover accessories, installed on the monopolar curved scissors instruments failed intra-operatively during da vinci robotic procedures and with 3 of the tip cover failures resulting in patient injuries. The following details concerning the second patient injury as indicated in the article is as follows: during a da vinci urological procedure, while performing a retroperitoneal lymphadenectomy using the monopolar curved scissors instrument with the mcs tip cover installed, the patient's small bowel was injured. Damage to affected area was recognized immediately and was repaired by reapproximation of the serosa using a silk suture. Inspection of the tip cover accessory by the site found that the tip cover exhibited a hole at the junction of the silicon and gray plastic shaft on the tip cover. Per the information indicated in the report, the tip cover accessory was inspected prior to use, was installed on the mcs instrument using the tip cover installation tool with no obvious mechanical imperfections observed in the tip cover. No other details concerning the patient were provided. Medwatch report 2955842-2008-01353 was submitted to the fda concerning the first patient injury and medwatch report 2955842-2012-01412 was submitted for the third report of a patient injury. | Isi has contacted the initial reporter concerning the reported event. A follow-up mdr will be submitted if additional information is provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2910970 | 2884101 | 2780637 | 2955842201201414 | 1 | NAY | Manufacturer | LiteratureHealth Professional | Physician | Initial | 2012-08-02 | 2012-12-21 | Yes | Yes | Health Professional | 400180 | No | Yes |  | No | Yes | Unkown | 2012
400180 |  | Injury | Required Intervention | On (b)(6) 2012, dr (b)(6) at (b)(6) hospital forwarded to intuitive surgical an article, titled robotic instrument insulation failure: initial report of a potential source of patient injury published by dr. Adam mues, geoffrey box and ronney abaza. The article alleges that from july 2008 - january 2009 12 mcs tip cover accessories, installed on the monopolar curved scissors instruments failed intra-operatively during da vinci robotic procedures and with 3 of the tip cover failures resulting in patient injuries. The following details concerning the second patient injury as indicated in the article is as follows: during a da vinci urological procedure, while performing a retroperitoneal lymphadenectomy using the monopolar curved scissors instrument with the mcs tip cover installed, the patient's small bowel was injured. Damage to affected area was recognized immediately and was repaired by reapproximation of the serosa using a silk suture. Inspection of the tip cover accessory by the site found that the tip cover exhibited a hole at the junction of the silicon and gray plastic shaft on the tip cover. Per the information indicated in the report, the tip cover accessory was inspected prior to use, was installed on the mcs instrument using the tip cover installation tool with no obvious mechanical imperfections observed in the tip cover. No other details concerning the patient were provided. Medwatch report 2955842-2008-01353 was submitted to the fda concerning the first patient injury and medwatch report 2955842-2012-01414 was submitted for the third report of a patient injury. | Isi has contacted the initial reporter concerning the reported event. A follow-up mdr will be submitted if additional information is provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2910975 | 2884107 | 2780643 | 2955842201201412 | 1 | NAY | Manufacturer | LiteratureHealth Professional | Physician | Initial | 2012-08-02 | 2012-12-21 | Yes | Yes | Health Professional | 400180 | No | Yes |  | No | Yes | Unkown | 2012
IS3000 | 2012-07-13 | Injury | Required Intervention | It was reported that during a da vinci si surgical procedure, the patient sustained damage to the colon and bladder. The damage required an additional surgery, which was completed successfully. No further details have been provided. | There is no allegation of malfunction of da vinci system, instruments or accessories. Furthermore system logs for the date of the surgery showed no system error codes were observed during the procedure. Isi clinical sales representative informed that the patient had swelling from a recent prior surgery, resulting in obstruction of the surgeon's view. The events are not entirely clear, however the patient's colon was injured. The patient required a cholecystectomy to repair the damage. The instruments and accessories user manual specifically states: - caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the patient. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2907217 | 2880580 | 2777116 | 2955842201200414 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-09 | 2012-12-19 | Yes | No | Health Professional | IS3000 | No | No | 2012-08-09 | No | Yes | Initial | 2012
42009308 | 2012-11-13 | Injury | Required Intervention | It was reported that during a da vinci si surgical procedure, the surgical staff reported that the prograsp forceps instrument would not regrasp the pulmonary vein, which caused major bleeding. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | Complaint of instrument not grasping was not confirmed. Instrument placed on system and driven. Recognition and engagement passed. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. Instrument was placed on ipt for testing and failed offset test. System logs for the da vinci si system were also reviewed and showed no errors occured during the surgery. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2907324 | 2880664 | 2777200 | 2955842201201378 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-11-19 | 2012-12-19 | Yes | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-11-19 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-03-09 | Injury | Required Intervention | The following patient story concerning a da vinci si hysterectomy procedure was found posted on the da vinci stories online. Had pop level 4 and hysterectomy, leaving cervix only, done robotically. Still have major abdominal pain and bowel issues. He remains to tell me everything ok. Had ct scan showing 9 mm kidney stone; said this is reason for pain. I'm having the same pain experienced before surgery and after vaginal exams. I feel bad all the time; no energy, bloating, and weight gain. If i take four advil, it helps, but if on my feet awhile, it gets worse. So i sit! i'm tired of complaining to dr and family. Not in my head; this is real. I'm frustrated with everyone. With the holidays coming, will have company for 1 week. Dreading this. I usually look forward to it. Everyone thinks i should be ok, but i'm not! went to family doctor; no answer. Has anyone else felt like this? i also have diverticulosis, but symptoms are different. I put pop as diagnosis, but was a total vaginal prolapse; uterus was out. So had mesh for bladder, rectocele plus all aforementioned. Can't stand pain and being tired anymore. I'm now (b)(6), but feel older, fatter, and lazy. I used to love working in yard, shopping, and cleaning; normal housekeeping. Had cleaning lady after surgery, but let her go because i wanted to do it myself, but it is hard to keep up. Just my husband and dog. Our children are grown. We have large home and yard. Just want an answer and to feel normal!! what else can i say to get this posted. Please, i need answers and can't get them. Is this how i am going to live or will it get better. Some say 1 year to 18 months for full recovery, but the doctor sure didn't say this; he never really said. | There is no allegation of device malfunction. System logs were reviewed for the system but no surgeries were logged for the date of the event ((b)(6) 2012). No other systems could be found for this particular account. If additional information is received, a follow-up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2903048 | 2876728 | 2773264 | 2955842201201336 | 1 | NAY | Manufacturer | Consumer | Other | Initial | 2012-11-18 | 2012-12-17 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-11-18 | No | No | Reuse | 2012
40018010 |  | Injury | Required Intervention | On (b)(6) 2012, dr (b)(6) provided the article titled electrocautery-associated vascular injury during robotic-assisted surgery published by dr. Beatrice cormier, farr nezhat, jason sternchos, yukio sonoda, and mario leitao. The article alleges that the mcs tip cover accessory, installed on the monopolar curved scissors instruments failed intra-operatively during da vinci robotic procedures. Three of the tip cover accessory failures were alleged to have resulted in patient injuries. The following details are related to the first patient injury as alleged in the article: during a da vinci hysterectomy procedure, a (b)(6) woman with a body mass index of (b)(6) was undergoing robotic-assisted hysterectomy, bilateral salpingo-oophorectomy, and pelvic lymph node dissection for endometrial cancer using the robotic surgical platform system. During retrieval of the right obturator lymph nodes, cutting-type current set at 35 w sparked through the intact protective sheath of the monopolar scissors onto an adjacent metallic suction irrigator that was retracting the external iliac vein. Conduction through the suction irrigator resulted in a vessel injury. After robotic suturing and attempted clipping of the defect was unsuccessful, the patient required emergency laparotomy with external iliac vein ligation. Estimated blood loss was 2,000 ml, and she received 2 units of packed red blood cells. Postoperatively, she was administered prophylactic enoxaparin 40 mg daily initially but required full anticoagulation after right lower extremity venous thrombosis was confirmed. She was discharged home on postoperative day 10 without any further sequelae. At 6-month follow-up, she was without evidence of disease and did not have any leg edema. Medwatch report were submitted for the second and third reports of patient injury. | Isi has contacted the initial reporter concerning the reported event. A follow-up mdr will be submitted if additional information is provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2901020 | 2874824 | 2771360 | 2955842201201314 | 1 | NAY | Manufacturer | LiteratureHealth Professional | Physician | Initial | 2012-08-02 | 2012-12-16 | Yes | Yes | Health Professional | 40018010 | No | Yes |  | No | Yes | Unkown | 2012
IS3000 A60P6 | 2012-11-14 | Other |  | It was reported that prior to starting a urological surgical procedure, using the da vinci si surgical system, while setting up the illuminator, it was noted that the light output from the illuminator was dim. The site checked the settings and reseated all parts; however, the issue persisted. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. No patient harm was reported. | The investigation conducted by the field service engineer (fse) found that the dim light issue experienced by the site was due to the incorrect vision settings on the surgeon side cart (ssc) and the brightness control on the illuminator was turned down. Isi has contacted the site several times to verify additional information concerning the status of the patient. However, no additional information has been provided. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2898890 | 2873194 | 2769730 | 2955842201201296 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-11-14 | 2012-12-13 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-11-14 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-11-13 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the scope was fogging up and the procedure had to be converted to open surgical techniques. No patient harm, adverse outcome or injury was reported. | The endoscope was returned and evaluated. The scope was received with mechanical damage to the distal window assembly hermetic seal, resulting in fluid invasion and subsequent minor damage to the optical components. The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2898931 | 2873221 | 2769757 | 2955842201201303 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-11-14 | 2012-12-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-11-14 | Device Not Retu | No | Unkown | 2012
IS1200 A43P9 | 2012-11-15 | Other | Other | It was reported that during a da vinci standard prostatectomy procedure, the camera icon was blinking, as the customer forgot to select an optic. Fifteen minutes later, the customer called back to report that the focus was not working from the surgeon side console (ssc). With the assistance of the technical support engineer (tse), the surgeon performed several troubleshooting steps as directed. The rear ssc connectors were checked, the focus was activated from the focus controller, the camera cable was checked to be sure it was tightened from both ends, and lastly, the camera cable was replaced with a new one. However, as the same problem persisted, the procedure was converted to open surgical techniques. As recommended by the engineer, the camera head was replaced by the head nurse the following day to resolve the reported problem. The system was working as expected after the exchange. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer were associated with the camera head. The camera head produces three dimensional images to the da vinci standard vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. On (b)(6) 2012 the customer provided additional information regarding the reported event. It was stated that the surgery was performed with open surgical techniques, and the patient was released from the hospital. The patient tolerated the procedure well, however, had a longer hospital stay. As of (b)(6) 2012, there have been no reported recurrences of the same issue at this hospital. The camera head has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | DA VINCI STANDARD SURGICALSYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2898936 | 2873240 | 2769776 | 2955842201201308 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-11-15 | 2012-12-13 | Yes | Yes | Health Professional | IS1200 A43P9 | No | Yes | 2012-11-15 | Device Not Retu | No | Reuse | 2012
| 2012-12-10 | Injury | Life ThreateningOther | I was using the davinci system for a hysterectomy on a (b)(6). The precise bipolar instrument was inserted and when i grasped tissue the current activated without the foot pedal being pressed. The audible power sound was heard and the grasped tissue which fortunately was a fallopian tube was cauterized. The entire crew in the operating room witnessed this event. I then grasped the uterus in several locations and the instrument again transmitted energy with the grasping of tissue without use of the foot pedal. The device was removed and the pk device inserted, tested and the case proceeded without incident. Exam of the bowel and surrounding tissue disclosed no injury and the pt was discharged the next morning. This only occurred 2 days ago however so potential for delayed injury still exists. A similar event had occurred 2 months ago using a fenestrated bipolar device on the same system and generator where upon insertion of the instrument energy was immediately delivered prior to foot pedal activation. The device was disconnected, reinserted and the problem did not recur on the second attempt. |  | DA VINCI SURGICAL SYSTEM | SURGICAL SYSTEM | 2910388 | 2881469 | 2778005 | MW5028218 | 1 | NAY | Voluntary |  | Physician | Initial | 2012-12-12 | 2012-12-12 | No | Yes | Service Personnel |  | Yes | Yes |  |  |  |  | 2012
| 2011-09-07 | Injury |  | It was reported that months after a da vinci si hysterectomy procedure a patient returned to a hospital with abdominal pain. A second, non-da vinci surgical procedure was performed and a foreign body was discovered and removed. The surgeon performing the second surgery indicated the foreign body may have originated from a da vinci instrument. | Conclusions: no da vinci instrument has been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On (b)(6) 2012, (b)(6) at the site of the original da vinci hysterectomy procedure, (b)(6) medical center, indicated that they had become aware of a subsequent event after the patient was admitted to a different hospital for abdominal pain. The surgeon performing the second surgery contacted (b)(6) medical center to inform them that he had found a metal piece that was approximately 1-2 cm in total length and had a 2-3 mm round metal end to it with wires coming from it. Neither the metal fragment pictures nor the fragment itself were shared or shown to (b)(6) medical center or intuitive surgical, inc. (b)(6) medical center indicated that currently, to the best of their knowledge, that patient is in good condition. On (b)(6), 2012 the surgeon performing the second diagnostic laparoscopic procedure, dr (b)(6), informed us that he did not have access to the pictures of the instrument fragment. He also did not have a procedure video of the removal of the fragment. He informed us that the patient is doing well. He also informed us that the patient was admitted for pain, and an x-ray was performed which identified a foreign body. Dr (b)(6) did not believe the fragment to be the source of the pain but rather attributed the patient's pain to scar tissue from the initial hysterectomy procedure. The fragment was found in the lateral belly wall and was not poking anything. Further attempts will be made to gather additional information. If additional information is received, a follow-up mdr will be submitted. | 8MM ENDOWRIST INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 2905423 | 2877142 | 2773678 | 2955842201201251 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-11-12 | 2012-12-11 | Yes | No | Health Professional |  | No | No | 2012-11-12 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-11-12 | Other | Other | It was reported that during a da vinci si surgical procedure, the site experienced a degraded image in the right eye of the surgeon side console. With the assistance of an isi technical support engineer (tse) the surgical staff changed the vision to 2d. The customer then called back to inform isi that the case would be converted to open surgery due to the lack of 3d image. At this time the surgical staff made the decision to convert the planned procedure to traditional open surgical techniques to successfully complete the procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the technical support engineer (tse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. On (b)(4) 2012, a follow up call was made with the reporter of this complaint. It was confirmed that the procedure was converted to open surgical techniques. Per customer, the patient was doing well and suffered no complications as a result. No injuries or adverse outcomes occurred. The patient has been released from the hospital. The staff has since ordered a new camera cable, and no similar issues have been observed. No further information was available at this time. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. (b)(4). | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2888394 | 2862918 | 2759454 | 2955842201201226 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-11-12 | 2012-12-10 | Yes | Yes | Health Professional | IS3000 A60P6 | No | Yes | 2012-11-12 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-11-06 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the patient sustained a brachial plexus injury due to how the patient was positioned. | Isi contacted the intuitive surgical clinical sales representative (csr) on (b)(4) 2012. The csr indicated that she met with the surgeon who performed the surgical procedure on (b)(6) 2012. The surgeon indicated that the patient was doing ok. The csr indicated that the surgeon advised her that she was unsure of what caused the brachial plexus injury sustained by the patient. The csr indicated that the surgeon did not provide any other details concerning the patient or the reported event. Intuitive surgical has contacted the site several times to obtain additional details concerning the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2884452 | 2858633 | 2755163 | 2955842201201180 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-11-06 | 2012-12-05 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-11-06 | No | No | Reuse | 2012
| 2012-07-10 | Malfunction |  |  |  |  | INSTRUMENT ROBOTIC SURGICAL HA | 2917024 | 2890524 | 2787060 | 2890524 | 1 | NAY | User Facility |  |  | Initial | 2012-11-07 | 2012-12-04 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2012
40018012 | 2012-10-30 | Injury |  | On (b)(6) 2012, clinical sales representative received information from the customer to inform that during a davinci si myomectomy procedure, while the surgeon was re-docking the system, the surgical staff forgot to install the mcs tip cover on the monopolar curved scissors instrument. The patient had a bowel burn and a general surgeon was called in to remove the burn piece of colon out. | On (b)(4) 2012, intuitive surgical contacted the initial reporter clinical sales representative, csr indicated that she was not present during the case, but according to the information provided to her by the customer, the nurse forgot to install the monopolar curved scissors (mcs) tip cover on the mcs monopolar curved scissors instrument. According to the information provided to the clinical sales representative, this was a onetime incident, the surgeon and the surgical staff at this account is an experienced robotic team. Clinical sales representative had contacted the davinci coordinator to initiate a training section with the surgical team to ensure that this incident does not re-occur. The myomectomy procedure was completed with the davinci robotic system and the patient status is stable. | 8MM MCSTIP COVER | ENDOSCOPIC ELECTROSURGICAL INS | 2879585 | 2853531 | 2750067 | 2955842201201146 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-11-02 | 2012-12-03 | No | Yes | Health Professional | 40018012 | No | No | 2012-11-02 | No | Yes | Unkown | 2012
410322 | 2012-12-03 | Malfunction |  | Blade on the vessel sealer was exposed upon initial use when it had not been activated to be exposed. Removed, no injury to pt. |  | DAVINICI ENDOWRIST ONE VESSEL SEALER | VESSEL SEALER | 2888924 | 2861382 | 2757918 | MW5027955 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-12-03 | 2012-12-03 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2012
400272 | 2012-12-03 | Malfunction |  | While using the harmonic ace curved shears during a robotic case, the tip of the shears burnt off and fell into the pt. Item retrieved and product removed from pt. No harm noted to pt. |  | DAVINCI HARMONIC ACE CURVED SHEARS INSERT | HARMONIC ACE FOR ROBOT | 2889956 | 2861342 | 2757878 | MW5027943 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-12-03 | 2012-12-02 | No | Yes | Health Professional | 400272 | Yes | Yes |  |  |  |  | 2012
| 2012-11-20 | Malfunction |  | During robotic-assisted laparoscopic myomectomy and uterine reconstruction, the pk dissecting forcep stopped working. The wire controlling the articulating jaw had frayed. Surgeon notified. Instrument sequestered. An x-ray was taken; radiologist found nothing in the pt's abdomen. According to robot console, this instrument was used eight times out of ten possible, meaning it had two "lives" left before discard. |  | DAVINCI PK DISSECTING FORCEP | DAVINCI PK DISSECTING FORCEP | 2903309 | 2875582 | 2772118 | 2875582 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2012-11-26 | 2012-11-29 | No | Yes | Health Professional |  | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2012
40018012 | 2012-10-30 | Injury | Other | On (b)(6) 2012, clinical sales representative received information from (b)(6), robotic coordinator, alleging that an monopolar curved scissors (mcs) tip cover used in conjuction with the monopolar curve scissors instrument failed. According to information provided by (b)(6), the patient had burns on her uterus. | Isi contacted initial reporter isi representative. According to the information provided to the isi representative, the customer noted that the patient's uterus was burned during the procedure. The customer noted blanch marks on the patient's uterus. According to the information provided to the isi representative, no additional procedure and/or treatment was necessary to repair the patient's uterus. The patient condition is stable; she was released from the hospital as planned. Isi has contacted robotic coordinator (b)(6) at (b)(6) medical center to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if a device is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2873386 | 2847052 | 2743587 | 2955842201201074 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-10-30 | 2012-11-27 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-10-30 | No | Yes | Unkown | 2012
400180 | 2012-11-16 | Malfunction |  | During uterine dissection utilizing the davinci robotic surgery system, the device had an arc between two instruments. In use at the time were the monopolar scissors and the precise bipolar instrument. Tiny holes were found in the protective sheath cover of the monopolar scissor. The cover was immediately changed and the surgery continued. No pt injury reported. |  | 8MM MCS TIP COVER | MONOPOLAR TIP COVER | 2884639 | 2856021 | 2752557 | MW5027869 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2012-11-23 | 2012-11-23 | No | Yes | Health Professional | 400180 | Yes | Yes |  |  |  |  | 2012
IS2000 A51P8 | 2012-06-27 | Injury | Required InterventionHospitalization | No allegations of deficieny have been made against da vinci system, instruments or accessories. System logs for the da vinci system were pulled for the date of surgery. A review of system logs showed no errors occured on the date of surgery. The root cause of the adverse events of this case cannot be determined. Attempts will be made to gain information from the account and if additional information is received, a follow-up mdr will be submitted to the fda. | On (b)(6) 2012, the console surgeon provided additional details on the event. The surgeon informed that there was no malfunction of da vinci system, instruments or accessories. The surgeon indicated that the first surgery was completed successfully without any patient complications. The surgeon indicated that the cause of the patient dehiscence after the surgery was the result of having vaginal intercourse 4. 5 weeks after the surgery. He had instructed the patient to abstain from intercourse for a period of 8 weeks, however this patient did not follow these guidelines. Dr. Barnes indicated this was the cause of the dehiscence. A second surgery was provided to mitigate the vaginal dehiscence and was also completed successfully without complications. The patient is currently doing well.The following patient story concerning a da vinci s hysterectomy procedure was submitted to intuitive surgical on (b)(6) 2012 through the (b)(4) website: my experience was horrific. I had a radical hysterectomy for cervical cancer. My family was unsupportive as well. I would not ever recommend da vinci. And i certainly would not recommend the hospital where my surgery was performed. I was not provided any information about what to expect. The admitting nurses did not complete all paperwork for proper admittance. I was given too much anesthesia and not fully recovered in the recovery room. While still groggy from surgery i was asked all kinds of medical questions and didn't know what was going on, which i think is just ridiculous of a medical staff. I was in icu for 1 day and there was a total lack of privacy for urinating once the catheter was removed. At 10 weeks post surgery, i had to have surgery again due to vaginal cuff dehiscence. My recovery time was double - my hospital bills were more than double. The vaginal cuff dehiscence was very scary. I was a great amount of pain for 6 days prior to the second surgery. The recovery times for both surgeries were much longer than advertised on the internet. I think the nursing staffs need to do a better job of setting expectations prior to surgery and following up with patients once they leave the hospital. I have 5 - 1 +long scares from the incisions. One incision became infected with staph. The most painful was the incision that went thru muscle. Also there was a great deal of pain in my neck from the anesthesia, after the first surgery. Since the second surgery to repair the first surgery, i have had a lot of swelling in my abdomen. I have lost over 20 pounds - you can't tell because of the swelling. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2869275 | 2842860 | 2739413 | 2955842201201043 | 1 | NAY | Manufacturer | Consumer | Other | FollowupInitial | 2012-10-24 | 2012-11-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-10-24 | Device Not Retu | No | Reuse | 2012
42020505 | 2012-10-22 | Injury |  | It was reported that during a da vinci si hysterectomy procedure, the patient possibly sustained burns at the port site incisions were a monopolar curved scissors instrument and fenestrated bipolar forceps instrument were installed. In addition, it was reported that the wire on the fenestrated bipolar forceps instrument was frayed. There were no missing or fallen pieces reported. Medwatch report 2955842-2012-01052 was submitted concerning the monopolar curved scissors instrument. | The instrument was returned and evaluated. Engineering evaluation was unable to confirm the customer's initial report. The cannula was not returned with the instrument. Microscopic inspection showed an exposed conductor wire. The instrument passed electrical continuity testing. The instrument showed no signs of charring or burn mark around the grip assembly. Engineering confirmed the customer's report of frayed wires. Both of the pitch cables were broken at the distal clevis hub and both cable segments that contain the crimp were still installed in the clevis. The clevis does not exhibit any damage or wear marks. On (b)(4) 2012, intuitive surgical contacted (b)(6) and she indicated the it is indeterminable if the damage to the patient's port site incisions were burns or bruises. (b)(6) indicated that the circumference of the port site were dark; however, the port sites did not exhibit any char marks. (b)(6) indicated that a valley lab force triad electrical surgical unit (esu) was used during the surgical procedure and there were no issues noted with the esu during the surgical procedure. (b)(6) indicated that the fenestrated bipolar forceps instrument was inspected prior to use and there was no damage to the instrument noted. (b)(6) indicated that there were no errors generated from the grounding pad used during the surgical procedure that indicated that it was defective. (b)(6) indicated that arcing during the surgical procedure was not observed; however, the fenestrated bipolar forceps instrument made contact with a monopolar forceps instrument that was also used during the surgical procedure, causing abrasions on one of the instruments. (b)(6) indicated that she is not sure which instrument was damaged. (b)(6) indicated that the planned surgical procedure was successfully completed and there were no intra-operative complications experienced by the patient. (b)(6) indicated that the port sites were excised and closed and to the best of her knowledge, the patient has not returned to the hospital due to experiencing any post surgical complications as are result of the reported event. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2869654 | 2843161 | 2739717 | 2955842201201053 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-24 | 2012-11-21 | Yes | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-10-24 | Device Not Retu | No | Reuse | 2012
42017910 | 2012-10-22 | Injury |  | It was reported that during a da vinci si hysterectomy procedure, the patient possibly sustained burns at the port site incisions were a monopolar curved scissors instrument and fenestrated bipolar forceps instrument were installed. In addition it was reported that the wire on the fenestrated bipolar forceps instrument was frayed. There were no missing or fallen pieces reported. Medwatch report 2955842-2012-01053 was submitted concerning the fenestrated bipolar forceps instrument. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. No sign of arcing was found at the clevis, nor the scissors of the instrument. The instrument was fully functional. On (b)(4) 2012 intuitive surgical contacted (b)(6) and she indicated the it is indeterminable if the damage to the patient's port site incisions were burns or a bruises. (b)(6) indicated that the circumference of the port sites were dark; however, the port sited did not exhibit any char marks. (b)(6) indicated that a valley lab force triad electrical surgical unit (esu) was used during the surgical procedure and there were no issues noted with the esu during the surgical procedure. (b)(6) indicated that monopolar curved scissors instrument was inspected prior to use and there was no damage to the instrument noted. (b)(6) indicated that there were no errors generated from the grounding pad used during the surgical procedure that indicated that it was defective. (b)(6) indicated that arcing during the surgical procedure was not observed; however, the monopolar curved scissors instrument made contact with a fenestrated bipolar forceps instrument that was also used during the surgical procedure, causing abrasions on one of the instruments. (b)(6) indicated that she is not sure which instrument was damaged. (b)(6) indicated that the planned surgical procedure was successfully completed and there were no intra-operative complications experienced by the patient. (b)(6) indicated that the port sites were excised and closed and to the best of her knowledge, the patient has not returned to the hospital due to experiencing any post-surgical complications as are result of the reported event. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2869631 | 2843184 | 2739739 | 2955842201201052 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-24 | 2012-11-21 | Yes | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-10-24 | Device Not Retu | No | Reuse | 2012
IS3000 A6094 | 2012-10-26 | Other |  | It was reported that approximately 30 minutes into a da vinci si partial nephrectomy procedure, the right eye image in the high resolution stereo viewer went black. Unable to resolve the issue, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2869694 | 2843239 | 2739794 | 2955842201201061 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-10-26 | 2012-11-21 | Yes | Yes | Health Professional | IS3000 A6094 | Yes | No | 2012-10-26 | Device Not Retu | No | Reuse | 2012
420194 | 2012-11-08 |  | Required Intervention | Instrument malfunctioned; no adverse outcome to pt. Will be returned to mfr. The device is certified to be reprocessed 10 times; the instrument had been reprocessed less than 10 times. Reprocessing was performed at (b)(6) hospital. |  | DAVINCI INTUITIVE | MEGA NEEDLE DRIVER | 2886136 | 2857752 | 2754288 | MW5027910 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2012-11-21 | 2012-11-21 | No | Yes | Health Professional | 420194 | Yes | Yes |  |  |  |  | 2012
420343 | 2012-09-24 | Malfunction |  | Da vinci thoracic grasper black protective coating ripped by metal da vinci trocar. Instrument inspected by two doctors and stated instrument coating appeared to be complete; wound inspected no coating pieces noted. |  | ENDOWRIST DA VINCI | SYSTEMSURGICALCOMPUTER CONTROL | 2878310 | 2852318 | 2748854 | 2852318 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-25 | 2012-11-19 | No | Yes | Physician | 420343 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42020706 | 2012-10-18 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the tenaculum forceps instrument installed on patient side manipulator (psm) arm #2 locked, could not be removed and the led indicator on the psm was green and yellow. Prior to contacting isi technical support engineering (tse) for troubleshooting assistance the surgeon made the decision to convert the planned surgical procedure to open surgical techniques, due to the patient bleeding. After the system was undocked from the patient, the tse instructed the site to restart the system and the site was able to remove the tenaculum forceps instrument and the system functioned normally. No other details concerning the patient was provided. | Conclusions: the investigation performed by the field service engineer (fse) was unable to replicate the issue experienced by the customer. When the fse arrived on site to evaluate the system, the site was setting up the system for the next procedure. The instrument and sterile adapter were returned for evaluation. Engineering was unable to replicate the issue experienced by the site during functional testing of the instrument on an in-house is3000 system with an in-house sterile adapter. The instrument passed recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly and there were no locking issued noted. The instrument was fully functional. The instrument was also tested on an in-house is3000 system using the returned sterile adapter and testing showed some intermittent engagement issues with the adapter input disc; preventing the instrument from moving intuitively; however, instrument locking was not observed. The fse indicated that during his conversation with (b)(6), she indicated that the surgeon's decision to convert the planned surgical procedure using open surgical techniques was unrelated to the da vinci surgical system; however, the specific reason was not provided. At this time the cause of the bleeding experienced by the patient is unknown. Isi has contacted the site to verify additional details; however, no additional information has been received. If additional information becomes available, a follow-up medwatch report will be submitted to the fda. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 2895078 | 2851029 | 2747584 | 2955842201200983 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-10-18 | 2012-11-17 | Yes | Yes | Health Professional | 42020706 | Yes | Yes | 2012-10-22 | Yes | No | Reuse | 2012
IS2000 A51P8 | 2012-10-18 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure the site lost the image in the right eye of the high resolution stereo viewer (hrsv) on the surgeon side console (ssc). With the assistance of an isi technical support engineer (tse), the surgeon inspected the camera cable and found that the camera cable's insulation sheath was damaged. The surgeon indicated to the tse that he was able to move the sheath. The site attempted to replace the sheath on to the connector, and restarted the camera control unit (ccu); however, the issue persisted. The patient was under anesthesia and one port incision had been created when the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. No patient harm was reported. | The investigation conducted by field service engineering determined that the customer reported problem was associated with the camera cable. Inspection of the camera cable by the fse found no visible damage to the camera cable; however, during functional testing of the camera cable, vision was lost in one eye when the cable was touched. The fse concluded that the camera cable malfunction was related to normal wear and tear. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2863470 | 2836351 | 2732907 | 2955842201200955 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-10-18 | 2012-11-16 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-10-18 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-10-23 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, when the scrub nurse inserted the monopolar curved scissors instrument and bipolar instrument into the patient, an audible sound from the electrical surgical unit (esu) occurred and, the esu activated when the circulating nurse connected the bipolar and monopolar cables to the esu, causing a burn to the fundus of the patient's uterus. The instruments were removed and then re-inserted into the patient and the surgeon proceeded with the procedure. The surgical staff did not experience any other issues with uncommanded cautery activation and the planned surgical procedure was completed. There was no report of an adverse outcome. | The investigation performed by the isi field service engineer was able to replicate the issue experienced by the site determined that the uncommanded cautery activation experienced by the site was caused by the site's valley lab force 2 electrical surgical unit (esu), which is not recommend for use with the da vinci si surgical system. The fse tested the site's esu using his da vinci truck stock valley lab esu cable, monopolor curved scissors instrument and a bipolar forceps instrument. The fse powered on the esu with the settings established at 30 for bipolar, 30 for monopolar cut, and 30 for monopolar coag, with a pure cut setting (the original cut setting is unknown). The fse connected the esu cables on both ends from the instrument. After approximately 15 minutes of testing, an audible beep was noted coming from the esu and the return electrode monitor (rem) light on the esu lit up. The fse then connected the esu to the site's da vinci surgical system and connected the esu cable to the back of the esu and pressed the monopolar foot pedals on the surgeon side cart. After activating the foot pedals, monopolar coag and cut both fired when pressed and stopped firing when released. The fse noted that an audible tone was present with each activation from the esu. The fse found that foot pedal recognition showed active for both monopolar coag and monopolar cut when pressed and there was no indication of active energy when released. The fse was also able to duplicate these results with the bipolar energy. While on site the fse consulted with (b)(6) in the biomedical department at (b)(6). They indicated to the fse that the rem indicator on the esu should not beep without foot pedal activation and that the rem light can come on without concern; however, if an audible noise is present without activation and not connected to the surgeon side cart it is associated with a fault to the esu. The fse advised the site that the valley lab force 2 electrical surgical unit is not recommended for use with the da vinci si surgical system and the fse provided the site with a list of the approved esus for the da vinci si surgical system. The fse indicated that the site reported that their investigation into the reported event found no issue with the monopolar curved scissors instrument or the maryland bipolar instrument and that they have continued to use the instruments. In addition, the fse advised that the site has sequestered the valley lab force 2 esu and is now using an approved esu with their da vinci si surgical system. The instruments and accessories user manual specifically states: caution: for the valleylab force 2 (approved for use with da vinci and da vinci s systems only), erbe icc 350 and erbe vio 300d, follow the specific configuration instructions provided in this section, to ensure that the desired energy is being activated from the da vinci surgeon console on (b)(4) 2012, isi contacted (b)(6) and she indicated that the fundus of the patient's uterus was burned; however, there was no negative impact to the patient and repair of the affected area was not required, since the patient's uterus was being removed as part of the planned surgical procedure. (b)(6) indicated that the procedure was successfully completed and no adverse outcome occurred. (b)(6) also indicated that the patient has not returned to the hospital due to experiencing any post-surgical complications. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2863586 | 2836473 | 2733029 | 2955842201200960 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-10-17 | 2012-11-16 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No | 2012-10-17 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-09-27 | Injury |  | It was reported that during a da vinci s prostatectomy procedure, while the surgeon was performing pedicle dissection, arcing from the mcs tip cover accessory installed on the monopolar curved scissors instrument occurred causing unintentional burning. The mcs instrument was removed and inspection of the tip cover accessory by the surgical staff found that it had a crack at the wrist. The planned surgical procedure was completed using a replacement tip cover accessory. No other information concerning the patient was provided. | The tip cover was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover has surface damage 0. 54 from the distal tip opening. The tip cover was found with a hole approximately 0. 03 in diameter and a slight burn mark was found on the inner wall of the tube. The hole opening on the outer wall is larger than the hole opening on the inner wall. On (b)(4) 2012, intuitive surgical contacted the initial reporter, (b)(6) (distributor). (b)(6) indicated that the burning as described in the initial report means that a spark was observed coming from the tip cover accessory. (b)(6) indicated that tissue was burned; however, it was minimal as the issue was identified early. (b)(6) indicated that repair of the affected area was not required as no perforation or bleeding occurred. (b)(6) indicated that a valley lab force fx was used and the coag setting was 50 watts (the mode was not provided). (b)(6) indicated that the mcs tip cover, mcs instrument and cannula were inspected prior to use and there were no abnormalities observed. Celeste indicated that a bipolar forceps instrument was also used during the surgical procedure; however, the mcs instrument did not make contact with bipolar forceps instrument. (b)(6) indicated that the mcs instrument was last used on (b)(6) 2012 and has been discarded and there is no video recording of the planned surgical procedure. (b)(6) indicated that the patient has not returned to the hospital due to experiencing any post-surgical complications. (b)(6) further indicated that she suspects that the damage to the tip cover occurred as a result of the surgeon's extreme manipulation of the mcs instrument during articulation. On july 12, 2012, intuitive surgical sent a product notification letter to the distributor (b)(6) advising them that a new version of the mcs tip cover (-12) is available. As of (b)(4) 2012, it is unknown why the site has continued to use the -10 version of the tip cover. Intuitive surgical has provided (b)(6) with a copy of the product notification letter that was sent on july 12, 2012. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2863583 | 2836485 | 2733041 | 2955842201200959 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-10-19 | 2012-11-16 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-10-19 | Device Not Retu | No | Unkown | 2012
42020505 | 2012-10-17 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the bipolar forceps instrument broke and half of the jaw fell into the patient. The surgeon was able to locate the broken jaw piece in the pelvis, the piece was retrieved and the planned surgical procedure was completed. | Isi contacted initial reporter, register nurse (b)(6). Based on the information provided by (b)(6), dual consoles were used for this procedure, surgeon dr (b)(6) was using one console and the resident surgeon was using the second console. She indicated that there was a collision between the fenestrated bipolar instrument used by dr (b)(6) and the prograsp instrument used by the resident surgeon. As the surgeons attempted to separate both instruments, they observed that the fenestrated bipolar forceps jaw broke. It took approximately one hour to locate and remove the broken piece from the patient. An x-ray had to be performed to locate the instrument piece; the piece was located in the patient's pelvis. Patient did not suffer any additional complications from this incident or from the exposure to additional anesthesia. Patient was released from the hospital as planned. The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one grip jaw broken at distal clevis grip base. A piece of approximately 0. 185 x 0. 209 had broken off from the instrument. There was no damage found on the cable or the wire. Instrument passed electrical continuity test. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2863691 | 2836577 | 2733133 | 2955842201200966 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2012-10-19 | 2012-11-16 | Yes | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2012-10-19 | Device Not Retu | No | Reuse | 2012
420205 | 2012-10-17 | Other | Required Intervention | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported a broken jaw on the fenestrated bipolar instrument. A piece of the jaw was reported breaking into the patient. The instrument fragment was retrieved with the assistance of a patient x-ray. No patient harm, injury or adverse outcome was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be confirmed a follow-up mdr will be submitted if additional information is received. On (b)(6) 2012, the initial reporter indicated that there was an instrument collision, which caused the tip break to occur. The instruments and accessories user manual specifically states: general precautions and warnings of handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2863702 | 2836582 | 2733138 | 2955842201200975 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-17 | 2012-11-16 | Yes | Yes | Health Professional | 420205 | No | No | 2012-10-17 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-10-04 | Injury | Other | It was reported that after a successful davinci si tongue base resection (tors) procedure, it was noted that the patient had sustained a burn to the upper lip. The surgeon, dr (b)(6), indicated that no additional surgical procedure was required to treat the patient; however, ointment was applied treat the affected area. Dr (b)(6) indicated that the patient has not returned to the site as a result of the reported event. Dr (b)(6) also indicated that he does not recall of traditional laparoscopic techniques were used during the surgical procedure. Dr (b)(6) also indicated that an 8mm cannula was used in conjunction with the 5mm instruments, and it is his belief that the burn sustained the patient was caused by the cannula. Dr (b)(6) indicated that the 5mm flared cannula is now being used for tors procedures at (b)(6). | Based on the information reported by dr (b)(6), no additional surgical procedure was required to treat the patient and to the best of his knowledge the patient has not returned to the hospital due to any post-surgical complications. The tors procedure reference guide recommends that the 5mm flared cannula be used in conjunction with the endowrist 5mm cautery instruments during da vinci si surgical procedures. Investigation has shown that it is possible for the electrosurgical unit (esu) current to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2863677 | 2836597 | 2733153 | 2955842201200967 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2012-10-17 | 2012-11-16 | Yes | Yes | Health Professional | IS3000 A60P7 | No | No | 2012-10-17 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-10-19 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the customer experienced a recoverable system error code 23005. The surgeon elected to convert to traditional open surgical techniques to finish the planned surgery, prior to speaking to an isi technical support representative. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23005, experienced by the site were associated with the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The field service engineer subsequently tested the system, and found it to be working to manufacturer's specifications with no defect or malfunction. The system error code 23005 appears when the da vinci si safety system determines a patient side manipulator (psm) arm has collided with another arm or something else in the environment. Upon determining these conditions, the safety systems put da vinci si in a recoverable safe state. Error code 23005 is a recoverable error. The system user manual recommends that the error code can be cleared by pressing the recover fault button on the patient-side vision cart touchscreen, or the recover button on the surgeon console touchpad. On (b)(6) 2012, isi contacted director of surgery, (b)(4), at (b)(4) hospital, and he indicated that the patient did not experience any surgical complications as a result of the reported event and the patient was released from the hospital. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2863693 | 2836599 | 2733154 | 2955842201200968 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2012-10-19 | 2012-11-16 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-10-19 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-10-18 | Other | Other | It was reported that while driving the patient side cart (psc), the site experienced system error code 816. The surgeon decided to convert the procedure to open prior to calling isi technical support. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineer concluded that system error code 816 experienced by the customer was associated with the patient side cart battery and battery box. The patient side cart battery and battery box were replaced to resolve the issue. The system was tested to isi specifications. The patient side cart battery was returned and evaluated. The battery failure was confirmed and the battery was found to be bulging at the top. The patient side cart battery box was returned and evaluated. The battery box was tested by charging and passed testing. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2863703 | 2836608 | 2733164 | 2955842201200976 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-18 | 2012-11-16 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-10-18 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-10-16 | Other | Other | It was reported that while preparing to start a da vinci si prostatectomy procedure, the site experienced a problem with psm1 not moving a full range of the insertion axis. With the assistance of an isi technical support engineer, the site attempted to use the instrument on other psms, the instrument advanced on the other two psms of the system. The site attempted to check the drape on psm1 and check for external/internal interference on psm1; however, this did not resolve the problem. The surgeon made the decision to abort the planned surgical procedure after ports had been placed and anesthesia had been administered. | The investigation conducted by field service engineering concluded that the insertion issue experienced by the customer was associated with a bent cover on the psm arm insertion path. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected suj harness. On (b)(4) 2012, the initial reporter indicated that the account investigation of the event revealed that the system had been moved the night prior to the surgery. During the move, the system psm arm was bumped against a door. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2860157 | 2832860 | 2729416 | 2955842201200933 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-16 | 2012-11-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-10-16 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-10-16 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the site experienced problems gripping with instruments located on psm1 and psm2 of the system. With the assistance of an isi technical support engineer, the site attempted to rotate sterile adapter discs; however, this did not resolve the problem. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the problem experienced by the customer was associated with sterile adapter. The instrument sterile adapter is part of the instrument arm drape and provides the point of connection between the instrument and the instrument arm. The psms are instrument arms located on the patient cart that provide the sterile interface for the endowrist instruments. The system was tested and performed to original equipment manufacturer specifications. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2860182 | 2832865 | 2729421 | 2955842201200932 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-16 | 2012-11-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-10-16 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-10-04 | Other |  | It was reported that post a successful davinci s prostatectomy procedure, the planned surgical procedure was completed without any problem and/or malfunction of the davinci robot system. The patient suffered an abdominal pain and had to be transferred by an ambulance the next day. The doctor observed that the patient had bleeding from the inferior epigastric artery. The doctor decided to perform an open hemostasis surgery. The patient recovered well enough to be able to be discharged as scheduled on the (b)(6). | Isi contacted initial reporter from (b)(6) university. They indicated there were no reported malfunctions of the davinci s system, instruments and/or accessories during the surgical procedure. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. A review of the system logs showed no system malfunction occurred during the procedure performed on (b)(6) 2012. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2857035 | 2829999 | 2726575 | 2955842201200901 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2012-10-15 | 2012-11-13 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2012-10-15 | No | No | Reuse | 2012
420110 | 2012-10-09 | Injury | Other | It was reported that during a da vinci si ureteral reimplantation procedure, the precise bipolar forceps instrument installed on patient side manipulator (psm) arm 2, nicked a hemotoma located on the patient's ovary. It was reported that when the issue occurred, the psm 2 was responding spastically. The site reported that prior to contacting intuitive surgical for technical support, the instrument and sterile adapter were reseated several times and at the time of their report to technical support engineering, the psm was functioning properly. The planned surgical procedure was completed. | The investigation conducted by field service engineering was unable to replicate the issue experienced by the site. The fse performed functional testing of the site's da vinci si system and it functioned to specification when tested with trunk stock sterile adapter and endowrist instrument. The fse investigation concluded that the spastical movement of the patient side manipulator was likely due to improper seating of the instrument on to the sterile adapter. The sterile adapter provides the sterile interface for the endowrist instruments. On (b)(4) 2012, isi contacted the da vinci coordinator, (b)(6) at (b)(6) medical center and she indicated that the hematoma on the patient's ovary was nicked while the psm arm was behaving spastically. (b)(6) indicated that the hematoma on the patient's ovary was a pre-existing condition and that repair of the affected area was not required. (b)(6) indicated that she discussed the reported event with the surgeon, and the surgeon indicated to her that damage to the patient's ovary would have likely occurred, due to the hematoma. (b)(6) indicated that she was unable to provide any other details concerning the patient since she was not in the or when the issue occurred. The precise bipolar forceps instrument was not returned to isi for failure analysis; however, the sterile adapter was returned and evaluated. Engineering evaluation was unable to duplicate the issue experienced by the site. The sterile adapter was found to function within specification. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The site has been re-trained by the fse on how to properly seat the endowrist instruments on to the sterile adapter. As of november 8, 2012, there have been no reported recurrences of the issue at this hospital. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2854992 | 2827892 | 2724468 | 2955842201200843 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-10-09 | 2012-11-12 | Yes | Yes | Health Professional | 420110 | No | No | 2012-10-09 | No | No | Unkown | 2012
IS3000 | 2012-10-11 | Injury | OtherRequired Intervention | It was reported that a patient sustained a bile duct injury during a da vinci si single site bile duct procedure. The patient underwent a second surgical procedure to have the damaged bile duct repaired at (b)(6) in (b)(6). No other information concerning the patient and event details were provided by the initial reporter. | Intuitive surgical's (b)(4) contacted dr (b)(6), the initial reporter of this event and he indicated that he is unable to provide any details concerning the hospital where the first da vinci si single site bile duct procedure was performed. Dr (b)(6) indicated that the patient is recovering well; however, he declined to provide any other details concerning the reported event. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2854949 | 2827939 | 2724515 | 2955842201200850 | 1 | NAY | Manufacturer | Company RepresentativeHealth Professiona | Physician | Initial | 2012-10-11 | 2012-11-12 | Yes | Yes | Health Professional | IS3000 | No | Yes | 2012-10-11 | No | No | Reuse | 2012
40018012 | 2012-10-11 | Injury | Required Intervention | On (b)(6) 2012, clinical sales representative received information from (b)(6), alleging that an mcs tip cover from a lot number recently involved in a medical device field action initiated by intuitive surgical, was used on a patient prior to the medical device field action. According to information provided by (b)(6), a patient acquired an infection, and the customer points to the possibility of an open package as the primary cause. | Isi has made multiple attempts to contact (b)(6); however, no response has been received as of the date of this report. A follow-up mdr will be submitted if additional information is received. (b)(4). Of those non-uniform seals, our tests indicate that in a small percentage of non-uniform seals, due to the presence of the pleats or creases, there may be a breach in the barrier integrity of that seal. (b)(4). | MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2855102 | 2828081 | 2724657 | 2955842201200863 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-10-18 | 2012-11-12 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-10-18 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-10-11 | Other |  | It was reported that during a da vinci s pyeloplasty procedure, the site experienced system error code 23. With the assistance of an isi technical support engineer, the site power cycled the system; however, system error code 297 occurred. The site then power cycled the system several more times; however, the issue persisted. The patient was under anesthesia and the port sites had been created when the surgeon made the decision to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation performed by the field service engineer (fse) concluded that the system error code 23 and system error code 297 experienced by the site was due to improper set up of the system. The fse found that patient side manipulator (psm) arm 1 on the patient side cart was in a distorted position, thus causing the system error codes experienced by the site. The fse homed the affected psm into the proper position and powered on the system with no error codes noted. Functional testing of the system found that it functioned to specification. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. System error code 297 indicates that the system detected that an electronic component was reporting an incorrect configuration. System error code 297 then puts the da vinci s system into a non-recoverable safe state. On (b)(4) 2012, isi contacted the initial reporter, (b)(6) and she indicated that the rescheduled procedure was performed on (b)(6) 2012. (b)(6) indicated that the patient tolerated the surgical procedure well and no patient harm, adverse outcome or injury occurred. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2855134 | 2828096 | 2724672 | 2955842201200869 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-10-11 | 2012-11-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-10-11 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-10-17 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced a non recoverable system error 5. With the assistance of an isi technical support engineer, and isi clinical sales representative the site attempted to troubleshoot the issue by restarting the system several times; however, this did not resolve the problem. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by field engineering concluded that the system error code 5 experienced by the customer was associated with the enclosure fan assembly. The fan assembly helps regulate temperature of electronics in the patient side cart. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 5 appears when the da vinci s safety system determines one or more fans are not moving as desired. Upon determining this condition the safety system puts da vinci in a recoverable safe state. As of (b)(4) 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2856532 | 2829497 | 2726073 | 2955842201200897 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-10-17 | 2012-11-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-10-17 | Device Not Retu | No | Reuse | 2012
IS3000 A60P7 | 2012-10-10 | Other | Other | It was reported that when the surgical staff went to dock, the system to the patient for a da vinci si benign hysterectomy procedure the blue fiber cable was damaged and got disconnected from the system. A spare cable could not be found. As a result, the decision was made to abort the planned procedure after ports had been placed and anesthesia had been administered. No patient harm was reported. | The investigation conducted by an intuitive surgical field service engineer concluded that the system issue experienced by the customer was due to a broken blue fiber optic cable. The cable passes video, audio, and data during system operation. The affected cable was removed and a replacement cable was provided to the hospital. As of (b)(4)2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2851268 | 2823377 | 2719953 | 2955842201200818 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-10 | 2012-11-07 | Yes | Yes | Health Professional | IS3000 A60P7 | Device Returned To Manufacturer | No | 2012-10-10 | Device Not Retu | No | Reuse | 2012
IS2000 A51P7 | 2012-10-09 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon experienced non-intuitive movement of the instrument wrist while articulating the master tool manipulators from the surgeon side cart. Unable to resolve the issue and prior to contacting isi technical support, the surgeon made the decision to abort the planned procedure after port had been placed and anesthesia adminstered. No patient harm was reported. | The investigation conducted by the field service engineer (fse) was unable to replicate the issue experienced by the site. The system tested to original equipment manufacturer specifications. The fse concluded the site's da vinci s system issue with non-intuitive motion of the instrument wrist experienced by the surgeon was due to the sterile adapter not being properly engaged. The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). No parts were replaced to resolve the issue, however, training was performed to avoid future occurrences. (b)(4). | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2851322 | 2823405 | 2719981 | 2955842201200831 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-09 | 2012-11-07 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2012-10-09 | Device Not Retu | No | Reuse | 2012
428053 | 2012-11-02 | No Answer Provided |  | Single site medium-large clip applier-both in set, held clips too loosely, clips were falling off, even though it hasn't expired yet. |  | DAVINCI | SYSTEMSURGICALCOMPUTER CONTROL | 2880284 | 2854238 | 2750774 | 2854238 | 1 | NAY | User Facility |  |  | Initial | 2012-11-07 | 2012-11-07 | No |  | Invalid Data | 428053 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2012-10-28 | Malfunction | Other | It was reported that during a davinci s partial nephrectomy procedure, the site experienced problems with power on the vision side cart. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the problem with power on the vision side cart experienced by the site was associated with fuses. The system was repaired by replacing the two affected 5amp fuses on the power supply of the vision side cart. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2849499 | 2821627 | 2718223 | 2955842201200802 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2012-11-29 | 2012-11-06 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2012-11-29 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-10-08 | Other | Other | It was reported that during a davinci si sacrocolpopexy procedure, one of the patient side manipulator (psm) arms was jumping when articulating. The surgical staff replaced the davinci instrument and drape prior to calling isi technical support engineer (tse). With the assistance of a tse the customer reseated the instrument and powered off and powered on the system; however, the psm arm problem reoccurred. Surgeon complained that she had limited wrist movement and that the instrument was locked. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The field service engineer visited the site to inspect the patient side manipulator (psm); he performed a test drive, cable tension and sine cycle on the psm with no issues. The psm was replaced and returned to intuitive surgical for further inspection. The patient side manipulator is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The psm arm was returned to intuitive surgical for failure analysis investigation. During internal evaluation, the arm passed (b)(4) switch tests with all circuits reporting, and normal mode with no issues. The psm was tested with in-house test instrument, as the original instruments used during the procedure were not returned. The in-house test instrument jaws were moving smoothly with full range of motion. The cannula mount was current and the wrist output post heights were under specifications. Visual inspection found nothing unusual. The flex cable and connectors were intact and undamaged. Review of the log error codes did not show any error codes that could have caused or contributed to this failure. As of nov 2 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2846768 | 2818808 | 2715404 | 2955842201200798 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2012-10-08 | 2012-11-05 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-10-08 | Device Not Retu | No | Reuse | 2012
42022703 | 2012-09-26 | Injury | Required Intervention | It was reported that at the end of a davinci si myomectomy procedure, the surgeon believes that the patient's bowel was burned. | Isi contacted initial reporter (b)(6), the manager of (b)(6) hospital. Based on the information provided by (b)(6), at the end of a robot-assisted laparoscopic myomectomy, several areas on the small and large bowel appeared concerning for potential thermal injury (unclear and unsure whether they are or not), along with several other lesions concerning for bowel endometriosis. Colorectal surgeon was called in, but we could not confirm whether it is a burn or not. The customer performed a biopsy of the area for frozen, but it was reported as 'chronic and acute inflammatory changes and fibrosis', so the customer could not confirm whether this was in fact a thermal injury to the bowel. The concerning areas were overseen as a precaution. Based on the information provided by the customer, it is indeterminable if the da vinci si system, instruments and/or accessories caused or contributed to the post-surgical complications. Isi has contacted the risk management department at (b)(6) to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if a device is returned (post engineering evaluation) or if additional information is received. According to davinci coordinator mark smith, there have been no reported recurrences of this issue at this hospital. | PK DISSECTING FORCEPS | ELECTROSURGICAL INSTRUMENT | 2846184 | 2818082 | 2714678 | 2955842201200790 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-10-03 | 2012-11-03 | Yes | No | Health Professional | 42022703 | No | No | 2012-10-03 | No | No | Reuse | 2012
420036 | 2012-10-02 | Injury | OtherRequired Intervention | It was reported that during a da vinci s cardiac procedure, the patient side manipulator (psm) arm 2 jumped causing the debakey forceps instrument installed on that psm to lacerate the patient's mammary artery. The planned surgical procedure was completed using open surgical techniques. | On (b)(4) 2012, the isi field service engineer (fse) contacted (b)(6), the initial reporter, and she indicated that during the surgical procedure, the site experienced difficulty installing and removing the endowrist instruments. (b)(6) indicated that no system error or alarms occurred. (b)(6) indicated that the touchscreen monitor seemed to flicker simultaneously when the patient side manipulator (psm) arm jumped. (b)(6) indicated that the sterile draped did not appear to have caused any binding during the surgical procedure, and there was no physical damage to the psm observed. On (b)(4) 2012, the fse arrived at the site to perform an evaluation of the site's da vinci system and the fse learned that (b)(6) was not in the operating room when the reported event occurred and that another nurse present during the surgical procedure observed a blur of the instrument tip passing in front of the endoscope on the touchscreen monitor. (b)(6) referred the fse to the surgical technician, (b)(6) (surname unknown) who was performing an instrument exchange during the surgical procedure when the event occurred. (b)(6) indicated to the fse that the issue occurred after installation of the debakey forceps instrument on psm 2 and while installing a clip applier instrument on psm 1. (b)(6) indicated that he experienced difficulty inserting and removing instruments through the cannula installed on psm 2 and that inspection of the drape installed on the psm, the drape did not appear to be in the way. (b)(6) indicated that the surgeon left the surgeon side console (ssc) multiple times to ensure that the psms were not colliding or rubbing together. (b)(6) also indicated to the fse that after conversion of the planned surgical procedure to open techniques, inspection of the cannulae used during the surgical procedure found no damage. The fse spoke to dr (b)(6), the surgeon who performed the surgical procedure and dr (b)(6) indicated that the planned surgical procedure was progressing smoothly. The surgeon indicated to the fse that he was positioning the mammary artery for clipping when he lost control of the debakey instrument installed on psm 2. The fse indicated that dr (b)(6) described the instrument as lunging forward and turned; but not a full 360 degrees. The surgeon indicated that the instruments were not manipulating correctly at times, and he would go over to the patient side cart to inspect the psms to ensure that there were no obstructions, as he felt that there was a lot of tension on psm 2. The investigation conducted by the fse was unable to replicate the issue experienced by the site. The fse performed functional testing using the endowrist instruments and cannulae that were used during the surgical procedure and the instruments functioned properly with no issues noted and the system functioned within specification. The fse also performed an inspection of the instruments and cannulae used during the surgical procedure and there was no damage found. Review of the system log by the for the procedure date found that system error code 31009 occurred thus indicating that the connection of the instruments used during the surgical procedure, to the system was intermittent. System error code 31009 is an informational error code and occurs when an instrument is initially detected, and then one or more sensors inputs indicate the presence of the instrument was missing for 3 seconds. The sterile adapters were not available for evaluation, as the site indicated that they were discarded. While attempting to replicate the incident, the isi in-house engineering was able to observe the following scenario: when holding the master tool manipulators (mtm) on the surgeon side cart in place, a translational force was applied to the patient side manipulators (psm) arms on the patient side cart (psc). Depending on the direction of the exerted force and the orientation of the mtm, the roll axis of the mtm moved by as much as 60 degrees. Engineering investigation has suggested that the injury sustained by the patient, may have been due to inadvertent bumping of the psm or the psc during the surgical procedure while exchanging instruments. Forces from such a bump could cause movement of the instrument, which would result in the movement of the mtm through force reflection, and then the roll axis could move, thus leading to instrument tip roll motion. Isi has contacted dr (b)(6) to verify the current status of the patient; however, no response has been received. If additional information is received a follow-up medwatch report will be submitted. As of (b)(6) 2012, the hospital has continued to use the system with no reported recurrences of the reported event. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2844355 | 2816230 | 2712826 | 2955842201200770 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2012-10-03 | 2012-11-01 | Yes | No | Health Professional | 420036 | Device Returned To Manufacturer | No | 2012-10-03 | Device Not Retu | No | Unkown | 2012
IS3000 A60P6 | 2012-09-10 | Other | Hospitalization | On (b)(6) 2012, a patient posted the following information on (b)(4): i had the davinci procedure for my hysterectomy. I was not really informed much about it until the day of the surgery. I learned that i was to be inverted upside down for this supposed 6 hr surgery. It turned out to be an 8 hr surgery. Once i came out of anesthesia, i was ok. The second day was absolute hell. I have never in my life felt this much pain. I was told i needed to try and get up to walk. I tried, but failed. I sat up on the edge of the bed and i got very dizzy and nauseated. I did not get to walk that day. The next day i made it out of bed and finally got to walk. The incisions were very painful. It felt like i ripped them open completely when i got up. I still feel extremely bloated and in pain. I was in the hospital for a week. I was told i should have gone home one day after surgery, but i was in agonizing pain. It has been 3 weeks since my surgery and i think now my intestines are twisted. Trying to pass bowels is so painful. The inner groin is numb now due to the lymph node removal. I am thankful for the surgery but i would never recommend this type of surgery for anyone. There is not enough information out there for women to make an informed choice. I do not think that being upside down for an 8 hr surgery is good. That surely can't be good for a person's body. Luckily when i came out of surgery my head and neck weren't swollen like they said it would be. My kids would have been beside themselves. Like i said before, i would never tell anyone that this type of surgery is good. I wouldn't ever do it again. | Isi has contacted dr (b)(6) at (b)(6) medical center and she indicated that she does not recall that any malfunction of the da vinci system, instrument and/or accessories occurred during the surgical procedure. Dr (b)(6) indicated that the patient was positioned in the trendelenburg position. Dr (b)(6) indicated that she has not received any report that the patient experienced any post surgical complications despite the patient's report that she was in the hospital for a week. Dr (b)(6) also indicated that she is unable to provide any other details concerning the patient without written consent from the patient. A follow-up mdr will be submitted if additional information is received. A review of the system logs showed no system malfunction occurred during the procedure performed on (b)(6) 2011. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2835699 | 2807765 | 2704361 | 2955842201200716 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-09-30 | 2012-10-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-09-30 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-09-18 | Injury | Required InterventionHospitalization | It was reported that 4 days after a da vinci si hysterectomy surgery the patient presented a distended abdomen and complained of pain. Imaging was performed and the surgical staff concluded the patient had a sbo (small bowel obstruction) and a port site hernia. The patient had a laparotomy and small bowel resection 1 week after the original surgery and was in the hospital 2-3 weeks after that. On (b)(6) 2012 the surgeon confirmed that patient is at home and doing well currently. This investigation is still ongoing and a follow-up mdr will be generated if additional information is received. | Several attempts have been made to gather additional information and the investigation is still pending. To date no allegation of deficiency of da vinci system, instruments or accessories has been made. Additional attempts are being made to follow-up with the physician to determine the cause of the surgical complications experienced by the patient. If additional information becomes available, a follow up mdr will be submitted. The root cause of the surgical complication experienced by the patient is also currently undetermined. On (b)(6) 2012, the surgeon indicated that the patient is at home now and is doing well. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2835091 | 2807212 | 2703808 | 2955842201200690 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-09-27 | 2012-10-27 | Yes | Yes | Health Professional | IS3000 A60P7 | No | No | 2012-09-27 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-10-02 | Other | Other | It was reported that after of a successful davinci s hysterectomy w/sacrocolpopex procedure, while the surgical staff was removing the trocars from the patient, it was noted that the patient sustained superficial burns at the port site incisions. The patient did not require any additional procedure or treatment for the burns. No additional patient harm, adverse outcome or injury was reported. | Isi contacted initial reporter, nurse (b)(6), and she indicated there were no reported malfunctions of the davinci s system, or instruments/accessories during the surgical procedure. (b)(6) indicated that all of the cautery instruments used during the surgical procedure were tested on (b)(6) 2012 to verify for a breach in the instrument's insulation. Testing found that there were no issues with the cautery instruments. (b)(6) also indicated that the grounding pad was positioned on the patient's left thigh and the electrosurgical unit used was a valley lab device; however, she does not recall what mode and power settings were used during the surgical procedure. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2835099 | 2807220 | 2703816 | 2955842201200706 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2012-10-02 | 2012-10-27 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2012-10-02 | No | No | Reuse | 2012
IS3000 A51P8 | 2012-08-31 | Injury | Required Intervention | On (b)(6) 2012, a patient posted the following information on (b)(4): my doctor recommended the da vinci for my hysterectomy. I am only (b)(6) but, also needed to have my ovaries removed. I went in for the surgery and came out about three hours later to find out that the doctor had missed my uterus and hit my rectum. Considered a mistake. So, now i have to have another surgery on (b)(6)2012, and they will be reopening my c-section scar. I think the da vinci is the worse thing out there people! the doctors just like to use it because, it is a cool new toy for them! i am hoping now that there is not to much damage from the da vinci. I have scars that i didn't need to have and had a surgery that i didn't need to have! i would not recommend anyone to get the da vinci. I have had a ton of pain and because of it, i have to have another surgery! i have two small children and trusted that my doctor knew what he was doing but, that was completely wrong! surgeons should not tell you that it is going to be ok and that this is a great new tool because, it doesn't work at all! i had to have another general surgeon come and check to make sure the da vinci did not do a ton of damage! doctors should be educated more on using machines to operate with. We are not practice dolls! my doctor finally admitted that i was only tenth person he had done this on. I am not happy about that! there needs to be a better paper out there explaining what could end up happening if surgery goes wrong! i only received a pamphlet on how great this machine was and it turned out to ruin my life!. | Intuitive surgical contacted the clinical sales representative (csr), as he was present during the procedure. According to the information provided by the csr, the event occurred when the uterine manipulator, an instrument not sold by intuitive surgical, was placed by the physician, the tip of the rumi perforated the cervix. Physician had placed the ports into the patient's abdomen to create the port incisions, but the surgeon made the decision to abort the procedure once he realized that the uterine manipulator was misplaced. The patient was brought back at later date for a tah (total abdominal hysterectomy) due to the patient's anatomy. According to the information provided to the csr, the patient was released from the hospital. Based on the information provided by the csr, no system malfunction occurred with the da vinci s system, instruments and/or accessories that could have caused or contributed to the patient's complications. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2835103 | 2807222 | 2703818 | 2955842201200702 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-09-26 | 2012-10-27 | Yes | No | Service Personnel | IS3000 A51P8 | No | No | 2012-09-26 | No | No | Reuse | 2012
42020505 | 2012-09-11 | Injury | Required Intervention | On (b)(6) 2012, intuitive surgical became aware of a bowel perforation that allegedly occurred during a davinci si hysterectomy procedure performed on (b)(6) 2012. According to the information provided to the clinical sales representative (csr), the surgeon believes that the fenestrated bipolar instrument activated by itself, without the surgeon pressing the foot pedal. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Several attempts have been made to gather additional information regarding this event without success. As of the date of this report, no additional information has been provided from the site. A follow-up medwatch report will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | FENESTRATED BIPOLAR FORCEPS | ELECTROSURGICAL INSTRUMENT | 2835133 | 2807226 | 2703822 | 2955842201200703 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-10-02 | 2012-10-27 | Yes | Yes | Health Professional | 42020505 | No | No | 2012-10-02 | No | No | Reuse | 2012
IS3000 A60P6 |  | Other | Required Intervention | It was reported that a few days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient experienced 'minor bleeding', the bleeding was controlled in the operating room and there were no further complications associated with this patient. As of the date of this report, no additional information has been provided from the site. | Based on the information provided by dr (b)(6), it was determined that the patient's bleeding was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's bleeding. Several attempts have been made to gather additional information regarding this event from this site without success. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826370 | 2798155 | 2694751 | 2955842201200610 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-09-25 | 2012-10-22 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-09-25 | No | No | Reuse | 2012
IS3000 A60P6 |  | Other | Required Intervention | It was reported that a few days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient experienced 'minor bleeding', the bleeding was controlled in the operating room and there were no further complications associated with this patient. As of the date of this report, no additional information has been provided from the site. | Based on the information provided by dr (b)(6), it was determined that the patient's bleeding was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's bleeding. Several attempts have been made to gather additional information regarding this event from this site without success. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826342 | 2798160 | 2694756 | 2955842201200611 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-09-25 | 2012-10-22 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-09-25 | No | No | Reuse | 2012
40018012 | 2012-09-04 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, arcing was observed on monopolar cautery scissors (mcs) tip cover accessory that was installed on the mcs instrument, and the patient's uterus was burned. The customer replaced the mcs tip cover, but the issue re-occurred. The customer used a third mcs tip cover to complete the planned surgical procedure. | Intuitive surgical contacted the initial reporter, registered nurse (b)(6), and she indicated that the patient's uterus was burned while it was being removed during a hysterectomy procedure. The tip cover was discarded and is not available for investigation; the root cause of the customer reported failure mode cannot be determined. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2826374 | 2798178 | 2694774 | 2955842201200615 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-25 | 2012-10-22 | Yes | Yes | Health Professional | 40018012 | No | Yes | 2012-09-25 | No | Yes | Initial | 2012
IS3000 A60P6 | 2012-09-20 | Death | Death | It was reported that two days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient expired due to uncontrollable post operative bleeding. | Based on the information provided by dr (b)(6), it was determined that the patient's demise was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's death. Several attempts have been made to gather additional information regarding this event from this site without success. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826424 | 2798203 | 2694799 | 2955842201200608 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-09-25 | 2012-10-22 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-09-25 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-09-19 | Other | Other | It was reported that while preparing to start a da vinci s surgical procedure, the site experienced error code 23003 after trying to home the system. The patient was under anesthesia with ports placed when the surgeon decided to abort the planned procedure. | The investigation conducted by field service engineering concluded that the system error code 23003 experienced by the customer was associated with an endoscopic control manipulator (ecm). The ecm is an endoscopic camera arm located on the patient side cart that provides the sterile interface for the endoscope. The system alarm (system generated fault code) functioned as designed. System error code 23003 occurs when either;an arm has restricted movement or is bumped while performing a startup homing sequence or; signifies an encoder or computational problem. The ecm was returned for evaluation and the system error was confirmed. Axis-3 would not home and visual inspection found the brake assembly was loose from its mount, with screws coming out. The brake assembly was bent at an angle preventing normal brake release preventing axis-3 from moving. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826417 | 2798223 | 2694819 | 2955842201200624 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-09-19 | 2012-10-22 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-09-19 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 |  | Other | Required Intervention | It was reported that a few days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient experienced 'minor bleeding', the bleeding was controlled in the operating room and there were no further complications associated with this patient. As of the date of this report, no additional information has been provided from the site. | Based on the information provided by dr (b)(6), it was determined that the patient's bleeding was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's bleeding. Several attempts have been made to gather additional information regarding this event from this site without success. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826430 | 2798234 | 2694830 | 2955842201200612 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-09-25 | 2012-10-22 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-09-25 | No | No | Reuse | 2012
40018012 | 2012-09-04 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, arcing was observed on monopolar cautery scissors (mcs) tip cover accessory that was installed on the mcs instrument, and the patient's uterus was burned. The customer replaced the mcs tip cover, but the issue re-occurred. The customer used a third mcs tip cover to complete the planned surgical procedure. | Intuitive surgical contacted the initial reporter, registered nurse (b)(6), and she indicated that the patient uterus was burned while it was being removed during a hysterectomy procedure. The tip cover was discarded and is not available for investigation; the root cause of the customer reported failure mode cannot be determined. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2826426 | 2798236 | 2694774 | 2955842201200614 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-25 | 2012-10-22 | Yes | Yes | Health Professional | 40018012 | No | Yes | 2012-09-25 | No | Yes | Initial | 2012
IS12000 A43P9 | 2010-09-27 | Death | Death | It was alleged that during a da vinci s partial nephrectomy procedure at (b)(6) hospital, the patient sustained cuts and lacerations, causing injury to the aorta, left renal artery, splenic vein and proximal renal vein. It is also alleged that during the surgical procedure the patient sustained perforation of the pancreas. The surgeon decided to convert the partial nephrectomy procedure to a radical nephrectomy procedure. A specialist was called in to assist with identification and repair of the damage organs and vessels. It is alleged that between (b)(6) 2010 as results of the various injuries and in particular the unrepaired and untreated injury to the pancreas, hemorrhagic debris and fluid accumulated in the contiguous residual cavity, causing and/or contributing to the development of a right pleural effusion and pericardial effusion. On september 28th, 2012, isi became aware that the patient expired on (b)(6) 2010. | At this time it is undetermined as to what caused the patient's death. Several attempts have been made to gather additional information regarding this event from this site without success. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2826487 | 2798286 | 2694882 | 2955842201200613 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-09-28 | 2012-10-22 | Yes | No | Health Professional | IS12000 A43P9 | No | No | 2012-09-28 | No | No | Reuse | 2012
42000606 | 2012-09-25 | Other | Required Intervention | It was reported that during a da vinci s gastrectomy procedure, the large needle driver instrument broke while the surgeon was suturing and an instrument piece fell into the patient. During retraction, a second and third piece came off. Two pieces were retrieved - however, a 1 mm/2 mm pin (part of the cabling system) was not recovered and remained in the patient. This was verified by x-ray. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Several attempts have been made to gather additional information regarding this event without success. As of the date of this report, no additional information has been provided from the site. A follow-up medwatch report will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2828250 | 2800417 | 2697013 | 2955842201200609 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-09-25 | 2012-10-22 | Yes | Yes | Health Professional | 42000606 | No | No | 2012-09-25 | No | No | Reuse | 2012
PK420227 | 2012-06-01 | Malfunction |  | Broken wire noted in davinci instrument after cleaning. No problems noted during surgery. |  | DA VINCI HOT SHEARS | INSTRUMENT SURGICAL ROBOTIC | 2912526 | 2886217 | 2782753 | 2886217 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-19 | 2012-10-19 | No | Yes | Invalid Data | PK420227 | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P4 | 2012-09-16 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, it was reported that the site experienced low light output from the illuminator. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques due to the vision issue and patient anatomical issues. No patient harm was reported. | The investigation conducted by field service engineering (fse) found no problems with the sites illuminator and was unable to replicate the vision issue experienced by the site. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. As a precaution, the fse replaced the site's illuminator. On (b)(4) 2012, da vinci coordinator (b)(6), indicated that she was not present during the surgical procedure and that she was unable to provide details concerning the patient anatomical issues that prevented the procedure from being completed using the da vinci surgical system. No other details concerning the patient were provided. The illuminator was evaluated by the original equipment manufacturer (oem). The oem was unable to confirm the customer reported failure mode. Functional testing performed found that the illuminator functioned to specification. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2824103 | 2795215 | 2691811 | 2955842201200584 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-09-17 | 2012-10-17 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2012-09-17 | Device Not Retu |  | Reuse | 2012
IS2000 A51P8 | 2012-09-13 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, while docking the system to the patient, the site experienced system error code 31017. The site contacted isi for technical support engineering (tse) assistance. With the assistance of the tse, the site performed an emergency power off (epo) and power on the patient side cart. The issue was resolved and the site was able to home the system and proceed with the procedure. After some time, the site contacted isi technical support engineering to report that they experienced non-recoverable system error code 15. With the assistance of the tse, the site was able to clear the error code; however, the error code re-occurred. The site asked if it was possible for the surgeon to continue with the case; however, the tse advised the site that non-recoverable system error code 15 may re-occur during the surgical procedure. Review of the site's system log by the tse found that the system error code 15 was associated with patient side manipulator (psm) arm. The port incisions were already created when the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. | The investigation conducted by the field service engineer (fse) was unable to replicate system error code 31017 and system error code 15 experienced by the site. However, the fse concluded that the system error codes were associated with the remote arm controller (rac) and/or a patient side manipular (psm) arm and replaced them as a precaution. The rac consists of five printed circuit assembly (pca) boards, which operate together to provide control of the system arms. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 31017 when there was a system fault that occurred during system shutdown, which was not cleared by the power cycle. System error code 15 occurs when the system detects that an internal message was corrupted during transmission. The rac and psm were returned and evaluation. Engineering was unable to duplicate the system error codes experienced by the site. Performance testing of the returned affected parts found that they functioned to specification. On oct 8 2012, isi contacted rn, (b)(6) and he indicated that while setting up for the planned surgical procedure, there were no issues noted with the system, however, the system error code 31017 occurred while docking to the patient and all of the patient side manipulator leds were observed to be red. (b)(6) indicated that there was no harm to the patient and the rescheduled procedure was performed a week later. As of oct 12 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2815900 | 2786464 | 2683059 | 2955842201200559 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-13 | 2012-10-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-09-13 | Device Not Retu | No | Reuse | 2012
|  | Injury | Other | On (b)(6) 2012, intuitive surgical received information from dr (b)(6), the (b)(6) at (b)(6) hospital, concerning a posting on an online urology forum called (b)(4) regarding an odd germ. The posting was dated (b)(6) 2012 and the author was dr. (b)(6). Dr. (b)(6) indicated that a patient, (b)(6) underwent a radical prostatectomy procedure using the da vinci surgical system. Dr. (b)(6) indicated that the patient was hospitalized for 8 days and there was an absence of any pathological findings after the surgery procedure. Dr. (b)(6) indicated that the patient's histology was pt2 n0 gl 7apsa 5. 7 ng. Dr. (b)(6) indicated in the posting after 11 weeks clinical infection, bacteremia with bacteroides fragilis in the blood culture (facultative pathogenic bacteria). Cause was an infected lymphocele, in the punctuation also bacteroides fr. Dr. (b)(6) also indicated in the posting therefore an unusual germ, prominent protracted development. Isolated case? connection with method? no other information concerning the reported incident was provided. | Isi has made multiple attempts to contact dr. (b)(6); however, no response has been received; therefore the rootcause cannot be determined. A follow-up mdr will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2816023 | 2786589 | 2683185 | 2955842201200566 | 1 | NAY | Manufacturer | ForeignHealth Professional | Physician | Initial | 2012-09-12 | 2012-10-12 | Yes | No | Health Professional |  | No | Yes | 2012-09-12 | No | No | Unkown | 2012
IS3000 A60P6 | 2011-12-16 | Injury | Required Intervention | On (b)(6) 2012 a patient posted the following information on (b)(4): the patient indicated that on (b)(6) 2012 that she underwent an exploratory examination by dr (b)(6) using the da vinci si surgical system due to a ovarian cyst that burst. The patient reported that she was to undergo a da vinci surgical procedure to remove the cyst, however, 3 days prior to the planned surgical procedure, the cyst burst. Post operatively the patient experienced pain and was told by the attending nurse that dr (b)(6) performed an appendectomy (due to endometriosis and chronic appendicitis) and a cystectomy. Reportedly the patient was not aware that she had any other cysts except for the cyst that resolved itself. The patient was prescribed an antibiotic, nausea suppositories and pain medications and released from the hospital on (b)(6) 2011. The patient indicated that she experienced extreme nausea and vomiting 2 days post op and on (b)(6) 2011, she contacted dr (b)(6) however, dr (b)(6) was not available. The patient indicated that on (b)(6) 2011 went to the (b)(6) emergency room and it was discovered that the patient was septic and was loosing blood. The patient indicated that on (b)(6) 2011 the patient underwent a 2nd surgical procedure to suction the blood out of her pelvic cavity and the surgeon discovered that the patient's right ovarian ligament was torn during the first surgical procedure. The patient indicated that the surgeon who performed the 2nd procedure indicated that there was no obvious cystectomy performed and that a bilateral ovarian drilling procedure was performed. | Based on the information provided by the patient, it is indeterminable if the da vinci si system, instruments and/or accessories caused or contributed to the post surgical complications. A review of the system logs showed no system malfunction occurred during the procedure performed on (b)(6) 2011. Isi has contacted the risk management department at (b)(6) medical center to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2816053 | 2786605 | 2683201 | 2955842201200564 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-09-14 | 2012-10-12 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-09-14 | No | No | Reuse | 2012
42017910 | 2012-09-12 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissors instrument got stuck in the robotic arm of the system. The surgeon did not wish to wait for isi sales representative's help and made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. The representative was 30 to 40 minutes away and was able to remove the instrument from the arm upon arriving. No injuries or adverse harm was reported. | The monopolar curved scissors instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2814003 | 2784591 | 2681187 | 2955842201200538 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-09-12 | 2012-10-11 | Yes | Yes | Health Professional | 42017910 | No | No | 2012-09-12 | No | No | Unkown | 2012
8MM | 2012-10-10 | Malfunction |  | While using the needle driver, it stopped working. It appeared the cable was loose at the tip of the instrument. |  | ENDOWRIST MEGA SUTURECUT | SYSTEMSURGICALCOMPUTER CONTROL | 2848236 | 2820202 | 2716798 | 2820202 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-11 | 2012-10-11 | No | Yes | Physician | 8MM | Yes | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 |  | Injury | Other | On (b)(6) 2012, an article titled posterior ischemic optic neuropathy after minimally invasive prostatectomy published in j neuro-ophthalmol, vol 27, no 4, 2007 was reviewed by an isi clinical project manager. Per the information in the article, post op a da vinci prostatectomy procedure the patient complained of purple vision and loss of inferior visual fields in both eyes. The article indicates that the patient's symptoms worsened when he was upright and on the same day a neuro-ophthalmologist recorded visual acuities of 20/25 in both of the patient's eyes. The article indicates that the patient did not have facial edema and the optic discs were described as of normal appearance with a cup/disc ratio of 0. 25. The article also indicated that automated humphrey visual fields demonstrated bilateral inferior altitudinal defects and the patient's hemoglobin level was 9. 7 g/dl. The article indicates that the patient's blood pressure and pulse were normal. The article indicates that the results of preoperative laboratory studies performed 1 week before the procedure were all within normal ranges, with hemoglobin of 14. 2 g/dl and that the patient's blood pressure on the morning of the surgery was 127/66 mm hg, and the patient had no history of hypertension. The article indicates that shortly after induction of anesthesia, the patient became hypotensive (90/45 mm hg) for about 1 hour with a nadir of 85/33 mm hg. The article indicates that no blood loss occurred. The article indicates the following concerning the surgical procedure: the patient was then placed into a steep trendelenburg tilt in a low lithotomy position. The total operative time was 6 hours 35 minutes, and for >5 hours, the blood pressures were maintained in a range of 90-135 mm hg systolic and 50-95 mm hg diastolic. Multiple intraoperative hemoglobin measurements ranged from 8. 2 to 11. 9 g/dl. Approximately 4 hours into the procedure, after the patient had received 4,300 ml of lactated ringer's solution and had lost approximately 550 ml of blood, t he patient was given a transfusion of 1 unit of packed red blood cells. The procedure was prolonged by the loss of one needle during bladder repair but concluded uneventfully. Total blood loss was 1,200 ml. The article indicates that the patient was observed for 3 days and discharged, during which time the patient's vision remained unchanged. The article further indicates that at a follow up examination performed 3 months later, found that the patient's visual acuities were 20/20 in both eyes, and a visual field examination confirmed stable loss of bilateral inferior fields. The article indicates that an ophthalmoscopy performed disclosed bilateral superotemporal optic disc pallor. The article indicates that the patient lost vision inferiorly, however the patient retained good foveal function with 20/20 vision. The article alleges that prolonged operative time and trendelenburg positioning of the patient may have played a causative role in the ischemic optic neuropathy developed by the patient. | On (b)(6) 2012, isi contacted dr. (b)(6), a neuro-ophthalmologist, who was referenced as one of the authors of the article and he indicated the patient underwent the a da vinci prostatectomy procedure over 5 years ago at (b)(6). Dr. (b)(6) indicated that he did not recall the name of the patient's urologist and advised me to contact dr. (b)(6), who was also referenced as an author in the article. On (b)(6) 2012, isi received an email response from dr. (b)(6) and he indicated that he was not the surgeon who performed the surgical procedure and that he no longer has the patient's records and he is unable to provide information concerning the patient without the patient's consent. If additional information is provided, a follow-up mdr will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2811143 | 2781479 | 2678075 | 2955842201200521 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2012-09-10 | 2012-10-10 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-09-10 | No | No | Reuse | 2012
400180 | 2012-08-07 | Injury | Required InterventionHospitalizationDisability | It was reported that after a successful da vinci si sacrocolpopexy procedure, the patient was discharged on day after her postoperative lleus resolved. On the 5th day the patient began experiencing abdominal pain and nausea. The next day the patient went in for emergency care, where she was identified as having a plain film of her abdomen showing free air. A ct scan was performed of the abdomen and pelvis which showed extravasation of contrast from the large bowel. General surgery returned the patient to the operating room to explore her abdomen and likely perform a colostomy. The patient required the mesh that was placed during the sacrocoplexy to be removed. During exploratory laparotomy, a 5mm perforation was identified in the sigmoid colon several centimeters distant from the area of operation. The defect in the colon was alleged to be consistent with a cautery injury. It is alleged that this injury is likely associated with arcing/cautery of the robotic instrument. | Neither the tip cover accessory nor the monopolar curved scissors have been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The root cause of the surgical complication experienced by the patient is also currently undetermined. On (b)(6) 2012 the initial reporter indicated that the physician involved with this incident had made the determination that arcing had caused the damage; however this determination was made approximately 30 days after the procedure and after the patient had returned complaining of pain. The operating room staff and team did not observe any arcing during the procedure and could not confirm the physician's observation. Additionally, since the physician delayed reporting this arcing, the team could not isolate which instrument was used during the case. Isi was available to provide which monopolar curved scissor instrument was used during that case and has made an attempt to have this instrument returned. The tip cover accessory used with the case was not saved as it is disposable. The initial reporter explained that the physician did not have much experience and had done only a few of these surgeries. The operating room staff could not collaborate the physician's story and nothing abnormal was observed during the original case by the team. While monopolar curved scissors tip cover accessory was disposed of and was not available for return, however there was nothing abnormal observed with the tip cover during the case either. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2811321 | 2781665 | 2678261 | 2955842201200526 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-10 | 2012-10-10 | Yes | Yes | Health Professional | 400180 | No | Yes | 2012-09-10 | No | Yes | Initial | 2012
42017207 | 2012-05-24 | Other | Other | The customer reported that during a da vinci si prostatectomy procedure the maryland bipolar forceps instrument was found faulty. The instrument was replaced with another maryland bipolar forceps instrument and the procedure was completed successfully. On (b)(6) 2012 isi received additional information indicating a patient x-ray was performed post operatively. The x-ray showed no foreign body and the patient did not develop any complications after the surgery. | The instrument was returned and evaluated. Customer initially reported faulty instrument. However for clarification, the defect was a broken pitch cable. The pitch cable was broken at the distal clevis hub. Cable segment that contains the crimp was still installed in clevis. Idler pulley was found to be physically damaged. Idler pulley exhibited indentations and ridges indicative of mishandling and may have caused pitch cable to break. Tube insulation also appears to have scratches and gouge marks with the surface removed. Evidence not conclusive, but damage is likely caused by mishandling/misuse. The instruments and accessories user manual specifically states: - proper care and handling is essential for satisfactory performance of surgical instruments and accessories. Examine the instrument or accessory - including all of its components - thoroughly before and after each use. If any abnormality is found, do not use it. Use the device for its intended purpose only. - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not used a damaged cannula or reducer. Damaged cannulae and reducers may abrade the instrument shaft and generate particles that may fall inside the patient. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2811328 | 2781669 | 2678265 | 2955842201200532 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-05-29 | 2012-10-10 | Yes | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-09-10 | Device Not Retu | No | Reuse | 2012
| 2012-09-04 | Malfunction |  | During the procedure, the surgeon noted the instrument was arcing unintentionally, and burned the uterus. The tip was changed out and it continued to arc. A new instrument was opened and functioned with no problems. It is unknown if the new instrument was of the same or a different lot. ======================manufacturer response for monopolar curved scissors, da vinci monopolar curved scissors (per site reporter). ======================they want the device back within a very short period of time or they reserve the right to cancel the rma without notice. |  | ENDOWRIST HOT SHEARS | MONOPOLAR CURVED SCISSORS | 2869004 | 2842683 | 2739238 | 2842683 | 1 | NAY | User Facility |  |  | Initial | 2012-10-10 | 2012-10-10 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2012-09-10 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, while draping the system, the site experienced multiple occurrences of system error code 23020 and system error code 23034. The site contacted isi for technical support assistance and the technical support engineer (tse) determined that the system error codes experienced by the site were associated with the port clutch on the endoscopic camera manipulator (ecm) arm. With the assistance of the tse, the site restarted the system several times, however, the issue recurred. With the assistance of the tse, the site performed an emergency power off of the system from the surgeon side cart, the patient side cart, pressed the release and port clutch button on the ecm several times, however, the issue persisted. The patient was under anesthesia, however, the port incisions were not yet created when the surgeon made the decision to abort the planned surgical procedure and perform the surgical procedure using open surgical techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 23020 and system error code 23034 were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. System error code 23020 appears when the da vinci s safety system determines one of the switches in a specific manipulator is showing inconsistent signals on two switch leads. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put da vinci s in a recoverable safe state. System error code 23034 appears when the da vinci s safety system determines that after a specified amount of time, a valid event has not been seen for one of the remote compute engine switches. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The ecm was returned and evaluated. Engineering was able to replicate the issue experienced by the site. Functional testing of the ecm found that one of the switches in the manipulator shows inconsistent signals on its two switch leads. Engineering concluded that the inconsistent signals was likely due to a faulty switch and/or switch wiring. The camera focus control module cap (cfc) switch, the embedded serializer for the camera cannula (escc) board and connecting flex cable were replaced to repair the ecm. As of october 9, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2808803 | 2779384 | 2675980 | 2955842201200508 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-10 | 2012-10-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-09-10 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-09-14 | Other | Other | It was reported that during a da vinci s surgical procedure, when the site attempted to place the instrument on the patient side maniluator (psm), the 'invalid cannula detected warning message appeared on the screen. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm, adverse outcome or injury was reported. Patient was discharged from the hospital on (b)(6) 2012. | The investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with patient site manipulator (psm). The patient side manipulator (psm) is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault message) functioned as designed and there was no injury to the patient. The invalid cannula detected warning message appears when the davinci s safety system can not detect the cannula or if the cannula is not seated correctly. Upon determining these conditions, the system displays the message to the user and prevents the instrument from being used or inserted past the cannula end. Investigation of this incident is still ongoing and the root cause of the patient side manipulator is currently undetermined. A follow-up medwatch report will be submitted when the investigation is completed. As of (b)(4) 2012 there have been no reported reoccurrence of the event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2805925 | 2776519 | 2673115 | 2955842201200500 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2012-09-14 | 2012-10-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-09-14 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-09-26 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced multiple instances of system error code 23021. The site contacted isi for technical support engineering (tse) assistance. Review of the site's system log by the tse found that the system error code 23021 was related the endoscopic camera manipulator(ecm) arm. With the assistance of the tse the site performed an emergency power off the surgeon side cart (ssc) and the patient side cart (psc)and restarted the system; however, the issue persisted. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer(fse) concluded that the system error code 23021 experienced by the site was associated with the endoscopic camera manipulator (ecm. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23021 appears when the da vinci s safety system determines that one of the switches on the manipulator has lost its ground sense. Upon determining these conditions, the system records the fault and transitions to a safe state. Investigation of this incident is still ongoing and the root cause of the customer reported failure mode is currently undetermined. A follow-up vigilance report will be submitted when the investigation has been completed. As of october 5, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2804443 | 2775001 | 2671597 | 2955842201200499 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-09-26 | 2012-10-05 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2012-09-26 | No | No | Reuse | 2012
FORCE FXC | 2012-09-11 | Death | OtherDeathHospitalizationRequired Intervention | At the beginning of the surgery, while exploring the patient's abdomen, the sound of the monopolar esu was heard by the surgeon, staff rn, and davinci representative. The surgeon stated that she wasn't depressing the bipolar foot pedal. Physician phone statement noted that this incident occurred at the very start of the case, and the foot pedal was not being activated when the esu was heard going off by itself; observe the davinci rep unplugging the foot pedals connections from the robot at the rear of the esu. Heard it activate again after this was done. Nurse who wrote up incident statement says she was on the other side of the room, but did hear the esu going off. Says she was the person who unplugged the cables from the front of the device. She was not quite sure if this happened at the beginning or the middle of the case. Nurse witness statement: two doctors were using the davinci robot console; they were using mono-polar setting at that moment, heard when physician say "why is the mono-polar still going," said that surgeon saw smoke in the field. Saw davinci rep un-plugged the cables from the back of the force fx esu, switched out force fx to the newer force triad esu, continue with case. As a result of what appears to be an unintended delivery of monopolar energy through a bipolar maryland clamp, the patient sustained a perforated bowel, and despite careful examination of the surgical area, it was not discovered until the patient came back to the hospital approximately a week later with signs of infection. This required an exploratory laparotomy, which revealed a perforated bowel. A portion of colon was removed and sent to surgical pathology. Patient subsequently developed bilateral pulmonary emboli and remains hospitalized to this day. |  | FORCE FX | ELECTROSURGICAL CUTTING COAGU | 2831892 | 2803906 | 2700502 | 2803906 | 1 | GEI | User Facility |  | Nurse | Initial | 2012-10-04 | 2012-10-05 | Yes | Yes | Physician | FORCE FXC | Yes | Yes |  |  |  |  | 2012
FORCE FXC | 2012-09-11 | Death | OtherDeathHospitalizationRequired Intervention | At the beginning of the surgery, while exploring the patient's abdomen, the sound of the monopolar esu was heard by the surgeon, staff rn, and davinci representative. The surgeon stated that she wasn't depressing the bipolar foot pedal. Physician phone statement noted that this incident occurred at the very start of the case, and the foot pedal was not being activated when the esu was heard going off by itself; observe the davinci rep unplugging the foot pedals connections from the robot at the rear of the esu. Heard it activate again after this was done. Nurse who wrote up incident statement says she was on the other side of the room, but did hear the esu going off. Says she was the person who unplugged the cables from the front of the device. She was not quite sure if this happened at the beginning or the middle of the case. Nurse witness statement: two doctors were using the davinci robot console; they were using mono-polar setting at that moment, heard when physician say "why is the mono-polar still going," said that surgeon saw smoke in the field. Saw davinci rep un-plugged the cables from the back of the force fx esu, switched out force fx to the newer force triad esu, continue with case. As a result of what appears to be an unintended delivery of monopolar energy through a bipolar maryland clamp, the patient sustained a perforated bowel, and despite careful examination of the surgical area, it was not discovered until the patient came back to the hospital approximately a week later with signs of infection. This required an exploratory laparotomy, which revealed a perforated bowel. A portion of colon was removed and sent to surgical pathology. Patient subsequently developed bilateral pulmonary emboli and remains hospitalized to this day. |  | FORCE FX | ELECTROSURGICAL CUTTING COAGU | 2831892 | 2803906 | 2700502 | 2803906 | 1 | GEI | User Facility |  | Nurse | Initial | 2012-10-04 | 2012-10-05 | Yes | Yes | Physician | FORCE FXC | Yes | Yes |  |  |  |  | 2012
FORCE FXC | 2012-09-11 | Death | OtherDeathHospitalizationRequired Intervention | At the beginning of the surgery, while exploring the patient's abdomen, the sound of the monopolar esu was heard by the surgeon, staff rn, and davinci representative. The surgeon stated that she wasn't depressing the bipolar foot pedal. Physician phone statement noted that this incident occurred at the very start of the case, and the foot pedal was not being activated when the esu was heard going off by itself; observe the davinci rep unplugging the foot pedals connections from the robot at the rear of the esu. Heard it activate again after this was done. Nurse who wrote up incident statement says she was on the other side of the room, but did hear the esu going off. Says she was the person who unplugged the cables from the front of the device. She was not quite sure if this happened at the beginning or the middle of the case. Nurse witness statement: two doctors were using the davinci robot console; they were using mono-polar setting at that moment, heard when physician say "why is the mono-polar still going," said that surgeon saw smoke in the field. Saw davinci rep un-plugged the cables from the back of the force fx esu, switched out force fx to the newer force triad esu, continue with case. As a result of what appears to be an unintended delivery of monopolar energy through a bipolar maryland clamp, the patient sustained a perforated bowel, and despite careful examination of the surgical area, it was not discovered until the patient came back to the hospital approximately a week later with signs of infection. This required an exploratory laparotomy, which revealed a perforated bowel. A portion of colon was removed and sent to surgical pathology. Patient subsequently developed bilateral pulmonary emboli and remains hospitalized to this day. |  | FORCE FX | ELECTROSURGICAL CUTTING COAGU | 2831892 | 2803906 | 2700502 | 2803906 | 1 | GEI | User Facility |  | Nurse | Initial | 2012-10-04 | 2012-10-05 | Yes | Yes | Physician | FORCE FXC | Yes | Yes |  |  |  |  | 2012
| 2012-09-26 | Malfunction |  | The robotic forcep device refused to open. The failure occurred after the last use of the device and no other device was needed for this case. The forcep was replaced. There was no patient harm. |  | PROGRASP | SYSTEMSURGICALCOMPUTER CONTROL | 2869124 | 2842748 | 2739303 | 2842748 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-05 | 2012-10-05 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-09-06 | Malfunction |  | Harmonic working, then error message and would not work. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 2824523 | 2795758 | 2692354 | 2795758 | 1 | NAY | User Facility |  |  | Initial | 2012-10-03 | 2012-10-03 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
REF410322 VER04 | 2012-09-19 | Malfunction |  | The doctor would only get a partial cut, the manufacturer rep suggested it was the device and a new device was given and it performed accurately. The new device was from the same lot number. Unfortunately the doctor had limited recollection of the event with the time that has passed since the incident. The davinci representative was present and replaced the device with the new vessel sealer which performed without difficulty. It is difficult to determine from the md notes exactly when the device failed during the procedure. There was no patient harm. There were 3 vessel sealer related issues at our facility. The davinci device is new to our facility and physician use would have been low at the time of the event. |  | INTUITIVE | SYSTEM SURGICAL COMPUTER CONTR | 2867858 | 2841249 | 2737805 | 2841249 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Physician | REF410322 VER04 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
410322 | 2012-09-07 | Malfunction |  | Cutting blade would not retract. |  | ENDOWRIST | SYSTEMSURGICALCOMPUTER CONTROL | 2869196 | 2842805 | 2739358 | 2842805 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Invalid Data | 410322 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42029003 | 2012-09-11 | Malfunction |  | Noted a hole in the internal package. |  | DA VINCI SI | SYSTEMSURGICALCOMPUTER CONTROL | 2869205 | 2842811 | 2739364 | 2842811 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Physician | 42029003 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42029003 | 2012-09-11 | Malfunction |  | Noted a hole in the internal package. |  | DA VINCI SI | SYSTEMSURGICALCOMPUTER CONTROL | 2869205 | 2842811 | 2739364 | 2842811 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Physician | 42029003 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42029003 | 2012-09-11 | Malfunction |  | Noted a hole in the internal package. |  | DA VINCI SI | SYSTEMSURGICALCOMPUTER CONTROL | 2869205 | 2842811 | 2739364 | 2842811 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Physician | 42029003 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
420309 | 2012-09-19 | Malfunction |  | The medical device would not open. The device was checked after the case, there were 2 uses still left on device. |  | MEGA SUTURECUT | SYSTEMSURGICALCOMPUTER CONTROL | 2869233 | 2842816 | 2739369 | 2842816 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-10-02 | 2012-10-02 | No | Yes | Invalid Data | 420309 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
40027202 | 2012-09-06 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, while using the harmonic curved shears instrument and harmonic ace insert, a piece of the insert broke and fell into the patient when the harmonic ace insert collided with the uterine manipulator. Multiple x-rays and ct scans were performed however no instrument fragments were found. The patient has recovered well with no post-surgical complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the curved blade had a broken piece at the distal end. The broken piece was not returned for evaluation, but it approximately measures. 400 x. 085. Blade was found to be fractured. 400 below the distal tip, well above the clamp arm pin. The fracture surface was not straight. Engineering concluded that damage was likely due to mishandling/misuse. No other damage was found. There are no components in the harmonic curved shears instruments or inserts that meet the definition of a medical sharp. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2794893 | 2766107 | 2662703 | 2955842201200478 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-09-06 | 2012-10-01 | Yes | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-09-06 | No | Yes | Unkown | 2012
40018012 | 2012-09-06 | Injury | DisabilityRequired Intervention | It was reported that during a da vinci si partial nephrectomy procedure, the tip cover accessory progressively tore at the distal tip. It is alleged that the edges of the torn material were sharp and cut the patient's renal vein. It is alleged that the damage to the renal vein required the entire patient kidney to be removed. | The tip cover accessory was returned and evaluated. Engineering confirmed the reported failure of a split on the top of the tip cover. During failure analysis the distal end of the tip cover showed a couple of splits in the silicone clear part. The splits are roughly 180 degrees apart, offset from the parting lines by. 025. One split is. 150 long and the other is. 035 long. The splits appeared to have initiated at the distal end opening and propagated down. The edges created by the splits are not abnormally sharp. Combination silicone/pellethane section of tip cover exhibits a couple of light scratch marks. No other damage was found. A procedure video was available and reviewed, during which the split at the distal end of the tip cover is visibly seen. The damage to the patient's renal vein appears to have occurred when the split in the tip cover is facing away from the location of the split. It is believed that the injury was caused by application of cautery and arcing at the base of the monopolar curve scissors where it transitions to the tip cover. The tip cover instructions for use (ifu) specifically states: general precautions and warnings inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. Warning: failure to follow these precautions will result in electrical arcs from the wrist and alternate site burns. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2794920 | 2766150 | 2662746 | 2955842201200479 | 1 | NAY | Manufacturer | ForeignHealth Professional | Physician | Initial | 2012-09-12 | 2012-10-01 | Yes | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | Yes | 2012-09-12 | No | Yes | Initial | 2012
IS3000 A61P6 | 2012-09-10 | Death | Death | It was reported that post a successful da vinci si mitral valve repair procedure, the patient was unable to stabilize and the surgeon decided to complete the procedure via a thoracotomy. It was reported that the patient expired shortly after the procedure. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | The surgeon reported that no system malfunction occurred during the procedure and that the patient's outcome was related to her state of health. Based on the information provided by the surgeon, it was determined that the da vinci si surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the patient's demise. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2794990 | 2766182 | 2662778 | 2955842201200480 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-09-10 | 2012-10-01 | Yes | No | Health Professional | IS3000 A61P6 | No | No | 2012-09-10 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-09-04 | Death | Death | It was reported that during a da vinci si partial nephrectomy procedure during tumor removal the patient started bleeding. It was noticed that the patient's vena cava was ruptured. The patient was supplied blood, however the patient expired. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. On (b)(4) 2012, isi contacted the clinical sales representative (csr) and the csr indicated that the account had deemed this death due to surgeon error. During the time of the bleeding the surgeon was reported having a slow response to mitigating the bleeding. The account has refused to discuss this case any further as they claim it is unrelated to the da vinci system. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2788357 | 2759090 | 2655704 | 2955842201200459 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-09-06 | 2012-09-25 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-09-06 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-08-21 | Injury | Required Intervention | After a successful da vinci s hysterectomy procedure the patient suffered bleeding from the drain. Then the doctor decided to confirm the bleeding by traditional laparoscopic procedure. The doctor observed that the bleeding had already stopped, but the doctor treated the suspected area with coagulation. The patient was doing well after this procedure, and discharged on schedule (b)(6). There was no reported malfunction of da vinci system, instruments or accessories. | No malfunctions of the da vinci s surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. On (b)(4) 2012, isi contacted the distributor and the distributor indicated that the patient was discharged from the hospital and that there have no been report of injury to the patient. The account has stated that no further information will be provided. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2788360 | 2759106 | 2655682 | 2955842201200460 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-09-04 | 2012-09-25 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-09-04 | No | No | Reuse | 2012
400272 | 2012-09-25 | Injury | Required Intervention | Harmonic ace shears insert tip broke off inside the pt during surgical intervention. Reason for use: robotic myomectomy tubal perfusion dx laparoscopy. |  | HARMONIC ACE CURVED SHEARS INSERT DA VINCI | HARMONIC CURVED SHEARS | 2799297 | 2768114 | 2664710 | MW5027084 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-09-25 | 2012-09-25 | No | Yes | Health Professional | 400272 | Yes | Yes |  |  |  |  | 2012
| 2012-08-31 | Malfunction |  | The patient was undergoing a robotic total laparoscopic hysterectomy with bilateral salpingectomy. The wires on the instrument broke. The fragments were retrieved from the patient's abdomen. |  | DAVINCI ENDOWRIST PK DISSECTING FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 2869907 | 2843522 | 2740058 | 2843522 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-09-25 | 2012-09-25 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2012
M1020502203 | 2012-09-14 | Malfunction |  |  |  | DAVINCI | SYSTEMSURGICALCOMPUTER CONTROL | 2797456 | 2768215 | 2664811 | 2768215 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-09-24 | 2012-09-24 | No | Yes | Physician | M1020502203 | No | No Answer Provided |  |  |  |  | 2012
| 2012-09-19 | Malfunction |  | The patient was undergoing a robotic-assisted total laparoscopic hysterectomy with bilateral salpingectomy. During the procedure, one of the wires to the fenestrated bipolar forceps broke. This was the fourth use of this instrument; they are designed to be used ten times before being replaced. |  | DAVINCI | SYSTEMSURGICALCOMPUTER CONTROL | 2798155 | 2769190 | 2665785 | 2769190 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-09-24 | 2012-09-24 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2012
40018012 | 2012-08-09 | Injury | Required Intervention | On (b)(4) 2012, isi contacted (b)(6) in risk management and she indicated that the tip cover accessory is not being returned for evaluation, as they would like to retain the tip cover for their records. (b)(6) indicated that she will try to obtain a picture of the affected tip cover to verify the tip cover version. (b)(6) also indicated that she will follow up with the surgical staff to verify if there were any issues noted prior to use of the tip cover accessory. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | On (b)(6) 2012, isi contacted the risk manager, (b)(6) and she indicated that the affected mcs tip cover accessory was from model 400180-12. (b)(6) indicated that she followed up with the clinical staff concerning the reported event and they indicated to her that there were no issues noted during installation of the tip cover on to the monopolar curved scissors (mcs) instrument and nothing unusual occurred during advancement and removal of the of the mcs instrument (with the tip cover installed) through the cannula. (b)(6) indicated that the tip cover is being held by the hospital's pathology department and the tip cover is being retained for the hospital's records. (b)(6) also indicated that there is no video recording available of the planned surgical procedure. (b)(6) indicated that she is unable to provide any other details concerning the patient's status. On (b)(4) 2012, isi received a photographic image of the reported mcs tip cover accessory. Review of the photograph by isi engineering found no noticeable damage on the surface of the tip cover.On date (b)(6) 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: the patient was undergoing robotic surgery for a subtotal hysterectomy with implantation of a sling. During the procedure, the tip cover accessory which covers the end of the robotic arm was not attached at the end of the procedure. Prior to closure, the physician conducted a visual sweep with the scope in order to locate the tip cover accessory and was unable to locate it. Six x-rays were taken in an attempt to locate the tip cover accessory. Due to the fact that the tip cover accessory is non radiopaque, the results of the x-rays were inconclusive. A ct scan was then conducted. The ct scan was positive for the cover tip accessory which was noted to be in the left rectum muscle. The patient underwent a second procedure for the successful removal of the tip cover accessory. More specifically, the tip cover accessory covered the monopolar scissor. This event apparently occurred midway through the procedure. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2782720 | 2753502 | 2650097 | 2955842201200445 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2012-10-08 | 2012-09-20 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-10-08 | Device Not Retu | Yes | Unkown | 2012
420194 | 2012-09-13 | Malfunction |  | Under direct visualization during robotic laparoscopic procedure, articulating wire snapped while mega driver in use. No fragments of wire were noted in surgical field. Device removed and examined. One wire noted to be broken. No pt harm or surgical delay. X-ray taken post-op. |  | INTUITIVE MEDICAL | MEGA NEEDLE DRIVER | 2843712 | 2809958 | 2706554 | 2809958 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-09-14 | 2012-09-19 | No | Yes | Health Professional | 420194 | Yes | Yes |  |  |  |  | 2012
IS3000 A60P6 | 2012-08-24 | Other | Other | It was reported that at the beginning of a da vinci si hysterectomy procedure, the site experienced system error code 23017. The site contacted isi technical support for troubleshooting assistance. The technical support engineer (tse) reviewed the site's system log and found that system error code 23017 was associated with the endoscopic camera manipulator (ecm) arm. With the assistance of the tse, the site exercised an axis on the ecm and power cycled the system. However, system error code 25588 occurred. The tse instructed the site to power cycle the system several more times; however, the issue persisted. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering found system error code 23017 and 25588 in the system error logs. It was determined that these errors were associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured electrical behavior did not match the internal simulation of the motor circuit. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. System error code 25588 appears when the da vinci s safety system determines that, at startup, the current to one motor cannot be controlled to the required precision. Upon determining this condition, the safety systems put da vinci in a non-recoverable safe state. The ecm was returned to isi for failure analysis investigation. There was a data acquisition fault detection (dafd) motor voltage tracking failure on the axis 4 motor/encoder, thus generating the error codes experienced by the site. Isi has made multiple attempts to contact the customer to verify additional information concerning the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted if additional information is received. As of (b)(6) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775165 | 2746106 | 2642702 | 2955842201200434 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-24 | 2012-09-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-08-24 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-08-27 | Death | Death | It was reported that approximately 2 weeks post a da vinci si hysterectomy procedure, the patient developed a pulmonary embolism and expired. | On (b)(6) 2012, intuitive surgical contacted the surgeon who performed the surgical procedure and he indicated that the patient was released from the hospital 2 days post op and that one week post op the patient returned to the hospital in good condition; however, it was discovered that the patient had a small lesion of the bladder diathermic. The surgeon indicated the patient got a cad and had no complaints anymore and was ready to be discharged from the hospital when she developed a massive pulmonary embolism. The surgeon indicated that they tried to resuscitate the patient without success and the patient expired on (b)(6) 2012. The surgeon indicated that there was no malfunction of the da vinci si system, instruments and/or accessories during the planned surgical procedure and that the procedure was not related to the pulmonary embolism that the patient developed. No other information was provided. If additional information is provided a follow-up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775305 | 2746211 | 2642807 | 2955842201200435 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-09-04 | 2012-09-14 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-09-04 | No | No | Reuse | 2012
IS3000 A60P7 | 2012-08-24 | Injury | Other | It was reported that post a transoral robotic surgery(tors) procedure, using the da vinci si surgical system, it was noted that the patient sustained a burn on the lip. The monopolar cautery instrument was inspected and it was noticed that there was a weak area in the black insulation. It is believed this area may have caused energy to be released, and may have caused the burn. The patient was treated using ointment and is expected to make a full recovery. | The monopolar cautery instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The root cause of the surgical complication experienced by the patient is also currently undetermined. A follow-up mdr will be submitted if additional information is received or if the instrument is returned. As of september 13, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775260 | 2746240 | 2642836 | 2955842201200439 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-24 | 2012-09-14 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-08-24 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-08-27 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the site experienced loss of image in the right eye of the surgeon side console. With the assistance of an isi technical support engineer the surgical staff converted the vision to 2d; however, the doctor stated he did not want to troubleshoot or work in 2d. At this time the surgical staff made the decision to convert the planned procedure to traditional open surgical techniques to successfully complete the procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of september 13, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775296 | 2746241 | 2642837 | 2955842201200438 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-27 | 2012-09-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-08-27 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-08-22 | Other | Required Intervention | It was reported that during a davinci si hysterectomy procedure, a bowel laceration occurred; the laceration was repaired. | On (b)(6) 2012, (b)(6), the robotics coordinator at (b)(6) medical center. He indicated that a 'little nick/cut' was made on the patient's bowel towards the end of the procedure. The affected area was repaired with a stitch suture. The patient didn't suffer any other injuries related with this incident. He indicated that he did not recall receiving any system error codes during the procedure. The patient was positioned in the dorsal lithotomy position. An intuitive surgical field service engineer (fse) had been contacted to visit the account to inspect the davinci system. A follow-up mdr will be submitted if a device is returned (post engineering evaluation) or if additional information is received. As of the date of this report, the customer has continued using the davinci si robot system with no issues. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775304 | 2746251 | 2642847 | 2955842201200442 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-22 | 2012-09-14 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2012-08-22 | No | No | Reuse | 2012
420093 | 2012-08-28 | Other | Other | It was reported that during a davinci si sacrocolpopexy procedure, the prograsp forceps instrument joint caught on the patient's bowel when it was used as a retractor. Patient is suspected of having incurred a bowel injury. | Per the isi clinical sales representative, when the surgeon was attempting to use the wrist of the progrash forceps instrument to retract the bowel, the metal part of the prograsp forceps instrument caught on the patient's bowel and surgeon believed that the bowel could have been damaged. No repair was necessary. There were no patient complications; the patient is reported stable and recovering as planned. The prograsp forceps instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2775262 | 2746252 | 2642848 | 2955842201200437 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-28 | 2012-09-14 | Yes | No | Health Professional | 420093 | No | No | 2012-08-28 | No | No | Unkown | 2012
IS2000 A51P8 | 2012-08-28 | Death | Death | It was reported that during a da vinci s prostatectomy procedure the surgeon experienced visualization problems due to the patient's anatomy on (b)(6) 2008. After several hours the surgeon decided to convert the procedure to traditional open surgical techniques to complete the prostatectomy. During the open portion of the prostatectomy procedure a rectal laceration was discovered, and a specialist was called in to perform an emergency repair. It was reported that blood loss during the procedure was estimated at 3500cc, fluid replacement was 7500ml crystalloid, four units of packed red cells, and one unit of autogous blood. It was reported that on (b)(6) 2008 the patient was transported by ambulance to an intensive care unit. On (b)(6) 2008 the patient was transferred to a rehabilitation center. On (b)(6) 2012 isi became aware that the patient expired from cardiac arrest on (b)(6) 2012. | The surgeon reported that no system malfunction occurred during the procedure and that the procedure was prolonged and converted due to the patient's anatomy. The initial autopsy report states that the patient had an enlarged heart, and death was likely due to cardiac arrest. Based on the information provided by the surgeon, it was determined that the da vinci s surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the patient's demise. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775306 | 2746253 | 2642849 | 2955842201200443 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-28 | 2012-09-14 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-08-28 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-07-29 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the patient's bladder was punctured with a needle. The affected area was repaired and the patient is reported to have recovered ok. Reportedly, damage to the patient's bladder was unrelated to the da vinci si system, instrument and/or accessories. | On (b)(4) 2012, isi contacted (b)(4), the risk manager at (b)(6) hospital and he indicated that approximately 4 or 5 days post a da vinci si hysterectomy procedure, the patient met with her primary care physician (pcp) and during her examination, the patient indicated to the pcp that she was experiencing abdominal pain and a catheter was placed. (b)(6) indicated that later that day the patient was admitted into good (b)(6)'s emergency room due to experiencing abdominal pain and that an exploratory laparoscopic procedure revealed that the patient's bladder was damaged. (b)(6) indicated that the patient's bladder was repaired that same day using traditional laparoscopic techniques. (b)(6) indicated that it is the hospital's belief that damage to the patient's bladder occurred while the surgeon was suturing the patient's vaginal cuff. Jim indicated that during the surgical procedure, no malfunction of the da vinci system, instruments or accessories occurred during the surgical procedure, no intra operative complications were experienced by the patient and the planned surgical procedure was successfully completed. (b)(6) indicated that (b)(6) had their da vinci si system evaluated by a third party company; however, he indicated that the evaluation results are not available. (b)(6) indicated that it is unknown if a video recording of the planned surgical procedure is available and that he is unable to provide isi with copies of the patient's operative report due to hippa regulations. Isi explained to (b)(6) that hippa regulations are not applicable to adverse events; however, (b)(6) indicated that he did not feel comfortable providing this information and (b)(6) will respond to requests for additional information directly from the fda. On (b)(4) 2012, an isi field service engineer (fse) was present during evaluation of the site's da vinci surgical system by the third party company to provide general information concerning the functionality of the system. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital and the site has continued to perform surgical procedures using their da vinci si surgical system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775319 | 2746266 | 2642862 | 2955842201200444 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-23 | 2012-09-14 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-08-23 | Device Not Retu | No | Reuse | 2012
| 2000-10-31 | Injury | Required Intervention | On (b)(6) 2012, intuitive surgical received a patient story posted on the internet concerning their da vinci hysterectomy surgical procedure: the patient indicated that on (b)(6) 2000 she underwent a da vinci hysterectomy procedure at (b)(6) hospital (location unknown) and during the surgical procedure her bladder was nicked. The patient indicated that prior to the surgical procedure; she was advised to undergo a hysterectomy with cystocele and rectocele repair. The patient indicated that her doctor advised her that her bladder during the surgical procedure was not addressed because it had descended to a level 3, her cervix was not removed and the rectocele was repaired. The patient indicated that 2 days post op she was released from the hospital and that she did not recover from the surgical procedure. The patient indicated for 7 days after the da vinci surgical procedure, her bladder would harden and she felt as though she was experiencing end stage labor and had to seek a specialist for her bladder. The patient indicated that examination by the specialist found that her bladder exhibited extensive scarring, muscle damage and there was no evidence of rectocele repair. The patient indicated that she underwent a pap smear, performed by a physician's assistant (pa) and the pa could not find the patient's cervix and that the patient exhibited serious inflammation. The patient also indicated that she experienced discomfort during the pap examination. The patient indicated that she requires another surgery and is currently under full disability. Late submission of this medwatch report is due to a data entry oversight by the responsible customer service representative. | Intuitive surgical is currently investigating the reported event. A follow up medwatch report will be submitted when additional information has been provided. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2775326 | 2746273 | 2642869 | 2955842201200446 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-07-09 | 2012-09-14 | Yes | No | Health Professional |  | No | No | 2012-07-09 | No | No | Reuse | 2012
| 2012-03-21 | Malfunction |  | During davinci hysterectomy, monopolar curved scissor was being used. End of plastic tip cover accessory burned/melted and arcing of cautery occurred - spark hit uterus. Instrument removed, new tip cover accessory put on but physician still saw arcing occur up instrument. New monopolar instrument opened and used. One monopolar had 1 life left and one had 2 lives left. No harm to patient. |  | HOT SHEARS | SYSTEMSURGICALCOMPUTER CONTROL | 2802655 | 2773130 | 2669724 | 2773130 | 1 | NAY | User Facility |  |  | Initial | 2012-09-06 | 2012-09-14 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2012
420309 | 2012-09-10 | Malfunction |  | During a myomectomy, a wire snapped intraoperatively on the robotic mega suture cut needle driver. No pieces observed to release into abdomen. No pt harm. Three 'lives' remaining on device. |  | INTUITIVE DAVINCI | MEGA SUTURE CUT NEEDLE DRIVER | 2862095 | 2834127 | 2730683 | 2834127 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2012-09-11 | 2012-09-14 | No | Yes | Health Professional | 420309 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2012
IS2000 A51P8 | 2012-08-16 | Other | Other | It was reported that during set up of a da vinci s prostatectomy procedure, the site experienced an intermittent black line in the left eye of the surgeon's console. The site decided to proceed with the planned surgical procedure, however, after starting the surgical procedure, the left image froze, the black line re-appeared and the error message left video lost was displayed. The site re-started the system and reseated the vision cable to resolve the vision issue, however, the issue recurred. The surgeon made the decision to complete the planned surgical procedure using traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the site was associated with the console vision cable as review of the site's system log found that system error code 45311 occurred. The console vision cable connects the vision cart to the surgeon side console (ssc) and passes the video signals between the ssc and the vision cart. The system was repaired by replacing the console vision cable. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 45311 appear when the da vinci s safety system detects a loss of one or both surgeon's video inputs. Upon determining this condition, the safety system puts da vinci s in a non-recoverable safe state. The system was repaired by replacing the affected camera cable. As of september 7, 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2773383 | 2744370 | 2640966 | 2955842201200432 | 1 | NAY | Manufacturer | DistributorUser facility | Health Professional | Initial | 2012-08-20 | 2012-09-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-08-20 | Device Not Retu | No | Reuse | 2012
8MM | 2012-09-11 | Malfunction |  | While surgeon was suturing mesh with the da vinci instrument, the wire was noted to be frayed. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 2903372 | 2876966 | 2773502 | 2876966 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-09-12 | 2012-09-12 | No | Yes | Physician | 8MM | Yes | No Answer Provided |  |  |  |  | 2012
| 2012-08-31 | Malfunction |  | While using tenaculum forceps during a robotic procedure,one piece of metal came apart into the patient's abdomen. Surgeon was able to retreive the piece from the patient's abdomen without any trauma to the patient. |  | DAVINCI ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 2778979 | 2749653 | 2646248 | 2749653 | 1 | NAY | User Facility |  |  | Initial | 2012-09-10 | 2012-09-10 | No | Yes | UNKNOWN |  | Yes | No Answer Provided |  |  |  |  | 2012
| 2012-09-06 | Malfunction |  | As md was using the needle driver, she noticed the jaws would not open correctly. It was noted that the wires at the tip of the instrument were broken. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 2862311 | 2835140 | 2731696 | 2835140 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2012-09-10 | 2012-09-10 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 |  | Death | Death | It was reported that post a successful da vinci si hysterectomy procedure the patient expired. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. On (b)(6) 2012, isi contacted the clinical sales representative (csr) and the csr indicated that he spoke to the surgeon and the surgeon indicated that the patient's demise was unrelated to the da vinci si system, instruments and/or accessories. The csr indicated that the surgeon stated that the patient was a large patient and during the patient's post surgical care the clinical staff had difficulty getting the patient to walk around. The surgeon indicated to the csr that multiple attempts by the surgical staff to get the patient to move around post operatively was difficult, as the patient reportedly preferred to relax instead of walking. The csr indicated that the patient may have developed an embolism, however, this has not been confirmed. Isi has contacted the surgeon several times concerning the reported event, however, no response has been received. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2762623 | 2733783 | 2630379 | 2955842201200422 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-20 | 2012-09-07 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-08-20 | Device Not Retu | No | Reuse | 2012
40027202 | 2012-08-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the harmonic ace curved shears instrument, the a piece from the harmonic ace curved shears insert accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the curved blade is broken at its midpoint and the curved blade cross section at the broken point has an uneven surface. The broken piece was also returned, and it matched the broken surface that remains on the insert. No additional pieces were missing. On (b)(4) 2012, isi contacted (b)(4) and she indicated that the planned surgical procedure was completed and the patient did not experience any complications. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2765132 | 2736198 | 2632794 | 2955842201200429 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-08-14 | 2012-09-07 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-08-14 | Device Not Retu | Yes | Unkown | 2012
42022703 | 2012-08-15 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, when the surgeon attempted to separate the patient's pedicles the jaws on the pk dissecting forceps instrument did not seem to meet and the instrument would not coagulate. Reportedly after coagulating the pedicles, the pedicles would bleed after being cut. Reportedly, significant charred tissue was observed on the jaws of the pk dissecting forceps instrument, which required cleaning numerous times. It is the surgeon's belief that the instrument did not seal well. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. It was additionally reported that prior to the start of the procedure the patient's hemoglobin level was 12. 5 gm/dl and at the end of the procedure, the patient's hemoglobin level was 10. 5 gm/dl. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. Engineering observed that the instrument was returned with the accumulation of dried biodebris on the yaw pulley at the grip base. However, there are no char marks or signs of arcing, as the biodebris was able to be removed and no localized melting of the yaw pulleys was observed. The conductor wires are not damaged and the instrument passed electrical continuity testing. The instrument also successfully passed a functional cautery test when using a wet towel as steam was created steam from the towel with activation of the cautery pedal. There was no physical damage to the instrument observed. On (b)(4) 2012 isi contacted (b)(6), at (b)(6) hospital and she indicated that inspection of the pk dissecting forceps instrument prior to use found that it was fine and that arcing from the instrument while in use not observed. (b)(6) indicated that the patient did not require a blood transfusion as a result of the reported event and that there is no video recording of the planned surgical procedure. No other information concerning the patient was provided. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2765136 | 2736200 | 2632796 | 2955842201200431 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2012-08-16 | 2012-09-07 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2012-08-15 | Device Not Retu | No | Reuse | 2012
IS2000 A51P7 | 2010-07-19 | Death | DeathRequired Intervention | On (b)(4) 2012 isi received a legal summons and complaint indicating on (b)(6) 2010, the patient underwent a da vinci s partial nephrectomy procedure at (b)(6) medical center. During the surgical procedure the patient sustained perforation of the duodenum. Between (b)(6) 2010, the patient developed an abscess, infection and sepsis. The patient deceased on (b)(6) 2010. | A review of the site's system log found that no system error codes were generated during the surgical procedure. At this time it is undetermined as to what caused damage to patient's duodenum. Requests for additional information have been made; however no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2765102 | 2736219 | 2632815 | 2955842201200425 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-08-13 | 2012-09-07 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2012-08-13 | No | No | Reuse | 2012
40018303 | 2012-06-28 | Malfunction |  | On (b)(6) 2012, medwatch uf/importer report 2201760000-2012-8010 was received: event desc: the hook broke into two pieces during surgery. The pieces were removed from the area and saved. Nothing was retained in the patient. There was not harm to the patient. What was the original intended procedure? robotic-associated laparoscopic right partial nephrectomy. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | Since the instrument was not be returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received or the instrument is returned. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2765109 | 2736221 | 2632817 | 2955842201200423 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-08-20 | 2012-09-07 | No | Yes | Health Professional | 40018303 | No | No | 2012-08-20 | No | No | Reuse | 2012
40027202 | 2012-08-17 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while using the harmonic ace instrument, the harmonic ace insert accessory broke and fell into the patient. The accessory fragment was retrieved and the surgery was completed successfully. No patient harm, injury or adverse outcome was reported. | The harmonic ace insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(4) 2012 there have been no reported recurrences of the issue at the hospital. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2765104 | 2736223 | 2632819 | 2955842201200426 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-17 | 2012-09-07 | No | Yes | Health Professional | 40027202 | No | No | 2012-08-17 | No | Yes | Unkown | 2012
IS2000 A51P8 | 2012-08-16 | Other | Other | It was reported that during a davinci s surgical procedure, the operating room staff experienced 'purple tint in the right eye' on the surgeon side console. The site experienced system error code 45312. The account informed the isi technical support engineering that they had a bad camera cable and could not locate their backup camera cable. At this time, the surgeon made the decision to convert the planned surgical procedure to a traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 45312 appear when the da vinci s safety system detects a loss of one or both surgeon's video inputs. Upon determining this condition, the safety system puts da vinci s in a non-recoverable safe state. The system was repaired by replacing the affected camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2765107 | 2736225 | 2632821 | 2955842201200424 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-16 | 2012-09-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-08-16 | Device Not Retu | No | Reuse | 2012
42018405 | 2012-08-27 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the permanent cautery spatula instrument was bent at 90 degree angle and was required to be unbent to remove the instrument from the patient. Upon removal of the instrument, a crack at the yellow plastic collar was noticed. It was determined that a piece of the permanent cautery spatula instrument had broken and fell into the patient. The fragment was retrieved, and the planned surgical procedure was completed successfully. No patient harm was reported. | The harmonic ace insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(4) 2012, there have been no reported recurrences of the issue at the hospital. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2765129 | 2736228 | 2632824 | 2955842201200427 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-27 | 2012-09-07 | No | Yes | Health Professional | 42018405 | No | Yes | 2012-08-27 | No | No | Invalid Data | 2012
42019408 | 2012-08-13 | Malfunction |  | It was reported to intuitive surgical that during a da vinci si hysterectomy procedure, the mega needle driver instrument tip insert fell into the patient. The insert was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. During failure analysis the instrument was found with carbide insert missing from one of the jaw grips. The grip tip surface was found to be clean with no adhesive residue. No other physical damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2765131 | 2736230 | 2632826 | 2955842201200428 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-13 | 2012-09-07 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-08-13 | Device Not Retu | No | Reuse | 2012
42029602 | 2012-08-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the large suturecut needle driver instrument would not close and an item fell into the patient. The item was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the blade edges have circular shaped dents along the edge. Engineering concluded that the damage is likely due to mishandling and misuse. The instrument was found to be intact and there were no missing components at the jaw grip. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Isi has made multiple attempts to contact the customer to verify additional information concerning the reported event. A follow-up mdr will be submitted if additional information is received. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2765134 | 2736232 | 2632828 | 2955842201200430 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-08-14 | 2012-09-07 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2012-08-14 | Device Not Retu | Yes | Reuse | 2012
420023 |  | Other | Other | It was reported that prior to starting a da vinci si surgical procedure, when the 8 mm blunt bladeless obturator was inserted into the patient during creation of a port incision, the surface of the patient's liver was scratched. The affected area was repaired using cautery and a little suture. It was reported that it is the surgeon's belief that damage to the patient's liver was a result of the length of the 8 mm blunt bladeless obturator when used with the regular cannula. It is unclear as to which cannula was used. The planned surgical procedure was completed and no post-surgical complications were experienced by the patient. | The obturator has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The instruments and accessories user manual specifically states: precautions and warnings - to minimize the risks associated with port placement, ensure the following: appropriate patient positioning to shift organs away from the port placement site. An adequate level of insufflation. Obturator tip is pointing away from major vessels, organs, and other anatomic structures. When possible, visualization of the entire insertion of the cannula using the endoscope is preferred. Moderate, controlled pressure is employed when placing the cannula and obturator. Potential complications associated with the use of the cannulae, obturators and reducers are the same as those associated with the use of surgical trocars and laparoscopic surgery in general. These potential complications include, but are not limited to, superficial lesions, injury to internal vessels, bleeding, hematoma, injury to the abdominal wall, infection, and peritonitis. | 8 MM BLUNT BLADELESS OBTURATOR | ENDOSCOPIC ACCESSORY | 2765155 | 2736251 | 2632847 | 2955842201200433 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-15 | 2012-09-07 | Yes | No | Health Professional | 420023 | No | No | 2012-08-15 | No | Yes | Unkown | 2012
IS3000 A60P7 | 2012-06-15 | Injury | Other | It was reported that post a sleep apnea surgical procedure, using the da vinci si surgical system and traditional laparoscopic techniques, it was noted that the pt sustained a burn on the lip and tongue. The isi clinical sales rep present during the surgical procedure indicated that tongue dissection using the da vinci surgical system went well, the system was undocked and the surgeon proceeded with the surgical procedure using traditional laparoscopic techniques; however, approx 10 minutes into the surgical procedure, while the surgeon was utilizing traditional laparoscopic instrumentation the isi csr noted that the pt's lip and tongue exhibited discoloration. At the end of the surgical procedure, the site determined that the discolorations observed by the csr were burns. The csr present during the surgical procedure inspected the da vinci endowrist cautery instruments, and found that the main tube exhibited chipping; however, the instruments did not exhibit damage that would have caused the burn injury sustained by the pt. | The da vinci surgical system is not approved for use for sleep apnea surgical procedures and the surgeon is aware that it is not an approved indication. However, the surgeon's preference was to use their da vinci si system to perform the tongue base resection surgical task, as the surgeon believes that the pt outcomes are better. On (b)(4) 2012, the isi clinical sales rep spoke to the surgeon who performed the surgical procedure, and he indicated that the hospital's investigation into the reported event did not determine the root cause of the burns sustained by the pt and that arcing during the surgical procedure was not noted. The surgeon indicated that the pt recovered well and did not experience any post-surgical complications as a result of the reported event. On (b)(4) 2012 isi contacted the da vinci coordinator (b)(6) at (b)(6) hospital indicated that the (b)(6) esu used during the surgical procedure was inspected and no issues were found and that the pt was grounded properly, however, (b)(6) indicated that she does not know the location of the grounding pad as she was not present during the surgical procedure. (b)(6) indicated that the surgeon indicated to her that no add'l surgical procedure was required to treat the pt and to the best of her knowledge the pt has not returned to the hosp due to any post-surgical complications. On (b)(4) 2012, isi contacted (b)(6) in the clinical quality risk mgmt dept at (b)(6) hosp and she indicated no pieces from the instrument fell into the pt and that the da vinci cautery instruments are not available for return to isi for eval as the hosp has made the decision to retain the instruments for their records. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hosp. | EDA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2765958 | 2733835 | 2630431 | 2955842201200421 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-15 | 2012-09-05 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-06-15 | Device Not Retu | No | Reuse | 2012
40018012 | 2012-08-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing from the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument was observed. The mcs instrument was removed and inspection of the tip cover accessory by the site found that it had a hole. The planned surgical procedure was completed with a replacement tip cover and no pt harm, adverse outcome or injury was reported. | Conclusion - the tip cover accessory and monopolar curved scissors (mcs) instrument were returned. Per the customer reported complaint, engineering observed that the hole in the silicone/pellethane on the tip cover exhibits localized melting indicating that an arcing event occurred. With the tip cover installed on returned mcs instrument, the position of the hole is in the general area of the distal idler pulleys and distal clevis ears on the instrument. The tip cover has a. 020" x. 030" oval shaped hole in the silicone/pellethane section and is located. 345" above the silicone/pellethane to sleeve interface. The tip cover exhibits some small cuts/scratches in the silicone/pellethane section, located away from the hole. No other damage was found. Engineering eval of the mcs instrument found that it did not exhibit any defect that likely or directly contributed to arcing through the tip cover accessory. As the distal end components on the instrument do not exhibit any burrs or abnormally sharp edges. On (b)(4) 2012, isi contacted the da vinci coordinator, (b)(6), and he indicated that the mcs instrument, cannula and tip cover were inspected prior to use and there was no damage observed. Larry indicated that the electrical surgical unit (esu) setting was 40 watts, however, he indicated that he does not recall the mode used during the surgical procedure. Larry indicated that there was no injury to the pt and the pt has been released from the hospital. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2807365 | 2742269 | 2638865 | 2955842201200420 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2012-08-10 | 2012-09-05 | No | Yes | Health Professional | 40018012 | Yes | Yes | 2012-08-10 | Yes | Yes | Unkown | 2012
42009308 | 2012-08-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a piece broke off of the prograsp forceps instrument and fell into the patient. The broken piece was not retrieved. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the axis pin missing from the distal end. Grips come assembled from supplier with the pin installed. Grip axis pin is the pivot point for the grips, so grips cannot open/close properly with this part missing. One grip has a small gouge on the outer surface, adjacent to the hole for the grip axis pin. Additional observation not reported by site is tube damage. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2753396 | 2725053 | 2621649 | 2955842201200415 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-08-14 | 2012-09-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-08-14 | Device Not Retu | No | Reuse | 2012
420179 | 2012-08-06 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the site experienced wristing breaking of the monopolar curved scissors instrument. The instrument was removed and replaced with another of the same type; however the wrist breaks continued to occur. A total of 4 instruments were broken before the customer no longer had anymore monopolar curved scissors. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. The investigation conducted by field service engineering concluded that it is likely that the surgeon was putting too much force on the instrument tip. All instruments failed at the same specific location (distal tip where the clevis meets the fiberglass shaft, directly below the flush port window). The surgeon was likely unaware of how much pressure was applied at the distal tip. Isi has made several attempts to verify if the reported instrument is available for return to isi for evaluation, however, no information has been provided. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2753397 | 2725055 | 2615177 | 2955842201200411 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-06 | 2012-09-04 | Yes | Yes | Health Professional | 420179 | No | Yes | 2012-08-06 | No | No | Unkown | 2012
42018405 | 2012-03-28 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a fragment from the permanent cautery spatula instrument broke and fell into the patient. The instrument fragment could not be located after searching and was not retrieved. No further information has been provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one distal clevis ear was broken off at its base. The clevis ear on the conductor wire side is broken off, resulting in dislodgment of the yaw pulley. The broken piece was no return, but is approximately 0. 285 x 0. 220 in size. Additional observation not reported is the missing conductor cap. Conductor cap likely detached as a result of clevis ear breakage. Yaw pulley boss feature that mates with cap is broken off. Evidence not conclusive, but breakage is likely due to overloading the tip. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2753398 | 2725056 | 2449038 | 2955842201200413 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-08 | 2012-09-04 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2012-08-08 | Device Not Retu | No | Reuse | 2012
42018405 | 2012-03-28 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a fragment from the permanent cautery spatula instrument broke and fell into the patient's abdomen. The instrument fragment was retrieved, no patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one distal clevis ear was broken off at its base. The clevis ear on the conductor wire side is broken off, resulting in dislodgment of the yaw pulley. The broken piece was no return, but is approximately 0. 285 x 0. 220 in size. Additional observation not reported by the site are a broken ceramic sleeve is missing a 0. 070 x 0. 075 piece. The sleeve exhibits light scratch marks below broken section. Conductor cap is missing from the distal end, exposing wire segment. Cap is likely detached as a result of clevis ear breakage. Yaw pulley boss feature that mates with the cap is broken off. The evidence is not conclusive, but breakage is likely due to overloading the tip. The instruments and accessories user manual specifically states: general precautions and warnings to handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2753401 | 2725065 | 2449038 | 2955842201200412 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-08 | 2012-09-04 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2012-08-08 | Device Not Retu | No | Reuse | 2012
420179 | 2012-08-06 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the site experienced wristing breaking of the monopolar curved scissors instrument. The instrument was removed and replaced with another of the same type; however the wrist breaks continued to occur. A total of 4 instruments were broken before the customer no longer had anymore monopolar curved scissors. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. The investigation conducted by field service engineering concluded that it is likely that the surgeon was putting too much force on the instrument tip. All instruments failed at the same specific location (distal tip where the clevis meets the fiberglass shaft, directly below the flush port window). The surgeon was likely unaware of how much pressure was applied at the distal tip. Isi has made several attempts to verify if the reported instrument is available for return to isi for evaluation, however, no information has been provided. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2753761 | 2725344 | 2615177 | 2955842201200409 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-06 | 2012-09-04 | Yes | Yes | Health Professional | 420179 | No | Yes | 2012-08-06 | No | No | Unkown | 2012
420179 | 2012-08-06 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the site experienced wristing breaking of the monopolar curved scissors instrument. The instrument was removed and replaced with another of the same type; however the wrist breaks continued to occur. A total of 4 instruments were broken before the customer no longer had anymore monopolar curved scissors. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. The investigation conducted by field service engineering concluded that it is likely that the surgeon was putting too much force on the instrument tip. All instruments failed at the same specific location (distal tip where the clevis meets the fiberglass shaft, directly below the flush port window). The surgeon was likely unaware of how much pressure was applied at the distal tip. Isi has made several attempts to verify if the reported instrument is available for return to isi for evaluation, however, no information has been provided. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2753760 | 2725360 | 2615177 | 2955842201200410 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-06 | 2012-09-04 | Yes | Yes | Health Professional | 420179 | No | Yes | 2012-08-06 | No | No | Unkown | 2012
870269 | 2012-08-15 | Malfunction |  | A (b)(6), woman underwent robotic assisted laparoscopic hysterectomy for endometrial cancer. Robot was being docked when camera clutch malfunctioned - the camera arm holder broke and the clutch stuck. "intuitive surgical" representative was called and ecm was replaced. Broken assy, ecm, extended, part 380189-02, serial number (b)(4) was replaced with assy, ecm, extended part 380189-02, serial number (b)(4). |  | DA VINCI SYSTEM | ROBOT | 2770223 | 2738884 | 2635479 | MW5026820 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-09-04 | 2012-09-04 | No | Yes | Health Professional | 870269 | Yes | Yes |  |  |  |  | 2012
870269 | 2012-08-15 | Malfunction |  | A (b)(6), woman underwent robotic assisted laparoscopic hysterectomy for endometrial cancer. Robot was being docked when camera clutch malfunctioned - the camera arm holder broke and the clutch stuck. "intuitive surgical" representative was called and ecm was replaced. Broken assy, ecm, extended, part 380189-02, serial number (b)(4) was replaced with assy, ecm, extended part 380189-02, serial number (b)(4). |  | DA VINCI SYSTEM | ROBOT | 2770223 | 2738884 | 2635479 | MW5026820 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-09-04 | 2012-09-04 | No | Yes | Health Professional | 870269 | Yes | Yes |  |  |  |  | 2012
IS3000 A60P6 | 2012-08-07 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the site experienced system error code 10025, system error code 23025, and system error code 23008. With the assistance of an isi technical support engineer (tse)the site back drove the left master tool manipulator (mtml) with the system powered on and powered off, however, the system error codes recurred. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | ~ the investigation conducted by the field service engineer concluded that system error code 10025, system error code 23025, and system error code 23008 experienced by the site were associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console (ssc). One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 10025 appears when the da vinci safety system detects a failure to enable amps. System error code 23008 appears when the da vinci si safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. System error code 23025 occurs when the remote arm controller hardware detects an encoder quadrature fault. Upon determining these conditions, the system records the faults and transitions to a safe state. The mtml was returned to isi for failure analysis. Engineering was able to replicate system error code 23025 experienced by the site. Functional testing of the mtml found that the axis 2 encoder malfunctioned, thus generating the system faults experienced by the customer. The axis 2 was replaced and retesting of the mtml found that it functioned within specification. On (b)(4) 2012, isi contacted rn, (b)(6) at (b)(6) medical center in (b)(6), and she indicated that the patient did not experience any surgical complications as a result of the reported event and the patient was released from the hospital. As of august 30, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2752234 | 2723698 | 2620287 | 2955842201200417 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-07 | 2012-09-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-08-07 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-08-09 | Other | Other | It was reported to intuitive surgical that during the davinci s hysterectomy procedure, the surgical staff reported 'red tint in one eye' on the surgeon side console. The account informed isi technical support engineering they had a bad camera cable and could not locate their backup camera cable. At this time, the surgeon made the decision to convert the planned surgical procedure to a traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The camera cable was returned to intuitive surgical for evaluation. Engineering was able to replicate the reported failure mode and observed that the conductors were damage. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2752291 | 2723760 | 2620349 | 2955842201200416 | 1 | NAY | Manufacturer | User facility | Health Professional | Initial | 2012-08-09 | 2012-09-01 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-08-09 | Device Not Retu | No | Reuse | 2012
42011704 | 2012-08-13 | Malfunction |  | It was reported that during a da vinci si ureteral re-implant procedure the tip of the 5 mm needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one needle driver grip broken off at the grip base. Engineering concluded that damage was likely due to excess force applied to the grip. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2012, isi followed up with the isi clinical sales representative present during the surgical procedure and she indicated that there was nothing unusual noted with the instrument prior to breakage of the needle drive instrument, however, the instrument was uninstalled multiple times during the surgical procedure. The csr indicated that the broken instrument piece was removed through an assistant port and there was no patient harm, adverse outcome or injury. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2752357 | 2723824 | 2620413 | 2955842201200419 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-08-13 | 2012-09-01 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2012-08-13 | Device Not Retu | No | Reuse | 2012
428053 | 2012-08-14 | No Answer Provided |  | It was reported that during a da vinci si single-site cholecystectomy procedure, the wek hem-o-lok medium-large clip installed on the 5 mm medium-large clip applier instrument fell into the patient's abdominal cavity. The clip was not retrieved. No patient harm, adverse outcome or injury was reported. | On (b)(4) 2012, the intuitive surgical clinical sales representative (csr) present during the surgical procedure indicated that when the surgeon took control of the medium-large clip applier instrument, to apply the wek hem-o-lok clip to the patient's cystic duct, the surgeon inadvertently closed the grips on the medium-large clip applier instrument, causing the clip to fall into the patient. The site irrigated and suctioned the surgical area to retrieve the clip; however, the misfired clip is believed to be free floating. The csr indicated that inspection of the instrument by the surgical staff found no damage and the planned surgical procedure was completed with the same instrument. | 5 MM MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2752363 | 2723834 | 2620423 | 2955842201200418 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-08-14 | 2012-09-01 | No | Yes | Health Professional | 428053 | No | No | 2012-08-14 | No | No | Unkown | 2012
42022703 | 2012-07-30 | Other | Other | It was reported that during a da vinci si myomectomy procedure, a cable and a plastic piece from the pk dissecting forceps instrument broke off and fell into the patient. Only one piece was found and retrieved. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The instrument has not been received for evaluation, therefore, the root cause of the customer reported failure cannot be determined. Multiple attempts have been made to obtain additional information, however as of the date of this report no response has been provided by the hospital. Intuitive surgical's clinical sales representative (csr) was present during the procedure and indicated the cause of the instrument breakage was due to intraoperative cleaning of the instrument using another instrument. The instruments and accessories user manual specifically states: general precautions and warnings: do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2749472 | 2720122 | 2616716 | 2955842201200408 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-31 | 2012-08-29 | Yes | Yes | Health Professional | 42022703 | No | No | 2012-07-30 | No | No | Unkown | 2012
| 2012-08-27 | Malfunction |  | The #2 robotic arm instruments not working properly. The instrument was changed with no effect. The instrument arm plate/drape changed with return of function to instruments. |  | DAVINCI ENDOWRIST | ROBOTIC SURGICAL ACCESSORY | 2807973 | 2778584 | 2675180 | 2778584 | 1 | NAY | User Facility |  |  | Initial | 2012-08-29 | 2012-08-29 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
40018012 | 2012-08-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a breach was noticed on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument and a spark was observed at the gap breach. Additionally, the orange portion of the instrument was visible and it appeared as if the tip cover had slid down the instrument during the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2747109 | 2717926 | 2614521 | 2955842201200393 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-08-03 | 2012-08-28 | No | Yes | Health Professional | 40018012 | No | Yes | 2012-08-03 | No | Yes | Initial | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an anonymous online survey concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The information received is as follows: some complications after surgery, resulted in internal leaking and return to hospital to have drain and catheter installed for further three weeks. On the positive side i had very good bladder control within hours of catheter and drain removal. | Due to the limited information provided in the survey, intuitive surgical is unable to follow up with the reporter to determine the root cause of the reported event. If additional information is received concerning the reported event, a follow-up medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2747800 | 2718569 | 2615144 | 2955842201200404 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No | Yes | 2012-08-03 | No |  | Unkown | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an anonymous online survey concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The information received is as follows: contacted (b)(6) from procedure. | Due to the limited information provided in the survey, intuitive surgical is unable to follow up with the reporter to determine the root cause of the reported event. If additional information is received concerning the reported event, a follow-up medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2747801 | 2718570 | 2615145 | 2955842201200400 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No | Yes | 2012-08-03 | No |  | Unkown | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an anonymous online survey concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The information received is as follows: three reasons: the recovery time was much greater than expected. There were complications two months after the surgery (internal bleeding). The surgery was not successful - rising psa immediately after surgery. | Due to the limited information provided in the survey, intuitive surgical is unable to follow up with the reporter to determine the root cause of the reported event. If additional information is received concerning the reported event, a follow-up medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2747799 | 2718571 | 2615146 | 2955842201200403 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No | Yes | 2012-08-03 | No |  | Unkown | 2012
42018406 | 2012-07-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, arcing occurred with the permanent cautery spatula instrument. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the conductor wire has a section at the wrist that is dislodged from the conductor cap and exposed outside the yaw pulley. Wire insulation is damaged near the yaw pulley exit, exposing bare wire. Boss feature of the yaw pulley is missing. Charring on the yaw pulley was evident. Yaw pulley exhibits charring/localized melting adjacent to the exposed bare wire. Bare wire likely created a path for unintended arcing to the yaw pulley. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2747812 | 2718574 | 2615149 | 2955842201200406 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-08-02 | 2012-08-28 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-08-02 | Device Not Retu | No | Reuse | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an anonymous online survey concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The information received is as follows: resulted in incontinence and erectile dis-function. | Due to the limited information provided in the survey, intuitive surgical is unable to follow up with the reporter to determine the root cause of the reported event. If additional information is received concerning the reported event, a follow-up medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2747810 | 2718580 | 2615176 | 2955842201200405 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No | Yes | 2012-08-03 | No |  | Unkown | 2012
420179 | 2012-08-06 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the site experienced wristing breaking of the monopolar curved scissors instrument. The instrument was removed and replaced with another of the same type; however the wrist breaks continued to occur. A total of 4 instruments were broken before the customer no longer had anymore monopolar curved scissors. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. The investigation conducted by field service engineering concluded that it is likely that the surgeon was putting too much force on the instrument tip. All instruments failed at the same specific location (distal tip where the clevis meets the fiberglass shaft, directly below the flush port window). The surgeon was likely unaware of how much pressure was applied at the distal tip. Isi has made several attempts to verify if the reported instrument is available for return to isi for evaluation, however, no information has been provided. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2747802 | 2718581 | 2615177 | 2955842201200407 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-08-06 | 2012-08-28 | Yes | Yes | Health Professional | 420179 | No | Yes | 2012-08-06 | No | No | Unkown | 2012
42018306 | 2012-07-20 | Malfunction |  | On (b)(6) 2012, it was initially reported that during a da vinci si hysterectomy procedure, the tip on the permanent cautery hook (pch) instrument was observed to be broken. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The instrument was returned and during failure analysis evaluation, engineering observed pieces missing from the instrument. The hospital was informed of the missing pieces and on (b)(6) 2012, the hospital indicated that a fragment from the instrument fell into the patient. The fragment was removed, and there were no remaining fragments left in the patient. The hospital also mentioned that no post-operative complications were experienced by the patient and there is no video recording available of the planned surgical procedure. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the hook is broken off at the distal end. The broken piece was not returned. The hook shank broke off within the yaw pulley, at the cross section of the drilled hole. Engineering evaluation also found that the entire ceramic sleeve is missing from the yaw pulley. Engineering concluded that the sleeve likely detached during/after the hook broke and that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2764789 | 2719590 | 2616184 | 2955842201200395 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-08 | 2012-08-28 | No | Yes | Health Professional | 42018306 | Yes | Yes | 2012-08-08 | Yes | No | Initial | 2012
42022703 | 2012-08-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure it was observed that the unable on the pk dissecting forceps instrument broke. After completion of the planned surgical procedure and upon inspection of the instrument by the surgical staff it was noted that a small piece of plastic on the tip of the pk dissecting forceps instrument was missing. The surgical staff cannot be certain if the missing piece fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2762335 | 2719625 | 2616221 | 2955842201200398 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2012-08-07 | 2012-08-28 | No | Yes | Health Professional | 42022703 | No | Yes | 2012-08-07 | Device Not Retu | No | Unkown | 2012
42022703 | 2012-07-31 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the pk dissecting forceps instrument was observed to be broken at the wrist and a fragment from the instrument was observed to be lying on the patient's tissue. The surgeon attempted to retrieve the fragment, however, it was lost in the patient's lower pelvic area. The intuitive surgical clinical sales representative (csr) present during the surgical procedure indicated that the procedure was a difficult case due to a significant number of dense and course fibroids the patient had. The csr also noted that there was significant torque applied on the pk dissecting forceps instrument by the surgeon. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Conclusion: the instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2012, intuitive surgical received an e-mail response from the (b)(6) at (b)(6) medical center and she indicated that she followed up with the surgeon who performed the surgical procedure. Per (b)(6), the surgeon indicated that when she attempted take hold of the instrument fragment, it slipped out into a pool of blood in the patient's cul de sac. The surgeon indicated to (b)(6) that the instrument fragment was the size of a spec and that it was removed using a suction irrigator. (b)(6) indicated that the planned surgical procedure was not video recorded and the pk dissecting forceps instruments was discarded. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2764733 | 2719637 | 2616233 | 2955842201200399 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-31 | 2012-08-28 | No | Yes | Health Professional | 42022703 | No | No | 2012-07-31 | Device Not Retu | No | Unkown | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an online survey from an unk source concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The info received is as follows: "the scars were more pronounced and did not heal as well as with regular laparoscopic surgery. In addition, the surgery caused a severe c-diff infection which forced me back to the hospital for three days. ". | Due to the limited info provided in the survey, intuitive surgical is unable to follow up with the reporter to determine the root cause of the reported event. If additional info is received concerning the reported event, a f/u medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2764047 | 2721072 | 2617668 | 2955842201200397 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No |  | 2012-08-03 | Device Not Retu |  | Unkown | 2012
IS3000 A60P7 | 2012-08-03 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced system error #22003 upon start up of the system. With the assistance of an isi technical support engineer, the site attempted to disable a patient side manipulator (psm) arm and performed a power cycle of the system, however, the system error code persisted. The patient was under anesthesia and the port incisions had been made when the surgeon decided to convert the planned procedure to traditional open techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code #22003 experienced by the customer was associated with a set up joint (suj) cable harness. The suj is an unpowered arm used to support and position the instrument arms for surgery. The setup-joint (suj) harness is an electrical cable which carries power and communication between the psm, the suj and the controlling processor. The system was repaired by replacing the affected suj cable harness. The system error code #22003 functioned as designed and there was no injury to the patient. The da vinci si safety system determined a position of one or more joints on the specified suj, as measured by that joint's primary sensor, disagreed with the measurement from secondary sensor by more than the specified tolerance. Upon determining these conditions, the safety systems put da vinci si in a recoverable safe state. As of august 24, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2762823 | 2721073 | 2617669 | 2955842201200402 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-03 | 2012-08-28 | Yes | Yes | Health Professional | IS3000 A60P7 | Yes | Yes | 2012-08-03 | Yes | No | Reuse | 2012
|  | Other | Other | On (b)(4) 2012, intuitive surgical received the following comments via an online survey from an unk source concerning the outcome of a surgical procedure that was performed using the da vinci surgical system. The info received is as follows: "experienced complications. Urinary stricture. Had to have 2nd surgery". | Due to the limited info provided in the surgery, intuitive surgical is unable to f/u with the reporter to determine the root cause of the reported event. If additional info is received concerning the reported event, a f/u medwatch report will be submitted to the fda. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2763875 | 2721087 | 2617683 | 2955842201200396 | 1 | GEI | Manufacturer | Other | Other | Initial | 2012-08-03 | 2012-08-28 | Yes | No | Health Professional |  | No |  | 2012-08-03 | Device Not Retu |  | Unkown | 2012
420205 | 2012-06-27 | Injury | Other | On date (b)(6) 2012, intuitive surgical received uf/importer report # (b)(4) from the fda. The event details are provided below: "during robotic hysterectomy, no one was present at console to activate power resource to activate pedal. However, upon advancing instruments there was smoke and a burn spot on the bipolar cautery instrument. There was also a burn on the pt's uterus. The robotic unit had instigated the electrical without anyone telling the panel to. The pt was having a hysterectomy so she was not harmed. ". | The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if add'l info is received. On (b)(4), 2012 intuitive surgical contacted (b)(6) the (b)(6) at (b)(6) hosp in (b)(6), and she indicated that during the da vinci si hysterectomy procedure the surgical staff inadvertently connected the cautery cable into the incorrect receptible on the electrical surgical unit (esu) being used during the surgical procedure. Ellen indicated that a connector was used in conjunction with the cautery cable. (b)(6) indicated that there was no malfunction of the da vinci system, instrument or accessories that caused the reported event. (b)(6) indicated that the planned surgical procedure was completed using the da vinci si surgical system and the pt has not returned to the hosp due to any post surgical complication as a result of the reported event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2789772 | 2758399 | 2654995 | 2955842201200401 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-08-01 | 2012-08-28 | Yes | No | Health Professional | 420205 | No | Yes | 2012-08-01 | Yes | No | Unkown | 2012
400180 | 2012-03-26 | Malfunction |  | Device involved was an intuitive surgical davinci monopolar tip cover insulator accessory. During robotic laparoscopic total hysterectomy, surgeon noted arcing coming from the robotic scissor tip cover from tip to the uterus; instrument removed and scissor tip cover replaced with new one; item returned to supervisor for evaluation by company. Procedure resumed with no further incident. Tip insulator part # 400180 that was used during the procedure was sent back to manufacturer for evaluation. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2793737 | 2764672 | 2661268 | 2764672 | 1 | NAY | User Facility |  |  | Initial | 2012-08-27 | 2012-08-27 | No | Yes | Physician | 400180 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-07-13 | Malfunction |  | During procedure, washer noted in surgical field via camera from da vinci; removed said washer and instrument, notified charge nurse and team, xray will be obtained. |  | DA VINCI S ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL BI | 2752158 | 2723053 | 2619649 | 2723053 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-23 | 2012-08-26 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-07-17 | Malfunction |  | While using the da vinci robot for a scheduled laparoscopic hysterectomy case, the surgeon was experiencing trouble with the laparoscopic scissors. After replacing the scissors three times, it was determined by the da vinci technician, that the scissors were not faulty. He recommended replacing the drape that incorporated a plastic adaptor plate between the instrument and the robotic arm controlling the instrument. When the drape was replaced, the instrument performed as designed. |  | DA VINCI ACCESSORIES | SYSTEM ROBOTIC SURGICAL INSTRU | 2780256 | 2751127 | 2647722 | 2751127 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-26 | 2012-08-26 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2012
|  | Malfunction |  | It was reported that during a da vinci si surgical procedure, clips fell into the patient when using the clip applier instrument. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | 8 MM ENDOWRIST CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2743979 | 2714368 | 2610964 | 2955842201200392 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-25 | 2012-08-24 | No | Yes | Health Professional |  | No | No | 2012-07-25 | No | No | Unkown | 2012
|  | Other | Other | On (b)(6) 2012, intuitive surgical received a product liability complaint where a pt is alleging that her da vinci hysterectomy procedure resulted in damage to her ureter and that she sustained injury to her vaginal cuff and bowel during the surgical procedure. The pt also alleges that she continues to suffer from abdominal pain, pelvic pain, dyspareunia, bloating, abdominal distention, fatigue, and decreased energy and stamina. | Despite multiple requests, no additional info has been received from the initial reporter. As a result, it cannot be determined as to whether or not the pt's alleged injuries were a result of the da vinci surgical system. A follow-up medwatch report will be submitted if additional info is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2755785 | 2717723 | 2614318 | 2955842201200394 | 1 | NAY | Manufacturer | Other | ATTORNEY | Initial | 2012-07-27 | 2012-08-24 | Yes | Yes | Health Professional |  | No | No | 2012-07-27 | Device Not Retu | No | Reuse | 2012
420207 | 2012-06-07 | Malfunction |  | During the operative procedure some of the robotic instruments, the grasper and tenaculum, had rubbed together inside the patient's abdominal cavity and had worn thin. Portions of the tenaculum were frayed off. Upon discovery of the instruments' damage the patient cavity was methodically searched and a copious amount of irrigation was utilized. The manufacturing company has been consulted to assist with in determining the causation of the instruments failing. Both the grasper and tenaculum had eight (8) uses. |  | DA VINCI ENDOWRIST TENACULUM FORCEPS | INSTRUMENT ROBOTIC SURGICAL TE | 2752161 | 2723133 | 2619729 | 2723133 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-24 | 2012-08-24 | No | Yes | Physician | 420207 | Yes | No Answer Provided |  |  |  |  | 2012
420207 | 2012-06-07 | Malfunction |  | During the operative procedure some of the robotic instruments, the grasper and tenaculum, had rubbed together inside the patient's abdominal cavity and had worn thin. Portions of the tenaculum were frayed off. Upon discovery of the instruments' damage the patient cavity was methodically searched and a copious amount of irrigation was utilized. The manufacturing company has been consulted to assist with in determining the causation of the instruments failing. Both the grasper and tenaculum had eight (8) uses. |  | DA VINCI ENDOWRIST TENACULUM FORCEPS | INSTRUMENT ROBOTIC SURGICAL TE | 2752161 | 2723133 | 2619729 | 2723133 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-24 | 2012-08-24 | No | Yes | Physician | 420207 | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 | 2012-07-06 | Other | Other | It was reported that after a da vinci si pulmonary wedge resection procedure was completed, a burn was observed at one of the port sites when the cannula accessory was removed from the instrument arm. The burnt skin was excised, however, the patient was burnt all the way through the tissue. | Despite multiple requests, no additional information has been received from the hospital regarding this event. Intuitive surgical investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the instrument arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded or if the grounding pad used by the surgical staff is faulty. As of (b)(4) 2012, the site has continued to use their da vinci si surgical system and there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2741723 | 2712015 | 2608611 | 2955842201200391 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-24 | 2012-08-23 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2012-07-24 | No | No | Reuse | 2012
40016907 | 2012-07-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the harmonic curved shears (hcs) insert accessory that was being used with hcs instrument broke into two parts. A broken piece fell into the patient and was retrieved. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, it was observed that the clamp arm assembly had detached from the distal end. The shaft features that mate with the clamp arm pins are bent backwards slightly, indicating hyperextension of the clamp arm. It was determined that the damage is likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2739941 | 2710078 | 2606674 | 2955842201200388 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-07-25 | 2012-08-22 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2012-07-25 | Device Not Retu | Yes | Initial | 2012
42017207 | 2012-07-13 | Malfunction |  | It was reported that during a da vinci si pulmonary right lower wedge resection procedure, the washer came off of the maryland bipolar forceps instrument and fell into the patient. The washer was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pulley had come off of the distal clevis. The detached piece was returned with the instrument. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2740206 | 2710261 | 2606857 | 2955842201200389 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-26 | 2012-08-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-07-26 | Device Not Retu | No | Reuse | 2012
42030901 | 2012-06-01 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure, the megasuturecut needle driver instrument broke and pieces from the instrument fell into the patient. The instrument fragments were removed and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported and the patient is doing well six weeks post-op. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2012, intuitive surgical contacted (b)(6) the clinical nurse specialist at (b)(6) medical center and she indicated that the patient is doing well and to the best of her knowledge the patient has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event. Per the information provided by the initial reporter, the broken instrument pieces were successfully retrieved from the patient. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2738652 | 2708330 | 2604926 | 2955842201200390 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2012-07-24 | 2012-08-21 | No | Yes | Health Professional | 42030901 | No | Yes | 2012-07-24 | No | No | Reuse | 2012
420227 VER 03 | 2012-08-06 | Malfunction | Other | The jaw of the pk forceps broke off during the procedure. The surgeon visualized the tip and removed it. The procedure went on without further incident. |  | INTUITIVE DA VINCI SURGICAL SERVICES | 8MM PK DISSECTING FORCEPS | 2745946 | 2715016 | 2611612 | MW5026634 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2012-08-21 | 2012-08-21 | No | Yes | Health Professional | 420227 VER 03 | Device Returned To Manufacturer | No |  |  |  |  | 2012
8MM | 2012-08-16 | Malfunction |  | During robotic case a small piece of foreign body was found loose in cavity. It was deemed to be a piece of the 'spatula' that was being used at the time. |  | DAVINCI PERMANENT CAUTERY SPATULA | SYSTEM SURGICAL COMPUTER CONTR | 2747015 | 2717876 | 2614472 | 2717876 | 1 | NAY | User Facility |  |  | Initial | 2012-08-21 | 2012-08-21 | No | Yes | Physician | 8MM | Yes | No Answer Provided |  |  |  |  | 2012
42030901 | 2012-07-30 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, fragments from the megasuturcut needle driver instrument broke off and fell into the patient. All of the fragments were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was broken at the distal idler with segments or strands of cable sticking out of the wrist. The segments obstruct the path of the grips from fully closing. No broken or loose fragments were returned from the customer. The idler pulley was not damaged. Per the information provided by the initial reporter, the broken instrument pieces were successfully retrieved from the patient. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2735821 | 2705485 | 2602081 | 2955842201200387 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-30 | 2012-08-20 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-07-30 | Device Not Retu | No | Reuse | 2012
IS3000 A60P5 | 2012-07-24 | Other | Other | It was reported that prior to starting a da vinci si gynecological procedure, the surgeon side console (ssc) would not power on. The site contacted isi for technical support engineering (tse) assistance and review of the site's system log by the tse found that the power issue experienced by the was related to the ssc power supply. The tse, instructed the site to cycle the power breaker on the ssc and the system powered on; however, after approximately 5 minutes, the issue recurred. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded and confirmed that the issue experienced by the customer was associated with the surgeon side console (ssc) power supply. The ssc power supply converts the ac voltage from the wall outlet to dc voltage for use by the system. The system was repaired by replacing the affected power supply. On (b)(4) 2012, isi followed up with (b)(6) the surgical technician and she indicated that during set up of the da vinci si surgical system, the surgeon side cart would not power on, however, they believed that the issue was related to a power issue within the or. (b)(6) indicated that they were able to restore power to the system and they continued to prep the patient for the planned surgical procedure. (b)(6) indicated that after the patient was placed under anesthesia and taken into the or, the power issue recurred. (b)(6) indicated that the planned surgical procedure has been rescheduled, there was no harm to the patient and the surgeon advised the patient's family that the procedure was rescheduled due to the system issue. As of august 10, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735372 | 2705141 | 2601737 | 2955842201200367 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-07-24 | 2012-08-18 | Yes | Yes | Health Professional | IS3000 A60P5 | Yes | No | 2012-07-24 | Device Not Retu | No | Reuse | 2012
IS3000 A60P7 | 2012-07-25 | Other | Other | It was reported that prior to starting a da vinci si hysterectomy procedure, the site experienced a system error code 252 while entering the surgeon's profile information using the touchpad on the surgeon side cart. With the assistance of an isi technical support engineer, the site power cycled the system and upon power-up of the system, system error code 30130 appeared. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No patient harm, injury or adverse event was reported. | The investigation conducted by the field service engineering concluded that the system error code 31030 was associated with the touchpad. The touchpad is located in the middle of the surgeon console armrest and provides the means for selecting various system functions. The system was repaired by replacing the affected touchpad. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 31030 appears when the da vinci si safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining this condition, the safety system puts da vinci si in a recoverable safe state. The touchpad was returned and evaluated. Engineering was able to replicate the issue experienced by the customer as the touchpad failed when installed on an in-house is3000 system for functional testing. On (b)(6), 2012, isi contacted the scrub technician (b)(4) and he indicated that the da vinci si surgical system was not docked to the patient when the touchpad issue occurred and that the surgeon performed the planned surgical procedure using traditional laparotomy techniques. (b)(4) indicated that the patient tolerated the laparotomy procedure well and that to the best of his knowledge, the patient did not experience any post-surgical complications. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINSI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735356 | 2705145 | 2601741 | 2955842201200386 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-25 | 2012-08-18 | Yes | Yes | Health Professional | IS3000 A60P7 | Device Returned To Manufacturer | No | 2012-07-25 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-07-25 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the site experienced multiple instances of system error code 23008. With the assistance of an isi technical support engineer, the site restarted the system, however, the system error code recurred. The tse instructed the site to perform an emergency power off and restart the system, however, the issue persisted. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. The patient tolerated the open procedure well, did not experience any post surgical complications and the patient has been released from the hospital. | The investigation conducted by the field service engineer concluded that the system error code 23008 experienced by the site was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The right master tool manipulator (mtmr) was returned to intuitive surgical for evaluation. Engineering was able to replicate the system error code 23008 experienced by the site during functional testing of the mtmr. A visual inspection found that an outer mechanical cable for axis 3 within the mtmr was broken, thus generating the system error code experienced by the site. The affected cable was replaced and retesting of the mtmr found that it functioned within specification. As of august 17, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735376 | 2705151 | 2601747 | 2955842201200385 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-07-25 | 2012-08-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-07-25 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-07-23 | Other | Other | It was reported that while starting a da vinci s surgical procedure, the site experienced a system error. With the assistance of an isi technical support engineer the site confirmed that the vision side cart (vsc) transformer had a blown fuse. The site attempted to replace the fuse, however the system issue continued to occur. The patient was under anesthesia with port incisions made, when the surgeon decided to abort the planned surgical procedure. | The investigation conducted by field service engineering concluded that system fault experienced by the customer was associated with the vision side cart (vsc) isolation transformer. The fuses on the vision side cart were replaced to repair the system. As of august 16, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735298 | 2705098 | 2601694 | 2955842201200380 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-07-23 | 2012-08-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-07-23 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-07-30 | Other | Other | Event Description: --/-- It was reported that during the beginning of a da vinci si surgical procedure, the site experienced error code 23025 on the endoscopic camera manipulator(ecm). With the assistance of an isi technical service engineer (tse), the site restarted the system; however this did not resolve the problem. No surgical tasks had been performed at this time; however, the patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for the following day. | Manufacturer Narrative: --/-- The investigation conducted by field service engineering found system error code 23025 was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23025 occurs when the signal from the primary sensor has made a suspect transition from one state to another at a rate higher than a specified threshold. The ecm is being returned to isi for failure analysis investigation. A follow-up report will be generated upon performing failure analysis of the ecm. The da vinci si surgical system user manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by field service engineering found system error code 23025 was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23035 is a pot-to-encoder comparison message used for diagnostics. Upon determining these conditions, the system records the faults and transitions to a safe state. The ecm was returned to intuitive surgical for failure analysis investigation. Engineering was unable to replicate the reported issue; however based on the error code determined that the axis-3 motor/encoder had malfunctioned. Axis-1 and axis-2 motor/encoder were also replaced as a precaution. The da vinci si surgical system user manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. Additional information received from the initial reporter indicated that this procedure was converted to open. The patient made a full recovery from the open procedure. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2735330 | 2705099 | 2601695 | 2955842201200384 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | InitialFollowup | 2012-07-30 | 2012-08-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-08-27 | Device Not Retu | No |  | 2012
42017910 | 2012-06-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure a piece of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. After the procedure the surgeon ordered a patient x-ray be taken, the x-ray did not show any foreign bodies. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension broken at the distal clevis joint, with a piece approximately 0. 26 x 0. 3 broken off the extension. Additionally failure analysis observed burns at grip tips and char marks. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2735327 | 2705109 | 2601705 | 2955842201200379 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-19 | 2012-08-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-07-19 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-07-24 | Other | Other | It was reported that while preparing to start a da vinci s surgical procedure, the site experienced a red image in the both eyes of the surgeon side console and touch screen display. With the assistance of an isi technical support engineer, the site restarted the system and reconnected all cables; however, this did not resolve the problem. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the site was associated the lamp module of the illuminator. The lamp module provides the light to the endoscope. The system was repaired by replacing the affected lamp module of the illuminator. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735297 | 2705110 | 2601706 | 2955842201200382 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-07-24 | 2012-08-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-07-24 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-07-27 | Other | Other | It was reported that during a da vinci si surgical procedure, the customer experienced a system error code 23. With the assistance of an isi technical support engineer the site applied emergency power off (epo), attempted to cycle the system circuit breakers, and reseat the fiber cables however the system error continued to occur. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No further information was provided. | The investigation conducted by field service engineering concluded that the system error code 23 experienced by the customer was associated with a setup-joint(suj). The setup-joint (suj) is an unpowered configurable arm which supports the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. As of august 16, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2735296 | 2705111 | 2601707 | 2955842201200383 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-07-27 | 2012-08-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-07-27 | Device Not Retu | No | Reuse | 2012
40027202 | 2012-07-23 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure on the 4th use the 5mm maryland dissector instrument, the instrument tip broke and fell inside the patient. An x-ray showed the location of the instrument tip. The instrument tip was retrieved laparoscopically and the procedure was completed successfully. No patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the curved blade broken at its midpoint. Evidence not conclusive but damage likely due to crack in the curved blade. No other physical damage was found. --/-- Additional Manufacturer Narrative: --/-- The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the device is returned or if additional information is received. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | Unkown | 2735300 | 2705114 | 2601710 | 2955842201200381 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2012-07-23 | 2012-08-17 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2012-08-13 | Device Not Retu | Yes |  | 2012
IS3000 A60P6 | 2012-07-17 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, there was no video displayed on the left side monitor of the surgeon console. An intuitive surgical technical support engineer attempted to troubleshoot the issue via telephone and had the surgical staff cycle system power, reset the vision cart and surgeon console breakers, and replace the camera and camera cable. These actions were unable to resolve the customer reported vision issue. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the left side vision issue was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The doco was returned to intuitive surgical for failure analysis investigation. During internal evaluation, engineering was able to replicate the reported failure mode. It was determined that the left ccu had malfunctioned on the doco. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2725135 | 2694495 | 2591091 | 2955842201200378 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-07-17 | 2012-08-13 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-07-17 | Device Not Retu | Yes | Reuse | 2012
40018012 | 2012-07-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument, fell into the patient. The tip cover was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The initial reporter also informed intuitive surgical that prior to use, electro lube was used to assist with the installation of the tip cover accessory onto the mcs instrument. | Despite multiple requests, the tip cover accessory has not been returned for evaluation. The instruments and accessories user manual provides instructions for applying electro lube after the tip cover accessory has been properly installed on the mcs instrument. The initial reporter informed intuitive surgical that the applicable staff members have been instructed to refrain from using electro lube when installing the tip cover accessory onto the mcs instrument. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2721414 | 2691033 | 2587630 | 2955842201200370 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-07-11 | 2012-08-10 | No | Yes | Health Professional | 40018012 | No | No | 2012-07-11 | No | Yes | Initial | 2012
400156 | 2012-07-09 | Malfunction |  | It was reported that during a da vinci s surgical procedure the 5mm cautery hook a-sure accessory installed on the 5mm monopolar cautery instrument, broke off and fell into the patient. The cautery hook was retrieved and the planned surgical procedure was completed. No patient injury, harm or adverse outcome was reported. | The accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The 5mm monopolar cautery instrument used in conjuction with the cautery hook a-sure accessory was returned to isi. The instrument was installed on an is2000 in-house test system and the instrument passed recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions and a new cautery tip accessory can be screwed onto the instrument indicating that the threading is within specification. The instrument also passed electrical continuity testing and was found to be fully functional. | 5MM CAUTERY HOOK ASURE ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2721585 | 2691204 | 2587800 | 2955842201200371 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-24 | 2012-08-10 | No | Yes | Health Professional | 400156 | No | No | 2012-07-24 | No | Yes | Unkown | 2012
IS3000 A60P6 | 2012-07-16 | Other |  | It was reported that during a da vinci si surgical procedure, the customer experienced a system error code 23. With the assistance of isi technical support engineering (tse) the site performed a power cycle and an emergency power off of the system, however, the issue recurred. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error code 23 experienced by the customer was associated with a patient side manipulator (psm) setup-joint harness. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The setup-joint (suj) harness is an electrical cable which carries power and communication between the psm, the suj and the controlling processor. The system was repaired by replacing the affected suj harness. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. On (b)(6) 2012, isi contacted da vinci coordinator, (b)(6) at (b)(6) hospital and she indicated that prior to starting a da vinci si ent procedure, when the surgical staff attempted to move the patient side cart (psc) towards the patient, a patient side manipulator arm would not moved. (b)(6) indicated that the patient had cancer; however, she indicated that she did not recall the specific procedure to be performed. (b)(6) indicated that she is unable to provide any additional information concerning the patient. As of (b)(6) 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2721608 | 2691207 | 2587803 | 2955842201200366 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-07-16 | 2012-08-10 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-07-16 | Device Not Retu | No | Reuse | 2012
42017908 | 2012-06-29 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the blade on the monopolar curved scissors instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned evaluated. Per the customer reported complaint engineering observed that the instrument was returned with the left blade detached from the instrument. The broken blade was also returned. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage found. On (b)(4) 2012, the returned affected instrument was reported to intuitive surgical indicating that the instrument was damaged during a da vinci s surgical procedure and that no pieces from the instrument fell into the patient. The event was not initially reported to the fda as the event as reported did not meet the code of federal regulation's guidelines of causing or contributing to a death, serious injury or likely causing or contributing to a death or serious injury if the malfunction were to recur. However, upon return and evaluation of the returned instrument, isi engineering found that the instrument was returned with one of the instrument blades broken off. On (b)(4) 2012, isi confirmed with the distributor representative that the broken instrument piece fell into the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2721609 | 2691208 | 2587804 | 2955842201200368 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-07-18 | 2012-08-10 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2012-07-18 | Device Not Retu | No | Reuse | 2012
420179 | 2012-06-13 | Malfunction |  | On date july 24, 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: "cautery machine made a noise while in use from 60 to 165 on cut mode. Rn turned machine off and turned it back on again. Machine went through self-test and rn was able to reset for approximately 15-20 minutes. Flame was noticed at connection point of cautery cord and laparoscopic scissors. Cautery cord and lap scissors were removed from field and replaced. No harm to pt. No other incident occurred. Cautery machine was removed from service and cautery cord and had a cut in it, causing a visual arc. Unit was tested to mfr specifications and output was within tolerance. Monopolar scissors given to material coordinator. Cord was also sequestered, but no numbers available. ". | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Isi has made several attempts to verify if the reported instrument is available for return to isi for evaluation, however, no information has been provided. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2721620 | 2691221 | 2527559 | 2955842201200369 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-07-24 | 2012-08-10 | No | Yes | Health Professional | 420179 | No | No | 2012-07-24 | No | No | Unkown | 2012
| 2012-07-11 | Other | Other | It was reported that during a da vinci si pulmonary wedge resection procedure the site experienced a bent jaw on a endowrist instrument. The site replaced this instrument with an instrument of the same type and experienced the same failure. The site had a 3rd instrument of the same type which was not sterilized. The site contacted a isi technical support engineer, to inquire about flash sterilizing this instrument; however the isi technical support engineer referenced a user manual warning to avoid flash sterilization of instruments. The site decided not to flash sterilize the instrument and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No other patient harm was reported. | The instruments involved in this case have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. | ENDOWRIST INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2721606 | 2691231 | 2587845 | 2955842201200374 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-07-11 | 2012-08-10 | Yes | Yes | Health Professional |  | No | Yes | 2012-07-11 | No | No | Invalid Data | 2012
| 2012-07-11 | Other | Other | It was reported that during a da vinci si pulmonary wedge resection procedure the site experienced a bent jaw on a endowrist instrument. The site replaced this instrument with an instrument of the same type and experienced the same failure. The site had a 3rd instrument of the same type which was not sterilized. The site contacted a isi technical support engineer, to inquire about flash sterilizing this instrument; however the isi technical support engineer referenced a user manual warning to avoid flash sterilization of instruments. The site decided not to flash sterilize the instrument and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No other patient harm was reported. | The instruments involved in this case have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. | ENDOWRIST INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2721628 | 2691249 | 2587845 | 2955842201200373 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-07-11 | 2012-08-10 | Yes | Yes | Health Professional |  | No | Yes | 2012-07-11 | No | No | Invalid Data | 2012
42014304 | 2012-07-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si thyroidectomy procedure the 5mm maryland dissector instrument made contact with the 5mm introducer, causing the instrument shaft to peel away. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the main tube insulation exhibited damage along the distal end. Insulation was found with several gouge and abrasive marks within the area. The marks were short in length and not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation and there are missing pieces of insulation. Engineering concluded that the damage was likely due to mishandling. On (b)(6) 2012 intuitive surgical received an email response from (b)(6), the surgical practitioner and robotics nurse at (b)(6) indicating that damage to the 5mm maryland dissector instrument was noted at the end of the planned surgical procedure and no fragment(s) from the instrument fell into the patient. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On august 2, 2012, isi received additional information from the site indicating that fragments from the shaft of the reported instrument fell into the patient. On august 8, 2012, isi contacted the clinical sales representative and he indicated that at the time of the initial report, there was no confirmation from the site that any fragments from the instrument fell into the patient, however, the site indicated that they were concerned about the potential of fragments falling into a patient. Isi contacted the customer on august 7, 2012 and august 8, 2012 to verify additional details concerning the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | 5MM MARYLAND DISSECTOR INSTRUMENT | Reuse | 2721875 | 2691467 | 2588063 | 2955842201200375 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-07-11 | 2012-08-10 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2012-08-15 | Device Not Retu | No |  | 2012
42014304 | 2012-07-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si thyroidectomy procedure the 5mm maryland dissector instrument made contact with the 5mm introducer, causing the instrument shaft to peel away. The planned surgical no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the black tube insulation is damaged at the distal end. The insulation is gouged on one side and has an. 080 x. 070 and. 040 x. 060 pieces missing approximately 4. 25 above the snake wrist. No other damage was found. On (b)(4) 2012 intuitive surgical received an email response from (b)(6), the surgical practitioner and robotics nurse at (b)(6) indicating that damage to the 5mm maryland dissector instrument was noted at the end of the planned surgical procedure and no fragment(s) from the instrument fell into the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2012, isi received additional information from the site indicating that fragments from the shaft of the reported instrument fell into the patient. On (b)(4) 2012, isi contacted the clinical sales representative and he indicated that at the time of the initial report, there was no confirmation from the site that any fragments from the instrument fell into the patient, however, the site indicated that they were concerned about the potential of fragments falling into a patient. Isi contacted the customer on (b)(4) 2012 to verify additional details concerning the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | 5MM MARYLAND DISSECTOR INSTRUMENT | Reuse | 2721887 | 2691480 | 2588076 | 2955842201200376 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-07-11 | 2012-08-10 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2012-08-15 | Device Not Retu | No |  | 2012
42014304 | 2012-07-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si thyroidectomy procedure the 5mm maryland dissector instrument made contact with the 5mm introducer, causing the instrument shaft to peel away. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. This report is being reported late due to an oversight in reporting by the responsible clinical sales representative. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the main tube insulation exhibited damage along the distal end. The insulation was found with several gouge marks within the area. Marks were short in length and not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation removed and pieces of tube insulation are missing. Engineering concluded that damage to the instrument was likely due to mishandling. On (b)(6) 2012 intuitive surgical received an email response from (b)(6), the surgical practitioner and robotics nurse at (b)(6) indicating that damage to the 5mm maryland dissector instrument was noted at the end of the planned surgical procedure and no fragment(s) from the instrument fell into the patient. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On august 2, 2012, isi received additional information from the site indicating that fragments from the shaft of the reported instrument fell into the patient. On august 8, 2012, isi contacted the clinical sales representative and he indicated that at the time of the initial report, there was no confirmation from the site that any fragments from the instrument fell into the patient, however, the site indicated that they were concerned about the potential of fragments falling into a patient. Isi contacted the customer on august 7, 2012 and august 8, 2012 to verify additional details concerning the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | 5MM MARYLAND DISSECTOR INSTRUMENT | Reuse | 2722013 | 2691611 | 2588207 | 2955842201200377 | 1 | NAY | Manufacturer | User facilityDistributor | Other | FollowupInitial | 2012-07-11 | 2012-08-10 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2012-08-15 | Device Not Retu | No |  | 2012
420177 | 2012-07-26 | Malfunction |  | Black coating is peeling off the instrument. These have been used less than 10 times. ======================manufacturer response for robotic bowel grasper davinci, bowel grasper robotic (per site reporter). ======================they would like the device to evaluate. |  | DAVINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL BO | 2734739 | 2704645 | 2601240 | 2704645 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-08 | 2012-08-08 | No | Yes | Physician | 420177 | Yes | No Answer Provided |  |  |  |  | 2012
42017910 | 2012-06-18 | Malfunction |  | On (b)(4) 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from university (b)(6) with the following information: during laparoscopic myomectomy, the robotic monopolar scissor broke off inside the patient. The instrument was immediately removed and disconnected from the robot, and the pieces retrieved from the patient. | The initial reporter indicated that the affected instrument will not be returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2716061 | 2685724 | 2582320 | 2955842201200361 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-07-09 | 2012-08-07 | No | Yes | Health Professional | 42017910 | No | No | 2012-07-09 | No | No | Unkown | 2012
400183 | 2012-06-28 | Malfunction |  | The hook broke into two pieces during surgery. The pieces were removed from the area and saved. Nothing was retained in the patient. There was no harm to the patient. |  | PERMANENT CAUTERY HOOK | SURGICAL INSTRUMENT COMPUTER C | 2718467 | 2688165 | 2584761 | 2688165 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-07 | 2012-08-07 | No | Yes | Invalid Data | 400183 | Yes | No Answer Provided |  |  |  |  | 2012
8MM | 2012-07-26 | Malfunction |  | As the surgeon was just completing a davinci total laparoscopic hysterectomy, the mega needle driver jaws failed while in use. |  | DAVINCI MEGA NEEDLE DRIVER | SURGICAL INSTRUMENT COMPUTER C | 2718665 | 2688568 | 2585164 | 2688568 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-07 | 2012-08-07 | No | Yes | Physician | 8MM | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-07-31 | Malfunction |  | Davinci monopolar scissor broke off intra-abdominally in patient. Instrument was unable to be detached from trocar. Trocar had to be removed from the patient to remove the instrument. New instrument applied and the procedure continued. There was no harm to the patient. |  | DAVINCI | SURGICAL INSTRUMENT COMPUTER C | 2718666 | 2688589 | 2585185 | 2688589 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-08-07 | 2012-08-07 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
8MM | 2012-08-02 | Malfunction |  | One 5mm robotic cannula was discovered bent and removed from case. One robotic maryland was discovered having shorted out and was removed from service. |  | DAVINCI MARYLAND BIPOLAR FORCEPS | SURGICAL INSTRUMENT COMPUTER C | 2723750 | 2693393 | 2589989 | 2693393 | 1 | NAY | User Facility |  |  | Initial | 2012-08-06 | 2012-08-06 | No | Yes | Physician | 8MM | Yes | No Answer Provided |  |  |  |  | 2012
42018306 | 2012-03-29 | Malfunction |  | On july 12, 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: the da vinci hook cautery was used during a robotic nissen fundoplication with a hernia repair and cholecystectomy surgery. The scrub tech noticed a piece of the instrument fell off in the patient. The piece was retrieved by the physician and given to the nurse along with the instrument. No harm was done to the patient. The instrument was sent down with the broken piece along with the lot number and other needed information. The resource nurse was notified. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the ceramic sleeve was broken at the base of the hook. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. On (b)(4) 2012, isi's clinical sales representative (csr) reported that during a da vinci si nissen fundoplication procedure a piece of plastic broke from the distal end of the permanent cautery hook (pch) instrument. On (b)(4) 2012, the csr reported that no piece from the reported affected instrument fell into the patient. Based on the additional information provided by isi csr, it was determined that the reported incident was not reportable to the fda, as the reported device malfunction did not meet the code of federal regulations' guidlines of causing or contributing to a death, serious injury or likely cause or contribute to a death or serious injury if the malfunction were to recur. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2712664 | 2681919 | 2578535 | 2955842201200365 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-07-12 | 2012-08-04 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-07-12 | Device Not Retu | No | Reuse | 2012
42018406 | 2012-07-09 | Malfunction |  | It was reported that during a da vinci si mediastinal tumor resection procedure, a piece of the permanent cautery spatula instrument broke and fell into the patient. The fragment was retrieved, and the planned surgical procedure was completed successfully. No patient harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the ceramic sleeve broken at the base of the spatula. Engineering concluded that damage to the instrument is likely due to direct impact on the sleeve. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2712463 | 2681741 | 2578357 | 2955842201200362 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-11 | 2012-08-03 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-07-11 | Device Not Retu | No | Reuse | 2012
| 2012-07-13 | Other | Other | It was reported that during a da vinci s nephrectomy procedure the cautery for the bipolar instrument arced causing damage to the patient anatomy. Anastomosis was used to repair the patient injury. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No other patient harm was reported. | On (b)(6), 2012 the initial reporter indicated the issue had been traced to user error, a bipolar instrument was plugged into a monopolar electrosurgical unit (esu). This bipolar instrument was used in conjunction with another monopolar instrument, when the two instruments were brought within proximity of each other energy was released. The initial reporter indicated the patient required anastomosis to repair the injury and the patient made a full recovery after the incident. As of august 2, 2012, there have been no reported recurrences of this issue at this hospital. | BIPOLAR ENDOWRIST INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2712499 | 2681770 | 2578386 | 2955842201200363 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-07-13 | 2012-08-03 | Yes | Yes | Health Professional |  | No | No | 2012-07-13 | No | No | Invalid Data | 2012
IS2000 A51P8 | 2012-07-06 | Other | Other | It was reported that during a da vinci s myomectomy procedure the surgical staff experienced system error code 603. The surgeon console was powered off and then back on, which cleared the error. The case was continued, however, system error code 603 returned. To resolve the issue, the surgeon console was again powered off, its breaker was reset, and the power cord was changed to another outlet. The system restarted successfully, however, system error code 603 recurred within 10 minutes of restart. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 603 was associated with the gemini user switch panel (gusp) printed circuit assembly board. The gusp is located on the surgeon console and is responsible for providing a mounting base for the leds, providing a mounting location for the 26 pin connector, and local power for board mounted leds. The system was repaired by replacing the affected gusp board. System error code 603 appears when the system detects that the ground for the panel has faulted. Upon determining this condition, the system records the fault and transitions to a safe state. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The gusp was returned and evaluated. The customer reported failure could not be replicated during internal testing, however, the switch power button was found to be sticky. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2712575 | 2681827 | 2578442 | 2955842201200364 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-07-06 | 2012-08-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-07-06 | Device Not Retu | No | Reuse | 2012
| 2007-09-28 | Injury | Hospitalization | Pt with history of pelvic pain had elective robotically assisted lap right oophorectomy/salpingectomy and appendectomy on (b)(6) 2007. On (b)(6) 2011, a diagnostic laparoscopy was conducted to remove the foreign body - an intuitive hot shears cover that was previously left behind. The cover had not been counted and apparently came off the hot shears when removed from trocar this was not noted by the surgical team. Facility notified on (b)(6) 2012. Retained plastic cover not thought to contributed to pt's pain or symptoms because it did not cause an inflammatory response in her body. |  | INTUITIVE HOT SHEARS COVER | HOT TIP COVER | 2715447 | 2682398 | 2579014 | MW5026371 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2012-07-31 | 2012-07-31 | Yes | No | Health Professional |  | No | Yes |  |  |  |  | 2012
40018012 | 2012-06-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a hole was detected on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. Upon detection of the hole the mcs tip cover accessory was removed and replaced with a new mcs tip cover accessory and the planned surgical procedure was completed successfully. The nurse did not observe any arcing from the instrument or accessory. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed damage on the side of the accessory. The mcs tip cover accessory was found with a hole approximately 0. 03 in diameter, with evidence of arcing found on the edge of the hole. Multiple cuts and tear damage was found near the hole, approximately 0. 06 in length, which is likely due to contact with other instruments. Multiple tears approximately 0. 11 in length was found at the distal end of the tip cover. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2703978 | 2673677 | 2570293 | 2955842201200357 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-28 | 2012-07-30 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-06-28 | Device Not Retu | Yes | Initial | 2012
40018012 | 2012-06-18 | Malfunction | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The damaged anatomy was stitched by a colorectal surgeon. No additional information was provided. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed holes in the tip cover accessory. The tip cover accessory exhibits localized melting around the holes, indicative of arcing. Holes are oblong shaped, roughly 0. 120 x 0. 020 and 0. 100 x 0. 040 and are centered from 0. 030 to 0. 115 below the combined pellethane/silicone section. Multiple holes indicate multiple arcing events. Position of holes correspond to proximal clevis section below ears where the tip cover is installed on the monopolar curved shears (mcs). | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2703994 | 2673688 | 2556252 | 2955842201200342 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-02 | 2012-07-30 | Yes | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-07-02 | Device Not Retu | Yes | Initial | 2012
UNKNOWN | 2012-07-06 | Injury | Required Intervention | It was reported that during a pediatric da vinci si colectomy and appendectomy procedures, both instruments could not be viewed at the same time. One instrument was not in the field of view, when the surgeon looked for this instrument it was discovered in the abdominal wall of the patient causing bleeding. The damage to the patient's abdominal wall was repaired using suture and the case was converted to laparoscopic surgery. The case was completed successfully and the patient made a full recovery. | Intuitive surgical's clinical sales representative (csr) was present during the procedure and indicated the reason the surgeon could not view both instruments during the procedure was because both the colectomy and appendectomy procedures were both performed using the same port placements. Per the csr, the surgeon did not want to place additional ports on the patient. Per the csr, the surgeon indicated he reviewed the video after the procedure and concluded this event was caused by user error resulting from not being able to see his instruments at all times. Isi's csr indicated that the surgery had no malfunction of the da vinci si system, instruments and/or accessories during the case. The surgeon reported the patient made a full recovery and the operation was completed successfully. The instruments and accessories user manual specifically states: - caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the patient. | 5MM ENDOWRIST INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2703998 | 2673692 | 2570308 | 2955842201200358 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-07-09 | 2012-07-30 | Yes | No | Health Professional | UNKNOWN | No | No | 2012-07-09 | No | No | Reuse | 2012
IS3000 A60P3 |  | Other | Hospitalization | On (b)(4) 2012, a patient who underwent a da vinci si hysterectomy procedure on an unknown date, provided intuitive surgical information regarding their postsurgical experience. The patient indicated that approximately 1 week after their procedure, on (b)(6) 2012, the patient became very ill and experienced pain and was admitted into (b)(6) medical center for a sleeping bowel. Per the patient, her belly was filled with fluid and required draining. Upon returning home the patient threw up, experienced a full belly, and continued to feel ill. The patient was advised her symptoms were a natural response to her procedure. On (b)(6) 2012, the patient was admitted into the er and a cat scan found fluid, however no drainage was performed. The attending physician released the patient and advised the patient that the symptoms were the result of an inflammatory condition. Per the patient, on (b)(6) 2012, one-liter of fluid was drained from her stomach and she was prescribed steroids to control the inflammation. Per the patient, dye tests performed revealed that her ureter was nicked. Per the patient a stent was placed in her ureter, an abdominal drain and catheter were placed to repair her ureter. No further details were provided. | Postoperative side effects are known and assessed normal risks of surgical procedures, and can develop regardless of the modality used to perform the surgery. Per the information provided, no system malfunctions were alleged. A request for additional information was made to the physician, however due to litigation involving this case, no further information was provided. To date, no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2704070 | 2673745 | 2570361 | 2955842201200360 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2012-06-21 | 2012-07-30 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2012-06-21 | No | No | Reuse | 2012
42017207 | 2012-06-22 | Malfunction |  | It was reported during a da vinci si hysterectomy that the maryland bipolar forceps instrument smelled of burning, sparks and smoke. The instrument was returned and evaluated and during failure analysis it was determined that arcing may have taken place. The account was contacted with this information and indicated there was no patient harm, adverse outcome or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on both yaw pulleys between the base grips. Charring is a sign of an arcing event. The instrument passed electrical continuity testing. It was determined that the charring was likely due to a breach in the instrument insulation, carbonized tissue which created a conductive path, or high generator settings. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2704096 | 2673790 | 2570406 | 2955842201200359 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-07-12 | 2012-07-30 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-07-12 | Device Not Retu | No | Reuse | 2012
42017910 | 2012-06-28 | Malfunction |  | It was reported that during a da vinci si low anterior resection procedure, the monopolar curved scissors instrument broke. After removing the instrument from the patient, a small piece was observed to be missing. The surgeon visually inspected the patient's abdominal cavity and was unable to locate the missing piece. The initial reporter indicated that the surgeon was satisfied that the missing instrument piece did not remain inside the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is broken and missing a piece at the proximal clevis interface. Additionally, the clevis is dislodged from the tube extension. It was determined that the tube extension likely broke as a result of excessive side loading. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2702966 | 2672601 | 2569217 | 2955842201200354 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-28 | 2012-07-27 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-06-29 | Device Not Retu | Yes | Reuse | 2012
IS3000 A60P7 | 2011-01-04 | Other | Hospitalization | On june 27, 2012, intuitive surgical received a letter regarding a patient who underwent a da vinci si prostatectomy (dvp) procedure at (b)(6) hospital (b)(6) begun on (b)(6) 2011, and completed on (b)(6) 2011. The patient alleges that the surgeon attempted urethrovesical anastomosis repeatedly, however, was unable to re-attach the urethra using the da vinci si surgical system. The surgeon's operative report stated that it was unclear why this could not be accomplished and it was felt perhaps related to the robotic arm and also perhaps related to the suture provided. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. Post-surgery, renal failure ensued and the patient spent four days in the intensive care unit. Since the dvp, the patient has had six additional surgical procedures performed at other hospitals to address recurrent bladder neck contracture. These procedures were performed on (b)(6) 2011. | When requesting additional information regarding this event from (b)(6) hospital, a hospital representative informed intuitive surgical that they are unwilling to discuss the patient or release any information without a signed request release from the patient. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2703147 | 2672816 | 2569432 | 2955842201200356 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-06-27 | 2012-07-27 | Yes | Yes | Health Professional | IS3000 A60P7 | No | No | 2012-06-27 | No | No | Unkown | 2012
IS3000 | 2012-06-28 | Other | Other | It was reported that prior to using the system to perform any tasks for a da vinci si sacrocolpopexy procedure, the surgeon experienced non-recoverable system error code 297. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the issue experienced by the customer was associated with the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. System error code 297 appears when the system detects that an electronic component was reporting an incorrect configuration. Upon determining this condition, the system records the fault and transitions to a safe state. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. As of july 26, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2703119 | 2672860 | 2569476 | 2955842201200355 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-28 | 2012-07-27 | Yes | Yes | Health Professional | IS3000 | Device Returned To Manufacturer | No | 2012-06-28 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-06-26 | Other | Other | Event Description: --/-- It was reported that during a da vinci si mitral valve repair procedure, the surgical staff experienced system error code 252. The system was power cycled and the same error recurred. While an isi technical support engineer attempted to troubleshoot the issue via telephone, the initial reporter indicated that water had been accidentally spilt on the vision cart prior to system error code 252 occurring. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- An isi field service engineer (fse) went onsite and was able to replicate the customer reported failure. The investigation conducted by the fse concluded that system error code 252 was associated with a video slice (vsl) printed circuit assembly board. The vsl is a video processing board in the system. The vsl takes in 8 video streams, conditions them, and blends them together into 4 blended outputs. It was determined that the water which was inadvertently spilt onto the vision cart by the surgical staff made contact with the vsl board, thus causing the vsl board to malfunction. The system was repaired by replacing the affected vsl board. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 252 appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the system records the fault and transitions to a safe state. The vsl was returned, however, reported failure experienced by the customer could not be replicated during internal evaluation. This may have been due the water moisture evaporating prior to evaluation. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012, intuitive surgical received an email response from (b)(6), indicating that the procedure was not converted to open as previously indicated. The mitral valve repaired surgical procedure was completed successfully with another davinci si robot system borrowed from the next operating room. This caused a 20 minutes delay, the procedure was completed successfully with no adverse event or affect to the patient. Customer provided additional information regarding this case. According to the customer a scope warming thermo solution fell from a mayo instrument stand near the vision side cart and some of the saline solution from the thermo was absorbed into the front of the vision side cart where the fan vents are located. The system warning error code 252 occurred and the customer was able to undock the system in the 10 seconds before it shutdown. | DA VINCI SI SURGICAL SYSTEM | Unkown | 2701648 | 2671323 | 2567939 | 2955842201200353 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | InitialFollowup | 2012-09-27 | 2012-07-26 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-09-27 | Device Not Retu | No |  | 2012
UNKNOWN | 2012-07-16 | Injury | Required Intervention | It was reported that while performing para-aortic lymph node dissection during a da vinci si hysterectomy procedure, the monopolar curved scissors (mcs) instrument wrist would rotate while the surgeon attempted to open/close the scissor blades. As a result of the unintended motion, the patient's superior vena cava was nicked and bleeding occurred. An instrument installed on another system arm was used to control the bleeding. The mcs instrument was removed and the instrument sterile adapter was reseated, however, unintended motion recurred. The instrument sterile adapter was replaced and the instrument arm resumed normal function. The patient's superior vena cava was repaired intra-operatively with suture and the planned surgical procedure was completed. The patient is reportedly in stable condition and recovering well. | The instrument sterile adapter is built into the sterile drape and manually attached to the instrument arm. When an instrument is installed onto the sterile adapter, the sterile adapter transfers motions from the instrument arm to the instrument. The instrument sterile adapter used during the procedure has not been returned for evaluation, therefore, it cannot be determined if the customer reported failure was caused by a defective instrument sterile adapter or user error when draping or installing the instrument. A follow-up medwatch report will be submitted if the instrument sterile adapter is returned or if additional information is received. | INSTRUMENT ARM DRAPE IS3000 | ENDOSCOPIC INSTRUMENT ACCESSOR | 2697339 | 2667011 | 2563627 | 2955842201200352 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-07-16 | 2012-07-24 | Yes | Yes | Health Professional | UNKNOWN | No | Yes | 2012-07-16 | No | Yes | Initial | 2012
420179 8MM | 2012-07-09 | Malfunction |  | During davinci, monopolar curved scissor was in arm 1. The instrument needed to be taken out but it was stuck. At the tip of the instrument, the instrument tip was turned and tip cover was curled in such a way the instrument could not be moved. To get instrument out, the whole trocar was removed. The tip cover was taken off so the tip of the instrument could be assessed. The orange paper the tip cover goes over was noted to be tearing and that probably caused tip cover to get stuck so the instrument could not be removed from trocar. The monopolar scissor had 2 lives left on it after this use. Instrument bagged and placed in office. No harm caused to patient. |  | DAVINCI MONOPOLAR CURVED SCISSOR | MONOPOLAR CURVED SCISSOR | 2705378 | 2674814 | 2571430 | 2674814 | 1 | NAY | User Facility |  |  | Initial | 2012-07-10 | 2012-07-24 | No | Yes | Physician | 420179 8MM | No | No Answer Provided |  |  |  |  | 2012
40018010 | 2012-06-22 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgeon observed a tear in the first mcs tip cover installed on the monopolar curved scissors instrument. The mcs instrument was removed and a replacement tip cover was installed. After some time of use, the second tip cover accessory developed a hole. A third tip cover was used and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The first and second tip cover accessories were returned and evaluated. Per the customer reported complaint, engineering observed that the first tip cover accessory has a tear approximately 0. 11 in length. Engineering concluded that damage to the tip cover is likely due to repetitive opening and closing of the mcs instrument blades. The tip cover does not exhibit evidence of arcing. Medwatch mfr report 2955842-2012-00351 was submitted to the fda concerning the second tip cover accessory. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2696443 | 2665803 | 2562401 | 2955842201200340 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-06-26 | 2012-07-23 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-06-26 | Device Not Retu | Yes | Unkown | 2012
40018010 | 2012-06-22 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgeon observed a tear in the first mcs tip cover installed on the monopolar curved scissors instrument. The mcs instrument was removed and a replacement tip cover was installed. After some time of use, the second tip cover accessory developed a hole. A third tip cover was used and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Both the first and second tip cover accessories were returned and evaluated. Per the customer reported complaint, engineering observed that the second tip cover has a hole approximately 0. 12 in length at the mid point of the tip cover. This tip cover does not exhibit evidence of arcing. Medwatch mfr report 2955842-2012-00340 was submitted to the fda concerning the first tip cover accessory. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2696451 | 2665815 | 2562401 | 2955842201200351 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-06-26 | 2012-07-23 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-06-26 | Device Not Retu | Yes | Unkown | 2012
IS3000 A60P6 | 2012-07-02 | Other | Required Intervention | It was reported that approximately 3. 5 hours into a da vinci si pancreatectomy procedure, the site experienced multiple instances of non-recoverable system error code 23 and the surgeon was unable to control the patient side manipulator (psm) arms. Unable to resolve the issue, and prior to contacting isi technical support for trouble-shooting assistance, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. With the assistance of the tse, the site reseated the blue fiber cable, however, the fiber cable led illuminated red. The patient experienced some bleeding, however, the surgeon was unable to control the patient's bleeding due to the system issue. | The investigation conducted by field service engineering (fse) was unable to replicate the customer reported issue. However, review of the site's system log by the fse found an occurrence of system error code 23 and it was determined that the system error code 23 was related to the blue fiber cable. The blue fiber cable carries communication over optical fibers between the patient side cart to the surgeon side console. As a precaution, the fse replaced the blue fiber cable. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults between two system components. In the event of these communication issues, the system records the fault and transitions to a safe state. On july 13, 2012 isi contacted the initial reporter and he indicated that he was not in the operating room when the reported issue occurred, however, he was called into the operating room to provide trouble-shooting assistance. The initial reporter indicated that he does not know what caused the bleeding experienced by the patient and that he was unable to provide any other details concerning the reported event. If additional information is received concerning the patient, a follow-up mdr will be submitted to the fda. As of july 20, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2694578 | 2663924 | 2560560 | 2955842201200341 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-02 | 2012-07-20 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-07-02 | Device Not Retu | Yes | Reuse | 2012
40015604 | 2012-06-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure the 5mm cautery hook a-sure accessory broke and fell into the patient. A 5mm port incision was created on the patient to retrieve the fragment. The accessory fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. As of july 18, 2012, there have been no reported recurrences of the incident. | 5MM CAUTERY HOOK ASUREACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2694675 | 2663982 | 2560618 | 2955842201200344 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-25 | 2012-07-20 | No | Yes | Health Professional | 40015604 | No | No | 2012-06-25 | No | Yes | Initial | 2012
IS1200 A43P9 | 2008-01-18 | Other | Hospitalization | On june 25, 2012, a patient who underwent a da vinci prostatectomy procedure on (b)(6) 2008, provided intuitive surgical information regarding their postsurgical experience. The patient indicated that their procedure was successful in removing cancer, as the patient is still cancer free 4. 5 years after the procedure. However, the patient alleges numerous complications through the case: the patient claims that during the surgery there was damage to his vein, a scratch on his right eye and damage to his penis causing a closed urethra. The patient reported that the procedure left him incontinent, and impotent. The patient also reported that the procedure has caused him to have pain during orgasm, and have pain in the legs and groin. The patient also alleges that a second procedure was required to repair a groin hernia. The patient alleges that he might not have underwent surgery had they seen the recommendation from (b)(4), which he claims, suggests men not get screened for prostate cancer partly because it results in overtreatment. | Post-operative incontinence and impotence is considered a known and assessed normal risk of prostatectomy procedures, and can develop regardless of the modality used to perform the surgery. Likewise, many complications listed: changes in orgasm, sterility, lymphedema, decrease in penis length, inguinal hernia are also potential complications from prostatectomy surgery. A review of system logs for the date of surgery indicates no system malfunctions occurred. Additionally there is no allegation of system, instrument or accessory malfunction. Complaints that the patient filed with the (b)(4) also determined, the care rendered in this case was consistent with the standard of practice in the community. A request for additional information has been made from the hospital. To date, no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2694670 | 2664013 | 2560649 | 2955842201200343 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2012-06-25 | 2012-07-20 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2012-06-25 | No | No | Reuse | 2012
40016907 | 2012-03-05 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci surgical procedure, the harmonic curved shears (hcs) insert installed on the harmonic curved shears instrument broke during the procedure. It was reported that a fragment broke off and fell into the patient. The fragment was retrieved and the planned procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the clamp arm with the white pad broke off. Evidence is not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The harmonic curved shears (hcs) insert has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be generated after the instrument is returned and evaluated. | HARMONIC CURVED SHEARS INSERT | Unkown | 2694676 | 2664017 | 2560653 | 2955842201200345 | 1 | NAY | Manufacturer | ForeignDistributor | Other | FollowupInitial | 2012-07-03 | 2012-07-20 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2012-08-01 | Device Not Retu | Yes |  | 2012
40018010 | 2012-06-25 | Malfunction |  | It was reported that approximately 58 minutes into a da vinci s surgical procedure, the surgeon observed that the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument had a tear. The surgeon made the decision to continue to use the tip cover accessory; however, after some time of use, the same tip cover accessory developed a hole. The tip cover accessory was replaced and the planned surgical procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the silicone part of tip cover has a. 120 long tear at the distal end. The tear is in the axial direction, through the entire thickness of the silicone. The silicone also exhibits a hole burned through the entire thickness. And exhibits localized melting around the hole, thus indicating that an arcing event occurred. The hole size is roughly. 020 in diameter and is located. 365 above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. The instruments and accessories user manual specifically states: general precautions and warnings o inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. Do not use another instrument to clean the tip cover accessory. To prevent damage to the tip cover accessory insulation, always straighten the instrument wrist under endoscopic visualization before removing the instrument through the cannula. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2694808 | 2664108 | 2560744 | 2955842201200348 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-27 | 2012-07-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-06-27 | Device Not Retu | Yes | Unkown | 2012
40018012 |  | Other | Other | It was reported that during a da vinci s prostatectomy procedure, while the surgeon was dissecting tissue, arcing to adjacent tissue occured from a split in the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was immediately removed and examination of the tip cover accessory by the surgical technician found that the silicon portion of the tip cover accessory had a char mark. A replacement tip cover was used to complete the planned surgical procedure. No patient harm, adverse outcome or injury occurred and no surgical intervention or repair was required as the affected tissue was tissue that was being removed as part of the surgical procedure. Per the initial reporter the tip cover accessory was discarded and to the best of his knowledge the patient has not returned to the hospital due to any post surgical complications as a result of the reported event. | The damaged tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover and/or instrument is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2694796 | 2664137 | 2560773 | 2955842201200349 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-21 | 2012-07-20 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-06-21 | No | Yes | Unkown | 2012
420110 | 2012-05-22 | Malfunction |  | Event Description: --/-- On (b)(4) 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: during robotic procedure, surgeon was using the precise bipolar. While in use he noticed that instrument was sparking at the tips. There was no harm to the patient. Instrument was immediately withdrawn and sequestered. --/-- Additional Event Description: --/-- It was reported that during a da vinci s surgical procedure, the surgeon observed a tear in the mcs tip cover installed on the monopolar curved scissors instrument. The mcs instrument was removed and a replacement tip cover was installed. After some time of use, the replacement tip cover accessory developed a hole. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2012, isi contacted (b)(6) the risk manager at (b)(6) and she indicated that the patient has not returned to the hospital due to experiencing any post surgical complications as a result of the reported event. This medwatch report was initially submitted with the incorrect event information. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover accessory has a hole approximately 0. 11 in length. Engineering concluded that damage to the tip cover is likely due to excess force applied at the tip opening, thus causing the tip cover to tear. The tip cover does not exhibit evidence of arcing. Medwatch report was submitted to the fda concerning the 1st tip cover accessory. | PRECISE BIPOLAR FORCEPS INSTRUMENT | Unkown | 2694805 | 2664147 | 2562401 | 2955842201200350 | 1 | NAY | Manufacturer | OtherCompany Representative | RISK MANAGER | InitialFollowup | 2012-06-26 | 2012-07-20 | No | Yes | Health Professional | 420110 | No | No | 2012-06-26 | No | No |  | 2012
40018012 | 2012-05-30 | Malfunction |  | On date (b)(6), 2012, intuitive surgical received (b)(6) report from the fda. The event details are provided below: multiple (5) version 12 tip covers did not adequately form to the tip of the monopolar hot shears, resulting in a loose fit and/or loss of integrity. In one case, the tip fell off into the patient's abdomen, but was retrieved. | (b)(4). The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, tip cover accessory was successfully retrieved from the patient. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2694798 | 2664150 | 2528228 | 2955842201200347 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-06-26 | 2012-07-20 | No | Yes | Health Professional | 40018012 | No | No | 2012-06-26 | No | Yes | Unkown | 2012
40018012 | 2012-06-18 | Malfunction |  | It was reported that while installing the monopolar curved scissors (mcs) tip cover accessory onto the mcs instrument for a da vinci si hysterectomy procedure, it was noted that the tip cover fit tightly and moved past the orange portion of the instrument. This caused the tip cover to stretch over the shaft of the instrument. It was also observed that the fit was tight when inserting the instrument into the trocar. During instrument removal after use, the tip cover fell into the patient. The tip cover was retrieved laparoscopically and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation despite multiple requests, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2689936 | 2659616 | 2556252 | 2955842201200337 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-06-18 | 2012-07-18 | No | Yes | Health Professional | 40018012 | No | No | 2012-06-18 | No | Yes | Initial | 2012
IS1200 A43P9 | 2012-06-19 | Other | Other | It was reported that during a da vinci hysterectomy procedure, damage to the patient's bowel may have occurred. No other information concerning the reported event has been provided. | The site has been contacted for additional information related to this event. However, no additional information has been provided, therefore, the root cause of the reported event cannot be determined. A follow-up mdr will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2690999 | 2660573 | 2557209 | 2955842201200338 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-19 | 2012-07-18 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2012-06-19 | No | No | Reuse | 2012
42017910 | 2012-06-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure,the monopolar curved scissors instrument was smoking during use. When the instrument was inspected by the surgical staff, it was observed to be hot and defective. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed arcing damage on the distal end of the main tube, with slight material melted and a hairline crack leading up to the damaged area. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2684712 | 2654944 | 2551560 | 2955842201200328 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-19 | 2012-07-13 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-06-19 | Device Not Retu | No | Reuse | 2012
42018306 | 2012-06-21 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, smoking was noticed at the end of the insulation of the permanent cautery hook instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed an arc trac at the base of the instrument yaw pulley with slightly melted materials. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2684724 | 2654955 | 2551571 | 2955842201200329 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-06-21 | 2012-07-13 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-06-21 | Device Not Retu | No | Initial | 2012
42018304 | 2012-06-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the permanent cautery hook (pch) instrument collided with another instrument. As a result, the plastic insulation seal of the pch instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal clevis was broken at the main tube connection joint and a piece of the clevis was missing. Additionally, the main tube fiber was broken off at the joint. It was determined that the instrument damage was likely caused by user handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2684754 | 2654965 | 2551581 | 2955842201200330 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-06-20 | 2012-07-13 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2012-06-20 | Device Not Retu | No | Reuse | 2012
IS3000 A60P7 | 2012-06-08 | Death | Death | It was reported that the patient underwent a hemicolectomy procedure on (b)(6) 2012 in which traditional laparoscopic techniques and the da vinci si system were used. An unrecognized bowel perforation occurred while performing traditional laparoscopic techniques during adhesiolysis. The patient expired on (b)(6) 2012. | On (b)(4) 2012, intuitive surgical's clinical sales representative (csr) contacted the surgeon and the surgeon indicated that it was his intent to perform the planned surgical procedure using traditional laparoscopic techniques and the da vinci si surgical system. Per the csr, the surgeon indicated that the patient's adhesions were dissected using traditional laparoscopic surgical techniques and that the da vinci si system was not docked to the patient during adhesiolysis. Per the csr, the surgeon believes that damage to the patient's bowel was introduced while performing adhesiolysis during the traditional laparoscopic portion of the surgical procedure. The surgeon indicated to isi's csr that the planned surgical was completed with the da vinci si surgical system and that no malfunction of the da vinci si system, instruments and/or accessories occurred. No other information concerning the reported event was provided. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2684781 | 2655010 | 2551626 | 2955842201200326 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-14 | 2012-07-13 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-06-14 | No | No | Reuse | 2012
371715 | 2012-06-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a flame was observed on the energy activation cable attached to the monopolar scissors curved instrument. The flame occurred at the instrument end of the cable and resulted in a char on the instrument drape and sterile adapter. Prior to the flame occurring, the surgeon was in cut mode with a power setting of 50. However, the surgical staff observed the cut power setting spike to 100 while the uterus was being morcellated. The power setting was changed back to 50 and the flame occurred during the next activation of energy. The portion of the procedure using the da vinci si surgical system was completed and per the hospital's standard process, the remaining part of the case was completed using traditional laparoscopic techniques. No patient harm was reported. | Investigation was conducted by an isi field service engineer (fse) who went onsite and evaluated the system. Testing was performed and the system was found to fully functional and performing to specifications. The fse concluded that the issue experienced by the site was associated with the energy activation cable which connects the device to the electrosurgical generator. As of july 13, 2012, there have been no reported recurrences of this issue at this hospital. | VALLEYLAB BIPOLAR ENERGY ACTIVATION CABLE | ENDOSCOPIC ELECTROSURGICAL INS | 2684920 | 2655133 | 2551749 | 2955842201200331 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-06-13 | 2012-07-13 | No | Yes | Health Professional | 371715 | No | No | 2012-06-13 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-06-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the site experienced system error code 23002. With the assistance of an isi technical support engineer the site back drove the master tool manipulators (mtm) with the system on and off, however, the system error code 23002 continued to occur. The tse then instructed the site to perform and emergency power off of the system, however, the issue persisted. The surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23002 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23002 system error code appears when the da vinci s safety system determines that the angular position of one of the joints, as measured by the primary sensor, is outside the range of acceptable values. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The mtmr was returned to isi for failure analysis investigation and the fault experienced by the customer was duplicated. The flex circuit assembly was replaced to address the system error code 23002. On (b)(6) 2012, isi contacted the initial reporter and she indicated that there was no patient harm as a result of the reported event, the open surgical procedure was successfully completed and the patient tolerated the open surgical procedure well. The initial reporter indicated that it is unknown if the patient has returned to the hospital due to experiencing any post surgical complications. As of july 13, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2685131 | 2655193 | 2551810 | 2955842201200333 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-15 | 2012-07-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-06-15 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-06-15 | Injury | OtherRequired Intervention | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument causing a burn to the patient's bowel. The affected area of the patient's bowel was repaired and the planned surgical procedure was completed. | Manufacturer Narrative: --/-- Corrected information was provided. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. | 8MM MCS TIP COVER ACCESSORY | Initial | 2685470 | 2655456 | 2552072 | 2955842201200335 | 1 | NAY | Manufacturer | User facilityOther | Nurse | FollowupInitial | 2012-06-18 | 2012-07-13 | Yes | Yes | Health Professional | 40018010 | No | Yes | 2012-06-18 | Device Not Retu | Yes |  | 2012
40018010 | 2012-06-18 | No Answer Provided | Other | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to gain additional information about the event, but no additional information has been received. If additional event details become available a follow up mdr will be submitted. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole on the transparent part of the tip cover. Engineering also observed cracks at the proximal end of the tip cover. No arcing damage was found at the hole and the cracks. Evidence is not conclusive, but damage is likely due to mishandling/misuse. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip.It was reported that during a da vinci s surgical procedure, smoke was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was observed to have a hole in the transparent portion of the tip cover accessory. The planned surgical procedure was completed by using a backup tip cover. No patient harm, injury or adverse outcome was reported. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2685471 | 2655457 | 2552073 | 2955842201200334 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | FollowupInitial | 2012-06-21 | 2012-07-13 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-07-13 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-06-18 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, arcing was observed originating from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument, causing a burn to the patient's uterus. The uterus was removed as a part of the planned surgical procedure. No patient harm or adverse outcome was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2685488 | 2655458 | 2552072 | 2955842201200332 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-18 | 2012-07-13 | Yes | Yes | Health Professional | 40018010 | No | Yes | 2012-06-18 | No | Yes | Initial | 2012
42014304 | 2012-06-18 | Malfunction |  | It was reported that during a da vinci si pyeloplasty procedure the 5mm maryland dissector instrument was found to have tube insulation peeling away with fragments having fallen into the patient. The instrument fragments were retrieved and the procedure was completed. No patient harm, injury or adverse outcome was reported. No further information was provided. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the black coating had peeled. The main tube insulation had scratches and gouge marks. Engineering concluded that damage was likely due to mishandling or misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2685472 | 2655459 | 2552075 | 2955842201200336 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-19 | 2012-07-13 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | Yes | 2012-06-19 | Device Not Retu | No | Reuse | 2012
428053 |  | Other | Required Intervention | It was reported that after a da vinci si single site cholecystectomy procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the hemolok ml clip applier instrument had opened. To address the bleeding, the surgeon performed a second surgery to repair the bleeding. No further information was provided. | Since the hemolok ml clip applier instrument was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. | HEMOLOK ML CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 2685469 | 2655472 | 2552088 | 2955842201200325 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-13 | 2012-07-13 | Yes | Yes | Health Professional | 428053 | No | No | 2012-06-13 | No | No | Unkown | 2012
400180 | 2010-12-29 | Injury | DisabilityHospitalization | Event Description: --/-- It was reported that after a successful da vinci si hysterectomy procedure, the pt experienced pain during recovery. The pt pain is allegedly attributed to a bladder injury caused by arcing from a hole in the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. The pt required prolonged hospitalization time and it was reported that the pt continues to suffer from the effects of the bladder surgery. No other adverse pt outcome was reported. --/-- Additional Event Description: --/-- It was reported that after a successful da vinci s hysterectomy procedure, the patient experienced pain during recovery. The patient pain is allegedly attributed to a bladder injury caused by arcing from a hole in the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. The patient required prolonged hospitalization time and it was reported that the patient continues to suffer from the effects of the bladder surgery. No other adverse patient outcome was reported. | Manufacturer Narrative: --/-- The tip cover accessory has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. The root cause of the surgical complication experienced by the pt is also currently undetermined. Multiple attempts have been made to gain additional info about the event, but the hosp has yet to provide further details. If additional event details become available a f/u mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- (b)(4). | 8MM MCS TIP COVER ACCESSORY | Initial | 2694341 | 2661145 | 2557781 | 2955842201200246 | 1 | NAY | Manufacturer | User facility | ATTORNEY | InitialFollowup | 2012-05-01 | 2012-07-13 | Yes | Yes | Health Professional | 400180 | No | No | 2012-05-01 | Device Not Retu | Yes |  | 2012
UNKNOWN | 2012-06-28 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument jaw was moving in the opposite direction from the surgeon's intended movement. An isi technical support engineer (tse) attempted to troubleshoot the issue via telephone and recommended the surgical staff reseat the instrument and sterile adapter. When attempting to remove the instrument, the instrument became stuck in the sterile adapter. In an effort to release the instrument, the instrument arm was undocked from the patient. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | A review of the system error logs by the tse showed multiple occurrences of system error code 31009. System error code 31009 was due to a loss of electrical connection between the instrument and system. This message occurs if, after full detection of an instrument, one or more sensors are not detected (the system uses multiple redundant sensors for instrument detection). This message is a diagnostic message only and does not interrupt surgery. The appearance of this message is consistent with a poor connection between the instrument and the sterile adapter. An isi field service engineer (fse) went onsite and evaluated the system. While testing the instrument arm using a reliable sterile adapter, the fse was unable to replicate the issue experienced by the customer. The instrument sterile adapter used during the procedure was discarded by the hospital, therefore, it cannot be determined if the customer reported failure was caused by a defective instrument sterile adapter or user error when draping or installing the instrument. The instrument sterile adapter is built into the sterile drape and is manually attached to the instrument arm. When an instrument is installed onto the sterile adapter, the sterile adapter transfers motions from the instrument arm to the instrument. | INSTRUMENT ARM DRAPE IS3000 | ENDOSCOPIC INSTRUMENT ACCESSOR | 2683810 | 2654063 | 2550680 | 2955842201200327 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-28 | 2012-07-12 | Yes | Yes | Health Professional | UNKNOWN | No | Yes | 2012-06-28 | No | Yes | Initial | 2012
VS 3000 | 2012-06-27 | Malfunction |  | During robotic hysterectomy, no one was present at console to activate power source to activate pedal. However, upon advancing instruments there was smoke and a burn spot on the bipolar cautery instrument. The active esu cord had been plugged into the esu monopolar output instead of the bipolar output. There was also a burn on the patient's uterus. The robotic unit had instigated the electrical without anyone telling the panel to. The patient was having a hysterectomy so she was not harmed. ======================manufacturer response for robotic surgical system, (brand not provided) (per site reporter). ======================error is probable when using singular banana plug receptacles. There is a disposable part that is configured to eliminate the possibility. The part number is e0512. Cord may have been plugged into monopolar port when it should have been in bipolar port. |  | DA VINCI | SYSTEM ROBOTIC SURGICAL | 2702047 | 2671515 | 2568131 | 2671515 | 1 | NAY | User Facility |  |  | Initial | 2012-07-12 | 2012-07-12 | No | Yes | Physician | VS 3000 | Yes | No Answer Provided |  |  |  |  | 2012
VS 3000 | 2012-06-27 | Malfunction |  | During robotic hysterectomy, no one was present at console to activate power source to activate pedal. However, upon advancing instruments there was smoke and a burn spot on the bipolar cautery instrument. The active esu cord had been plugged into the esu monopolar output instead of the bipolar output. There was also a burn on the patient's uterus. The robotic unit had instigated the electrical without anyone telling the panel to. The patient was having a hysterectomy so she was not harmed. ======================manufacturer response for robotic surgical system, (brand not provided) (per site reporter). ======================error is probable when using singular banana plug receptacles. There is a disposable part that is configured to eliminate the possibility. The part number is e0512. Cord may have been plugged into monopolar port when it should have been in bipolar port. |  | DA VINCI | SYSTEM ROBOTIC SURGICAL | 2702047 | 2671515 | 2568131 | 2671515 | 1 | NAY | User Facility |  |  | Initial | 2012-07-12 | 2012-07-12 | No | Yes | Physician | VS 3000 | Yes | No Answer Provided |  |  |  |  | 2012
IS1200 A43P9 | 2012-06-07 | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure the surgical staff experienced a non-recoverable system fault. The system was restarted multiple times, however, the system fault continued to recur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer concluded that the non-recoverable fault was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned for evaluation. The reported failure was replicated during internal testing and it was determined that the axis-6 potentiometer had malfunctioned, thus generating the system fault experienced by the customer. As of july 4, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2674211 | 2643755 | 2540368 | 2955842201200317 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-07 | 2012-07-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2012-06-07 | Device Not Retu | No | Reuse | 2012
IS3000 A60P7 | 2012-06-27 | Other |  | Event Description: --/-- It was reported that the surgeon console would not power on for use for a da vinci si surgical procedure. Multiple outlets were tried and the power cable was replaced, however, the issue could not be resolved. The patient had already been administered anesthesia when the surgeon made the decision to abort the planned surgical procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by an intuitive surgical field service engineer determined that the power issue experienced by the customer was associated with the medical grade power supply (mgps). The mgps converts the ac power from the wall socket (through the power cord) and converts it to dc power for use on the system. The system was repaired by replacing the affected mgps. As part of the da vinci surgical system pre-surgery inspection, the customer is instructed to power on the system before the patient is under anesthesia. This is to ensure the proper functioning of the system prior to use. As of july 4, 2012, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012, intuitive surgical was informed that the planned surgical procedure was completed using traditional laparoscopic techniques and not aborted post anesthesia as initially reported. Per the information provided above, intuitive surgical determined that this event does not require reporting via a medical device report as no adverse event occurred and the recurrence of the reported failure mode is not expected to likely cause or contribute to a death, serious injury or serious deterioration in the state of health of a patient. | DA VINCI SI SURGICAL SYSTEM | Initial | 2674521 | 2644079 | 2540692 | 2955842201200318 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | InitialFollowup | 2012-06-27 | 2012-07-06 | Yes | Yes | Health Professional | IS3000 A60P7 | Yes | No | 2012-07-06 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-06-13 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, using dual surgeon side consoles (ssc) the surgeon working from ssc 1 experienced resistance while manipulating the master tool manipulators (mtm) and when the surgeon let go of the mtms on the ssc1 the patient's abdominal wall was damaged. The site contacted isi technical support engineering (tse) for trouble-shooting assistance. With the assistance of the tse, the site switched from dual ssc setup to single ssc setup to proceed with the planned surgical procedure. Resistance on the ssc 2 mtms was not experienced by the surgeon. The tse asked the initial reporter if the surgeon working from ssc1 had removed head from the high resolution stereo viewer (hrsv) prior to releasing the his hands from the mtms; however, the surgical staff was unable to provide this information. Review of the site's system log by the tse found recoverable system error code 23002 occurred on the mtmr at the time of the customer's report. The planned surgical procedure was completed. | The field investigation conducted by the isi fse concerning this event was unable to replicate the issue experienced by the site. However, after manipulating the mtm on ssc1 the fse was able to induce the recoverable system error code 23002 and concluded that the system error code was related to the mtmr; however, he did not note any resistance on the mtmr or the patient side manipulators. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The reported resistance from the surgeon and the 23002 are both consistent with an inaccurate calibration, since the gravity compensation algorithm on the master relies on an accurate measurement of the arm position. Gravity compensation is provided for the comfort of the surgeon and is not intended to allow them to release their grip on the master while in following. As a precaution, the fse calibrated the mtms on ssc 1. The patient injury most likely occurred as a result of the surgeon's removal of his hands from the mtms prior to removal of his head from the high resolution stereo viewer. The da vinci si user manual warning(s) specifically states: once in following, the surgeon console operator must not remove is or her hands from the masters until removing his or her head from the surgeon console viewer-thereby taking the system out of following mode. Failure to do so may result in uncontrolled movement of the masters, resulting in serious harm to the patient. The previous calibration of the mtm occurred on june 12, 2012 during a routine preventative maintenance (pm) on the site's system. The fse reported issues with his laptop while calibrating the master tool manipulator (mtms) on ssc 1. With the assistance of isi dvstat support, it was determined that the fse's laptop had malfunctioned and the calibration files were restored to site's system. A sine cycle and test drive of the system was performed by the fse found that the system performed within specification. Isi has made several attempts to verify the status of the patient. However, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is provided. As of july 6, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2674587 | 2644122 | 2540735 | 2955842201200319 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2012-06-13 | 2012-07-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-06-13 | Device Not Retu | No | Reuse | 2012
42018406 | 2012-04-13 | Malfunction |  | It was reported to intuitive surgical on (b)(4) 2012 that during a da vinci si thymectomy procedure, the permanent cautery spatula (pcs) instrument stopped working. The report as initially reported to intuitive surgical was not reported to the fda as no adverse event or incident was alleged to have occurred by the site. However, failure analysis of the returned pcs instrument found that the instrument had a missing piece. On (b)(4) 2012, intuitive surgical contacted the supply chain coordinator and she indicated that the surgical staff confirmed that a piece from the pcs instrument broke off and fell into the patient. The broken instrument piece was successfully retrieved and no patient harm, adverse outcome or injury occurred. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument drive cables are intact and undamaged; however, the conductor wire is broken at the yaw pulley exit. The instrument failed electrical continuity testing, making cautery non-functional. No signs of arcing were observed. Engineering observation also found that a conductor cap is broken. One section of the cap is broken off with a missing piece. The wire segment is exposed as a result of breakage. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2674599 | 2644134 | 2540747 | 2955842201200320 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2012-06-15 | 2012-07-06 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-06-15 | Device Not Retu | No | Reuse | 2012
40018012 | 2012-06-12 | Other | Other | It was reported that during a da vinci s urological procedure, while the surgeon was dissecting tissue using the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed, arcing to adjacent tissue being removed occurred. The mcs instrument was removed and a replacement tip cover was installed on the mcs instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, intuitive surgical contacted surgical technician, (b)(6) at (b)(6) health system and he indicated that he was present during the surgical procedure and that arcing from a split in the tip cover accessory to adjacent tissue occurred, however, no intervention or repair was required as the affected tissue was being removed. Per (b)(6), there was no patient harm, adverse outcome or injury and to the best of his knowledge the patient has not returned to the hospital due to any post operative complications as a result of the event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2674600 | 2644135 | 2540748 | 2955842201200321 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-12 | 2012-07-06 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-06-12 | No | Yes | Unkown | 2012
40018012 | 2012-06-08 | Other | Other | It was reported that during a da vinci s gynecological procedure, while the surgeon was dissecting tissue using the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed, arcing to adjacent tissue being removed occurred. The mcs instrument was removed and a replacement tip cover was installed on the mcs instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4), intuitive surgical contacted rn (b)(6) at (b)(6) health system and she indicated that the tip cover accessory was discarded and that she was unable to provide any additional information concerning the reported event. On (b)(4), 2012 intuitive surgical contacted surgical technician, (b)(6) at (b)(6) health system and he indicated that he was present during the surgical procedure and that arcing to adjacent tissue did not require any intervention or repair as the affected tissue was being removed. Per (b)(6), there was no patient harm, adverse outcome or injury and to the best of his knowledge the patient has not experienced any post operative complications. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2674602 | 2644137 | 2540750 | 2955842201200323 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-12 | 2012-07-06 | Yes | Yes | Health Professional | 40018012 | No | No | 2012-06-12 | Device Not Retu | Yes | Unkown | 2012
40018012 | 2012-06-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the arcing from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed was observed. The mcs instrument was immediately removed and examination of the tip cover accessory by the surgical staff found that the tip cover had a hole. The planned surgical procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome, or injury was reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, the isi clinical sales representative indicated that he also examined the tip cover accessory and the tip cover had a hole located in the gray portion of the tip cover below the silicon portion of the tip cover. On (b)(4) 2012, isi contacted rn (b)(6) and she indicated that it was noted that during the surgical procedure, the monopolar curved scissors had made contact with a suction probe installed in an assistant port. (b)(6) indicated that the electrical surgical unit (esu) used during the procedure was a valley lab force fx, however, she does not recall what the esu settings/mode were at the time of the procedure. (b)(6) indicated that a video recording of the surgical procedure is not available for review by intuitive surgical and the mcs tip cover accessory was discarded. (b)(6) indicated that no patient harm, adverse outcome, or injury occurred. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2674604 | 2644139 | 2540752 | 2955842201200322 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-11 | 2012-07-06 | No | Yes | Health Professional | 40018012 | No | No | 2012-06-11 | No | Yes | Unkown | 2012
40018012 | 2012-06-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Since the tip cover accessory will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2674621 | 2644176 | 2540789 | 2955842201200324 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-07 | 2012-07-06 | No | Yes | Health Professional | 40018012 | No | No | 2012-06-06 | No | Yes | Initial | 2012
IS1200 A43P9 | 2012-06-13 | Other | Other | The investigation conducted by the field service engineer concluded that the vision issue experienced by the site was associated the high resolution video (hrsv) monitor. The hrsv monitor provides the video image for the surgeon side console (ssc). The system was repaired by replacing the affected high resolution video (hrsv) monitor. As of july 2, 2012, there have been no reported recurrences of the issue at this hospital. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the site was associated the high resolution video (hrsv) monitor. The hrsv monitor provides the video image for the surgeon side console (ssc). The system was repaired by replacing the affected high resolution video (hrsv) monitor. The hrsv was returned to intuitive surgical for failure analysis investigation. During evaluation, the reported failure mode was replicated and the left display was not powering on. The hrsv was repaired by replacing both right and left monitors.It was reported that while preparing to start a da vinci prostatectomy procedure, the site experienced a black bar in the left eye of the surgeon side console, which eventually turned completely black. With the assistance of an isi clinical sales representative, the site restarted the system; however this did not resolve the problem. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2671528 | 2641219 | 2537856 | 2955842201200315 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | InitialFollowup | 2012-06-13 | 2012-07-05 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2012-07-18 | Device Not Retu | No | Reuse | 2012
420179 |  | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The arcing occurred onto coopersurgical koh-efficient system uterine manipulator. No patient harm, adverse outcome or injury was reported. | The instrument and mcs tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Intuitive surgical sales representative was informed that coopersurgical investigation of the complaint revealed that this was the result of spark gap, bringing the instrument too close to the uterine manipulator. If additional event details become available, a follow up mdr will be submitted. A follow-up mdr will be submitted, if the instrument and/or tip cover accessory is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2671463 | 2641260 | 2537896 | 2955842201200314 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-05 | 2012-07-05 | No | Yes | Health Professional | 420179 | No | No | 2012-06-05 | No | No | Unkown | 2012
IS3000 A60P6 | 2012-06-13 | Other | Other | It was reported that while preparing to start a da vinci si surgical procedure the surgeon experienced reverse (non-intuitive) motion of instruments installed on patient side manipulator (psm) arm 1 and patient side manipulator 2. Patient side manipulator arm 3 was working normally. With the assistance of an technical service engineer, the site re-draped the system and verified the camera was not backwards; however this did not resolve the problem. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. The planned procedure was completed and no adverse outcome was reported. | The investigation conducted by the field service engineer (fse) was able to replicate the motion issues experienced by the surgeon. The system was repaired by reprogramming and calibrating the set-up joints on psm1 and psm2. The motion issues experienced by the customer were attributed to incorrect maintenance performed to the system prior to the surgery. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2671513 | 2641265 | 2537901 | 2955842201200316 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-13 | 2012-07-05 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-06-13 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2010-03-01 | Other |  | On june 5, 2012, intuitive surgical received a legal complaint regarding a patient who underwent a da vinci s hysterectomy procedure at (b)(6) medical center ((b)(6)) on or about (b)(6) 2010. The patient alleges that due to improper positioning prior to and/or during the procedure, she sustained a fractured coccyx and lumbosacral plexopathy, causing permanent pain and disability. | Despite multiple requests, the surgeon who performed the patient's surgical procedure has not provided intuitive surgical with any additional information regarding this event. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2669372 | 2639097 | 2535754 | 2955842201200313 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-06-05 | 2012-07-03 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-06-05 | No | No | Reuse | 2012
371498 | 2012-05-31 | Other | Other | It was reported that during a da vinci s hysterectomy procedure the cautery for the bipolar instrument stopped working. The patient had been under anesthesia for approximately 45 minutes when the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by an isi field service engineer (fse) who went onsite and evaluated the system. The fse concluded that the issue experienced by the site was associated with the bipolar energy activation cable which connects the device to the electrosurgical generator. As of june 29, 2012, there have been no reported recurrences of this issue at this hospital. | VALLEYLAB BIPOLAR ENERGY ACTIVATION CABLE | ENDOSCOPIC ELECTROSURGICAL INS | 2666331 | 2636096 | 2532752 | 2955842201200308 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-31 | 2012-07-02 | Yes | Yes | Health Professional | 371498 | Yes | No | 2012-05-31 | Device Not Retu | No | Invalid Data | 2012
IS2000 A51P8 | 2012-05-30 | Other | Other | It was reported that during system setup for a da vinci s surgical procedure the surgical staff experienced video loss on the right side image of the surgeon console. Via telephone, an isi technical support engineer instructed the staff to power down the system, reseat all power cords on the vision cart, swap the ends of the blue system cable, and restart the system. The system powered up and homed successfully and the staff continued with setup. Immediately after the system was docked to the patient, the same issue occurred when the surgeon sat down at the surgeon console. The patient was under anesthesia and the port incisions had been made when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm was reported. | Investigation was conducted by an isi field service engineer (fse) who went onsite and evaluated the system. The fse concluded that the vision issue experienced by the site was associated with the blue system cable. The blue system cable passes video signals between the surgeon console and vision cart. The system was repaired by replacing the blue system cable. As of june 28, 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2666336 | 2636101 | 2532757 | 2955842201200310 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-30 | 2012-07-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-05-30 | Device Not Retu | No | Reuse | 2012
IS2000 A51P7 | 2012-06-04 | Other | Other | It was reported that approximately 1 hour and 45 minutes into a da vinci s prostatectomy procedure, the surgeon experienced limited range of motion and drift while articulating the master tool manipulators. Prior to contacting isi technical support engineering (tse) for trouble-shooting assistance, the site performed an emergency stop, however, the issue persisted. With the assistance of an isi tse, the site swapped out the mcs instrument and re-draped the psm, the issue continued to recur. Review of the site's system log by the tse did not find any system errors related to the issue. The tse instructed the site to change the surgeon scaling from fine to normal; however, the surgeon indicated that the issue appeared worse. Unable to resolve the issue, the surgeon made the decision to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) was unable to replicate the limited range of motion experienced by the surgeon. The fse inspected the cables of the affected patient side manipulator (psm) and the cable tensions were found to be within specification. The fse concluded that the limited range of motion experienced by the customer was likely due to a faulty sterile adapter. The fse was able replicate a slight drift on the right master tool manipulator (mtmr). The instrument sterile adapter is part of the instrument arm drape and provides the point of connection between the instrument and the instrument arm. The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The mtmr was calibrated to repair the system. On (b)(4) 2012, the isi clinical sales representative (csr) present during the surgical procedure indicated that the surgeon had completed an extensive lysis of adhesion and prior to starting dissection at the vas deferens, the surgeon noticed the drift. The csr indicated that the drift experienced by surgeon occurred when the surgeon removed his hands from the mtms while his head remained in the high stereo resolution viewer. The csr indicated that the instruments installed on the psm functioned when the surgeon manipulated the mtms. Per the csr, the patient is recovering for 2 to 3 weeks and will return to the hospital to have the procedure completed. The csr indicated that no harm, adverse outcome or injury to the patient occurred and the patient was released from the hospital. The csr indicated that he followed up with the surgeon and the surgeon reported on (b)(4) 2012, and the patient was doing fine and was awaiting the next viable surgery date. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2666352 | 2636117 | 2532773 | 2955842201200307 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-04 | 2012-07-02 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2012-06-04 | Device Not Retu | No | Reuse | 2012
42017910 | 2012-05-24 | Malfunction |  | It was reported that during a da vinci si surgical procedure a piece of the monopolar curved scissors instrument blade fell into the patient. The detached piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, both scissor blades are intact and undamaged. However, the distal end of the tube extension is broken on one side, adjacent to one of the keys. A. 195 x. 225 detached piece of the tube extension was returned with the instrument. The tube extension exhibits cracking directly below the broken section. It was determined that the damage is likely due to mishandling/misuse. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2666353 | 2636119 | 2532775 | 2955842201200309 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-05-30 | 2012-07-02 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-05-30 | Device Not Retu | No | Reuse | 2012
40018012 | 2012-05-29 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument came off of the instrument and fell inside the patient. The tip cover was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- Since the tip cover accessory will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- Complaint of instrument tip cover accessory falling off of the instrument could not be confirmed. Tip cover accessory was installed on original instrument and placed on the system for testing, but the failure could not be duplicated. No physical damage was found on the tip cover. | 8 MM MCS TIP COVER ACCESSORY | Initial | 2666622 | 2636382 | 2533038 | 2955842201200311 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2012-06-01 | 2012-07-02 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-07-05 | Device Not Retu | Yes |  | 2012
40018010 | 2012-04-30 | Malfunction |  | It was reported to intuitive surgical on (b)(4) 2012 that during a da vinci s hysterectomy procedure, a char was observed on the monopolar curved scissors instrument when it was removed from the patient. The distal tip of the instrument was also observed to be melted and have a crack. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(4) 2012 intuitive surgical received med watch report for patient identifier (b)(4) from medical center (b)(6) indicating the following: at approximately 0900 circulator and scrub noticed an electrical burning odor while the surgeon was using monopolar scissor 311 and notified surgeon. Assistant removed trocar with arm attached and discovered scissor arm damaged and melted. Case converted to paparascopic. Tissue on uterus charred. Case was a robotic assist hyst. | On june 4, 2012 intuitive surgical received the returned mcs tip cover accessory used in conjunction with the monopolar curved scissors (mcs) instrument. Engineering evaluation found that the tip cover has a. 030 x. 050 oval shaped hole in the silicone. The hole is located. 500 above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. With tip cover installed on mcs instrument the hole position is near grip cable and distal clevis ear. Engineering evaluation of the mcs instrument found that the main tube was found broken located slightly below the proximal clevis. The broken main tube can be rotated 360 degrees; however, it still remains intact with the instrument. The broken area shows signs of arcing due to localized melting. Engineering concluded that the damage was likely due to excessive side loading or other type of misuse. The instrument also passed electrical continuity testing. Engineering evaluation of the mcs instrument also found black char marks on the blade tips. The surface of the blades exhibit burnt material, indicative that unintended arcing occurred. Tube reinforcement ring also shows peeling off of surface material. On (b)(4) 2012, intuitive surgical contacted (b)(6) in risk management at medical center (b)(6), and she indicated that during the da vinci s hysterectomy procedure, while the surgeon was dissecting the patient's vaginal cuff using the monopolar curved scissors (mcs) instrument arcing to the patient's uterus occurred, causing charring to the patient's uterus, which was being removed. (b)(6) indicated that to the best of her knowledge, the surgical team did not experience any issues with the da vinci s system prior to the arcing event; however, the planned surgical procedure was completed using laparoscopic techniques due to the damaged condition of the mcs instrument and it was unknown at the time of the event if any harm or injury to the patient occurred. Melissa also provided the following additional details: - the monopolar curved scissors (mcs) instrument and mcs tip cover were inspected prior to use and there was no damage found. - the monopolar curved scissors instrument made contact with another instrument used during the surgical procedure. - the cannula was inspected prior to use. - the esu generator used during the surgical procedure was a valley lab, force, fx and the settings were coag 90, cut 90. On (b)(4) 2012 the site's da vinci s system was inspected by an isi field service engineer (fse) and found no issues. Functional testing of the system found that it functioned within specification. Isi's regional fse manager also met with (b)(6) and members of the surgical staff present during the surgical procedure and they indicated that the patient involved had a tight anatomy and that the surgical staff observed that the mcs instrument exhibited excessive torquing. Based on the additional information provided by the site, the regional fse manager has concluded that excessive torquing of the mcs instrument caused the instrument to break, which allowed energy to escape and arc to the patient's uterus, which was being removed. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has submitted manufacturer report 2955842-2012-00245 concerning the monopolar curved scissors instrument. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2668069 | 2637842 | 2534498 | 2955842201200312 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-04 | 2012-07-02 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-06-04 | Device Not Retu | No | Unkown | 2012
IS3000 A60P6 | 2012-05-25 | Death |  | It was reported that after undergoing a da vinci si cholecystectomy procedure on (b)(6) 2012, the patient later expired on (b)(6) 2012. During the surgical procedure, the surgeon decided to remove the patient's gall bladder using traditional laparoscopic techniques and without using a bag. An isi representative in attendance during the surgical procedure indicated that no system malfunctions occurred during the case and the planned procedure was completed without any issues. | Despite multiple requests, the surgeon who performed the patient's da vinci surgical procedure has not provided intuitive surgical with any additional information regarding this event. As stated, an isi representative in attendance during the surgical procedure indicated that no system malfunctions occurred during the case and the planned procedure was completed without any issues. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2664879 | 2634699 | 2531355 | 2955842201200300 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-29 | 2012-06-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-05-29 | No | No | Reuse | 2012
40018012 | 2012-05-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the monopolar curved scissors (mcs) tip cover accessory, that was installed on the mcs instrument, fell into the patient as the instrument was being pulled out of the trocar. The tip cover was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Visual inspection shows no damage to the tip cover. A fit check was performed by installing the tip cover on a mcs instrument, which revealed a proper fit against the tube extension with no slippage between the mating parts. The customer reported failure could only be replicated when the proximal end of the tip cover is partially installed over the tube extension shoulder and the instrument is removed from the cannula accessory at a high speed/force. The instruments and accessories user manual specifically states: the tip cover accessory is not properly installed if any part of the orange surface is visible, it is also not properly installed if installed beyond the orange surface and over the shaft. This causes a bulge on the shaft and may prevent it fitting through the cannula. | 8 MM MCS TIP COVER ACCESSORY | Unkown | 2664856 | 2634733 | 2531389 | 2955842201200299 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-05-29 | 2012-06-29 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-08-20 | Device Not Retu | Yes |  | 2012
|  | Other | Other | The following patient comments were found posted on intuitive surgical's marketing feedback patient survey website concerning the outcome of their da vinci surgical procedure: i didn't recover as quickly as i thought. I wasn't fully prepared for the incontinence. I also developed a left hip pain that lasted for a month, maybe more because i am still in recovery. Doctors aren't sure if the hip pain is related to the procedure. No other information was provide. Late submission of this medwatch report is due to an oversight during review of the survey by isi marketing. | Based on the limited information provided in the survey, intuitive surgical was unable to follow up with the initial reporter of this event; therefore, the root cause of the reported event cannot be determined. A follow-up mdr will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2665261 | 2635087 | 2531743 | 2955842201200301 | 1 | NAY | Manufacturer | Consumer | Other | Initial | 2012-05-07 | 2012-06-29 | Yes | No | Health Professional |  | No | No | 2012-05-07 | No |  | Unkown | 2012
IS3000 A60P5 | 2012-05-29 | Other | Other | It was reported that prior to starting a da vinci si hysterectomy procedure, the site experienced video loss in the left eye of the surgeon side console. With the assistance of an isi technical support engineer, the site attempted to troubleshoot by checking cable connections and restarting the system; however, the vision problem persisted. The patient was under anesthesia for 45 minutes when the surgeon decided to convert the planned procedure to traditional open techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) concluded the 12v power supply to the left camera controller unit (ccu) was defective. The camera cable was also found to be shorted. Based on the system error logs, it was determined that the system error experienced by the site may be associated with defective doco units. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected docos and replacing the camera cable. The docos were returned to isi for failure analysis and the customer reported complaint was confirmed, troubleshooting revealed the problem was caused by the left ccus. As a result it was determined that the system error code was observed with the doco. As of june 28, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2665376 | 2635166 | 2531822 | 2955842201200302 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-29 | 2012-06-29 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2012-05-29 | Device Not Retu | No | Reuse | 2012
42805307 | 2012-06-08 | Malfunction |  | It was reported that during a da vinci si single-site cholecystectomy procedure, a clip installed on the medium-large clip applier instrument fell into the patient when the instrument was advanced through the cannula. The site made three additional attempts to advance the instrument with replacement clips, however, each time the instrument was advanced, the clips fell into the patient. All 4 of the clips were removed laparoscopically from the patient and the surgeon made the decision to complete the surgical task using a traditional laparoscopic clip applier instrument through an assistant port. The planned surgical procedure was successfully completed with the da vinci si surgical system and no patient harm, adverse outcome or injury occurred. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. A visual inspection of the instrument found no damage to the grips. An in-house test clip loaded onto the instrument grips without difficulty. The instrument was placed on an in-house (b)(4) system for functional testing and the clip did not fall out while the instrument was driven in roll nor during partial opening/closing of the grips. A clip was also successfully applied to a piece of suture. No physical damage was found. | 5 MM MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2665822 | 2635621 | 2532277 | 2955842201200304 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-06-08 | 2012-06-29 | No | Yes | Health Professional | 42805307 | Device Returned To Manufacturer | No | 2012-06-08 | Device Not Retu | No | Reuse | 2012
42805307 |  | Malfunction |  | It was reported that during a da vinci si single-site cholecystectomy procedure a clip installed on the medium-large clip applier instrument fell into the patient. The clip was removed laparoscopically and the planned surgical procedure was completed with a replacement medium-large clip applier instrument and clip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. A visual inspection of the instrument found no damage to the grips. An in-house test clip loaded onto the instrument grips without difficulty. The instrument was placed on an in-house is3000 system for functional testing and the clip did not fall out while the instrument was driven in roll or during partial opening/closing of the grips. A clip was also successfully applied to a piece of suture. No physical damage was found. | 5 MM MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2665844 | 2635636 | 2532292 | 2955842201200305 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-30 | 2012-06-29 | No | Yes | Health Professional | 42805307 | Device Returned To Manufacturer | No | 2012-05-30 | Device Not Retu | No | Reuse | 2012
42805307 |  | Malfunction |  | It was reported that during a da vinci si single-site cholecystectomy procedure, a clip installed on the medium-large clip applier instrument fell into the patient. The clip was removed laparoscopically and the planned surgical procedure was completed with a replacement medium-large clip applier instrument and clip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. A visual inspection of the instrument found no damage to the grips. An in-house test clip loaded onto the instrument grips without difficulty. The instrument was placed on an in-house is3000 system for functional testing and the clip did not fall out while the instrument was driven in roll nor during partial opening/closing of the grips. A clip was also successfully applied to a piece of suture. No physical damage was found. | 5 MM MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2665849 | 2635641 | 2532297 | 2955842201200303 | 1 | GCJ | Manufacturer | Company Representative | Other | Initial | 2012-05-30 | 2012-06-29 | No | Yes | Health Professional | 42805307 | Device Returned To Manufacturer | No | 2012-05-30 | Device Not Retu | No | Reuse | 2012
42018306 | 2012-06-05 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the permanent cautery hook (pch) instrument broke and a piece from the instrument fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that instrument was returned with the hook and ceramic missing from the yaw pulley. The hook shank was not found within the yaw pulley. The yaw pulley exhibits a crack in the axial direction. Engineering concluded that the instrument pieces likely broke off as a result of mishandling/misuse. On june 12, 2012 isi contacted robotics coordinator, (b)(6) at hospital and she indicated that while removing the permanent cautery hook (pch) instrument from the trocar, the hook on the pch instrument caught onto the trocar, causing the tip to break. (b)(6) indicated the patient has not retained any instrument piece(s) and there was no harm or injury to the patient. The instruments and accessories user manual specifically states: general precautions and warnings; handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2665826 | 2635645 | 2532301 | 2955842201200306 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-06-05 | 2012-06-29 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | Yes | 2012-06-05 | Device Not Retu | No | Unkown | 2012
| 2012-03-29 | Malfunction |  | The davinci hook cautery was used during a robotic nissen fundoplication with a hernia repair and cholecystectomy surgery. The scrub tech noticed a piece of the instrument fell off in the patient. The piece was retrieved by the physician and given to the nurse along with the instrument. No harm was done to the patient. The instrument was sent down with the broken piece along with the lot number and other needed information. The resource nurse was notified. |  | HOOK L CAUTERY 8MM | PERMANENT CAUTERY HOOK | 2678869 | 2648434 | 2545050 | 2648434 | 1 | KNS | User Facility |  | Invalid Data | Initial | 2012-06-29 | 2012-06-29 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
40018010 | 2012-05-28 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgeon observed smoke coming from the tip of the monopolar curved (mcs) scissors instrument during energy activation. When the surgeon inspected the instrument, a hole was found on the mcs tip cover accessory that was installed on the instrument. The tip cover was removed and replaced with a new tip cover. Two hours later, the same issue occurred. Since it was towards the end of the procedure, a new tip cover was not used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2012-00298 was submitted for the other tip cover accessory used during the surgical procedure. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt though the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The hole sizes range from. 010 diameter to. 050 long with varying shapes (round and oblong) and are located. 035 to. 165 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has various mechanical tears at the edge of one hole and in the general vicinity of the holes. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2663247 | 2633162 | 2529822 | 2955842201200297 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-05-28 | 2012-06-28 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-05-28 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-05-28 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgeon observed smoke coming from the tip of the monopolar curved (mcs) scissors instrument during energy activation. When the surgeon inspected the instrument, a hole was found on the mcs tip cover accessory that was installed on the instrument. The tip cover was removed and replaced with a new tip cover. Two hours later, the same issue occurred. Since it was towards the end of the procedure, a new tip cover was not used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2012-00297 was submitted for the other tip cover accessory used during the surgical procedure. | (b)(4). The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt though the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole size is approximately. 020 in diameter and located. 365 above the silicone to sleeve interface. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2663271 | 2633200 | 2529860 | 2955842201200298 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-05-28 | 2012-06-28 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-05-28 | Device Not Retu | Yes | Initial | 2012
40018012 | 2012-05-29 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the site was removing the monopolar curved scissors instrument with the mcs tip cover installed, from the cannula, the mcs tip cover accessory fell onto the patient's bowel. The tip cover accessory was removed laparoscopically and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | (b)(4). The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012 isi contacted surgical technologist, at (b)(6) and she indicated that it was noted that the tip cover fit loosely on the monopolar curved scissors instrument and was easily turned. A visual inspection of the cannula did not find any damage. (b)(6) also indicated that an inspection of the cannula using an obturator found that the cannula was round. No patient harm, adverse outcome or injury occurred. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2661646 | 2631532 | 2528228 | 2955842201200296 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-29 | 2012-06-27 | No | Yes | Health Professional | 40018012 | No | No | 2012-05-29 | No | Yes | Unkown | 2012
| 2012-06-26 | Malfunction |  | Upon use, it was noticed the da vinci forcep had tissue attaching to the operating end. Staff questioned if the forcep was dull and draggin against the tissue. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL BI | 2669887 | 2639796 | 2536452 | 2639796 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-06-27 | 2012-06-27 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 |  | Injury | Other | It was reported that a patient sustained a bowel injury while undergoing a da vinci s surgical procedure. | Intuitive surgical has requested additional information regarding this event from the hospital. No additional information has been provided to date, therefore, no conclusion can be drawn as to what may have caused or contributed to the patient's bowel injury. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2657284 | 2626773 | 2523469 | 2955842201200290 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2013-01-08 | 2012-06-22 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-05-23 | No | No | Reuse | 2012
42034301 | 2012-05-17 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the shaft of the thoracic grasper instrument bent, causing damage to the casing and the endowrist to break. It is unknown whether shavings from the instrument broke off and fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported failure mode. The grip cable was found broken at the endowrist. The main tube insulation appears to have scratches and gouge marks. Engineering concluded that the damage to the instrument was likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2657276 | 2626797 | 2523492 | 2955842201200292 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-24 | 2012-06-22 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-05-23 | Device Not Retu | No | Reuse | 2012
40018012 | 2012-05-24 | Malfunction |  | It was reported that during a da vinci si right colectomy procedure the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument came off of the instrument and fell into the patient after another instrument rubbed against it. The tip cover was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Since the tip cover accessory has not been returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2657282 | 2626803 | 2523498 | 2955842201200293 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-24 | 2012-06-22 | No | Yes | Health Professional | 40018012 | No | No | 2012-05-24 | No | Yes | Initial | 2012
IS2000 A51P8 |  | Injury | Other | It was reported that a patient sustained a bowel injury while undergoing a da vinci s surgical procedure. | Intuitive surgical has requested additional information regarding this event from the hospital. No additional information has been provided to date, therefore, no conclusion can be drawn as to what may have caused or contributed to the patient's bowel injury. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2657286 | 2626805 | 2523500 | 2955842201200291 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2013-01-08 | 2012-06-22 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-05-23 | No | No | Reuse | 2012
42018406 | 2012-05-23 | Malfunction |  | It was reported that while dissecting the lymph node between the lobes during a da vinci s lobectomy procedure, energy was observed leaking through the distal wires of the permanent cautery spatula instrument. The instrument was in use for 1. 5 hours when the leaking was noticed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed an arc track at the base of the spatula. It was determined that arcing likely occurred while the spatula was perpendicular to the distal clevis. Slightly melted material was also found on the yaw pulley. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2657285 | 2626807 | 2523503 | 2955842201200294 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-05-23 | 2012-06-22 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-05-23 | Device Not Retu | No | Reuse | 2012
42018306 | 2012-05-21 | Malfunction |  | It was reported that during a da vinci si lower anterior resection procedure, sparkling was observed on the wrist of the permanent cautery hook instrument when using electrocautery. The instrument was immediately removed from the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, an arc trac and melted materials were observed on the yaw pulley at the base of the hook. As a result of arcing, the instrument conductor wire was damaged and broken. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2657307 | 2626810 | 2523505 | 2955842201200295 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-05-25 | 2012-06-22 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-05-25 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-05-22 | Other | Other | It was reported that when starting a da vinci si prostatectomy procedure, the surgical staff experienced system error code 22005 and restarted the system. After restart, the system would not complete the homing process. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer concluded that the system error codes experienced by the customer were associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 22005 appears when the da vinci safety system determines the system did not self calibrate correctly during setup. Upon determining this condition, the system records the fault and transitions to a safe state. The mtmr was returned and evaluated. The error code experienced by the customer was confirmed in the system logs, however, engineering was unable to replicate the reported failure during internal testing. Based on the system error logs, it is believed that the mtmr arm may have been physically obstructed in the operating room. As of june 15, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2654039 | 2623363 | 2520059 | 2955842201200285 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-05-22 | 2012-06-21 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-05-22 | Device Not Retu | No | Reuse | 2012
42020505 | 2012-05-18 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, a cable broke on the fenestrated bipolar forceps instrument and a fragment fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- Despite multiple requests, the instrument has not been returned for evaluation. As a result, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, the pitch up cable was observed to be broken at the distal clevis hub and the pitch down cable is frayed at the clevis hub. The cable segment that contains the crimp is still installed in the clevis. The yaw pulley is also missing a. 175 x. 060 piece opposite of the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 2655198 | 2624879 | 2521574 | 2955842201200288 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2012-05-22 | 2012-06-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-07-16 | Device Not Retu | No |  | 2012
42020505 | 2012-05-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a thin strand was noticed coming from the wrist of the fenestrated bipolar forceps instrument. The thin strand appeared to be a frayed cable. Upon instrument removal, the frayed spot on the cable was evident, however, the strand was not. The strand was later observed on the shaft of the replacement instrument, where it appeared to have adhered to the inside of the instrument cannula and was reintroduced when the replacement instrument was inserted into the patient. The frayed cable was removed using a laparoscopic grasper. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable is frayed at the distal clevis hub and the frayed strands stick out at the wrist. Two dark colored strands, both approximately. 420 in length, were returned with the instrument. The detached strands are darker in color than the gray colored tungsten pitch down cable. Additionally, the detached strands measure approximately. 003 in diameter, compared to. 001 diameter for the individual tungsten strands of the pitch down cable. As a result, it cannot be determined if the detached strands came from the tungsten pitch down cable. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2655205 | 2624889 | 2521544 | 2955842201200287 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-22 | 2012-06-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-05-22 | Device Not Retu | No | Reuse | 2012
42030901 | 2012-05-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the tip on the mega suturecut needle driver instrument broke and fell into the patient's abdomen. The surgeon attempted to find the tip robotically, but the tip could not be located. An x-ray of the patient was taken. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Since the instrument will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. On june 20, 2012, (b)(6) hospital informed intuitive surgical that the patient has not developed any complications. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2655585 | 2625279 | 2521975 | 2955842201200289 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-22 | 2012-06-21 | No | Yes | Health Professional | 42030901 | No | No | 2012-05-22 | No | No | Invalid Data | 2012
420207 VER05 | 2012-06-20 | No Answer Provided |  | During a laparoscopic with robotic procedure the da vinci 5mm tenaculum forcep's pulley snapped making the instrument unusable. The surgeon switched to the da vinci cautery hook and 30 minutes into the procedure the hook collided with another instrument (pk dissecting forceps, da vinci also) and a piece of the plastic insulation seal of the monopolar cautery hook near the base of the shaft broke off and landed on the operative site. The small plastic piece was retrieved by the surgeon and a thorough inspection of the site was done and no other pieces remained in the patient. Da vinci rep was present the whole time the incident happened. Rep witnessed the whole thing. Manufacturer was notified of the incident. Products (tenaculum and hook) will be returned to them for analysis. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL TE | 2661825 | 2631711 | 2528407 | 2631711 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-06-20 | 2012-06-21 | No |  | Physician | 420207 VER05 | Yes | No Answer Provided |  |  |  |  | 2012
420207 VER05 | 2012-06-20 | No Answer Provided |  | During a laparoscopic with robotic procedure the da vinci 5mm tenaculum forcep's pulley snapped making the instrument unusable. The surgeon switched to the da vinci cautery hook and 30 minutes into the procedure the hook collided with another instrument (pk dissecting forceps, da vinci also) and a piece of the plastic insulation seal of the monopolar cautery hook near the base of the shaft broke off and landed on the operative site. The small plastic piece was retrieved by the surgeon and a thorough inspection of the site was done and no other pieces remained in the patient. Da vinci rep was present the whole time the incident happened. Rep witnessed the whole thing. Manufacturer was notified of the incident. Products (tenaculum and hook) will be returned to them for analysis. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL TE | 2661825 | 2631711 | 2528407 | 2631711 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-06-20 | 2012-06-21 | No |  | Physician | 420207 VER05 | Yes | No Answer Provided |  |  |  |  | 2012
42017910 | 2012-04-19 | Injury |  | On (b)(6) 2012, intuitive surgical received uf/importer report # (b)(4) from (b)(6) in (b)(6). The event details are provided below: "pt in operating room for robotic laparoscopic hysterectomy; surgeon documented grossly enlarged uterus with multiple fibroids. As surgeon was bivalving the uterus, he saw a flash of blood coming from the left side of the pt and immediately recognized vascular injury. The robot was immediately undocked and procedure was converted to open procedure. The vascular surgeon responded immediately to repair the pt's left external iliac artery, which he described as "charred" due to the cautery injury. Operating room staff reported small tear in the plastic that covered the metal end of the monopolar scissor. The pt was stabilized after surgical repair and administration of blood products and transferred to the intensive care unit post-operatively. She remains stable and was transferred to the surgical unit the following day. The pt was discharged home (b)(6) 2012". | The mcs instrument was returned and evaluated. Per the customer reported complaint (with the following clarification) " the defect was a tube extension dislodged from the main tube". The entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. The surface of the reinforcement tube extension exhibited burnt molded material, indicating unintended arcing may have occurred. Engineering concluded that the dislodged tube at the distal end is likely due to overloading at the distal end which led to a path for unintended arching. The mcs tip cover accessory was not returned. On (b)(4) 2012, isi contacted risk mgr, (b)(6) at (b)(6) and she indicated that the pt had a large uterus which required that the surgeon cut the pt's uterus in half to allow for removal. She indicated that the surgeon did not observe arcing from the instrument however; there was an excessive amount of smoke in the surgical area when the injury to the pt occurred. Since she stated that a vascular surgeon was immediately called to repair the affected vessel and during repair of the vessel, the vascular surgeon noted that it exhibited charring, thus it was determined that an arcing event had occurred. Per (b)(6), the surgeon did not observe that the mcs instrument broke during the surgical procedure, however, examination of the tip cover accessory by the surgical staff found that it had a hole; however, she does not know the status of the tip cover accessory. (b)(6) indicated that the pt is recovering well and has not returned to the hospital due to experiencing any post-operative complications. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2665332 | 2632716 | 2529412 | 2955842201200276 | 1 | NAY | Manufacturer | OtherHealth Professional | RISK MANAGER | Initial | 2012-05-22 | 2012-06-21 | Yes | Yes | Health Professional | 42017910 | Yes | Yes | 2012-05-22 | Yes | No | Reuse | 2012
| 2012-06-13 | Malfunction |  | Cautery machine made a noise while in use on da vinci procedure. The setting went from 60 to 165 on cut mode. Rn turned machine off and turned it back on again. Machine went through self test and rn was able to reset settings. Dr continued to work with cautery for appoximately 15-20 minutes. Flame was noticed at connection point of cautery cord and laparoscopic scissors. Cautery cord and lap scissors were removed from field and replaced. No harm to pt. No other incident occurred. Cautery machine was removed from service and checked by biomed. They stated that "the cautery cord had a cut in it, causing a visual arc. Unit was tested to mfr specifications and output was within tolerance. " monopolar scissors given to materials coordinator. Cord was also sequestered, but no numbers available. |  | FORCE FX ELECTROSURGICAL GENERATOR | ELECTROSURGICAL GENERATOR | 2661111 | 2630863 | 2527559 | 2630863 | 1 | GEI | User Facility |  |  | Initial | 2012-06-20 | 2012-06-20 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2012-05-23 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, the site experienced error message patient cart not connected, check red cable. With the assistance of isi technical support engineering (tse), the site checked the power cable and found no issues with the cable. The tse instructed the site to perform an emergency power off (epo) of the surgeon side cart (ssc) and the patient side cart (psc) and the issue was resolved. A few minutes later, the site contacted isi technical support to report that error message reoccurred. While the tse attempted to provide troubleshoot assistance the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the error message experienced by the site was associated with a defective fiber cable. The fiber cable connects the patient side cart to the surgeon console and passes the video, audio and data during system operation. The system was repaired by replacing the fiber cable. On (b)(4) 2012, the isi clinical sales representative (csr) indicated that the site reported that the aborted planned surgical procedure was a da vinci s prostatectomy with lymph node dissection. Per the csr, the site indicated that the rescheduled procedure was performed on (b)(6) 2012, and the patient was released from the hospital on (b)(6) 2012. As of (b)(4) 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2650698 | 2620044 | 2516740 | 2955842201200278 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-23 | 2012-06-18 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2012-05-23 | No | No | Unkown | 2012
42027804 | 2012-04-15 | Malfunction |  | On date may 25, 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: during a robotic assisted laparoscopic left renal cyst decortication and lysis of the intra-peritoneal adhesions, the davinci grasping forcep broke off into the patient. It was successfully removed without any harm. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | GRAPTOR RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2650699 | 2620059 | 2516755 | 2955842201200277 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2012-05-25 | 2012-06-18 | No | Yes | Health Professional | 42027804 | No | No | 2012-05-25 | No | No | Unkown | 2012
42000606 | 2012-04-10 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, a piece from the large needle driver instrument broke off and fell into the patient. The broken instrument piece was retrieved from the patient; however, it fell onto the or floor and was not located by the surgical staff. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one carbide insert missing from its grip. The grip faces do not make contact when the grips are closed and the brazing surface of the grip exhibits poor surface coverage of filler materials. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2650780 | 2620156 | 2516852 | 2955842201200273 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-29 | 2012-06-18 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2012-05-29 | Device Not Retu | No | Reuse | 2012
420179 | 2012-05-22 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissors (mcs) broke. The surgeon made the decision to complete the planned surgical procedure due to the patient's anatomy. No patient harm was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, the isi clinical sales representative (csr) present during the planned surgical procedure indicated that the monopolar curved scissors (mcs) instrument broke during the procedure, causing energy to escape the instrument; however, no patient harm or injury occurred. The csr indicated that the surgeon reported that the patient had many abdominal adhesions that made a continued attempt to proceed with the planned surgical procedure using minimally invasive techniques unsafe. Per the csr, the surgeon indicated that although the mcs instrument malfunction did influence his decision to convert the procedure to open, it was not the primary reason for the conversion and without the patient anatomical challenges, he would have moved forward with the procedure robotically. The surgeon reported to the csr that the patient was discharged from the hospital on (b)(6) 2012, is doing well and did not experience any post-operative complications. On june 15, 2012, isi contacted the (b)(6) concerning the mcs instrument and he indicated that he was not present during the surgical procedure when the instrument broke. (b)(6) indicated that he will follow up with the surgical staff to verify the status of the mcs instrument and that he will contact isi to advise if the instrument is available for return for evaluation by isi. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2650827 | 2620168 | 2516864 | 2955842201200284 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-22 | 2012-06-18 | Yes | Yes | Health Professional | 420179 | No | No | 2012-05-22 | Device Not Retu | No | Unkown | 2012
IS2000 |  | Other | Other | It was reported that in the (b)(6) 2008, a patient underwent a da vinci s mitral valve repair procedure. The recovery was quick and easy, however, the patient alleges that the mitral valve repair was not successful and he still has a heart murmur and regurgitation of blood through the valve. | Despite multiple requests, the surgeon who performed the patient's da vinci surgical procedure has not provided intuitive surgical with any additional information regarding this event. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2650845 | 2620217 | 2516913 | 2955842201200283 | 1 | NAY | Manufacturer | Consumer | Other | Initial | 2012-05-20 | 2012-06-18 | Yes | Yes | Health Professional | IS2000 | No | No | 2012-05-20 | No | No | Reuse | 2012
40018012 | 2012-05-18 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument came off of the instrument and fell into the patient's pelvis. The tip cover was retrieved with a laparoscopic grasper and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. The tip cover was installed on an in-house mcs instrument with no issues observed. The tip cover fit properly on the instrument and did not slide off. The customer reported failure could not be replicated. It was also observed that the distal end of the tip cover has a. 115 long tear through the silicone part. Additionally, cuts were found on the sleeve part of the pellethane insert, which do not go through the thickness of the pellethane. Cuts were also found on the silicone overmold, near the sleeve shoulder, and do not go through the combined silicone/pellethane wall thickness. No other damage was found. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2648980 | 2618548 | 2515244 | 2955842201200279 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-18 | 2012-06-17 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-05-18 | Device Not Retu | Yes | Initial | 2012
42018306 | 2012-05-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the permanent cautery hook (pch) instrument sparked and started to emit smoke from the tip of the instrument. The planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed car marks and charring on the yaw pulley. Additionally, engineering observed one grip close cable is derailed at the distal idler pulley. Grips can still open and close, but movement may not be precise. Cable derailment is likely due to cable losing contact with pulley during wrist articulation. Fleet angle between proximal and distal pulleys may contribute to derailment. The banana plug is also bent sideways. The instrument passed electrical testing. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2649062 | 2618630 | 2515326 | 2955842201200280 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2012-05-18 | 2012-06-17 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-05-18 | Device Not Retu | No | Unkown | 2012
IS2000 A51P7 | 2012-05-17 | Other | Other | It was reported that prior to starting a da vinci s beating heart revascularization procedure, the site experienced a recoverable fault led to light up. Isi technical support engineering (tse) checked the system log and found error message 45310. With the assistance of isi technical support engineering (tse), the error message was cleared and the issue was resolved. The surgeon then reported having scope fogging issues and made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation conducted by the field service engineer could not replicated the reported failure mode. The field service engineer viewed and inspected the endoscope and found it to be working as intended. Error code 45310 occurs when the display coordinator (dc) detects a loss of primary video inputs. The site's system was tested and performed to specifications. The site informed that the surgeon was very impatient and wanted to convert to open procedure without trying a different endoscope. As of june 15, 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2649118 | 2618688 | 2515384 | 2955842201200281 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-17 | 2012-06-17 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2012-05-17 | Device Not Retu | No | Reuse | 2012
420172 | 2012-04-06 | Malfunction |  | It was reported that during da vinci si surgical procedure, while using the maryland bipolar forceps instrument, a piece broke off and fell into the patient. No further information was provided concerning if the piece was retrieved. No patient injury, harm or adverse event was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to gain additional information about the event, but the hospital has yet to provide further details. If additional information is received, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2649194 | 2618761 | 2515457 | 2955842201200282 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-18 | 2012-06-17 | No | Yes | Health Professional | 420172 | No | No | 2012-05-18 | No | No | Unkown | 2012
420227 | 2012-05-23 | Malfunction |  | During procedure, the pk cable started to break. No apparent harm to pt. |  | DAVINCI S | PK DISSECTING FORCEPS | 2661182 | 2629701 | 2526397 | 2629701 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2012-06-08 | 2012-06-16 | No | Yes | Health Professional | 420227 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2012
IS1200 A43P9 | 2012-05-16 | Other | Other | Event Description: --/-- It was reported that prior to using the system to perform any tasks for a da vinci prostatectomy procedure, the surgeon experienced no video signal at the surgeon console. The connections were reconnected and then the system and synchronizer were restarted, however, the issue continued to recur. The patient was under anesthesia and the port incisions had been made when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was related to the synchronizer. The synchronizer coordinates the left eye and right eye video streams so that events are displayed simultaneously to each of the surgeon's eyes. The system was repaired by replacing the affected synchronizer. As of june 14, 2012, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- The synchronizer was returned to the original equipment manufacturer (oem) for evaluation. The oem was unable to replicate the customer reported failure during testing and found the synchronizer to be fully functional and working to specification. | DA VINCI SURGICAL SYSTEM | Reuse | 2648798 | 2618337 | 2515033 | 2955842201200274 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2012-05-16 | 2012-06-15 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2012-06-14 | Device Not Retu | No |  | 2012
IS3000 A60P6 |  | Death | Death | It was reported that a patient expired after undergoing a da vinci si prostatectomy procedure. | The hospital was contacted for additional information, however, the hospital informed intuitive surgical that they are unwilling to release any details regarding this event. As a result, no conclusion can be drawn by intuitive surgical regarding the cause of the patient's death. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2648803 | 2618353 | 2515049 | 2955842201200275 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-18 | 2012-06-15 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-05-18 | No | No | Reuse | 2012
400180 |  | Malfunction |  | It was reported that during da vinci si procedure, arcing from a hole in the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument was observed. The instrument was disposed of. No further details were provided. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to gain additional information about the event, but the hospital has yet to provide further details. If additional event details become available a follow up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2647170 | 2617061 | 2513757 | 2955842201200271 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-05-15 | 2012-06-14 | No | Yes | Health Professional | 400180 | No | No | 2012-05-15 | No | Yes | Initial | 2012
IS3000 A60P6 | 2012-03-06 | Other | Other | On (b)(6) 2012, a patient underwent a da vinci si surgical procedure on his stomach at (b)(6) medical center, which reportedly took 6. 5 hours to complete. The patient also had his hiatal hernia repaired. The patient alleges that post surgery, he continues to have pain in his stomach and issues with keeping his food down. Late submission of this report is due to a data entry oversight by customer service. | The surgeon who performed the patient's da vinci si surgical procedure indicated that the system did not malfunction during the case. Despite multiple requests, the surgeon has not provided intuitive surgical with any more information. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2645101 | 2615077 | 2511773 | 2955842201200270 | 1 | NAY | Manufacturer | Consumer | Other | Initial | 2012-05-12 | 2012-06-13 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-05-12 | No | No | Reuse | 2012
400180 | 2012-05-30 | Malfunction |  | Multiple (5) version 12 tip covers did not adequately form to the tip of the monopolar hot shears, resulting in a loose fit and/or loss of integrity. In one case, the tip fell off into the patient's abdomen, but was retrieved. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL TI | 2652429 | 2621805 | 2518500 | 2621805 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-06-07 | 2012-06-13 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2012
40018012 | 2012-05-09 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument came off of the instrument and fell inside the patient. The tip cover was retrieved and the planned surgical procedure was completed using a replacement tip cover of the same type. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Visual inspection shows no damage to the sleeve. A fit check was performed using an in-house mcs instrument and installation tool. The tip cover was able to be fully installed on the instrument without any issues observed. The tip cover did not slip off when the instrument was positioned with the tip pointing vertically downward. Removal tests were performed using a force gage and it was determined that the average removal force required to remove th tip cover with pure axial loading (no twisting) was 17. 0 lbs. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2634281 | 2604749 | 2501436 | 2955842201200265 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-09 | 2012-06-08 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-05-09 | Device Not Retu | Yes | Initial | 2012
42001105 | 2012-04-06 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the cannula bent, causing tiny shavings from the thoracic grasper instrument to break off and fall into the patient. The fragments were not retrieved. Tiny shavings from a 2nd thoracic grasper instrument used during the procedure also broke off and fell into the patient. The fragments were not retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Mdr 2955842-2012-00260 was submitted for the lst thoracic grasper instrument and mdr 2955842-2012-00261 was submitted for the 2nd thoracic grasper instrument. Late submission of this report is due to a data entry oversight that occurred after receipt of the complaint information by customer service. | The cannula has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the it is returned (post engineering evaluation). On may 15, 2012, isi contacted the (b)(6) medical center and she indicated that the initial reporter of this event is on leave and that she was not aware of the reported event and that she will have to perform an investigation into the reported event and she follow up with isi with additional information upon completion of her investigation. On may 17, 2012 isi contacted the surgeon assisting at the patient side cart and he indicated that he was not aware that any shavings from the thoracic grasper instruments used during the surgical procedure fell into the patient and he does not recall the number of times the thoracic grasper instrument(s) were changed out during the surgical procedure, however, the planned surgical procedure was completed and there was no patient harm, adverse outcome or injury. On may 24, 2012, isi followed up with (b)(6) and she indicated that her investigation into the reported event is still in progress, however, there was no report by the surgical staff concerning the reported event and that once her investigation was complete she would provide additional information. Isi has made several attempts to follow -up with (b)(6) concerning the status of her investigation, however, however, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. | 5MM CANNULA | ENDOSCOPIC ACCESSORY | 2634306 | 2604751 | 2501438 | 2955842201200259 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-06 | 2012-06-08 | No | Yes | Health Professional | 42001105 | No | Yes | 2012-04-06 | No | No | Unkown | 2012
42034301 | 2012-04-06 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the cannula bent, causing tiny shavings from the thoracic grasper instrument to break off and fall into the patient. The fragments were not retrieved. Tiny shavings from a 2nd thoracic grasper instrument used during the procedure also broke off and fell into the patient. The fragments were not retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Mdr 2955842-2012-00259 was submitted for the cannula and mdr 2955842-2012-00260 was submitted for the 1st thoracic grasper instrument. Late submission of this report is due to a data entry oversight that occurred after receipt of the complaint information by customer service. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported failure mode. The main tube insulation appears to have scratches and gouge marks. Engineering concluded that the damage to the instrument was likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On may 15, 2012, isi contacted (b)(6) medical center and she indicated that the initial reporter of this event is on leave and that she was not aware of the reported event and that she will have to perform an investigation into the reported event and she follow up with isi with additional information upon completion of her investigation. On may 17, 2012 isi contacted the surgeon assisting at the patient side cart and he indicated that he was not aware that any shavings from the thoracic grasper instruments used during the surgical procedure fell into the patient and he does not recall the number of times the thoracic grasper instrument (s) were changed out during the surgical procedure, however, the planned surgical procedure was completed and there was no patient harm, adverse outcome or injury. On may 24, 2012, isi followed up with (b)(6) and she indicated that her investigation into the reported event is still in progress, however, there was no report by the surgical staff concerning the reported event and that once her investigation was complete she would provide additional information. Isi has made several attempts to follow -up with (b)(6) concerning the status of her investigation, however, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2634308 | 2604755 | 2501443 | 2955842201200261 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-04-06 | 2012-06-08 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-04-06 | Device Not Retu | No | Reuse | 2012
42034301 | 2012-04-06 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the cannula bent, causing tiny shavings from the thoracic grasper instrument to break off and fall into the patient. The fragments were not retrieved. Tiny shavings from a 2nd thoracic grasper instrument used during the procedure also broke off and fell into the patient. The fragments were not retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Mdr 2955842-2012-00259 was submitted for the cannula and mdr 2955842-2012-00261 was submitted for the 2nd thoracic grasper instrument. Late submission of this report is due to a data entry oversight that occurred after receipt of the complaint information by customer service. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported failure mode. The black tube insulation is peeling, however, no metal material is exposed. Engineering concluded that the damage is likely due to misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings; handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On may 15, 2012, isi contacted (b)(6) medical center and she indicated that the initial reporter of this event is on leave and that she was not aware of the reported event and that she will have to perform an investigation into the reported event and she follow up with isi with additional information upon completion of her investigation. On may 17, 2012 isi contacted the surgeon assisting at the patient side cart and he indicated that he was not aware that any shavings from the thoracic grasper instruments used during the surgical procedure fell into the patient and he does not recall the number of times the thoracic grasper instrument (s) were changed out during the surgical procedure, however, the planned surgical procedure was completed and there was no patient harm, adverse outcome or injury. On may 24, 2012, isi followed up with (b)(6) and she indicated that her investigation into the reported event is still in progress, however, there was no report by the surgical staff concerning the reported event and that once her investigation was complete she would provide additional information. Isi has made several attempts to follow -up with (b)(6) concerning the status of her investigation, however, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2634289 | 2604756 | 2501443 | 2955842201200260 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-04-06 | 2012-06-08 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-04-06 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-05-14 | Other | Other | It was reported that during a da vinci si lobectomy procedure the patient side cart (psc) power transitioned to the back-up battery power. With the assistance of isi technical support engineering (tse) the site replugged the psc into 3 different outlets, however, the issue persisted. Unable to resolve the issue the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the power issue experienced by the customer was associated with the patient side cart medical grade power supply (mgps), power cord and the battery pack module. The mgps converts the ac power from the wall socket (through the power cord) and converts it to dc power for use on the system. The battery pack module provides a backup source of dc power to expedite the undocking of the system in situations where ac power is lost. The battery pack module was returned to the original equipment manufacturer (oem) for evaluation. The oem observed that gas gage board required replacement. Isi contacted the initial reporter on june 4, 2012 and she reported that conversion of the planned surgical procedure was due to the battery issue. She indicated that there was no harm to the patient due to the system issue and the patient tolerated the open procedure well and that it is unknown if the patient has returned to the hospital due to experiencing any post operative complications. No other information was provided. As of june 8, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2634326 | 2604758 | 2501445 | 2955842201200262 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-14 | 2012-06-08 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-05-14 | Device Not Retu | No | Reuse | 2012
42018406 | 2012-05-11 | Malfunction |  | It was reported that during a da vinci si pulmonary wedge resection procedure, the insulation on the tip of the permanent cautery spatula instrument broke and fell into the patient. The fragment was retrieved, however, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed. The distal end of the ceramic sleeve is broken off. The broken pieces were not returned. The intact part of sleeve exhibits light scratching on one side. When the instrument was installed on an isi in-house system, the system showed that the instrument was not used. No other damage was found. Engineering concluded that damage to the instrument is likely due to mishandling/misuse. However, the site may have reported and returned the incorrect instrument, as the remaining instrument lives show that the instrument was not used. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On may 16, 2012, or technician, (b)(6) indicated that the broken instrument piece was retrieved from the patient laparoscopically; however, she indicated that conversion of the planned surgical procedure to open was because the lobe being removed was too large to remove laparoscopically. (b)(6) indicated that to the best of her knowledge, the patient tolerated the open procedure well. No other details concerning the patient were provided. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2634336 | 2604785 | 2501472 | 2955842201200264 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-05-11 | 2012-06-08 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-05-11 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-04-25 | Other | Other | It was reported that post-op a successful da vinci s mitral valve repair procedure, performed on (b)(6) 2012, the patient experienced mild dysphagia, dysarthria and a strange tongue. A ct scan performed found that the patient developed a right brain infarction and was treated with the drug radicut. Examination of the patient by a neurological physician confirmed that the patient experienced a brain infarction, which resulted in the patient developing right side tongue weakness. The patient side surgeon has reported that there was no relationship between the reported event and the da vinci s system, instrument and/or accessories, as there was no malfunction of the da vinci s system, instrument and/or accessories during the surgical procedure. The site has indicated that the brain infarction experienced by the patient is a potential complication of intracardiac surgery with arrested heart. | On (b)(4) 2012, the patient side surgeon reported that the patient was discharged from the hospital on (b)(6) 2012, and at the time of discharge, the dysphagia and tongue deviation was not observed and the site has concluded that the patient's symptoms are in remission. The patient side surgeon commented that from the preoperative, postoperative cerebral infarction, and it can happen as a complication of surgery performed under cardiac arrest, but not limited to robotic surgery. The surgeon also indicated that the cause of the brain infarction is unknown. No other information concerning the patient or reported event was provided. As of june 8, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2634342 | 2604786 | 2501473 | 2955842201200267 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-11 | 2012-06-08 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-05-11 | No | No | Unkown | 2012
40018012 | 2012-05-10 | Malfunction |  | It was reported that during a da vinci si endometriosis resection procedure the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument came off of the instrument and fell inside the patient. The tip cover was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Since the tip cover accessory will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2634363 | 2604802 | 2501495 | 2955842201200266 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-05-10 | 2012-06-08 | No | Yes | Health Professional | 40018012 | No | No | 2012-05-10 | No | Yes | Initial | 2012
42034301 | 2012-05-04 | Malfunction |  | It was reported that during a da vinci si mediastinal tumor resection procedure, tiny shavings from the thoracic grasper instrument broke off of the instrument. It is not known whether any shavings fell into the patient. No further information was provided. | Since the instrument has not be returned for evaluation, the root cause of the customer reported failure mode cannot be determined. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received or if the instrument is returned for evaluation. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2634365 | 2604803 | 2501496 | 2955842201200268 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-05-10 | 2012-06-08 | No | Yes | Health Professional | 42034301 | No | No | 2012-05-10 | No | No | Unkown | 2012
IS3000 A60P6 | 2012-05-18 | Other | Other | It was reported that during a da vinci si partial nephrectomy procedure, while the surgeon was grasping a renal artery, the vision side cart (vsc) shut down. The site contacted isi technical support engineering (tse) for trouble-shooting assistance and found that the vsc was accidentally unplugged from the outlet, however, prior to resolution of the issue and due to time constraints the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation performed by isi technical support engineering determined that the issue experienced by the site was related to a loss of power to the vision side cart (vsc), as the site confirmed that one of the operating room staff accidentally tripped over the powercord, causing the powercord to unplug from the wall outlet. The vsc houses the system's central processing and vision equipment. Power was restored to the site's vsc after re-plugging the powercord into the wall outlet. On june 7, 2012 isi contacted the initial reporter at (b)(6) and she indicated that there was no harm to the patient as a result of the reported event and the patient tolerated the open procedure well. As of june 8, 2012 there have been no other report recurrence of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2634407 | 2604809 | 2501502 | 2955842201200263 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2012-05-18 | 2012-06-08 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-05-18 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-05-15 | Other | Other | It was reported that during a da vinci s ureterectomy procedure, the site experienced system error code 25513. With the assistance of an isi technical support engineer (tse) the site performed multiple emergency power offs of the system, however, the issue persisted. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation performed by isi field service engineering determined that the system error code 25513 experienced by the site was associated with a remote arm controller (rac)pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25513 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put davinci in a recoverable safe state. On (b)(4) 2012, the initial reporter (surgeon) at (b)(6) indicated that while he was performing dissection using the da vinci s surgical system a vesicovaginal fistula occurred. The surgeon indicated that his intention was to repair the fistula after removal of the patient's tumor, however, when the system malfunction occurred, he no longer had sufficient exposure to adequately close the fistula. The surgeon indicated that it is his belief that the fistula likely occurred due to the aggressiveness of the patient's cancer and adhesions from prior surgery. Per the surgeon, the patient was discharged 4 days post op and did experience post operative complications as a result of the vesicovaginal fistula. A foley catheter was placed to treat the patient. No other information concerning the patient was provided. The remote arm controller (rac) pca board was returned to isi for failure analysis investigation. Engineering found that when the rac was installed on an isi in-house system the rac failed at start up. As of june 8, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2636316 | 2606579 | 2503275 | 2955842201200269 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-05-15 | 2012-06-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-05-15 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-02-28 | Injury | Other | It was reported that when removing the trocar sleeves from the patient's abdomen at the end of a da vinci si hysterectomy procedure, one trocar site was noted to have a burn area at the edge of the incision. Afterwards, the surgical staff observed that the monopolar curved scissors (mcs) tip cover accessory (used with the mcs instrument during the procedure) had a hole in it. No additional patient harm was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 190 x. 065 oval shaped hole in the silicone. The hole center is located. 595 above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicating that an arcing event occurred. It was determined that the arcing event may have caused the patient port site burn if the area of the breached insulation was up inside or close enough to the cannula tip, allowing stray current to transfer. The mcs instrument was also returned and evaluated. The instrument was visually inspected and did not exhibit char marks or any burrs or sharp edges that would damage the tip cover. No evidence was observed on the instrument that would be a main contributor to an arcing event. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2633103 | 2603694 | 2406798 | 2955842201200251 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-08 | 2012-06-07 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-05-08 | Device Not Retu | Yes | Initial | 2012
IS2000 A51P8 | 2012-05-08 | Other | Other | It was reported that prior to starting a da vinci s prostectomy surgical procedure, the site could not power on the system. With the assistance of an isi technical support engineer the site applied emergency power off (epo) before powering on the system. The site attempted to use different power outlets and a different power cord, however the system would not power on. The patient was under anesthesia with no port incisions made, when the surgeon decided to abort the planned surgical procedure. | The investigation conducted by field service engineer concluded that system fault experienced by the customer was associated with the power supply. The power supply was replaced. As of june 7, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2633296 | 2603898 | 2500605 | 2955842201200252 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-05-08 | 2012-06-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-05-08 | Device Not Retu | No | Reuse | 2012
IS2000 | 2010-01-14 | Other | Hospitalization | On may 8, 2012, intuitive surgical received a legal complaint that alleged the following: on (b)(6) 2010, a patient underwent a da vinci s hysterectomy procedure with a left salpingo-oophorectomy at (b)(6). There were no records of any complications during the surgery and the patient was discharged on (b)(6) 2010. On (b)(6) 2010, the patient experienced severe pain in her lower abdomen and fever and was subsequently re-admitted to (b)(6). The patient was immediately started on triple intravenous antibiotic therapy including rocephin, gentamicin, and fagyl. A diagnosis of pelvic abscess along the vaginal cuff was made. Secondary to this, the patient had radiology tests done, including a ct scan of the abdomen and pelvis with oral and intravenous contrast, that confirmed a large mass in the left lower quadrant and pelvis having inflammatory-type features and containing air most suggestive of an abscess. The mass was measured to be 7. 3 x 5. 1 centimeters in size. The reading radiologist also noted that this could also be a necrotic type mass which communicates with bowel. The patient was noted to have an elevated white blood cell count with a left shift consistent with sepsis. On (b)(6) 2010, the decision was made to proceed with surgery to evaluate the abscess. Finding at surgery confirmed a large pelvic abscess along the vaginal cuff. A large amount of pus was encountered on manipulating the bowel over the vaginal cuff. Another surgeon was called in to evaluate the colon and it was felt that there was no discrete bowel injury. A drainage tube (jp drain) was then placed after a copious amount of irrigation with a bacitracin solution. Following the (b)(6) 2010, surgery, the patient developed a post-operative ileus. Her cultures revealed a positive e. Coli bacterial infection sensitive to levaquin. The patient was discharged on (b)(6) 2010, on oral antibiotics. The patient continued to have abdominal and pelvic pain and bowel discomfort and was referred for a follow-up ct scan on (b)(6) 2010. The study showed a peripherally enhancing fluid collection which measured approximately 3. 6 cm. On (b)(6) 2010, due to persisting lower abdominal pain, the patient had a third ct scan done which revealed a 2. 9 cm peripherally enhancing fluid collection superior to the bladder. On (b)(6) 2010, due to persistent left lower quadrant abdominal pelvic pain and bloating, the patient had an additional ct scan of the abdomen and pelvis performed which revealed residual inflammatory scarring along the vaginal cuff. The patient further alleges that due to the injuries sustained to her vaginal cuff and bowel during the da vinci s hysterectomy procedure, the patient had to have multiple additional medical tests and physician consultations and has suffered pain, loss of function, emotional distress, and permanent injury. | Multiple attempts have been made to obtain additional information regarding this event from the surgeon who performed the da vinci s hysterectomy with a left salpingo-oophorectomy procedure. To date no response has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2633290 | 2603935 | 2500628 | 2955842201200253 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-05-08 | 2012-06-07 | Yes | No | Health Professional | IS2000 | No | No | 2012-05-08 | No | No | Reuse | 2012
42018406 | 2012-05-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing occurred with the permanent cautery spatula instrument. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a section of wire with damaged insulation and charring. The charring indicates that an arcing event occurred. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2633300 | 2603945 | 2500638 | 2955842201200257 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-05-14 | 2012-06-07 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-05-14 | Device Not Retu | No | Reuse | 2012
42034403 | 2012-05-09 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure the curved bipolar dissector instrument was not cauterizing solely within the area of the tips, but throughout the whole area from the metal tips to the metal connection on the instrument shaft. Upon noticing cautery in the wristed area, the surgeon immediately removed the instrument from the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both yaw pulleys exhibit charring and localized melting at the grip base, between the grips. The charring indicates that an arcing event occurred. No other damage was found. | CURVED BIPOLAR DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2633301 | 2603946 | 2500639 | 2955842201200254 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-09 | 2012-06-07 | No | Yes | Health Professional | 42034403 | Device Returned To Manufacturer | Yes | 2012-05-09 | Device Not Retu | No | Reuse | 2012
42034403 | 2012-05-09 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure the curved bipolar dissector instrument was not cauterizing within the area of the tips, but sparks were occurring at the neck of the instrument and the energy isolation material appeared to have melted. Upon noticing the sparks, the surgeon immediately removed the instrument from the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both yaw pulleys exhibit charring and localized melting at the grip base, between the grips. The charring indicates that an arcing event occurred. No other damage was found. | CURVED BIPOLAR DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2633241 | 2603947 | 2500639 | 2955842201200255 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-09 | 2012-06-07 | No | Yes | Health Professional | 42034403 | Device Returned To Manufacturer | Yes | 2012-05-09 | Device Not Retu | No | Reuse | 2012
42018306 | 2012-05-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the permanent cautery hook instrument began smoking and then caught on fire shortly after use. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. It was later observed that part of the instrument was missing, presumably melted or burnt off. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. The initial reporter has been contacted multiple times to request additional information regarding the product problem, however, no response has been received to date. A follow-up medwatch report will be submitted if the device is returned or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2633304 | 2603949 | 2500642 | 2955842201200258 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-05-10 | 2012-06-07 | No | Yes | Health Professional | 42018306 | No | No | 2012-05-10 | No | No | Unkown | 2012
42020702 | 2012-05-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, black fragments from the tenaculum forceps instrument fell inside the patient. All fragments were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation despite multiple requests, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned with the entire grip assembly and clevis broken off at the main tube joint. The proximal clevis and distal clevis were not returned and the main tube was found with fiber missing at the clevis connection joint. It was determined that the instrument damage is likely due to excess force applied to the clevis and main tube. | TENACULUM FORCEPS INSTRUMENT | Reuse | 2633242 | 2603960 | 2500607 | 2955842201200256 | 1 | GEI | Manufacturer | User facility | Other | InitialFollowup | 2012-05-11 | 2012-06-07 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | No | 2012-06-28 | Device Not Retu | No |  | 2012
IS3000 A60P6 |  | Death | Death | It was reported that after a da vinci partial nephrectomy was completed, it was determined that the patient required another procedure. The patient then underwent an open procedure, for a total nephrectomy. Afterwards, the patient was transferred to (b)(6), where the patient expired on an unknown date. The delay in reporting is due to an administrative oversight where the aware date was inadvertently not updated by a customer service representative. | Several attempts have been made to gather additional information regarding this event from this site without success. On (b)(6) 2012, (b)(6) informed intuitive surgical that no further information about this case would be provided. If additional event details become available a follow up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2629774 | 2600178 | 2496885 | 2955842201200249 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-04-12 | 2012-06-06 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-04-12 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-05-08 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, when the surgeon articulated the master tool manipulators (mtms) the patient side manipulator (psm) arms would freeze up. With the assistance of isi technical support engineering (tse) the site rebooted the system and the issue was resolved. However, due to patient anatomical issues and the port placements the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by isi field service engineering (fse) was unable to replicate the issue experienced by the surgeon. The fse performed multiple test drives of the site's da vinci si system and found that it functioned within specification. Upon follow up with the surgeon by the fse, the surgeon indicated that the master tool manipulators (mtms) were locking up in place and when he clutched the mtm grips, the mtms unlocked. The surgeon indicated to the fse that he does not recall if any icons were present when the issue occurred. The fse concluded that the freezing up of the patient side manipulators (psm) experienced by the surgeon was likely due to the surgeon removing his head from the high resolution stereo viewer while manipulating the mtms. The hrsv provides the video image for the ssc. The fse advised the surgeon to keep his head in the hrsv when manipulating the mtms to prevent the psms from freezing up during da vinci surgical procedures. The da vinci si user manual specifically states: caution: the infrared head sensors performs a safety function by preventing movement of the patient cart arms when the surgeon's head is not in the viewer. Do not defeat this safety feature by intentionally blocking the sensors. On (b)(4) 2012, the isi representative present during the surgical procedure indicated that it is his belief that the planned surgical procedure could have been completed using the da vinci si system and that the surgeon's decision to complete the procedure using open surgical techniques was likely due to difficulty intubating the patient prior to starting the procedure, patient anatomical issues and the surgeon's experience using the da vinci si surgical system. Isi has made multiple attempts to obtain additional information from the surgeon concerning the patient and reported event. No information has been forthcoming. If additional information is received, a follow-up mdr will be submitted to the fda. As of june 6, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2631707 | 2602287 | 2498974 | 2955842201200250 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-08 | 2012-06-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-05-08 | Device Not Retu | No | Reuse | 2012
41029902 | 2012-04-27 | Malfunction |  | On (b)(6) 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda, which contained the following information: event desc: end of instrument fell off in the patient. It was retrieved from the patient. What was the original intended procedure? laparoscpic robotic bowel resection this product problem was initially reported to intuitive surgical by the user facility on (b)(6) 2012. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal tip had detached from the snake wrist. The distal tip was also returned, with the pebax tubing still attached. The tip had dislodged from the four cable crimps at the distal end of the snake wrist. The snake wrist discs are not dislodged. The base of the crimps appear to have adhesive residue. A couple of crimp holes in the distal tip exhibit cracking near the edge of the hole. | SUCTION IRRIGATOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2625140 | 2595934 | 2492627 | 2955842201200248 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-05-02 | 2012-06-01 | No | Yes | Health Professional | 41029902 | Device Returned To Manufacturer | No | 2012-05-02 | Device Not Retu | Yes | Initial | 2012
IS3000 A60P6 | 2012-04-03 | Other | Other | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the illuminator shut off multiple times. Unable to resolve the issue and due to other surgical reasons, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. This report is late due to a data entry error by the distributor. | Manufacturer Narrative: --/-- On (b)(6) 2012, intuitive surgical received additional information from the (b)(6) and he indicated that the surgeon reported that the excessive bleeding experienced by the patient was due the patient's advanced stage of cancer (t3). Per (b)(6), the surgeon indicated that the patient's bleeding was located at the prostate and rectum upper wall and that he was able to resolve the issue by performing coagulation and suturing of the patient's rectum. Per (b)(6), it is the surgeon's belief that the bleeding experienced by the patient was unrelated to the vision issue but rather to patient anatomical issues. Per (b)(6), the surgeon indicated that the open surgical procedure was completed successfully and the patient is healing fine. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by field service engineering concluded that the issue experienced by the customer was associated with the illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. The illuminator was returned to isi for failure analysis investigation. Engineering evaluation was unable to replicate the issue experienced by the customer. Engineering concluded that an obstruction of air intake at the site likely caused the thermal switch to trip, causing the illuminator to shut off. On may 30, 2012, isi received additional information from the initial reporter of this incident indicating that the planned surgical procedure was converted to open due to uncontrolled bleeding experienced by the patient that occurred as a result of an injury that the patient sustained. The initial reporter indicated that there was no harm or adverse outcome to the patient due to conversion of the surgical procedure to open and the patient was released from the hospital on (b)(6) 2012, and the patient did not experience any post operative complication. Isi has made several requests for additional information concerning the injury sustained by the patient; however, no additional information has been provided. Should additional information be provided, a follow-up medwatch report will be submitted to the fda. As of may 31, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2624305 | 2595109 | 2491823 | 2955842201200247 | 1 | NAY | Manufacturer | Distributor | Other | InitialFollowup | 2012-06-08 | 2012-05-31 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-06-08 | Device Not Retu | No |  | 2012
42017910 | 2012-04-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, a char was observed on the monopolar curved scissors instrument when it was removed from the patient. The distal tip of the instrument was also observed to be melted and have a crack. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- Since the instrument has not been returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012, the following medwatch report was received from (b)(6): at approximately 0900 circulator and scrub noticed an electrical burning odor while the surgeon was using monopolar scissor 311 and notified surgeon. Assistant removed trocar with arm attached and discovered scissor arm damaged and melted. Case converted to laparoscopic. Tissue on uterus charred. Case was a robotic assist hyst. Product will remain with hospital risk management pending further review. On (b)(6) 2012, (b)(6) in risk management at (b)(6) indicated that during the da vinci s hysterectomy procedure, while the surgeon was dissecting the patient's vaginal cuff using the monopolar curved scissors (mcs) instrument arcing to the patient's uterus occurred, causing charring to the patient's uterus, which was being removed. (b)(6) indicated that to the best of her knowledge, the surgical team did not experience any issues with the da vinci s system prior to the arcing event; however, the planned surgical procedure was completed using laparoscopic techniques due to the damaged condition of the mcs instrument and it was unknown at the time of the event if any harm or injury to the patient occurred. (b)(6) also provided the following additional details: the monopolar curved scissors (mcs) instrument and mcs tip cover were inspected prior to use and there was no damage found. The monopolar curved scissors instrument made contact with another instrument used during the surgical procedure. The cannula was inspected prior to use. The esu generator used during the surgical procedure was a valley lab, force, fx and the settings were coag 90, cut 90. The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube was found broken located slightly below the proximal clevis. The broken main tube can be rotated 360 degrees; however, it still remains intact with the instrument. The broken area shows signs of arcing due to localized melting. Engineering concluded that the damage was likely due to excessive side loading or other type of misuse. The instrument also passed electrical continuity testing. Engineering evaluation of the mcs instrument also found black char marks on the blade tips. The surface of the blades exhibit burnt material, indicative that unintended arcing occurred. Tube reinforcement ring also shows peeling off of surface material. On (b)(6) 2012 the site's da vinci s system was inspected by an isi field service engineer (fse) and found no issues. Functional testing of the system found that it functioned within specification. Isi's regional fse manager also met with (b)(6) and members of the surgical staff present during the surgical procedure and they indicated that the patient involved had a tight anatomy and that the surgical staff observed that the mcs instrument exhibited excessive torquing. Based on the additional information provided by the site, the regional fse manager has concluded that excessive torquing of the mcs instrument caused the instrument to break, which allowed energy to escape and arc to the patient's uterus, which was being removed. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012 intuitive surgical received the returned mcs tip cover accessory used in conjunction with the monopolar curved scissors (mcs) instrument. Engineering evaluation found that the tip cover has a. 030 x. 050 oval shaped hole in the silicone. The hole is located. 500 above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. With tip cover installed on mcs instrument the hole position is near grip cable and distal clevis ear. Intuitive surgical has submitted manufacturer report 2955842-2012-00312 concerning the tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2622311 | 2592853 | 2489546 | 2955842201200245 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2012-06-04 | 2012-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-06-04 | Device Not Retu | No |  | 2012
| 2012-05-29 | Other | Other | Davinci robotic instruments are difficult to clean devices. When i contacted intuitive surgical to clarify cleaning and decontamination instructions, i was given erroneous information by several people. An infection control nurse called in attempt to get clarification and was told incorrect information. Finally, when speaking with a sterilization specialist. The comment was made that they had some training issues and he had 395 employees and couldn't train them all. I am a member of (b)(6). I know what needs to be done. Everyone says follow the manufacturer's ifu. But the manufacturer is not giving out accurate information. |  | INTUITIVE DAVINCI ROBOT | DAVINCI ROBOT | 2631739 | 2599733 | 2496426 | MW5025634 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-05-30 | 2012-05-30 | No | No | Service Personnel |  |  | Yes |  |  |  |  | 2012
40018010 | 2012-02-01 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone and ultem sleeve exhibit localized melting around the holes, indicative of arcing. The hole sizes range from. 015 to. 040 with varying shapes (round and oblong) and are located from 0 to. 215 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has mechanical tears at the edge of one of the holes and at the distal end opening. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620089 | 2590881 | 2381233 | 2955842201200237 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-02-07 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The hole sizes range from. 010 to. 025 with varying shapes (round and oblong) and are located from. 030 to. 415 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620110 | 2590882 | 2487575 | 2955842201200238 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-03-26 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing occurred with the the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The holes are approximately. 020 in diameter and are located from. 220 to. 355 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has mechanical tears at the edges of the holes. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620112 | 2590883 | 2487589 | 2955842201200240 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-04-05 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. All of the tip cover damage is contained within one-half of the silicone, between the two parting lines. The hole size is. 015 x. 010 and is located. 350 above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the hole and at the distal end opening. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620121 | 2590887 | 2487602 | 2955842201200242 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-03-27 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone. The silicone exhibits melting around the hole, indicative of arcing. The hole size is approximately. 010 in diameter and is located. 220 above the silicone to sleeve interface. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620111 | 2590896 | 2487589 | 2955842201200239 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-04-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole size is approximately. 025 in diameter and is located. 130 above the silicone to sleeve interface. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620122 | 2590906 | 2487599 | 2955842201200244 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-04-05 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The hole sizes range from. 010 to. 060 with varying shapes (round and oblong) and are located from. 060 to. 490 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620105 | 2590908 | 2487602 | 2955842201200243 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-04-02 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing occurred with the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is approximately. 025 in diameter and is located. 495 above the silicone to sleeve interface. The tip cover does not exhibit mechanical tears in the general vicinity of the hole, but has a small tear at the distal end opening. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2620123 | 2590909 | 2487602 | 2955842201200241 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-30 | 2012-05-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-30 | Device Not Retu | Yes | Initial | 2012
42017910 | 2012-04-26 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed with the monopolar curved scissors instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury reported was reported. | The instrument was returned and evaluated. Visual inspection shows no sign of arcing on the main tube, tube extension, or metal components at the distal end. The distal end does not have any components or burrs or abnormally sharp edges that would contribute to arcing. The instrument was found to be fully functional with no physical damage observed. The customer reported failure could not be confirmed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2617740 | 2588726 | 2485419 | 2955842201200234 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-30 | 2012-05-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-04-30 | Device Not Retu | No | Reuse | 2012
400156 | 2012-04-26 | Malfunction |  | It was reported that while dissecting tissue during a da vinci si pyeloplasty procedure the 5 mm cautery spatula tip accessory that was installed on the 5 mm monopolar cautery instrument detached and fell into the patient. The tip accessory was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip accessory has not be returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip accessory is returned or if additional information is received. On may 17, 2012, the initial reporter indicated that no post-surgical complications have been experienced by the patient. | 5 MM CAUTERY HOOK TIP ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2617743 | 2588729 | 2485422 | 2955842201200233 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-26 | 2012-05-25 | No | Yes | Health Professional | 400156 | No | No | 2012-04-26 | No | Yes | Unkown | 2012
IS2000 A51P8 | 2012-04-17 | Death | Death | It was reported that during a da vinci s prostatectomy procedure, minor bleeding occurred. The bleeding was regarded as normal bleeding from minor vessels. The surgery was completed successfully and the patient went into recovery. The patient seemed fine during recovery; however, at an unknown time after the surgery, the patient expired. No further information was provided. | A review of system logs showed that no system faults occurred during surgery. There was no report of system, instrument or accessory malfunction related to this procedure. The root cause of death is thought to be either intubation or the result of bleeding. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. As of may 25, 2012, no adverse events have been experienced with the site's system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2618377 | 2589305 | 2485998 | 2955842201200235 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-04-27 | 2012-05-25 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-04-27 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-05-09 | Other | Other | It was reported that while docking the system for a da vinci s hysterectomy surgical procedure, the site experienced system error code 23. With the assistance of an isi technical support engineer (tse) the site restarted the system, which cleared the error. However system error code 23 reappeared while re-docking. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that system fault experienced by the customer was associated with the red fiber cable. The fiber cable provides communication from the surgeon side console to the patient cart. The input/output panel provides the connection point for the fiber cable into the surgeon console. The system was repaired by replacing the red fiber cable. As of may 21, 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2618379 | 2589307 | 2486000 | 2955842201200236 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-05-09 | 2012-05-25 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-05-09 | Device Not Retu | No | Reuse | 2012
400180 | 2012-03-28 | Injury | Required Intervention | Event Description: --/-- On (b)(6) 2012, intuitive surgical received a copy of (b)(4) report (b)(4) from (b)(6) on behalf of (b)(4). The report was submitted by (b)(6) located in (b)(6), concerning injuries sustained by the patient during a da vinci s periaortic lymphadectomy procedure. The following event description was provided: please note: there are 2 incidents - both occurred to the same client during the same surgery. Incident 1: robot arm 1 holding scissors (monopolar curved scissors) moved by itself and punctured the aorta causing 2 holes. Surgeon's head was in position and was opening the mouth of the scissors when the arm moved down toward the patient's aorta. Surgeon repaired the aorta and resumed robotic surgery. Incident 2: small hole found on the tip accessory cover of the scissors that created a third hole in the aorta from heat transfer. Hole was repaired and robotic surgery continued. Report submitted by (b)(6) dated (b)(6) 2012, the da vinci s system was verified by a technical support representative ((b)(6)), to determine if the problems come from poor functioning of the da vinci s system. Please see letter from (b)(6). Mfr report 2955842-2012-00227 has been submitted concerning the 2nd incident that occurred during the surgical procedure. | Manufacturer Narrative: --/-- On (b)(4) 2012, intuitive surgical received additional information from (b)(6). (b)(6) indicated that he followed up with (b)(6) hospital and the site reported that the patient is healthy and has resumed normal activities. Per (b)(6), the site indicated that damage to the patient's aorta was repaired with sutures. (b)(6) indicated that there is no video recording available of the planned surgical procedure. Per (b)(6), the site reported that an inspection of the mcs tip cover and monopolar curved scissors (mcs) instrument was performed by the surgical staff prior to use, a con med system 5000 esu was used during the surgical procedure with the settings between 35 watts and 45 watts; however, the mode setting was not provided. Per (b)(6), the site did not indicate if the mcs instrument made contact with another instrument during the surgical procedure and it is the surgeon's belief that the arcing event was caused due to a hole in the tip cover accessory. (b)(6) indicated that the site reported that the tip cover and mcs instrument are not available for return as they were discarded. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by the isi distributor field service engineer was unable to replicate the reported issue experienced by the site. Functional testing of the site's da vinci s system found that the system functioned within specification and a review of the site's system log revealed that there were no error codes generated during the surgical procedure. Several follow-up attempts have been made to verify additional information concerning the reported event, however, no response has been received. A follow-up mdr will be submitted if additional information is received. | 5MM MCS TIP COVER ACCESSORY | Unkown | 2614910 | 2586003 | 2482696 | 2955842201200228 | 1 | NAY | Manufacturer | OtherForeign | Other | FollowupInitial | 2012-05-01 | 2012-05-24 | Yes | Yes | Health Professional | 400180 | Yes | No | 2012-06-11 | Device Not Retu | Yes |  | 2012
400272 | 2012-03-08 | Malfunction |  | On april 30, 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda, which contained the following information: event desc: using harmonic scalpel in robot case, the metal tip of the harmonic insert broke off into the patient. Surgeon tried to retrieve it, but due to uterus size was unable to do so immediately. Patient was opened to find it and finish case. X-ray and c-arm were used to assist with finding piece. Counts were correct and patient was stable upon leaving the or. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the device is returned or if additional information is received. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2616557 | 2587814 | 2484507 | 2955842201200229 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-30 | 2012-05-24 | No | Yes | Health Professional | 400272 | No | Yes | 2012-03-08 | No | Yes | Unkown | 2012
42011704 | 2012-04-24 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure while using the needle driver instrument, the insulation broke off and the fell into the patient. The planned surgical procedure was completed; however, the hospital has not communicated if the broken instrument piece was removed from the patient. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube insulation was coming off. The main tube insulation appears to have deep scratches and gouge marks, and a large piece of insulation, approximately 0. 75 in length was removed from the distal end of the main tube. There are several other small pieces of insulation removed. Engineering concluded that mishandling/misuse likely caused the damage. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On may 14, 2012, isi contacted the initial reporter and he indicated that there was no harm to the patient and the patient is doing well; however, due to hippa regulations, he declined to confirm if the patient has retained any fragments from the instrument. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2616567 | 2587829 | 2484522 | 2955842201200231 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-24 | 2012-05-24 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2012-04-24 | Device Not Retu | No | Reuse | 2012
40016907 | 2011-08-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, part of the jaw of the harmonic curved shears (hcs) insert accessory (installed on the hcs instrument) broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned and / or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the clamp arm detached and missing from the distal end. Shaft features that mate with clamp arm pins are bent slightly backwards, consistent with hyperextension of the clamp arm. Evidence is not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | Unkown | 2616564 | 2587836 | 2484529 | 2955842201200230 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2012-05-03 | 2012-05-24 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2012-06-01 | Device Not Retu | Yes |  | 2012
42011704 | 2012-04-24 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure while using the needle driver instrument, the insulation broke off and the fell into the patient. The planned surgical procedure was completed; however, the hospital has not communicated if the broken instrument piece was removed from the patient. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube insulation was coming off. The main tube insulation appears to have deep scratches and gouge marks, and a large piece of insulation, approximately 1. 4 in length was removed from the distal end of the main tube. There are several other small pieces of insulation removed. Engineering concluded that mishandling/misuse likely caused the damage. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2012, isi contacted the initial reporter and he indicated that there was no harm to the patient and the patient is doing well; however, due to hippa regulations, he declined to confirm if the patient has retained any fragments from the instrument. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2616604 | 2587874 | 2484522 | 2955842201200232 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-24 | 2012-05-24 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2012-04-24 | Device Not Retu | No | Reuse | 2012
| 2012-05-22 | Malfunction |  | During robotic procedure, surgeon was using the precise bipolar. While in use he noticed that instrument was sparking at the tips. There was no harm to the patient. Instrument was immediately withdrawn and sequestered. |  | DA VINCI ENDOWRIST PRECISE | INSTRUMENT ROBOTIC SURGICAL BI | 2620566 | 2591378 | 2488071 | 2591378 | 1 | NAY | User Facility |  |  | Initial | 2012-05-24 | 2012-05-24 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 |  | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the surgeon experienced non-intuitive movement with the instrument and had to reseat the instrument sterile adapter and drape. Post-surgery, the patient developed an infection caused by a hole found in her bowel. As a result, the patient's bowel was resected. | The cause of the hole in the patient's bowel is currently unknown. Intuitive surgical is in the process of obtaining additional information regarding this event from the hospital. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2614744 | 2585804 | 2482497 | 2955842201200224 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-23 | 2012-05-23 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2012-04-23 | No | No | Reuse | 2012
42014504 | 2012-04-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si bilateral ureteral reimplantation procedure, the jaw on the debakey forceps instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found the jaws broken at the base. Engineering concluded that the damage was likely due to force applied to the jaw. The broken surface indicated slight bending prior to the breakage. No damage was found on the snake wrist. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | DEBAKEY FORCEPS INSTRUMENT | Reuse | 2614808 | 2585893 | 2482586 | 2955842201200225 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2012-04-30 | 2012-05-23 | No | Yes | Health Professional | 42014504 | Device Returned To Manufacturer | No | 2012-05-24 | Device Not Retu | No |  | 2012
400180 | 2012-03-28 | Injury | Required Intervention | Event Description: --/-- On may 1, 2012, intuitive surgical received a copy of (b)(6) report from (b)(6) on behalf of (b)(4). The report was submitted by (b)(6) hospital, concerning injuries sustained by the patient during a da vinci s periaortic lymphoidectomy procedure. The following event description was provided: please note: there are 2 incidents - both occurred to the same client during the same surgery. Incident 1: robot arm 1 holding scissors (monopolar curved scissors) moved by itself and punctured the aorta causing 2 holes. Surgeon's head was in position and was opening the mouth of the scissors when the arm moved down toward the patient's aorta. Surgeon repaired the aorta and resumed robotic surgery. Incident 2: small hole found on the tip accessory cover of the scissors that created a third hole in the aorta from heat transfer. Hole was repaired and robotic surgery continued. Report submitted by (b)(4) dated 2012/03/29 the da vinci s system was verified by a technical support representative ((b)(4)) to determine if the problems come from poor functioning of the da vinci s system. Mfr report 2955842-2012-00228 has been submitted concerning the 1st incident that occurred during the surgical procedure. | Manufacturer Narrative: --/-- On (b)(6) 2012, intuitive surgical received additional information from (b)(6) at (b)(6) (distributor). (b)(6) indicated that he followed up with (b)(6) and the site reported that the patient is healthy and has resumed normal activities. Per (b)(6), the site indicated that damage to the patient's aorta was repaired with sutures. (b)(6) indicated that there is no video recording available of the planned surgical procedure. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Several attempts have been made to the site to verify additional information concerning the reported event. However, no response has been received. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | Unkown | 2614888 | 2585943 | 2482636 | 2955842201200227 | 1 | NAY | Manufacturer | OtherForeign | Other | InitialFollowup | 2012-05-01 | 2012-05-23 | Yes | Yes | Health Professional | 400180 | No | No | 2012-06-11 | Device Not Retu | Yes |  | 2012
IS2000 A51P8 | 2012-04-30 | Death | Death | It was reported that 1 day after a da vinci s hysterectomy procedure was successfully performed, the patient expired. | On (b)(6) 2012, isi's clinical sales representative (initial report) indicated that he spoke to dr. (b)(6), the surgeon who performed the surgical procedure and dr. (b)(6) reported that there was no malfunction of the da vinci s surgical system, instruments and/or accessories during the surgical procedure. Dr. Harden reported that the planned surgical procedure was completed after approximately 3 hours and intra-operatively the patient tolerated the planned surgical procedure well. Dr. Harden indicated that approximately 3 to 4 hours post- op the patient's foley catheter was removed, the patient was able to walk around and appeared to be doing well and the next day at approximately 6:30 am, she received a call advising that the patient expired. Per dr. (b)(6), an autopsy was performed and there was no evidence of hemorrhaging or pulmonary embolism and it is believed that the patient's demise was a result of consuming too many narcotics. On (b)(6) 2012, isi contacted dr. (b)(6) and she indicated that the patient underwent the da vinci s hysterectomy procedure due to dysfunctional bleeding and that the patient was taking anti-depressants due to depression and that the patient's sister contacted her after the patient expired and advised dr. (b)(6) that the patient consumed alcohol while taking anti-depressant medications. A toxicology report is currently pending. Dr. (b)(6) indicated that there was no malfunction of the da vinci system, instruments and/or accessories, no patient complications occurred and the patient tolerated the planned surgical procedure well. She is unable to provide the patient's records; however, it is her belief that the patient's demise is unrelated to the da vinci system, instruments and/or accessories. On (b)(6) 2012, isi contacted (b)(6) in the patient safety department at (b)(6) and she indicated that she is unable to provide any additional information concerning the patient due to hippa regulations. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2614857 | 2585957 | 2482650 | 2955842201200226 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-05-01 | 2012-05-23 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-05-01 | No | No | Reuse | 2012
42019408 | 2012-04-20 | Malfunction |  | It was reported that while the surgeon was applying hemostatsis agent on the vaginal cuff during a da vinci si hysterectomy procedure, the mega needle driver instrument insert separated from the tip and fell into the patient. The insert was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found both of the carbide inserts to be broken off of one grip, only one insert was returned. The returned insert was separated from the instrument. No damage was found to the instrument clevis or cable. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2609509 | 2580369 | 2477068 | 2955842201200220 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-20 | 2012-05-18 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2012-04-20 | Device Not Retu | No | Reuse | 2012
IS1200 A43P9 | 2012-04-19 | Other | Other | It was reported that at the beginning of a da vinci cystectomy procedure, the site experienced a non-recoverable system error code 21008. With the assistance of an isi technical support engineer, the site removed the instrument and sterile adapter installed on a patient side manipulator (psm) arm, manually manipulated an axis on the psm, reinstalled the sterile adapter and restarted the system, however, the error code reoccurred. The tse instructed the site to un-install the sterile adapter and reboot the system and the psm system homed. Due to the repeated occurrence of system error code 21008, as a precaution the surgeon made the decision to abort the planned surgical procedure and reschedule for a later date. The patient was under anesthesia and the port incisions were made. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer concluded that the non-recoverable system error code 21008 experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The error code 21008 appeared when the da vinci safety system determined the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The psm was returned for failure analysis investigation. Engineering was not able to replicate the customer reported failure mode, however based on the system error logs the axis potentiometer and motor/encoder were replaced as a caution. Dr. (b)(6) with (b)(6) hosp was contacted on may 17th and they indicated that the planned surgical procedure was successfully completed (b)(6) 2012. As of may 18, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2609530 | 2580370 | 2477069 | 2955842201200219 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-04-19 | 2012-05-18 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2012-04-19 | Device Not Retu | No | Reuse | 2012
42019408 | 2012-04-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the fenestration blade portion of the blade on the tip of the mega needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the carbide insert is broken off from the instrument tip. The insert was not returned with the instrument. Per the information provided by the isi's clinical sales representative, the broken instrument piece were successfully retrieved from the patient. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2609547 | 2580395 | 2477088 | 2955842201200218 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-20 | 2012-05-18 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2010-08-05 | Device Not Retu | No | Reuse | 2012
42018406 | 2012-04-20 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the tip on the permanent cautery spatula (pcs)instrument broke off and fell into the patient while advancing the pcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found that the instrument was returned with the spatula broken off along with the spatula fragment. Engineering concluded that damage to the instrument was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2609615 | 2580451 | 2477144 | 2955842201200221 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-20 | 2012-05-18 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-04-20 | Device Not Retu | No | Reuse | 2012
400180 | 2012-04-19 | Injury | HospitalizationLife ThreateningRequired Intervention | Event Description: --/-- On (b)(4) 2012, medwatch uf/importer report (b)(4) was received: patient in or for robotic laparoscopic hysterectomy; surgeon documented grossly enlarged uterus with multiple fibroids. As surgeon was bivalving the uterus, he saw a flash of blood coming from the left side of the patient and immediately recognized vascular injury. The robot was immediately undocked and procedure was converted to opened procedure. The vascular surgeon responded immediately to repair the patient's left external iliac artery, which he described as charred due to cautery injury. Operating room staff reported small tear in the plastic that covered the metal end of the monopolar scissor. The patient was stabilized after surgical repair and administration of blood products and transferred to the intensive care unit postoperatively. She remained stable and was transferred to the surgical unit the following day. The patient was discharged home (b)(6) 2012. | Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The mcs instrument used in conjunction with the tip cover has not been returned for evaluation. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The monopolar curved scissors (mcs) instrument used in conjunction with the mcs tip cover accessory was returned and evaluated. The instrument was tested with the system, the system recognized the instrument. The instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. Scissors cut cleanly through 0. 006 latex. No signs of arcing were observed. Electrical continuity test passed. The instrument was found to be fully functional. The tip cover accessory has not been returned for evaluation. | 8MM MCS TIP COVER ACCESSORY | Initial | 2609712 | 2580502 | 2477201 | 2955842201200217 | 1 | NAY | Manufacturer | Company Representative | Other | FollowupFollowup | 2013-01-08 | 2012-05-18 | Yes | Yes | Health Professional | 400180 | No | No | 2012-04-19 | Device Not Retu | Yes |  | 2012
42027501 | 2012-04-18 | Other | Other | On may 15, 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda with the following information: the physician completed the part of the procedure that required the shears. He disconnected the cord that powers the pedal to activate the shears and moved the robot away from the patient. The team heard the shears/scalpel activate despite being disconnected. The patient sustained a small burn. What was the original intended procedure? robotic hysterectomy device usage problem: device malfunction - that is, the device did not do what it was supposed to do the user facility had initially reported this issue to intuitive surgical on (b)(4) 2012. | Additional information was provided by the initial reporter, who indicated that after the surgical procedure was successfully completed with the da vinci si surgical system, the harmonic ace curved shears instrument with the harmonic ace curved shears insert accessory installed was placed on the patient's abdomen. The instrument was laying across the patient's abdomen when it suddenly energized and caused a minor burn. The patient's burn was treated with neosporin and no post-surgical complications have been reported. At the time of the event, the instrument was completely detached from the system. Although the instrument was connected to the ethicon gen11 electrosurgical unit (esu), the esu was not connected to the system either. Therefore, the initial reporter further indicated that it us unknown how the instrument became energized. The initial reporter further indicated that the da vinci si surgical system did not cause or contribute to the event as it was not connected to the instrument nor was it connected to the esu. The esu manufacturer was notified to inspect their energy source. Despite multiple requests, the device has not been returned for evaluation. A follow-up medwatch report will be submitted if the instrument and insert are returned or if additional information is received. | HARMONIC ACE CURVED SHEARS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2609717 | 2580506 | 2477205 | 2955842201200222 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-18 | 2012-05-18 | Yes | Yes | Health Professional | 42027501 | No | Yes | 2012-04-18 | No | Yes | Reuse | 2012
IS3000 | 2012-04-20 | Other | Other | Event Description: --/-- It was reported that during a da vinci si surgical procedure the surgical staff experienced recurring system error code 20013. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. --/-- Additional Event Description: --/-- It was reported that during a da vinci surgical procedure the surgical staff experienced recurring system error code 20013. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by the field service engineer determined that system error code 20013 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of may 18, 2012, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- Additional information received from the account indicated there was no complications as a result of converting to open surgery. | DA VINCI SURGICAL SYSTEM | Reuse | 2609740 | 2580527 | 2477226 | 2955842201200223 | 1 | NAY | Manufacturer | ForeignUser facility | Other | InitialFollowup | 2012-05-22 | 2012-05-18 | Yes | Yes | Health Professional | IS3000 | Device Returned To Manufacturer | No | 2012-04-20 | Device Not Retu | No |  | 2012
| 2012-04-16 | Malfunction |  | During a robotic assisted laparoscopic left renal cyst decortication and lysis of intra-peritoneal adhesions, the da vinci grasping forcep broke off into the patient. It was successfully removed without harm to the patient. |  | DA VINCI ENDOWRIST GRAPTOR | INSTRUMENT ROBOTIC SURGICAL GR | 2611231 | 2582154 | 2478846 | 2582154 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-05-18 | 2012-05-18 | No | Yes | Physician |  | No Answer Provided | No Answer Provided |  |  |  |  | 2012
IS2000 A43P9 | 2010-09-01 | Other | Hospitalization | The following patient story concerning a da vinci prostatectomy procedure was found posted on the (b)(4): i was diagnosed with a t1cnxmx gleason 6 adenocarcinoma, 2 of 6 biopsies in (b)(6) 2010. Thought robotic surgery would be the best type of surgery available. The first urologist bent over backwards to sell the robotic surgery and my wife and i agreed based on the fast healing time as all other outcomes would be the same as open surgery. Based on my diagnosis, i thought i would have an excellent chance of cure and a fast recovery. Da vinci website and the surgeon alluded to great expectations, i had the surgery in (b)(6) 2010. Outcome was atrocious, urinary tract damage, uti, bladder and kidney infection, 12 months of perineal pain and physio to help treat it and still require physio occasionally (the physio man thought damage was caused by nerve damage related to postural issues during surgery), incontinence, impotence and at lastly an incisional hernia from a very large oversized port above my navel (6cm long or 2. 5 inches) da vinci website says 1/2 inch (12mm) ports. Now my stomach looks very distorted from having a very large mesh inserted to repair the hernia and i now sport a 125mm (5 inch) scar horizontally above my navel. Time off work to date to initially have the surgery and follow up surgery for the hernia is approximately 14 weeks. Medical out of pocket costs now exceed (b)(6) plus lost income for 14 weeks. None of my expectations were met by this surgery and the surgeon didn't know and couldn't logically explain why it all happened. He was at least very honest and i appreciate that. The outcome is devastating and now it seems my psa is increasing albeit it is now at. 03ug/l whereas for 12 months it was. 01ug/l. Now i am being told it is possible circulating kallikreins. Funny it wasn't there in the first 12 months. My quality of life is low, drug controlled continence and ed issues linger. I continually regret the day i decided to have robotic surgery. | The system logs of the da vinci system were reviewed and no errors were observed which indicate that there were issues with the system. The root cause of the post surgical complications experienced by the patient is currently undetermined. Additional information has been requested about the event, but the customer has yet to provide further details. As of may 16, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2608165 | 2579296 | 2475995 | 2955842201200215 | 1 | NAY | Manufacturer | Consumer | Other | Initial | 2012-04-22 | 2012-05-17 | Yes | No | Health Professional | IS2000 A43P9 | No | No | 2012-04-27 | No | No | Reuse | 2012
40018010 | 2012-04-17 | Injury | Other | Event Description: --/-- It was reported that during a da vinci s hysterectomy procedure during lymph node dissection, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The arcing created a 3mm perforation along the medial external iliac vein, resulting in bleeding. The surgeon used the maryland bipolar forceps to clamp the vein and repair it robotically. The planned surgical procedure was completed and no further patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On may 27, 2012, intuitive surgical received an phone response from rn (b)(6), indicating that the patient is recovering well and has not developed any post-operative complications. --/-- Additional Manufacturer Narrative: --/-- | 8MM MCS TIP COVER ACCESSORY | Initial | 2608168 | 2579299 | 2475998 | 2955842201200216 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-01-08 | 2012-05-17 | Yes | Yes | Health Professional | 40018010 | No | No | 2012-04-17 | No | No |  | 2012
IS3000 A60P6 |  | Other | Other | The following patient story concerning a da vinci si hysterectomy procedure performed on (b)(6) 2012, at (b)(6) was found posted on the (b)(6) by intuitive surgical on (b)(6) 2012. After trying other measures failed to stop the bleeding and pain of fibroids a hysterectomy was necessary. The procedure was simple and scaring and bleeding were minimal. The first 2 weeks went by smoothly and the pain was gone within the first few days. The biggest inconvenience was not being able to drive for 2 weeks. Then i had a turn for the worse. The day after my 2 week follow up appointment i started to bleed heavily. I was soaking a pad an hour. The on-call surgeon did not seem to think it was serious and advised to wait 4 hours and go to the er if it continued. The er doctor did and physical exam and did not see any signs of a rupture. He too consulted the on-call surgeon who suggested i be released and to contact my doctor the following morning. On the way home from the er i began to experience intense pain in my lower back and cramps equivalent to contractions. I felt a huge gush of blood and fainted. The ambulance was called and i was taken back to the er by this time i had been bleeding for 8 hours and lost 2 pints of blood. After an ultra sound it was decided that i would need to undergo a second surgery. The sutures had dissolved and the cuff had opened causing the bleeding. This complication also caused me to become anemic, but fortunately i did not have to undergo a blood transfusion. Its now 7 weeks. Since the initial procedure and all is well. My doctor in his 20 years experience had never seen this complication. During my initial 2 week recovery i did all as instructed. I was even ready to go back to work the very next day when the bleeding started. I look forward to no longer having pain or discomfort from the fibroids or menstruation. | Based on the information provided by the patient, it is indeterminable if the da vinci si system, instruments and/or accessories caused or contributed to the post surgical complications experienced by the patient. Isi has contacted the surgeon who performed the surgical procedure several times to obtain additional information concerning the reported event, however, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2605972 | 2576917 | 2473629 | 2955842201200211 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2012-04-16 | 2012-05-16 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-04-16 | No | No | Reuse | 2012
42018404 | 2012-04-20 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy with salpingo-oophorectomy procedure, the permanent cautery spatula instrument (pcs) broke and a fragment from the instrument broke off and fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with a. 175 x. 320 piece missing from one side of the proximal clevis and most of the missing section is between the clevis ears. Engineering has concluded that the damage was likely due to overloading at the instrument tip while the instrument wrist is pitched. Engineering observation also found that the pitch cable is broken and the cable crimp is missing. The pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2012 isi contacted the initial reporter at (b)(6) hospital and she indicated that no additional surgical procedure was required to remove the instrument fragment as the surgeon was able to remove the fragment using a laparoscopic grasper through one of the assistant ports. An x-ray of the affected area was not performed however, irrigation of the affected area was performed to ensure that there were no remaining instrument piece(s). She further indicated that no harm or injury to the patient as a result of the reported event. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2606399 | 2577280 | 2473994 | 2955842201200213 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-04-20 | 2012-05-16 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2012-05-22 | Device Not Retu | No |  | 2012
IS2000 A51P8 | 2011-06-01 | Other | Required Intervention | On april 16, 2012 isi received a legal summons indicating on (b)(6) 2011, the patient underwent a da vinci s hysterectomy procedure at (b)(6). During the surgical procedure the patient sustained injuries to her left ureter and bladder. The planned surgical procedure was converted to open laparotomy surgery with repair of the bladder injury. The patient was given two units of packed red blood cells due to blood loss during surgery. The patient was discharged from the hospital on (b)(6) 2011. On (b)(6) 2011, the patient began having severe pain in her back and noticed that she was passing urine through her vagina. Secondary to this, the patient went back to the hospital and had radiology tests performed which, confirmed that the patient had a ureteral injury requiring an immediate diversionary nephrostomy tube to be placed by a radiologist. A percutaneous nephrostomy tube was placed and the stent was able to pass freely into the bladder. The patient was outfitted with a nephrostomy bag to collect her urine. The patient was then taken back to the operating room and the ureteral stent in the left ureter was internalized. This allowed for the discontinuation of the nephrostomy bag. On (b)(6) 2011, the patient was taken back to the operating room for a cystoscopy, left ureteral stent removal, bilateral ureteroscopy, and right retrograde pyelogram. Despite all of this treatment, the patient had recurrent difficulty with pelvic pain and bladder function and continued to experience involuntary loss of urine through her vagina. The patient had to undergo an additional operation on (b)(6) 2011, consisting of cystoscopy and fulguration of bladder lesion for a diagnosis of vesicovaginal fistula. Despite this treatment, the patient continues to suffer from nocturea, pelvic pain, dyspareunia, and continued leakage of urine vaginally. | A review of the site's system log found that no system error codes were generated during the surgical procedure. At this time it is indeterminable as to what caused damage to patient's ureter and bladder. Requests for additional information have been made to the hospital's risk management department and to the surgeon. To date, no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2606389 | 2577302 | 2474014 | 2955842201200212 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-04-16 | 2012-05-16 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-04-16 | No | No | Reuse | 2012
42030901 | 2012-04-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si sacrocolpopexy procedure, the mega suturecut needle driver instrument broke into pieces and a piece of wire from the instrument fell into the patient. The fragment was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. This report is being reported late due to late entry of the complaint information by the responsible customer service representative. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the physician, the broken instrument piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, the grip cable is broken at the distal idlers. The idler pulley spins freely and does not exhibit damage. The cable segment sticks out at wrist. The piece of cable broken off and retrieved by surgeon was not returned with the instrument. No other damage was found. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Initial | 2606407 | 2577334 | 2474048 | 2955842201200214 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2012-04-02 | 2012-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2012-05-14 | Device Not Retu | No |  | 2012
420172 | 2012-05-13 | Malfunction |  | When getting ready to use the maryland dissector, surgeon noticed that tips were not aligned. The instrument was removed from the sterile field by the scrub tech and a new instrument was given to the surgeon. The instrument was given to the manufacturer's rep. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL MA | 2609701 | 2580460 | 2477153 | 2580460 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-05-14 | 2012-05-14 | No | Yes | Physician | 420172 | No | No Answer Provided |  |  |  |  | 2012
VERSION 8 | 2012-04-13 | Malfunction |  | Upon getting ready to use the da vinci prograsp graspers, the scrub tech noticed a wire sticking out of the instrument. The instrument was taken off the sterile field and replaced with a new prograsp. |  | DA VINCI ENDOWRIST PROGRASP | INSTRUMENT ROBOTIC SURGICAL FO | 2609702 | 2580461 | 2477154 | 2580461 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-05-14 | 2012-05-14 | No | Yes | Invalid Data | VERSION 8 | No | No Answer Provided |  |  |  |  | 2012
VERSION 7 | 2012-04-13 | Malfunction |  | Da vinci maryland dissectors needed to be replaced during surgery, tips not aligning. No harm to patient, instrument removed from sterile field and to be returned to manufacturer. |  | DA VINCI ENDOWRIST | INSTRUMENT ROBOTIC SURGICAL BI | 2610165 | 2581066 | 2477759 | 2581066 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-05-14 | 2012-05-14 | No | Yes | Physician | VERSION 7 |  | No Answer Provided |  |  |  |  | 2012
400180 | 2012-04-10 | Injury | Other | Event Description: --/-- It was reported that during a da vinci si hysterectomy with bilateral salpingoophorectomy procedure, arcing from a hole in the mcs tip cover accessory installed on the monopolar curved scissors(mcs) instrument occurred, causing injury to the patient's colon/bowel. Damage to the patient's colon/bowel was repaired using a 3. 0 silk suture. No other adverse patient outcome was reported. | Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. However, the mcs instrument used in conjunction with the tip cover was returned for evaluation and engineering observation found no evidence of arcing. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burned through the silicone. Silicone exhibits localized melting around the holes, indicative of arcing. Hole sizes range from 0. 015 to 0. 025 with varying shapes (round and oblong), and are located from 0. 200 to 0. 480 above the silicone to sleeve interface. Multiples holes indicate multiple arcing events occurred. Tip cover also has various mechanical tears at the edge of the two of the holes and more tearing in the general vicinity of the holes. The tip cover holes line up with the edge of the distal clevis ears when installed on the monopolar curved shears (mcs) instrument. | 8MM MCS TIP COVER ACCESSORY | Initial | 2601256 | 2571951 | 2465967 | 2955842201200207 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupFollowup | 2013-01-08 | 2012-05-11 | Yes | Yes | Health Professional | 400180 | Device Returned To Manufacturer | No | 2012-04-11 | Device Not Retu | Yes |  | 2012
IS2000 A51P8 | 2012-04-12 | Malfunction |  | It was reported that during a da vinci s surgical procedure, a lens was observed to be missing from the endoscope. No efforts were made to retrieve the missing lens. No patient harm, adverse outcome or injury was reported. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment manufacturer (oem) for evaluation. Per the customer reported complaint, the oem confirmed that an endoscope lens was missing. It was determined that the endoscope was exposed to a high temperature load caused by unsuitable cleaning and sterilization, which allowed fluid to infiltrate the endoscope. As a result, the glue is washed out and the optical components were damaged. Sterility is the responsibility of the person/institution performing the sterilization. Intuitive surgical's reprocessing instructions for cleaning, disinfection, and sterilization information for reusable instruments, accessories and endoscopes used with da vinci, da vinci s and da vinci si systems lists intuitive surgical's recommendations for methods of sterilization and the parameters. The da vinci s surgical system user's manual specifically states: caution: do not autoclave the endoscope. Autoclave cycles introduce high temperatures and sudden temperature changes, which will cause damage to the endoscope. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2601258 | 2571953 | 2468666 | 2955842201200209 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-04-12 | 2012-05-11 | No | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-04-12 | Device Not Retu | No | Reuse | 2012
400180 | 2012-04-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingoophorectomy procedure, arcing from a hole in the mcs tip cover accessory installed on the monopolar curved scissors(mcs) instrument occurred. No patient harm or adverse outcome was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. However, the mcs instrument used in conjunction with the tip cover was returned for evaluation and engineering observation found no evidence of arcing. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2601259 | 2571954 | 2468667 | 2955842201200208 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-13 | 2012-05-11 | No | Yes | Health Professional | 400180 | No | No | 2012-04-10 | No | Yes | Initial | 2012
42805306 | 2012-03-31 | Malfunction | Required Intervention | It was reported that during a da vinci si single-site cholecystectomy procedure the jaws of the 5 mm medium-large clip applier instrument would not open after gripping tissue. The surgical staff tried using a wrench to open the jaws, but was unsuccessful. The tissue within the instrument jaw grips was then cut and removed with the instrument. A replacement instrument of the same type was used to complete the planned procedure, with a 30 minute delay. No additional patient harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grips are stuck in the open position. The grip input disc is very loose and the grips do not open and close properly. The housing was removed and the grip axis input shaft was found to be broken below the top chassis bearing. The input gear no longer meshes with the sector gear, so the grips cannot open and close properly. The instrument was disassembled to separate the grip assembly from the drive train. The grips are able to open and close when the push pull rod is manually actuated. The sector gear will not rotate and appears to be seized on the dowel pin. It was determined that the seized gear may have contributed to the breakage of the input shaft. | 5 MM MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2601261 | 2571955 | 2468668 | 2955842201200210 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-04-12 | 2012-05-11 | Yes | Yes | Health Professional | 42805306 | Device Returned To Manufacturer | No | 2012-04-12 | Device Not Retu | No | Reuse | 2012
400180 | 2012-04-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) with a mcs tip cover accessory installed causing a burn to fat located near the patient's bowel. The mcs instrument was immediately removed and inspection of the tip cover accessory by the surgical staff revealed that the mcs tip cover had a hole. The planned surgical procedure was completed with a replacement mcs instrument tip cover accessory. No patient harm, adverse outcome or injury was reported. | The site reported that the tip cover accessory will not be returned to intuitive surgical for failure analysis, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(4) 2012 intuitive surgical contacted the initial reporter and she indicated that there were no intra operative or immediate post operative complications experienced by the patient as a result of the reported event and no patient harm, adverse outcome or injury occurred. As of (b)(4) 2010 she does not know if the patient has returned to the hospital due to experiencing any post operative complications. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2598457 | 2569234 | 2465947 | 2955842201200206 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-10 | 2012-05-10 | No | Yes | Health Professional | 400180 | No | Yes | 2012-04-10 | No | Yes | Unkown | 2012
40018010 | 2012-04-10 | Other | Other | It was reported that during a da vinci si hysterectomy with pelvic-aortic lymphadenectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used, causing a burn to the patient's bowel. The affected area of the patient's bowel was repaired and the planned surgical procedure was completed using a 4th tip cover accessory. Medwatch mfr reports 2955842-2012-00203 & 2955842-2012-00204 were submitted for the first and third occurrences of arcing. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the tip cover was returned with various mechanical tears and a hole burned though the silicone. The silicone exhibits localized melting around the hole, indicating that an arcing event occurred. The hole is roughly. 015 in diameter, and is located. 160 above the silicone to sleeve interface. The tip cover also has a couple of mechanical tears in the general vicinity of the holes. When installed on the instrument returned with the tip cover, the hole in the tip cover is in the area of the proximal clevis. Engineering evaluation of the monopolar curved scissors (mcs) instrument used in conjunction with the returned tip cover accessory found that the distal end of the instrument shows no char marks or other signs of arcing. The metal components at wrist do not have abnormally sharp edges that would potentially damage the tip cover. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2598441 | 2569254 | 2465967 | 2955842201200205 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-10 | 2012-05-10 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-04-10 | Device Not Retu | Yes | Unkown | 2012
| 2012-01-06 | Malfunction |  | While the surgeon was using the mega suturecut needle driver on the patient, the suture cut control cable broke. A new needle driver was obtained. It is unknown whether the new device was of the same or a different lot. |  | DA VINCI ENDOWRIST MEGA SUTURECUT | INSTRUMENT ROBOTIC SURGICAL SU | 2609700 | 2580459 | 2477152 | 2580459 | 1 | NAY | User Facility |  |  | Initial | 2012-04-10 | 2012-05-10 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
40018010 | 2012-04-10 | Other | Other | It was reported that during a da vinci si hysterectomy with pelvic-aortic lymphadenectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used, causing a burn to the pt's bowel. The affected area of the pt's bowel was repaired and the planned surgical procedure was completed using a 4th tip cover accessory. Medwatch mfr reports #2955842-2012-00205 & 2955842-2012-00204 were submitted for the first and second occurrences of arcing. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the tip cover was returned with various mechanical tears and a hole burned though the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is roughly. 015" in diameter, and is located. 175" above the silicone to sleeve interface. The cover also has a couple of mechanical tears in the general vicinity of the holes. When installed on the returned instrument, the hole in the tip cover is in the area of the proximal clevis. Engineering eval of the monopolar curved scissors (mcs) instrument used in conjunction with the returned tip cover accessory found that the distal end of the instrument shows no char marks or other signs of arcing. The metal components at wrist do not have abnormally sharp edges that would potentially damage the tip cover. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2623593 | 2582933 | 2465967 | 2955842201200203 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2012-04-10 | 2012-05-10 | Yes | Yes | Health Professional | 40018010 | Yes | No | 2012-04-10 | Yes | Yes | Unkown | 2012
40018010 | 2012-04-10 | Other | Other | It was reported that during a da vinci si hysterectomy with pelvic-aortic lymphadenectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used, causing a burn to the pt's bowel. The affected area of the pt's bowel was repaired and the planned surgical procedure was completed using a 4th tip cover accessory. Medwatch mfr reports #2955842-2012-00203 & 2955842-2012-00205 were submitted for the first and second occurrences of arcing. | The tip cover accessory was returned and evaluated. Engineering was unable to confirm the customer reported complaint, as a visual inspection of the tip did not find any evidence of arcing and there were no holes and or tears observed in the silicone part of the tip cover. No physical damage was found. Engineering eval of the monopolar curved scissors (mcs) instrument used in conjunction with the returned tip cover accessory found that the distal end of the instrument shows no char marks or other signs of arcing. The metal components at wrist do not have abnormally sharp edges that would potentially damage the tip cover. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2623626 | 2582936 | 2465967 | 2955842201200204 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2012-04-10 | 2012-05-10 | Yes | Yes | Health Professional | 40018010 | Yes | No | 2012-04-10 | Yes | Yes | Unkown | 2012
40018010 | 2012-02-29 | Other | Other | Event Description: --/-- On april 10, 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda, with the following information: event desc: during case, monopolar scissors with cautery for the davinci, caused arc during procedure. Md stated that the arc had happened and then noticed a hole in the protective sheath that is placed on the scissors. Md noted there was damage to the right external iliac artery. The scissors were removed and the protective sheath was replaced before surgery continued. What was the original intended procedure? robotic radical laparoscopic hysterectomy, bilateral salpingo-oophorectomy with parametrial resection, pelvic and common iliac lymph node dissection. Device usage problem: device malfunction - that is, the device did not do what it was supposed to do. | Manufacturer Narrative: --/-- On (b)(6) 2012, a representative from the hospital's risk management department indicated that a vascular surgeon was called in during the surgical procedure to assess the patient burn. The vascular surgeon determined the burn to be superficial and no actions were required to address the burn. The hospital representative also indicated that no post-surgical complications have been reported by the patient. --/-- Additional Manufacturer Narrative: --/-- The hospital has been contacted multiple times for additional information regarding this event, however, no additional information has been provided to date. The tip cover accessory has not been returned for evaluation despite several requests, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | Unkown | 2596428 | 2567645 | 2464358 | 2955842201200201 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | InitialFollowup | 2012-04-10 | 2012-05-09 | Yes | Yes | Health Professional | 40018010 | No | No | 2012-05-11 | Device Not Retu | Yes |  | 2012
IS3000 A60P6 | 2012-04-09 | Death | Death | It was reported that approximately 45 minutes into a da vinci si procedure when the assistant at the patient side cart (psc) performed a tool change to switch the instruments in the patient side manipulator (psm) arms, the patient's aorta was punctured. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned surgical procedure. The patient was reported to be in stable condition right after the open surgical procedure, however, the patient expired the following day. The site has been contacted for additional information, however, no additional information has been forthcoming. | The isi clinical sales representative followed up with the assistant rn and the rn indicated that the patient side assistant had not received training by the site on how to perform a guided tool change. Based on the information provided, the puncturing of the patient aorta is the result of an incorrect instrument change. The surgeon had requested that instruments located on the patient side manipulators (psm) arms be swapped, such that the instrument on psm arm 1 would be switched with the instrument located on psm arm 2. It is not clear whether the assistant nurse used a guided tool change to safely change instruments or whether the nurse used the arm clutch to insert the instrument manually. It is unclear whether any malfunction of the system occurred and the exact mechanism which resulted in puncturing of the aorta. The site has been contacted for additional information and once this information has been received, a follow up report will be sent to the fda. The da vinci si instructions for use (ifu) specifically states: warning: the instrument may not be immediately visible when being moved from the cannula into the patient. Use appropriate caution when manually inserting instruments into the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2596655 | 2567858 | 2464571 | 2955842201200202 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-04-09 | 2012-05-09 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-04-09 | No | No | Reuse | 2012
IS3000 A60P6 | 2010-02-24 | Death | Death | Event Description: --/-- On (b)(6) 2012, intuitive surgical received a legal summons indicating that during a da vinci si hysterectomy procedure on (b)(6) 2012, while the surgeon was performing node dissection, the surgeon cut the patient's iliac artery and vein causing bleeding from those vessels. A hemo-clip was applied to the affected area, however, the clip tore through the area where the clip was applied causing significant hemorrhage from the pelvic vessels. The surgeon made the decision to convert the planned surgical procedure to open laparotomy techniques. The patient expired on (b)(6) 2012. | Manufacturer Narrative: --/-- A review of the site's system log found no errors for surgeries performed on the date of the reported event ((b)(6) 2010). --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2012, isi contacted (b)(6) in risk management at (b)(6) to obtain additional information concerning the reported event, however, she declined to provide any information concerning the reported event due to litigation constraints. If additional information is received, a follow-up mdr will be submitted to the fda. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2592098 | 2563461 | 2460174 | 2955842201200198 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2012-04-05 | 2012-05-08 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-04-05 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-04-05 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the site's electrical surgical unit would not function while connected to the system and the site was unable to take control of 2 of the patient side manipulator (psm) arms. Unable to resolve the issues and prior to contacting isi for technical support assistance, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the cautery issue experienced by the site was associated personality module energy device (pmed) and the issue with the patient side manipulator (psm) arms were associated with the camera arm setup joint (suj) potentiometer. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from its own joint positions to the remote arm controller. The system was repaired by replacing the affected suj. The pmed is designed to interface and to control functions within various cautery generators or esu's (electro-surgical unit), which are activated by a foot pedal switch. The pmed provides enhanced control, enables multiple pedal inputs, multiple or expanded generator control and system visibility of foot pedal stimulus. The system was repaired by replacing the affected pmed and potentiometer. As of may 4, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2592102 | 2563483 | 2460196 | 2955842201200199 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-04-05 | 2012-05-08 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-04-05 | Device Not Retu | No | Reuse | 2012
42019408 | 2012-04-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a piece of the mega needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal pin was dislodged from the proximal clevis. The proximal pin is broken and a. 100 long segment of the pin that contains the pin head was also returned. The pin fractured just before it enters the distal clevis. One ear of the distal clevis exhibits scratch marks next to the pin hole. As a result of the pin breakage, three proximal pulleys are missing. One pulley was returned. One grip close cable was also found broken at the distal idlers. The idler pulley spins freely, however, the outer edge exhibits minor damage. The other cables at the wrist are not damaged. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2591304 | 2562495 | 2459208 | 2955842201200200 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-12 | 2012-05-07 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-04-12 | Device Not Retu |  | Reuse | 2012
| 2012-03-28 | Malfunction |  | Permanent cautery spatula instrument broke inside of patient. A small plastic piece broke off and we were able to locate and remove it from the abdomen. Instrument was removed from service. |  |  | INSTRUMENT ROBOTIC SURGICAL PE | 2590956 | 2562171 | 2458884 | 2562171 | 1 | NAY | User Facility |  |  | Initial | 2012-05-06 | 2012-05-06 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
42807203 | 2012-03-26 | Malfunction | Required Intervention | It was reported that during a da vinci si single-site pyeloplasty procedure, the surgeon noticed an instrument moving strangely due to the 5 x 250 mm curved cannula arm 2 accessory becoming unsoldered between the shaft and bowl. This resulted in extra rotation of the instrument shaft and consequently, the wrong orientation of the instrument. The hospital did not have a backup 5 x 250 mm curved cannula arm 2 accessory, therefore, the surgeon made the decision to convert to da vinci multi-port techniques to complete the planned procedure. No pt harm was reported. | The cannula accessory was returned and evaluated. The cannula was visually inspected and no damage was observed. The cannula was put on a system with an instrument installed and driven. The cannula was found to be fully functional with no issues detected. The customer reported failure could not be confirmed. | SINGLESITE 5 X 250 MM CURVED CANNULA ARM 2 | ENDOSCOPIC INSTRUMENT ACCESSOR | 2629317 | 2591810 | 2488503 | 2955842201200196 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-04-02 | 2012-05-04 | Yes | Yes | Health Professional | 42807203 | Yes | No | 2012-04-02 | Yes | No | Unkown | 2012
IS2000 A51P8 | 2011-04-18 | Other | DeathRequired Intervention | On april 4, 2012 intuitive surgical was contacted by the patient's grand-daughter that the patient underwent a kidney biopsy procedure using the da vinci s surgical system on (b)(6) 2011. Reportedly, during the biopsy procedure, the patient's renal artery was damaged. The planned procedure was converted to open surgical techniques to repair the affected vessel and the patient required a partial nephrectomy as a result of the damage. The patient lost 13 pints of blood. On (b)(6) 2012, the patient expired due to pneumonia, multiple organ failure and renal failure. | Several follow up attempts were made to the initial reporter to obtain the patient's medical records and a video recording of the planned surgical procedure, however, the information has not yet been provided. Should additional information be received concerning this report, a follow-up mdr will be submitted to the fda. On april 13, 2012, intuitive surgical contacted the surgeon, (b)(6) and he indicated that the intended planned procedure was a da vinci s partial nephrectomy procedure, however, he is unable to provide any other information concerning the reported event and that i should contact the hospital's risk management department. On april 20, 2012, intuitive surgical was contacted by attorney (b)(6) and she indicated that due to hippa regulations she is unable to provide additional information concerning the reported event to 3rd parties. However, the hospital will comply with any request for additional information directly from the fda. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2588470 | 2559247 | 2455960 | 2955842201200197 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2012-04-04 | 2012-05-03 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2012-04-04 | No | No | Reuse | 2012
IS2000 A51P8 | 2012-02-21 | Other | Other | It was reported that after a da vinci s prostatectomy surgical procedure the patient developed a postsurgical complication. The patient had abscess in the pelvis. The site has advised an anastomosis insufficiency caused the patient injury. The delay in reporting is due to vague event description and root cause determination by the user facility. As indicated, an investigation is ongoing, and until further root cause determination can be confirmed, this event shall be reported. | An investigation by hygiene specialist at (b)(6) tested the instruments used in the reported case, the test concluded to find no malfunction of intuitive surgical instruments. The instruments were also tested for bacterial contamination, the test revealed no contamination of intuitive surgical instruments. The hygiene specialist could not make a definite root cause determination, but did conclude after extensive analysis of the reported event, that the post-surgical complications of the patient are not a result of any potentially contaminated instruments. Microbiological findings of the post-surgery examinations and an anastomosis insufficiency in the operating field confirmed this. The hygiene specialist further noted that the infection causing bacteria in the pelvic cavity was determined to be a mixed flora, a. O. E. Coli, but the account confirmed there was no intestine leakage. Information recently received from (b)(6) stated that the abscess experienced by this site was found to be related to leak in anastomosis in which they have reported as a known complication for prostatectomy procedures performed and not a deficiency of the da vinci surgical system or the associated instruments or accessories. Our clinical research supports this finding and indicates that insufficiencies in anastomosis can occur regardless of technique used in order to achieve anastomosis in robotic assisted laparoscopic procedures. Please see cited article: williams, s. B. , m. Alemozaffar, et al. (2010). Randomized controlled trial of barbed polyglyconate versus polyglactin suture for robot-assisted laparoscopic prostatectomy anastomosis: technique and outcomes. European urology 58(6): 875-881. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2586045 | 2557008 | 2453721 | 2955842201200170 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-02-21 | 2012-05-02 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2012-02-21 | Device Not Retu | No | Reuse | 2012
42018306 | 2012-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff noticed a melted hook and possible crack in the wrist of the permanent cautery hook instrument. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring on the yaw pulley and localized melting. The metal hook is not melted. The charring is a sign of unintended arcing. Engineering also observed a disconnected wire. The instrument failed electrical testing, indicating damage to the conductor wire. Disassembly of the instrument shows the wire was disconnected from the hook at the yaw pulley exit. This likely created a path for arcing. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2586050 | 2557043 | 2453756 | 2955842201200191 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-03 | 2012-05-02 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-04-03 | Device Not Retu | No | Initial | 2012
40027202 | 2012-02-09 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure the housing of the harmonic ace curved shears instrument would not hold the insert accessory, allowing the insert to break and fall into the patient. The broken piece was retrieved. The same issue occurred with two additional inserts used with the same instrument. All broken pieces were retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. Mdrs 2955842-2012-00192, and 295842-2012-00193 were submitted for the second and third insert breakages. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 050 x. 025 piece of the teflon pad broken off at the distal tip. The broken piece was not returned. The curved blade was not damaged. No other damage was found. The damage is likely the result of mishandling/misuse. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. The harmonic ace curved shears instrument was also returned and evaluated. The reported failure could not be confirmed. Visual inspection shows all backend components are intact. A fit check with the returned insert and a known good insert passed. Instrument was installed on in-house system with the same insert and engagement and driving of insert was successful with no issues encountered. No physical damage was found on the instrument. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2586074 | 2557058 | 2453824 | 2955842201200194 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-04-05 | 2012-05-02 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-04-05 | Device Not Retu | Yes | Initial | 2012
40027202 | 2012-02-09 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure the housing of the harmonic ace curved shears instrument would not hold the insert accessory, allowing the insert to break and fall into the patient. The broken piece was retrieved. The same issue occurred with two additional inserts used with the same instrument. All broken pieces were retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. Mdrs 2955842-2012-00194, and 295842-2012-00193 were submitted for the first and third insert breakages. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 050 x. 025 piece of the teflon pad broken off at the distal tip. The broken piece was not returned. The curved blade was not damaged. No other damage was found. The damage is likely the result of mishandling/misuse. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. The harmonic ace curved shears instrument was also returned and evaluated. The reported failure could not be confirmed. Visual inspection shows all backend components are intact. A fit check with the returned insert and a known good insert passed. Instrument was installed on in-house system with the same insert and engagement and driving of insert was successful with no issues encountered. No physical damage was found on the instrument. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2586080 | 2557096 | 2453824 | 2955842201200192 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-04-05 | 2012-05-02 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-04-05 | Device Not Retu | Yes | Initial | 2012
40027202 | 2012-02-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si nissen fundoplication procedure the housing of the harmonic ace curved shears instrument would not hold the insert accessory, allowing the insert to break and fall into the patient. The broken piece was retrieved. The same issue occurred with two additional inserts used with the same instrument. All broken pieces were retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. Mdrs 2955842-2012-00194, and 295842-2012-00192 were submitted for the first and second insert breakages. | Manufacturer Narrative: --/-- A correction is required to patient identifier, previously it was referenced as (b)(6), it should be referenced as (b)(6). This was originally submitted on mfr report 2955842-2012-00193. --/-- Additional Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 050 x. 025 piece of the teflon pad broken off at the distal tip. The broken piece was not returned. The curved blade was not damaged. No other damage was found. The damage is likely the result of mishandling/misuse. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. The harmonic ace curved shears instrument was also returned and evaluated. The reported failure could not be confirmed. Visual inspection shows all backend components are intact. A fit check with the returned insert and a known good insert passed. Instrument was installed on in-house system with the same insert and engagement and driving of insert was successful with no issues encountered. No physical damage was found on the instrument. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | Initial | 2586093 | 2557111 | 2453824 | 2955842201200193 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2012-04-05 | 2012-05-02 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-04-05 | Device Not Retu | Yes |  | 2012
42018404 | 2012-04-03 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the customer experienced difficulty removing the permanent cautery spatula (pcs) instrument from the cannula and pieces from the instrument broke off and fell into the patient. The instrument fragments were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal clevis is detached from the main tube with damage at the distal end of the main tube with pieces removed and missing. Engineering concluded that damage to the instrument is likely due to excess force applied to the clevis. Clevis. The instruments and accessories user manual specifically states: general precautions and warnings - please refer to the individual esu manufacturer's operator's manual for operating instructions. Set the esu ti bipolar output. Set the power as low a possible to achieve adequate hemostasis - do not use bipolar instruments with a moopolar source output as this may cause damage to the instrument and harm to the patient or medical personnel. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2586450 | 2557314 | 2454027 | 2955842201200195 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-04-03 | 2012-05-02 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2012-04-03 | Device Not Retu | No | Reuse | 2012
| 2012-04-27 | No Answer Provided |  | End of instrument fell off in the patient. It was retrieved fromthe patient. |  | ENDOWRIST ONE | INSTRUMENT ROBOTIC SURGICAL SU | 2590245 | 2561626 | 2458339 | 2561626 | 1 | NAY | User Facility |  |  | Initial | 2012-05-01 | 2012-05-02 | No | No | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS1200 A43P9 | 2012-04-03 | Other | Other | It was reported that during a da vinci hysterectomy procedure the image appeared dark and grainy. Prior to contacting intuitive surgical technical support for assistance, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by an intuitive surgical field service engineer (fse) who went onsite and evaluated the system. System verification testing was performed with no issues observed and the system was found to be performing to specification. The fse also evaluated the hospital's six endoscopes, three light guide cables, and two illuminators. No issues were detected with any of these items. As of april 26, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2584140 | 2554995 | 2451708 | 2955842201200186 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-03 | 2012-05-01 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2012-04-03 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-04-03 | Other | Other | It was reported that while preparing for a da vinci si prostatectomy procedure the steering mechanism on the patient cart was not working and the surgical staff was unable to guide the patient cart to the patient for docking. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by an intuitive surgical field service engineer (fse) who went onsite and evaluated the system. The fse observed that the pin on the tiller handle of the patient cart was sheared off, thus causing the steering issue experienced by the customer. The system was repaired by replacing the tiller assembly. As of april 26, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2584141 | 2554996 | 2451709 | 2955842201200185 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-04-03 | 2012-05-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-04-03 | Device Not Retu | No | Reuse | 2012
| 2012-04-18 | Malfunction |  | The physician completed the part of the procedure that required the shears. He disconnected the cord that powers the pedal to activate the shears and moved the robot away from the patient. The team heard the shears/scalpel activate despite being disconnected. The patient sustained a small burn. |  |  | INSTRUMENT ROBOTIC SURGICAL HA | 2590241 | 2561624 | 2458337 | 2561624 | 1 | NAY | User Facility |  |  | Initial | 2012-05-01 | 2012-05-01 | No | Yes | Invalid Data |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
40018010 | 2012-03-26 | Malfunction |  | It was reported that during a da vinci si total hysterectomy procedure, a spark from the monopolar curved scissors (mcs) tip cover accessory was observed. The mcs instrument was immediately removed and a replacement tip cover was installed to complete planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone and mechanical tears. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole size is approximately. 010 in diameter and is located 0. 175 from the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2581107 | 2552320 | 2449033 | 2955842201200187 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2012-03-28 | 2012-04-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2012-03-28 | Device Not Retu | Yes | Initial | 2012
420184 | 2012-03-28 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, pieces of the permanent cautery spatula instrument broke off and fell into the patient. One broken instrument piece was retrieved, however, another broken piece could not be located. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2581135 | 2552325 | 2449038 | 2955842201200188 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-28 | 2012-04-29 | No | Yes | Health Professional | 420184 | No | Yes | 2012-03-28 | No | No | Unkown | 2012
40027202 | 2012-03-29 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hernia repair procedure a piece of white scalpel from the harmonic ace curved shears insert accessory (installed on the harmonic ace curved shears instrument) fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the insert accessory is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the entire white teflon pad is missing from the clamp arm. The pad was not returned with the insert. The clamp arm slot that holds the teflon pad is not damaged and the width measures within specification. There was no physical damage observed at the distal end. Device manufacture date: 01/2012. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | Unkown | 2581103 | 2552345 | 2449058 | 2955842201200189 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2012-03-29 | 2012-04-29 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-05-09 | Device Not Retu | Yes |  | 2012
IS3000 A60P6 | 2012-03-28 | Other | Other | It was reported that during a da vinci si gastric bypass procedure the surgeon was unable to move patient side manipulator (psm) arm 2 from the surgeon console. After restarting the system, psm arm 2 worked fine. After 15 minutes of operation the surgeon lost control of psm arm 1, resulting in damage to the patient's stomach. No further injuries were reported. | The system logs were reviewed by an isi technical support engineer and no errors were observed which indicate that there were issues with the psm arms. The psms are instrument arms located on the patient cart that provide the sterile interface for the endowrist instruments. The root cause of the surgical complication experienced by the patient is currently undetermined. Multiple attempts have been made to gain additional information about the event, but the customer has yet to provide further details. If additional event details become available a follow up mdv will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2581123 | 2552349 | 2449062 | 2955842201200184 | 1 | NAY | Manufacturer | Foreign | Other | Initial | 2012-03-28 | 2012-04-29 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No | 2012-03-28 | Device Not Retu | No | Reuse | 2012
400180 | 2007-09-28 | Other | Other | It was reported that the patient underwent a da vinci s appendectomy and oophorectomy procedure on (b)(6) 2007, to alleviate pelvic pain. Post-operatively the patient's pelvic pain became increasingly worse. A ct scan of the patient's pelvic area revealed that the patient had retained a foreign body and a diagnostic laparoscopy procedure performed on (b)(6) 2011, to remove the foreign body revealed that it was an mcs tip cover accessory used in conjunction with a monopolar curved scissors instrument that was used during the da vinci s appendectomy and oophorectomy procedure. | The mcs tip cover accessory has not been returned for evaluation, therefore, the root cause of the reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, intuitive surgical followed up with the surgeon, dr. (b)(6) at (b)(6) health services and he indicated that during the surgical procedure, a malfunction of the da vinci s system, instruments and accessories was not noted and it is unknown how the mcs tip cover accessory came off of the monopolar curved scissors (mcs) instrument during the surgical procedure. In addition, it is unknown, if the tip cover accessory and mcs instrument were inspected prior to use and if an installation tool was used to install the tip cover onto the mcs instrument, however, there were no issues noted by the surgical staff while installing the tip cover accessory onto the mcs instrument. The planned surgical procedure was completed and no patient intra-operative complications were appreciated at the time of the surgical procedure. Review of the patient's operative report provided by the initial reporter found that an endo gia stapler was also used during the surgical procedure and a slight malfunction of the device occurred which required reloading. A review of the site's system log found that no malfunction of the site's da vinci s surgical system occurred. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 2581126 | 2552353 | 2449066 | 2955842201200190 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-03-29 | 2012-04-29 | Yes | Yes | Health Professional | 400180 | No | No | 2012-03-29 | No | No | Unkown | 2012
42030901 | 2012-03-12 | Malfunction | Required Intervention | On (b)(4) 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from (b)(6) medical center with the following information: cable on robotic instrument frayed and broke while in use. Under visualization, the surgeon was not able to locate any missing portion. The instrument was removed and replaced. X-ray required to rule out retention of frayed/broken cable. No foreign body detected. No harm to patient. Remaining lives: (b)(6). | The instrument has not been returned for evaluation despite several requests, therefore, the root cause of the reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2576700 | 2547980 | 2444673 | 2955842201200180 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-03-26 | 2012-04-25 | Yes | Yes | Health Professional | 42030901 | No | No | 2012-03-26 | No | No | Reuse | 2012
42019408 | 2012-03-30 | Malfunction |  | It was reported that while performing a da vinci si hysterectomy procedure, a cable on the mega needle driver instrument broke and a piece fell into the patient. The surgeon was immediately able to retrieve the piece and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable was broken at the distal idlers. The idler pulley spins freely and was not damaged. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2577432 | 2548769 | 2445492 | 2955842201200181 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-04-02 | 2012-04-25 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-04-02 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-03-26 | Death | Death | It was reported that during a da vinci si ovarian transposition procedure, a complication occurred and the patient started to bleed. The surgeon attempted to control bleeding by applying pressure with an instrument grasper and applying laparoscopic ligaclips. The ligaclips failed to control the bleeding and the surgeon attempted to suture the vessel, but this was also unsuccessful at controlling bleeding. The surgeon made the decision to convert to traditional open surgical techniques. The surgical staff undocked the camera arm and one of the instrument arms from the patient but kept the other two arms with the instruments in the patient. The instruments arms were applying pressure to the site of bleeding. The surgeon attempted to find the source of bleeding. It is unclear on the sequence of events after this. The patient was confirmed to have expired. It is unknown whether the patient expired during or after the procedure. No further information was provided. | The system was tested by a field service engineer and performed to specifications. The account stated the case complications are unrelated to the da vinci system. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. As of (b)(4) 2012, no adverse events have been experienced with the site's system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2577408 | 2548815 | 2445528 | 2955842201200182 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-03-26 | 2012-04-25 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No |  | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-03-26 | Other | Other | It was reported that prior to starting a da vinci s hysterectomy procedure, the surgical staff reported seeing a pink image in the left eye in the surgeon console and touch screen. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue. The site switched their camera cable with a back up; however this did not resolve the problem. The patient was under anesthesia and the port incisions were made when the surgical staff made the decision to abort the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer concluded that the customer reported issue was associated with both the initial camera cable used and the backup camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cables. One of the camera cables was returned to intuitive surgical for failure analysis investigation. During evaluation, the reported failure mode was replicated and the camera cable was found to have a broken connection. As of april 24, 2012, there have been no reported recurrences of the incident. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2577520 | 2548914 | 2445607 | 2955842201200183 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-26 | 2012-04-25 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-03-26 | Device Not Retu | No | Reuse | 2012
IS2000 A60P6 | 2009-02-09 | No Answer Provided | OtherRequired InterventionHospitalization | The following patient story concerning a da vinci s prostatectomy procedure was found posted on the (b)(4) by intuitive surgical on (b)(6), 2012: i was diagnosed with prostate cancer and was convinced the davinci robotic assisted prostatectomy was the correct method. I was able to get in to see a new urologist in the clinic, who had trained on the davinci. I usually require a surgeon who has performed at least a minimum of 250 or the procedures. The equipment was fairly new, so no one seemed to have the experience. I was very positive going into the procedure. I awoke from the procedure to find it had not gone well. It had lasted approximately over an hour longer than expected. He explained the equipment had failed during the procedure and they had spent an hour trying to restart and reprogram the davinci. He then explained they elected to do a radical prostatectomy, and he had accidently cut two blood vessels that were bloated from the wait. They had given 6 units of blood due to the amount of blood loss according to what he told me. My bp dropped to 50/30 that night and i awoke to find an icu team just leaving my room. One person stayed with me to monitor me to be sure i did not need further assistance. I then went into affib and they considered taking me to shock me into the correct fibrolation when it corrected itself. Then i was told i may have possible kidney/liver or other problems, than proved to be false. Then i was told my blood sugar had become elevated. Meantime the doctor had to leave town on personal business and his associate would see me plus the hospitalist. They were both very competent and reassuring. I was sent home with orders to begin taking insulin for diabetes. That also did not stay elevated. The doctor saw me the next week and apologized multiple times and it was the last i saw or heard from him. I returned for my next visit to find he had left town. Be sure your surgeon has at least 250 successful procedures. I still recommend the davinci procure. I still recommend (b)(6) and dr (b)(6) that did my follow up. I just had to have an artificial sphincter installed to control my incontinence. The patient also reported that he was hospitalized for 5 days, experienced multiple post-operative complications in the hospital as a result of the surgical procedure performed by the surgeon, his recovery time took 9 months and he was able to return to normal activity after 6 months. | On (b)(6), 2012, isi contacted (b)(6) in the quality department at (b)(6) medical center in (b)(6) and she indicated that due to hippa regulations, she is unable to provide any information concerning the reported event. (b)(6) explained to (b)(6) that the hippa regulations are not applicable to investigations related to adverse events; however, she reiterated that she is unable to provide any information. At this time, (b)(6) is unable to provide additional information concerning the reported event. A follow-up mdr will be submitted to the fda if additional information becomes available. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2575882 | 2547370 | 2444073 | 2955842201200179 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-03-26 | 2012-04-24 | Yes | No | Health Professional | IS2000 A60P6 | No | No | 2012-03-26 | No | No | Reuse | 2012
42018306 | 2012-03-23 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the cables at the tip of the permanent cautery hook instrument frayed and a piece fell into the patient. The fragmented piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2571503 | 2543469 | 2440244 | 2955842201200173 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-23 | 2012-04-20 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-03-23 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-03-23 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff experienced non-recoverable system error code 2. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 2 was associated with the left master tool manipulator (mtm). The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 2 appears when the da vinci s safety system determines a reference voltage was out of range. Upon determining this condition, the system records the fault and transitions to a safe state. The mtml was returned to intuitive surgical for failure analysis investigation. Engineering was unable to replicate the reported failure during internal testing, however, an axis potentiometer and motor were replaced as a precaution. As of april 19, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2571517 | 2543497 | 2440229 | 2955842201200174 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-03-23 | 2012-04-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-03-23 | Device Not Retu | No | Reuse | 2012
420230 | 2012-03-21 | Death | Death | It was reported that after a da vinci si partial lobectomy procedure, while the patient was in recovery, the patient's blood pressure dropped. The drop in blood pressure was due to significant bleeding. The patient underwent emergency open surgery and the clip attached to the pulmonary artery was found to have come off. It was believed that the clipped tissue may have been sustained damage due to previous cancer treatment which included radiation exposure. The patient was provided an unknown amount of blood and the surgical staff attempted to control the bleeding without success. The patient expired. | This investigation is currently ongoing. Isi has made several attempts to gain additional information concerning the reported event, however, no additional information has been provided. If additional information is received a follow-up mdr will be sent to the fda. | HEMOOLOCK CLIP APPLIER LARGE | ENDOSCOPIC INSTRUMENT | 2571557 | 2543538 | 2440271 | 2955842201200176 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-21 | 2012-04-20 | Yes | No | Health Professional | 420230 | No | No | 2012-03-23 | No | No | Unkown | 2012
420230 | 2012-03-21 | Other | Required InterventionHospitalization | It was reported that after a da vinci si partial lobectomy procedure, while the patient was in recovery, the patient's blood pressure dropped. The drop in blood pressure was due to significant bleeding. The patient underwent emergency open surgery and the clip attached to the pulmonary artery was found to have come off. The bleeding was able to be controlled and the patient received an unknown amount of blood. The patient was reported as recovering from the event with no additional injuries reported. | This investigation is currently ongoing. Isi has made several attempts to gain additional information concerning the reported event, however, no additional information has been provided. If additional information is received a follow-up mdr will be sent to the fda. | HEMOOLOCK CLIP APPLIER LARGE | ENDOSCOPIC INSTRUMENT | 2571579 | 2543539 | 2440271 | 2955842201200175 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-21 | 2012-04-20 | Yes | No | Health Professional | 420230 | No | No | 2012-03-23 | No | No | Unkown | 2012
42018305 | 2012-03-16 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a small round plastic piece broke off of the permanent cautery hook (pch) instrument and fell into the patient. The surgeon was able to retrieve the broken instrument piece and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with the distal clevis ear broken and with the conductor wire cap detached from the clevis. Engineering concluded that excess force was likely applied to the clevis, causing the clevis ear to break. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2572116 | 2543946 | 2440677 | 2955842201200178 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-03-21 | 2012-04-20 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2012-03-21 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-03-23 | Injury | Other | It was reported that during a da vinci si pulmonary lobectomy procedure, the patient's pulmonary artery was lacerated, requiring hemostasis of the affected vessel using a cadiere forceps instrument to allow for completion of the planned surgical procedure using open surgical techniques. After obtaining adequate exposure, the site attempted to manually open the grips on the cadiere forceps instrument using several different allen wrenches, however, the instrument grips would not release. The site then attempted to release the instrument using the master tool manipulors (mtm), however, the error message unable to align masters was observed. Unable to resolve the issue, the site contacted isi technical support engineering (tse) for troubleshooting assistance. After performing an emergency stop/recover of the system, the grips on the cadiere forceps instrument opened. No other information concerning the reported event was provided. | On (b)(6) 2012, isi's supervisor of technical support engineering contacted the surgeon, dr. (b)(6) at (b)(6) to discuss the difficulties experienced while attempting to release the cadiere forceps instrument. Dr. (b)(6) indicated that after damage to the patient's pulmonary artery occurred, he used a cadiere forceps instrument to clamp down on the vessel to stop the blood flow to allow for conversion of the procedure to open. Dr. (b)(6) made the decision to release the cadiere forceps instrument from the patient bedside, instead of from the surgeon side cart. However when the surgical staff attempted to use an allen wrench to release the instrument, the instrument would not release. The tse reiterated that the instructions in the system user manual states that the emergency grip release mechanism facilitates removal of an instrument in the event of a system fault. Dr. Abbas was advised to perform an emergency stop in these situations to remove instruments from the patient side without a system error. Several follow up attempts with the site have been made concerning the patient's status, however, no additional information has been provided. A follow-up mdr will be submitted to the fda if additional information becomes available. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2572147 | 2543963 | 2440695 | 2955842201200177 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-23 | 2012-04-20 | Yes | Yes | Health Professional | IS3000 A60P6 | No | Yes | 2012-03-23 | No |  | Reuse | 2012
|  | Injury |  | Event Description: --/-- On (b)(6) 2012, intuitive surgical was provided a letter sent to (b)(6) from (b)(6), alleging prolonged procedures up to 11 hours can result in affects to the patient such as thrombosis, embolism, compartmental syndrome and hematomas. Additionally, an allegation was made that a patient experienced an abscess after a procedure. The letter also alleged general problems with proper sterilization of instruments. No site specific information was provided within the body of this letter. Several attempts have been made to contact the author of the letter, (b)(6) and staff, both via email and phone; however, as of the date of this report no returned communication has been received. | Manufacturer Narrative: --/-- Intuitive surgical made multiple attempts to verify the information as indicated in the letter received from (b)(4) and based on the investigation intuitive surgical was able to verify information concerning only one of the allegations, which concerned a patient that developed an abscess post-op a da vinci prostatovesiculectomy procedure. Intuitive surgical's investigation is as follows: on (b)(6) 2012 intuitive surgical contacted (b)(6) who initially distributed the dztim letter to (b)(4) who forwarded the letter to intuitive surgical. (b)(6) was out of office. On (b)(6) 2012 intuitive surgical contacted professor (b)(6) at the clinic in (b)(6) who is the author of the letter. No response was received. On (b)(6) 2012, intuitive surgical contacted (b)(6) at the (b)(6) and she indicated that she was not aware of a patient-abscess report and referred intuitive surgical to (b)(6) also at (b)(6). (b)(6) indicated that she did not have any information concerning the clinic where the abscess case occurred and she recommended that intuitive surgical proactively contacted all intuitive surgical customers in the area and to contact (b)(6). Multiple attempts to reach professor (b)(6) were made from (b)(6) 2012 through (b)(6) 2012, however, no response has been received. However, on (b)(6) 2012 intuitive surgical received an email from a (b)(4) attorney indicating that intuitive surgical should contact (b)(4) concerning the incident. Intuitive surgical responded back to (b)(4) indicating that there would be difficulty identifying the incident with (b)(4) without information concerning where the incident occurred. On (b)(4) 2012, intuitive surgical received an email from the attorney at (b)(4) indicating that the affected patient underwent a da vinci prostatovesiculectomy procedure in july 2011 at (b)(6) and post surgical complications experienced by the patient (abscess) was treated at clinics for urology of the (b)(6) in october 2011. On (b)(4) 2012 intuitive surgical sent an email to (b)(6). On (b)(4) 2012, intuitive surgical sent a reminder email to (b)(6). On (b)(4) 2012, intuitive surgical received an email response from (b)(6) the chief physician of robotic urology at (b)(6) indicating that review of their records found a da vinci prostatectomy procedure performed in july 2011 by (b)(6) that fit the patient description as indicated in the initial communication to intuitive surgical dated (b)(4) 2012. (b)(6) indicated that based on their hospital records concerning the surgical procedure, there were no unusual conditions that occurred during or post -op the surgical procedure. (b)(6) indicated that (b)(6) had no prior knowledge that the patient was being treated at (b)(6) 3 months post-op due to post surgical complications and had not received any communication from (b)(6) concerning the incident. (b)(6) indicated that due to the unusual germ spectrum, (b)(6) reviewed their sterilization documentation and review of the documentation did not reveal the source of the contamination. (b)(6) indicated that due to patient confidentially, he is unable to provide any other details concerning the patient. --/-- Additional Manufacturer Narrative: --/-- Intuitive surgical has made several attempts to contact the initial author of the letter and staff members of this (b)(6) to gain additional information regarding these allegations; however, no response has been received. A follow-up mdr will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | Unkown | 2570424 | 2542327 | 2439080 | 2955842201200172 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2012-03-20 | 2012-04-19 | Yes | No | Service Personnel |  | No | Yes | 2012-06-15 | Device Not Retu | No |  | 2012
| 2012-01-12 | Malfunction |  | During the procedure, the surgical first assistant noted a spring on the needle driver was not in its expected position, and was poking out of the device. The uses left on this particular device were 8 of 10. The needle driver was taken out of the field. |  | MEGA SUTURECUT NEEDLE DRIVER | INSTRUMENT ROBOTIC SURGICAL NE | 2573176 | 2544891 | 2441623 | 2544891 | 1 | NAY | User Facility |  |  | Initial | 2012-04-10 | 2012-04-19 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 |  | Death | Death | Event Description: --/-- It was reported that during a da vinci si bilateral salpingo-oophorectomy procedure, during placement of the uterine manipulator, the patient's bowel was ruptured. The surgeon repaired the patient's bowel, which took approximately 3 hours to repair. Two days post op the patient experienced cardiovascular complications and complications and expired. The surgeon is attributing the patient's demise to a t-berg. No other information was provided. | Manufacturer Narrative: --/-- Isi has made multiple attempts to contact the physician (initial reporter) at (b)(6), for additional information concerning the reported event, however, no additional information has been provided. If additional information is received a follow-up mdr will be sent to the fda. At this time, there has been no report by the site that during the planned surgical procedure that a malfunction of the da vinci si system, instruments and/or accessories occurred. --/-- Additional Manufacturer Narrative: --/-- Additional information received from the account indicates the patient death was the result of a heart attack and was not related to the surgery in this event. No further information provided. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2568339 | 2539834 | 2436587 | 2955842201200168 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2012-03-19 | 2012-04-18 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-04-19 | Device Not Retu | No |  | 2012
400180 | 2012-03-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure an arc was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was removed and upon removal, a small hole was noticed in the corner of the tip. A replacement tip cover was installed and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The hospital indicated that the tip cover accessory was discarded. Since the affected device is not returning for evaluation, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2568322 | 2539841 | 2436594 | 2955842201200166 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-19 | 2012-04-18 | No | Yes | Health Professional | 400180 | No | No | 2012-03-19 | No | Yes | Unkown | 2012
400180 | 2012-03-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure an arc was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The arc caused a slight burn to the patient's uterus, however, the uterus was being removed as part of the planned procedure. The tip cover was removed and upon removal, a round hole was observed. A replacement tip cover was installed to complete the procedure. No additional patient harm, adverse outcome or injury was reported. | The hospital indicated that the tip cover accessory was discarded. Since the affected device is not returning for evaluation, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2568353 | 2539847 | 2436600 | 2955842201200167 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-19 | 2012-04-18 | No | Yes | Health Professional | 400180 | No | No | 2012-03-19 | No | Yes | Unkown | 2012
IS3000 A60P6 |  | Death | Death | Event Description: --/-- It was reported that approximately one year ago, a patient expired after undergoing a da vinci si gynecological procedure. | Manufacturer Narrative: --/-- On (b)(6) 2012, a hospital representative informed intuitive surgical that the patient death had nothing to do with the use of the da vinci si surgical system and the patient issue would have occurred regardless of how the surgical procedure was performed. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012, the initial reporter informed intuitive surgical that the patient death was not related to surgical error, nor was it related to the anesthesia administered to the patient. The initial reporter also indicated that the da vinci si surgical system did not malfunction during the surgical procedure. Due to hipaa regulations, the initial reporter was unwilling to provide any further information. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2568363 | 2539857 | 2436610 | 2955842201200169 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2012-03-19 | 2012-04-18 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-05-04 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-03-19 | Other | Other | It was reported that during a da vinci si prostatectomy procedure the site experienced system error code 31059 after the surgeon knocked a bottle of betadine solution onto the patient side cart (psc) base, causing the betadine solution to flow into the base of the psc. With the assistance of an isi technical support engineer (tse) the site performed several troubleshooting techniques, however, the system continued to error. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineering concluded that the system error coded 31059 was associated with the remote arm controller (rac). The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 31059 occurs when one part of the internal system communication network fails in the middle of sending a message. Upon determining these conditions, the safety systems put da vinci si in a recoverable safe state. The rac was returned to isi for failure analysis investigation. Per the customer reported complaint, engineering found that the rac was contaminated with fluid and was able to replicate the system error code 31059 using an isi in-house pca test system. As of april 18, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2568650 | 2540145 | 2436898 | 2955842201200171 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2012-03-19 | 2012-04-18 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-03-19 | Device Not Retu | No | Reuse | 2012
| 2012-04-16 | Malfunction |  | Bipolar da vinci instrument shorted out while in use. Instrument removed from field six of ten uses left. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL BI | 2573146 | 2544881 | 2441613 | 2544881 | 1 | NAY | User Facility |  |  | Initial | 2012-04-18 | 2012-04-18 | No | Yes | Physician |  | No Answer Provided | No Answer Provided |  |  |  |  | 2012
IS1200 A43P9 | 2012-03-26 | Other | Other | It was reported that when starting a da vinci prostatectomy procedure, the surgical staff experienced non-recoverable system error code 2 and restarted the system. After restart, the system would not complete the homing process and a message informing the staff to remove the instrument was displayed, even though there was no instrument installed. The instrument sterile adapter was replaced to try and resolve the issue, however, the issue remained. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) determined that system error code 2 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 2 appears when the da vinci safety system determines a reference voltage was out of range. Upon determining this condition, the system records the fault and transitions to a safe state. During evaluation of the affected psm, the fse found dried disinfectant residue on the psm instrument carriage contacts. A hospital employee indicated that liquid disinfectant was used to clean the psm prior to installing the instrument sterile adapter. It was concluded that the residue likely caused a short circuit, thus generating the system error code experienced by the customer. The system was repaired by cleaning the psm contacts. As of april 16, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2564972 | 2537165 | 2433918 | 2955842201200165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-03-26 | 2012-04-16 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2012-03-26 | Device Not Retu | No | Reuse | 2012
410299 | 2012-03-14 | Malfunction |  | Upon using the da vinci endowrist one suction irrigator the tip broke/bent. No pieces broke off into the patient. Patient remained stable and was transported to recovery after the procedure. Taken to sterile processing department (spd). |  | DA VINCI ENDOWRIST ONE | INSTRUMENT ROBOTIC SURGICAL SU | 2569531 | 2541229 | 2437982 | 2541229 | 1 | NAY | User Facility |  |  | Initial | 2012-04-16 | 2012-04-16 | No | Yes | Physician | 410299 | Yes | No Answer Provided |  |  |  |  | 2012
42018406 | 2012-03-14 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, the brown coating on the permanent cautery hook instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the ceramic sleeve broken at the spatula base. A small piece approximately 0. 2 by 0. 4 in size was broken off the ceramic sleeve. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2562510 | 2535263 | 2432016 | 2955842201200149 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-14 | 2012-04-13 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2012-03-14 | Device Not Retu | No | Initial | 2012
IS2000 A51P8 | 2012-03-14 | Other | Other | It was reported that during a da vinci s surgical procedure, the surgical staff experienced system error code 23008. The system was restarted multiple times, however, system error code 23008 continued to occur. Unable to resolve the issue and prior to contacting isi for technical support assistance, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineering concluded that system error code 23008 experienced by the customer was associated with the left master tool manipulator (mtm), as review of the system log found multiple instances of system error code 23008. The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of april 10, 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2562503 | 2535296 | 2432049 | 2955842201200148 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-03-14 | 2012-04-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-03-14 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-03-14 | Other | Other | It was reported that during a da vinci s hysterectomy procedure an instrument arm shut off when being handled during repositioning. Prior to contacting intuitive surgical technical support for assistance, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by an intuitive surgical field service engineer (fse) who went onsite and evaluated the system. The fse reviewed the system error logs, which showed that system error code 23002 occurred and instead of recovering from the error code, the surgeon inadvertently disabled the instrument arm. The fse concluded that system error code 23002 had been generated by vigorous movement of the instrument arm during repositioning. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23002 appears when the da vinci s safety system determines that a joint went beyond its mechanical limit, as measured by that joint's primary control sensor (encoder). Upon determining this condition, the system records the fault and transitions to a safe state. The fse performed system verification testing. No issues were found and the system performed to specification. As of april 12, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2562508 | 2535298 | 2432051 | 2955842201200163 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-14 | 2012-04-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-03-14 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-03-14 | Other | Other | It was reported that 10 minutes into a da vinci s hysterectomy procedure, the surgical staff observed that the endoscope moved freely while positioning it in the endoscopic camera manipulator (ecm) arm. Unable to resolve the issue, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation conducted by intuitive surgical's distributor concluded that the customer reported problem was associated with the endoscopic camera manipulator (ecm) cannula mount. The ecm cannula mount is designed for the camera arm and attaches to a disposable cannula that serves as a port of entry for the endoscope. The ecm cannula mount was broken in half. The system was repaired by replacing the ecm cannula mount. As of april 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2562511 | 2535301 | 2432054 | 2955842201200150 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-03-14 | 2012-04-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-03-14 | Device Not Retu | No | Reuse | 2012
42000606 | 2012-03-16 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the cable broke on the large needle driver instrument and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable broken at the distal idlers. The idler pulley spins freely and is not damaged. The broken cable segment sticks out at the wrist and the other cables at the wrist are not damaged. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2562517 | 2535305 | 2432058 | 2955842201200164 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-16 | 2012-04-13 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2012-03-16 | Device Not Retu | No | Reuse | 2012
ACE | 2012-03-08 | Malfunction |  | Using harmonic scalpel in robot case, the metal tip of the harmonic insert broke off into the patient. Surgeon tried to retrieve it, but due to uterus size was unable to do so immediately. Patient was opened to find it and finish case. X-ray and c-arm were used to assist with finding piece. Counts were correct and patient was stable upon leaving the or. |  | DA VINCI HARMONIC SHEARS | INSTRUMENT ROBOTIC SURGICAL HA | 2567292 | 2538975 | 2435728 | 2538975 | 1 | NAY | User Facility |  |  | Initial | 2012-04-13 | 2012-04-13 | No | Yes | Physician | ACE | Yes | No Answer Provided |  |  |  |  | 2012
ACE | 2012-03-29 | Malfunction |  | During the middle of the case it was noticed that the intuitive device (harmonic shears) were starting to crack and break, resulting in a piece of the disposable item falling off and having to be recovered from the abdominal cavity. We requested that all of this lot number be removed and replaced with a new lot number, which was done. |  | DA VINCI HARMONIC SHEARS | INSTRUMENT ROBOTIC SURGICAL HA | 2569444 | 2541182 | 2437935 | 2541182 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-04-13 | 2012-04-13 | No | Yes | Physician | ACE | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
ACE | 2012-03-29 | Malfunction |  | While using the intuitive harmonic scalpel during a case, the instrument tip melted and started to separate while in the patient's abdominal cavity. We requested that this lot be removed from the facility and replaced with a new lot, which was done. |  | DA VINCI HARMONIC SHEARS | INSTRUMENT ROBOTIC SURGICAL HA | 2569510 | 2541186 | 2437939 | 2541186 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-04-13 | 2012-04-13 | No | Yes | Physician | ACE | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2010-07-05 | Injury | Required Intervention | Event Description: --/-- It was reported that after undergoing a da vinci s hysterectomy procedure on (b)(6) 2010, the patient suffered from necrosis and devascularization allegedly from the electrocoagulation process, which prevented proper healing. Post surgery, the patient experienced continuous pain and drainage. During a postsurgical visit, it was discovered that the patient's bowel was protruding into her vagina and surgery was performed on (b)(6) 2010, to address this issue. | Manufacturer Narrative: --/-- On april 6, 2012, additional information was provided by the physician who performed the surgical procedure. The surgeon informed intuitive surgical that the patient had failed to follow postoperative instructions. The surgeon also reported that no malfunctions occurred with the da vinci s surgical system during the procedure. No additional information was provided. A follow-up medwatch report will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- Initial mdr was inadvertently sent with incorrect information: type of reportable event. | DA VINCI S SURGICAL SYSTEM | Reuse | 2560062 | 2533111 | 2429845 | 2955842201200161 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-01-07 | 2012-04-12 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-03-13 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-03-13 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, when the pk dissecting instrument was installed on a patient side manipulator (psm) arm, the instrument grips would not close properly and the surgeon experienced non-intuitive movement. The site installed a replacement pk dissecting forceps instrument; however, the issue persisted. The site contacted isi technical support engineering (tse) for trouble shooting assistance and was instructed by the tse to remove the instrument and reseat the sterile adapter. The site indicated that the issue had improved and was able to proceed with the procedure. Approximately 20 minutes later, the site contacted isi technical support to report that the issue had returned, the issue was also noted on a second psm. The tse instructed the site to remove all of the instruments installed and perform a power cycle of the system. After powering on the system, the site confirmed with the technical support engineer that the system was functioning with no issues noted; however, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer (fse) was unable to replicate the issue experienced by the customer. However, the fse found that the patient side manipulator (psm) arms cable tensions were low. The fse concluded that the issue experienced by the customer was likely due to a combination of the psm arm low cable tension and a poorly seated or bad sterile adapter. The sterile adapter is part of the disposable drape isolating the non-sterile psm from the sterile field. It attaches to the psm on one side and holds the sterile instrument on the other side. Four disks on the sterile adapter allow motors on the psm to control motion of the instrument wrist. If the connection between the adapter and either the psm or the instrument is incorrect, it can lead to non-intuitive motion. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The fse readjusted the psm arm cable tensions to repair the system. The fse performed system diagnostics and verification and found that the system is functioning within specification. On (b)(6) 2012 isi contacted rn, (b)(6) (surname unknown) at (b)(6) hospital and he indicated that he was present during the surgical procedure and that the site was able to resolve the issue the second time with the assistance of isi technical support engineering, however, after some time the issue reoccurred and the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. There was no harm to the patient, the patient tolerated the open procedure well and to the best of his knowledge, the patient has not returned to the hospital due to experiencing any post-operative complications. As of april 12, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2560783 | 2533805 | 2430559 | 2955842201200162 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-13 | 2012-04-12 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-03-13 | Device Not Retu | No | Reuse | 2012
420003 | 2011-06-10 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | Investigation performed by an intuitive surgical representative found that during the time period that the event occurred, the site was using clips that were not compatible with the small clip applier instrument. Since this event, the surgical staff has been retrained on the proper clips to use. The instruments and accessories user manual specifically states: the small clip applier instrument is designed for use with weck hemoclip small titanium clips (weck product code (b)(4)). | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555224 | 2528171 | 2424924 | 2955842201200155 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
420003 | 2011-09-29 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | Investigation performed by an intuitive surgical representative found that during the time period that the event occurred, the site was using clips that were not compatible with the small clip applier instrument. Since this event, the surgical staff has been retrained on the proper clips to use. The instruments and accessories user manual specifically states: the small clip applier instrument is designed for use with weck hemoclip small titanium clips (weck product code (b)(4)). | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555251 | 2528176 | 2424929 | 2955842201200153 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
42020706 | 2012-01-09 | Malfunction |  | It was reported that after a da vinci si surgical procedure, during a post-op inspection it was noticed the upper left dial on the tenaculum forceps instrument was spinning freely. As a result of the problem with the instrument the account decided to take an x-ray of the patient to determine if any fragments had fallen into the patient. The x-ray did not show a foreign body in the patient. No patient harm, injury or adverse event was reported. | The instrument was returned for evaluation. Per the customer reported complaint the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555261 | 2528180 | 2424933 | 2955842201200151 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-03-13 | 2012-04-11 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-03-13 | Device Not Retu | No | Reuse | 2012
420003 | 2011-04-14 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | Investigation performed by an intuitive surgical representative found that during the time period that the event occurred, the site was using clips that were not compatible with the small clip applier instrument. Since this event, the surgical staff has been retrained on the proper clips to use. The instruments and accessories user manual specifically states: the small clip applier instrument is designed for use with weck hemoclip small titanium clips (weck product code (b)(4)). | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555268 | 2528215 | 2424964 | 2955842201200152 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
420003 | 2011-08-25 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | Investigation performed by an intuitive surgical representative found that during the time period that the event occurred, the site was using clips that were not compatible with the small clip applier instrument. Since this event, the surgical staff has been retrained on the proper clips to use. The instruments and accessories user manual specifically states: the small clip applier instrument is designed for use with weck hemoclip small titanium clips (weck product code (b)(4)). | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555305 | 2528232 | 2424981 | 2955842201200154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
420003 | 2012-02-13 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported complaint cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555286 | 2528248 | 2424997 | 2955842201200156 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
420003 | 2012-02-27 | Other | Required Intervention | It was reported that after a vinci si cardiac procedure was completed, the patient had to be brought back to the operating room due to bleeding. It was discovered that the clip used with the small clip applier instrument had popped off of the patient's internal mammary artery (ima). To address the bleeding, the surgeon increased the port incision that was used for the system camera and manually clipped the patient's ima. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported complaint cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2555329 | 2528256 | 2425025 | 2955842201200157 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-12 | 2012-04-11 | Yes | Yes | Health Professional | 420003 | No | Yes | 2012-03-12 | No | No | Unkown | 2012
42017910 | 2012-03-12 | Malfunction |  | It was reported that during a da vinci si rectopexy procedure, arcing from the side of the monopolar curved scissors (mcs) instrument with the mcs tip cover installed was observed. The mcs instrument was immediately removed and the planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated, however the tip cover accessory was not returned. Engineering was unable to confirm the customer reported failure mode as visual inspection of the instrument did not show any signs of arcing or melting on the distal end components or on the main tube and tube extension. The instrument was placed on an in-house is3000 system and driven. The instrument passed recognition and engagement testing and the instrument moved intuitively with a full range of motion in all directions. The instrument also passed a latex cut test and electrical continuity testing. No physical damage to the instrument was found. On april 4, 2012, isi contacted the scrub nurse at roper hospital and she indicated that inspection of the tip cover used in conjunction with the mcs instrument prior to use and after the event occured did not reveal any damage. The cannula was not inspected; however, there was no noted difficulty in advancing or removing any instruments from the cannula. On april 5, 2012 isi contacted the surgeon who performed the surgical procedure at roper hospital and he indicated that the patient did not sustain any injury due to the reported event. The patient was released from the hospital 4 days post-op, is recovering well and has not returned to the hospital due to experiencing any post operative complications. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2555778 | 2528678 | 2425431 | 2955842201200158 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-12 | 2012-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-03-12 | Device Not Retu | No | Initial | 2012
42018304 | 2012-03-12 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure when the customer attempted to remove the permanent cautery hook (pch) instrument the instrument became stuck on the cannula causing pieces from the instrument's insulation to break off and fall into the patient. A fragment from the instrument was retrieved, however, the surgical staff cannot be certain that there are no remaining fragments inside of the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the proximal clevis is broken at the main tube connection point. A piece approximately 9. 3 x 8. 7 in size broke off the clevis. The main tube distal tip also was damaged with some fibers removed. Engineering concluded that damage was likely due to excessive force applied to the clevis. On april 3, 2012 isi contacted the surgeon and she indicated that the fibroid being removed weighed 385 grams and force on the instrument was used to separate the fibroid from the patient's uterus. One piece from the instrument was recovered, however, she cannot be certain that there is no residual material remaining in the patient's abdomen. No additional surgical procedure was performed to locate the broken instrument pieces, however, it took 30 minutes to locate the broken instrument piece that was retrieved. Per the surgeon, as of april 3, 2012 the patient is recovering well and has not returned to the hospital due to experiencing post-operative complications related to the event. On april 10, 2012 the isi clinical sales representative (csr) present during the surgical procedure indicated that during the surgical procedure, the permanent cautery hook instrument (pch) became non-responsive. The csr instructed the surgical assistant to release and reseat the pch instrument, however, when the surgical assistant pressed the release levers on the instrument, they did not remove the instrument in time with the movement insertion axis on the patient side manipulator, causing the pch instrument to re-engage and catch onto the cannula tip. The force generated from the insertion axis retraction caused the instrument to break. The csr has instructed the surgical assistant on how to properly remove instruments to prevent future occurrences of the reported event. The da vinci si user manual specifically states: general notes and caution * the instrument arm insertion axis will automatically retract when the instrument is removed. If the instrument is not removed in time, the instrument sterile adapter may re-engage the instrument during the insertion axis retraction. If this happens, simply remove the instrument from the instrument arm. Any lateral pressure on the instrument during removal may damage the instrument. As of april 11, 2012 there have been no reported recurrences of the issue at this hospital. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2555874 | 2528777 | 2425546 | 2955842201200159 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-03-12 | 2012-04-11 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2012-03-12 | Device Not Retu | No | Reuse | 2012
42020504 | 2012-03-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the fenestrated bipolar instrument froze and during removal of the instrument, pieces broke off and fell into the patient. The broken instrument pieces were removed and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found that the tip assembly is detached from the distal end. The cables and wires have been cut below the proximal clevis. The proximal clevis exhibits various scratch/gouge marks. Engineering concluded that the likely failure was a dislodged clevis due to overloading the tip and the tip was most likely cut off to remove the instrument from cannula. Engineering observation also found that the main tube interface wall is broken on one side and has missing pieces below the wall. It is likely that some of the main tube damage occurred during the dislodgment and some occurred after initial failure, as the tip was being cut off. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2555789 | 2528791 | 2425524 | 2955842201200160 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-13 | 2012-04-11 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | Yes | 2012-03-13 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-03-12 | Other | Other | It was reported that during a da vinci si procedure the insertion axis on the camera arm would not move. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the staff undock and undrape the camera arm to check for binding, however, no binding was found. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system issue experienced by the customer was associated the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The ecm was returned for evaluation and engineering was able to replicate the reported failure during testing. While performing a visual inspection, engineering found that one of the five bearings attached to the carriage was loose, thus causing the ecm to become stuck. As of april 5, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2546350 | 2520851 | 2417604 | 2955842201200140 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-12 | 2012-04-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-03-12 | Device Not Retu | No | Reuse | 2012
IS2000 | 2012-03-08 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgeon experienced very low light conditions. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was performed by field service engineering, who concluded that the vision issue experienced by the customer was associated the lamp module located inside the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator lamp module. The da vinci s surgical system user's manual specifically states: the illuminator lamp should be replaced every 500 hours. Symptoms indicating the need for replacement include: - repeated ignition attempts without the lamp going into operation. - low light output evident with dark picture at high light level adjustment. In a report provided to the hospital for a previous service activity, the hospital was informed that the lamp module needed replacement. As of april 5, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2546353 | 2520852 | 2417605 | 2955842201200139 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-03-08 | 2012-04-06 | Yes | Yes | Health Professional | IS2000 | Yes | No | 2012-03-08 | Device Not Retu | No | Reuse | 2012
42009308 | 2012-03-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the pin that holds the prograsp instrument tip dislodged and the wheels fell into the patient. The wheels were retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2546375 | 2520864 | 2417617 | 2955842201200141 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-03-12 | 2012-04-06 | No | Yes | Health Professional | 42009308 | No | No | 2012-03-12 | No | No | Unkown | 2012
42018404 | 2012-03-07 | Malfunction |  | Event Description: --/-- On (b)(6) 2012, intuitive surgical received a copy of uf medwatch report (b)(4) from the fda with the following information: event desc: during a robotic surgery, the tip of the robotic spatula instrument broke off while being removed from the robotic arm. Surgeon explored the wound thoroughly and no foreign objects were found in the patient. X-rays were not taken because missing pieces are not radioopaque. The broken pieces were not found. Surgeon aware. Intuitive surgical rep was present and witnessed the incident. This particular instrument was purchased in 2008 and is a limited life instrument that is resterilized in our facility. This particular instrument can be used for 10 times before replacement is required. This device had been used 7 times. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). This product problem had previously been reported to intuitive surgical on (b)(6) 2012. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a break from the proximal clevis to the main tube interface. As a result of this, the clevis is dislodged from the main tube. The proximal clevis and main tube are both missing pieces. It was determined that the damage was likely due to excess force applied to the instrument. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2547610 | 2521938 | 2392040 | 2955842201200142 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | InitialFollowup | 2012-03-07 | 2012-04-06 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2012-03-20 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-03-09 | Injury | Other | Event Description: --/-- It was reported that during da vinci si surgical procedure, the patient was found to have a burn located at port incision site. Intuitive surgical technical service engineer (tse) provided the customer with potential causes that may contributed to the burns at the port incision site. No additional patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- Intuitive surgical investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. Several attempts have been made to gather additional information regarding this event from this site without success. If additional information is received, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by isi field service engineering (fse) was unable to replicate the issue experienced by the site as functional testing of the site's da vinci si system found that the system functioned within specification. The fse also tested the site's electrical surgical unit (esu) in conjunction with the system and it was found to be functioning normally. On (b)(4) 2012 intuitive surgical received uf/importer report numbers (b)(4) and (b)(4) from (b)(6) with the following event details concerning instruments used during the surgical procedure: a thermal burn was noted around the trocar site in the patient's left lower abdomen, which was superficial. The insulation on all instruments used was checked and found to be intact, this includes the needle drivers. The triad device was checked with the ground pad and found to be operating within normal manufacturer specifications. All electrical outlets in the room were checked and functioning properly. Per the surgeon, the port was not touched with the monopolar device. On (b)(4) 2012, intuitive surgical contacted (b)(6) in the risk management department at (b)(6) and she indicated that uf importer report numbers (b)(4) and (b)(4) were submitted to capture the instruments that were installed in the patient side manipulator (psm) arm of the affected port site during the surgical procedure. The issue occurred with one patient and one port site was affected. The patient sustained a very superficial thermal burn 2mm in size. The affected area was excised and no additional treatment of the affected area was required. There was no malfunction of the site's da vinci system, instruments and/or accessories during the surgical procedure. There were no post-operative complications experienced by the patient and as of (b)(6) 2012 the patient has not returned to the hospital due to experiencing post-operative complications as a result of the burn. As of (b)(6) 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2547561 | 2521978 | 2418731 | 2955842201200144 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-08 | 2012-04-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-05-04 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-01-27 | Injury | Required InterventionHospitalizationDisability | Event Description: --/-- On (b)(6) 2012, medwatch uf/importer report (b)(4) was received: event desc: the patient safety concern regarding the davinci robot is a possible design issue. The patient suffered a compartment syndrome of the right lower extremity (rle) after the robotic-assisted laparoscopy (with procedures). It is felt that the robotic arm, while docked on the right side of the patient, either bumped and/or rested on the patient's rle. This was a 7 hour procedure. The patient required emergency, 4-compartment fasciotomy followed by another surgery for skin graft (plus wound vac) and wound closure. What was the original intended procedure? da vinci robotic-assisted laparoscopy with lysis of adhesions, resection of endometriosis, bilateral salpingectomy, cystoscopy, and pelvic floor botox injection. | Manufacturer Narrative: --/-- The system was tested by a field service engineer and performed to specifications. Multiple attempts for additional information from the account were made. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- Several attempts have been made to gather additional information regarding this event from this site without success. The investigation is ongoing and an isi field service engineer is attempting to schedule a site visit to inspect the system and to conduct further investigations. If additional information becomes available, a follow up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2547605 | 2521980 | 2418733 | 2955842201200145 | 1 | NAY | Manufacturer | Other | Nurse | InitialFollowup | 2012-03-08 | 2012-04-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-03-08 | Device Not Retu | No |  | 2012
IS3000 A60P5 |  | Other | Other | It was reported that 2 days post-op a da vinci si gastric bypass procedure, the patient developed septic shock after the jejunum (jj) and gastrojejunum (gj) staple lines and suture closures opened. There was no report of a malfunction of the da vinci system, instruments and/or accessories and it is the surgeon's belief that opening of the jj and gs staple lines and suture closures was related to the patient's anatomy. No other information was provided. | On april 4, 2012 the intuitive surgical contacted (b)(6) hospital's risk manager and the risk manager indicated that she is unable to provide any information concerning the reported event without a patient name and/or date of the surgical procedure. To date, no information has been received which suggests that the da vinci si surgical system may have caused or contributed to the patient's post surgical complications. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2547607 | 2521982 | 2418735 | 2955842201200146 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-09 | 2012-04-06 | Yes | No | Health Professional | IS3000 A60P5 | No | No | 2012-03-09 | No | No | Reuse | 2012
42017910 | 2011-11-04 | Malfunction |  | It was reported that during a da vinci si partial hysterectomy procedure, the surgical staff observed arcing coming from the shaft of the monopolar curved scissors instrument. The instrument was replaced to successfully complete the procedure with no patient harm reported. This record was submitted late due to an error by the complaint analyst during the fda esubmission process. | The instrument was returned and during evaluation was found to have cracks along the main tube near the tube extension connection, likely caused leakage of energy through the main tube. No additional issues noted. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2547629 | 2522009 | 2418762 | 2955842201200147 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-04 | 2012-04-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2011-11-04 | Device Not Retu | No | Reuse | 2012
| 2012-03-28 | Malfunction |  |  |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL CA | 2562979 | 2535536 | 2432289 | 2535536 | 1 | NAY | User Facility |  |  | Initial | 2012-04-06 | 2012-04-06 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 | 2012-03-06 | Other | Other | It was reported that prior to starting a da vinci si partial nephrectomy procedure, the surgical staff reported seeing a yellow tint in the right eye in the surgeon console and touch screen. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue. The site switched their camera head and cable with a back up; however this did not resolve the problem. The patient was under anesthesia and the port incisions were made when the surgical staff made the decision to abort the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer concluded that the customer reported issue was associated with both the initial camera cable used and the backup camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cables. The camera cables were returned to intuitive surgical for failure analysis investigation. During internal evaluation, engineering was able to replicate the reported failure mode. Both camera cables failed during testing. One camera cable also showed bulging at the connector and had scratched insulation. As of april 2, 2012, there have been no reported recurrences of the incident. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2544240 | 2518600 | 2415354 | 2955842201200138 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-06 | 2012-04-05 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-03-06 | Device Not Retu | No | Reuse | 2012
| 2012-02-29 | Malfunction |  | During case, monopolar scissors with cautery for the davinci, caused arc during procedure. Md stated that the arc had happened and then noticed a hole in the protective sheath that is placed on the scissors. Md noted there was damage to the right external iliac artery. The scissors were removed, and the protective sheath was replaced before surgery continued. |  | DA VINCI S TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2543066 | 2517594 | 2414347 | 2517594 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-04-03 | 2012-04-03 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 | 2012-02-18 | Death | Death | It was reported that a patient expired 5 days after undergoing a da vinci si tongue base resection procedure on (b)(6), 2012. No issues occurred during the surgical procedure and the patient was released 3 days post-op. The patient was doing well at the time of discharge, however, he was discovered to have passed away two days later at his home. | Additional information was provided by the physician who performed the surgical procedure. The surgeon confirmed that no system malfunctions occurred during the case. The surgeon also does not believe that the use of the da vinci si surgical system had any direct relationship to the patient's death. The autopsy results are still pending, therefore, the cause of death is still undetermined. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2536563 | 2511072 | 2407823 | 2955842201200131 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-03-01 | 2012-03-30 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-03-01 | No | No | Reuse | 2012
42000306 | 2012-03-02 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, one of the jaws of the small clip applier instrument broke off and fell into the patient. The broken piece was not retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2536562 | 2511106 | 2407819 | 2955842201200134 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-03-02 | 2012-03-30 | No | Yes | Health Professional | 42000306 | No | Yes | 2012-03-02 | No | No | Unkown | 2012
IS1200 A43P9 | 2012-03-20 | Other | Other | It was reported that while homing the system prior to starting a da vinci prostatectomy procedure, the surgical staff experienced non-recoverable system error code 212. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the system records the fault and transitions to a safe state. The ecm was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported failure during testing and determined that the axis-1 motor had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2536568 | 2511108 | 2407821 | 2955842201200132 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-03-20 | 2012-03-30 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2012-03-20 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-03-06 | Other |  | It was reported that during a da vinci si hysterectomy procedure, system error code 31059 occurred. The system was restarted, however, error code 31030 appeared upon restart. With the assistance of an isi technical support engineer via telephone, the surgical staff attempted to resolve the issue by pressing the emergency power off button and reseating the fiber cable multiple times, however, the issue could not be resolved. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes 31059 and 31030 were associated with a remote arm controller (rac). The rac consists of five printed circuit assembly boards which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 31059 is generated when the da vinci safety system determines that a component experienced a communication failure in the middle of sending a message. System error code 31030 appears when the da vinci safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining these conditions, the system records the faults and transitions to a safe state. The rac was returned for evaluation. Engineering found evidence of contamination due to fluid intrusion, thus generating the system error codes experienced by the customer. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2536770 | 2511290 | 2408023 | 2955842201200133 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-03-06 | 2012-03-30 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-03-06 | Device Not Retu | No | Reuse | 2012
40018402 | 2012-02-22 | Malfunction |  | It was reported that during a da vinci myomectomy procedure, electric sparks were observed coming from the permanent cautery spatula instrument. The site completed the planned procedure with a replacement instrument. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed melted material and an arc track at the base of the distal yaw pulley assembly. Additionally broken pieces of the conductor wire cap were detected. The damage is indicative of an arcing event. Since this report was originally received, several attempts have been made to gather additional information without success. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2537464 | 2511926 | 2408657 | 2955842201200135 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-03-02 | 2012-03-30 | No | Yes | Health Professional | 40018402 | Device Returned To Manufacturer | No | 2012-03-02 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-02-29 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci si prostatectomy procedure, after installing an instrument into a patient side manipulator (psm) arm, the site experienced an error message that the instrument was incompatible with the cannula. With the assistance of an isi technical support engineer, the site was instructed to install a replacement cannula and re-installed the instrument to the psm while simultaneously closing/squeezing the cannula mount, however, the error message reoccurred. The tse instructed the site to restart the system and reseat the cannula, however, the issue persisted. The patient was under anesthesia and the port incisions had been made, when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was report. | The investigation conducted by the isi field service engineer (fse) concluded that the issue experienced by the customer was associate with the patient side manipulator (psm) arm cannula mount accessory. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The cannula mount accessory is attached to the psm and designed to hold the cannula securely in place, outside of the patient cavity. The system was repaired by replacing the affected psm cannula mount accessory. Isi has made several attempts to contact the initial reporter to obtain additional information concerning the patient's current status, however, no additional information has been provided. If additional information is received a follow up medwatch report will be submitted. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2537483 | 2511944 | 2408675 | 2955842201200136 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-29 | 2012-03-30 | Yes | Yes | Health Professional | IS3000 A60P6 |  | Yes | 2012-02-29 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2011-11-23 | Injury | Other | It was reported that during a da vinci s nissen fundoplication procedure, the surgical staff observed the bipolar instrument to have energized without the footpedal being pressed resulting in a burn to the patient's liver. The bipolar instrument was removed and the procedure was completed using replacement instruments. The site reported that the patient is doing fine; however, it is unknown if or how the patient was treated. An error in the initial review of this event resulted in a delay of the submission of this report. | The investigation concluded that the patient injury was related to a setup issue with the electrical surgical unit (esu) which is not supplied by intuitive surgical. The surgical staff advised that the root cause of this event was related to the bipolar cable being removed from the esu unit. The esu was set incorrectly to auto bipolar mode and when the bipolar cable was re-connected the instrument began to active as would be intended with these esu settings. Further testing from the site determined that the da vinci s surgical system was working to specifications and did not contribute to this event. As of (b)(4), 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2537530 | 2511990 | 2408721 | 2955842201200137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-23 | 2012-03-30 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2011-11-27 | Device Not Retu | No | Reuse | 2012
IS3000 | 2012-02-28 | Other | Other | It was reported that during the beginning of a da vinci si paraesophageal hernia repair procedure, the surgeon could move the camera, but not the patient side manipulator (psm) arms. As a result, the surgeon made the decision to abort the planned procedure. The patient had already been administered anesthesia when the decision was made. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was likely due to user error. The assisting nurse was not experienced in the instrumentation set up which led to the patient side manipulator (psm) arm difficulty as reported by the surgeon. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The field engineer worked with the surgical staff to verify system performance, and found the system to be working as intended. As of (b)(4) 2012, there have been no reported recurrences of the event. | DA VINSI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2534901 | 2509623 | 2406356 | 2955842201200129 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-02-28 | 2012-03-29 | Yes | Yes | Health Professional | IS3000 | Yes | Yes | 2012-02-28 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-02-28 | Other | Other | It was reported that after removing the cannula accessories after a da vinci si hysterectomy procedure was completed, charring was observed at one of the port incision sites. The surgeon applied a topical ointment to treat the affected area. Arcing during the surgical procedure was not observed by the site, however, the site inspected the monopolar curved scissors (mcs) instrument and tip cover accessory used during surgical procedure and found that the tip cover had a hole. Insulation testing performed on the mcs instrument and tip cover by the site's instrument coordinator observed that the insulator tester indicated that there was an electrical breach. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(6) 2012, intuitive surgical received an email response from rn (b)(6), at (b)(6) hospital indicating that the patient tolerated the planned surgical procedure well and the patient's hospitalization was not prolonged, as the patient was released from the hospital on (b)(6) 2012. The patient is recovering well and has not returned to the hospital due to any post-operative complications. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2535532 | 2510065 | 2406798 | 2955842201200130 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-28 | 2012-03-29 | Yes | Yes | Health Professional | 40018010 | No | Yes | 2012-02-28 | No | No | Unkown | 2012
IS3000 A60P6 | 2012-02-22 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, during startup of the da vinci si system the site noticed the system error code 23013 occurred. With the assistance of an isi technical support engineer (tse) the site performed an emergency power off of the system; however, system error code 23013 reoccurred. The tse asked the site if they would like to disable one of the patient side manipulators (psm) arms and continue with the procedure using only two psms, the surgeon declined. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for the following day. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 23013 appears when the da vinci safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2528309 | 2502605 | 2399334 | 2955842201200126 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-02-22 | 2012-03-23 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-02-22 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-02-22 | Injury | Other | It was reported that after removing the cannula accessories after a da vinci si hysterectomy procedure was completed, charring was observed at one of the port incision sites. No additional harm was reported. | There were no reported malfunctions with the system during the surgical procedure. However, intuitive surgical investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2528308 | 2502606 | 2399335 | 2955842201200127 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | InitialFollowup | 2013-01-08 | 2012-03-23 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-02-22 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-02-23 | Other | Other | It was reported that when the surgical staff went to dock the system to the patient for a da vinci si benign hysterectomy procedure a staff member stepped on a blue fiber cable causing it to break. As a result, the decision was made to convert to traditional open surgical technique to complete the planned procedure. No patient harm was reported. | The investigation conducted by an intuitive surgical field service engineer concluded that the system issue experienced by the customer was due to a broken blue fiber optic cable. The cable passes video, audio, and data during system operation. It was determined that the cable broke when a surgical staff member accidentally stepped on it during docking. The affected cable was removed and a replacement cable was provided to the hospital. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2528329 | 2502607 | 2399336 | 2955842201200128 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-02-23 | 2012-03-23 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-02-23 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-02-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was opening the vaginal cuff, a spark from the monopolar curved scissors (mcs) tip cover accessory, installed on the mcs instrument was observed when the mcs instrument made contact with a cooper rumi uterine manipulator also being used during the surgical procedure. The mcs instrument was immediately removed and a replacement tip cover was installed. After some time of use a spark was observed coming from the replacement mcs tip cover accessory when the mcs instrument made contact with the cooper rumi uterine manipulator. The mcs instrument was immediately removed and the planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. The surgical staff contacted isi's technical support engineering for trouble shooting assistance. The tse instructed the site to inspect the tip cover accessories for damage; however, inspection of the tip cover accessories and mcs instruments by the surgical staff and isi representative present did not find any damage to the tip covers or instruments. No patient harm, adverse outcome or injury was reported. Medwatch report 2955842-2012-00125 has been submitted for the 1st mcs tip cover. | The investigation conducted by the field service engineer determined that the cautery issue experienced by the customer was unrelated to the da vinci system as the fse was unable to replicate the issue experienced by the customer. Functional testing of the site's system by the fse found that it functioned within specification and review the site's system log did not find any error codes associated with cautery issues. As a precaution, the fse recommended that the site have their electrical surgical unit (esu) inspected and replace the esu cable. On (b)(4) 2012, isi contacted the da vinci coordinator (b)(6) at (b)(6) hospital and the following additional information concerning the reported event was provided: - the grounding pad and surgical site were found to be normal. - the esu was a valley lab force fx and the cut/coag mode settings were both 30 watts. - the site's biomedical department inspected the site's entire monopolar curved scissors instrument (mcs) inventory, cautery cable inventory, and esu; however no abnormalities were found. - the mcs instruments did not make contact with any other instruments during the surgical procedure. A response received on (b)(4) 2012, from the surgeon, dr. (b)(6) at (b)(6) hospital indicated that during the surgical procedure, the mcs instrument was in contact with the tissue being dissected and the characteristics of the sparks observed were bright flashes of light and that sparking occurred only when the vaginal cuff was being opened. She initially thought that a harmonic cup instead of a monopolar cup was being used, since it was generally when they were working on the cup. However, the surgical team confirmed that the correct cup was being used. No information was provided concerning what the surgeon believes may have caused the sparking. Dr. (b)(6) also indicated that there was no patient injury noted during the surgical procedure as a result of the sparking and the patient is recovering well. The tip cover accessory was returned and evaluated. Engineering visually inspected the tip cover under magnification and found no evidence of arcing on the silicone or ultem. No mechanical tears or physical damage were found. The customer reported failure cannot be confirmed. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2528331 | 2502608 | 2399372 | 2955842201200124 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-23 | 2012-03-23 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-02-23 | Device Not Retu | Yes | Initial | 2012
40018010 | 2012-02-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was opening the vaginal cuff, a spark from the monopolar curved scissors (mcs) tip cover accessory, installed on the mcs instrument was observed when the mcs instrument made contact with a cooper rumi uterine manipulator also being used during the surgical procedure. The mcs instrument was immediately removed and a replacement tip cover was installed. After some time of use a spark was observed coming from the replacement mcs tip cover accessory when the mcs instrument made contact with the cooper rumi uterine manipulator. The mcs instrument was immediately removed and the planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. The surgical staff contacted isi's technical support engineering for trouble shooting assistance. The tse instructed the site to inspect the tip cover accessories for damage; however, inspection of the tip cover accessories and mcs instruments by the surgical staff and isi representative present did not find any damage to the tip covers or instruments. No patient harm, adverse outcome or injury was reported. Medwatch report 2955842-2012-00124 has been submitted for the 2nd mcs tip cover accessory. | The investigation conducted by the field service engineer determined that the cautery issue experienced by the customer was unrelated to the da vinci system as the fse was unable to replicate the issue experienced by the customer. Functional testing of the site's system by the fse found that it functioned within specification and review the site's system log did not find any error codes associated with cautery issues. As a precaution, the fse recommended that the site have their electrical surgical unit (esu) inspected and replace the esu cable. On (b)(4) 2012, isi contacted the da vinci coordinator (b)(6) at (b)(6) hospital and the following additional information concerning the reported event was provided: - the grounding pad and surgical site were found to be normal. - the esu was a valley lab force fx and the cut/coag mode settings were both 30 watts. - the site's biomedical department inspected the site's entire monopolar curved scissors (mcs) instrument inventory, cautery cable inventory, and esu; however no abnormalities were found. - the mcs instruments did not make contact with any other instruments during the surgical procedure. A response received on (b)(4) 2012, from the surgeon, dr. (b)(6) at (b)(6) hospital indicated that during the surgical procedure, the mcs instrument was in contact with the tissue being dissected and the characteristics of the sparks observed were bright flashes of light and that sparking occurred only when the vaginal cuff was being opened. She initially thought that a harmonic cup instead of a monopolar cup was being used, since it was generally when they were working on the cup. However, the surgical team confirmed that the correct cup was being used. No information was provided concerning what the surgeon believes may have caused the sparking. Dr. (b)(6) also indicated that there was no patient injury noted during the surgical procedure as a result of the sparking and the patient is recovering well. The tip cover was returned and evaluated. Engineering visually inspected the tip cover under magnification and found no evidence of arcing on the silicon or ultem. No mechanical tears or physical damage were found. The customer reported failure cannot be confirmed. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2528330 | 2502639 | 2399372 | 2955842201200125 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-02-23 | 2012-03-23 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-02-23 | Device Not Retu | Yes | Initial | 2012
400180 | 2012-03-12 | Malfunction |  | During robotic assisted total hysterectomy with right salpingo-oophorectomy, the pa reported that an arc was seen coming from the da vinci monopolar curved scissors tip cover accessory. No flame seen. Instrument removed and tip cover replaced with a new one. Procedure resumed. No further incidents. Qa performed a fda maude database search for similar reported problems across the united states. We found many reports regarding tears and holes in the monopolar tip cover insulator accessory part # 400180 used with the curved scissors which resulted in some situations arcing or energy being delivered to unintended nearby tissue. Intuitive surgical had sent a customer safety notice letter in october 2011 to customers with safety recommendations and suggestions to prevent this situation from occurring. The main bullet points from the letter to minimize this risk are:*avoid instrument collisions. *carefully install the tip cover. *straighten wrist prior to removal. *inspect cannula prior to use. *do not exceed maximum monopolar cautery settings. *only use validated esus. *be aware of critical anatomy in contact with the instrument. *change tip cover in cases with extended cautery use. Each of bullet items listed above have detailed explanations in the letter. This information has been shared with surgical services management and administration. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2530466 | 2504742 | 2401475 | 2504742 | 1 | NAY | User Facility |  |  | Initial | 2012-03-23 | 2012-03-23 | No | Yes | Physician | 400180 | No | No Answer Provided |  |  |  |  | 2012
400180 | 2012-03-12 | Malfunction |  | During procedure monopolar curved scissors (hard shears) #420179 was used and electrical arc was observed to uterus by pa and physician. Scissor and tip were removed. A round hole was seen by scrub technician and circulator in tip cover. We put new tip cover accessory on scissor, no hole was observed and we continued with procedure, without difficulty (slight burn to uterus from arc, but uterus removed in scheduled surgical procedure), so no harm was done to patient. Qa performed a fda maude database search for similar reported problems across the united states. We found many reports regarding tears and holes in the monopolar tip cover insulator accessory part # 400180 used with the curved scissors which resulted in some situations arcing or energy being delivered to unintended nearby tissue. Intuitive surgical had sent a customer safety notice letter in october 2011 to customers with safety recommendations and suggestions to prevent this situation from occurring. The main bullet points from the letter to minimize this risk are:*avoid instrument collisions. *carefully install the tip cover. *straighten wrist prior to removal. *inspect cannula prior to use. *do not exceed maximum monopolar cautery settings. *only use validated esus. *be aware of critical anatomy in contact with the instrument. *change tip cover in cases with extended cautery use. Each of bullet items listed above have detailed explanations in the letter. This information has been shared with surgical services management and administration. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2530467 | 2504745 | 2401478 | 2504745 | 1 | NAY | User Facility |  |  | Initial | 2012-03-23 | 2012-03-23 | No | Yes | Physician | 400180 | No | No Answer Provided |  |  |  |  | 2012
42018405 | 2012-02-16 | Malfunction |  | Event Description: --/-- It was reported that while performing a da vinci si total benign hysterectomy procedure, pieces of the permanent cautery spatula instrument fell into the patient. The surgeon was immediately able to retrieve one piece, the second piece was found after irrigation. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off and one conductor cap broken off. The damage was likely caused by mishandling. Engineering also observed a crack on the proximal clevis. It was determined that the damage was likely due to excess force applied to the instrument. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2521357 | 2495756 | 2392489 | 2955842201200123 | 1 | NAY | Manufacturer | User facilityOther | Nurse | FollowupInitial | 2012-02-18 | 2012-03-19 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2012-03-26 | Device Not Retu | No |  | 2012
42030901 | 2012-02-20 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a pin and small disc on the mega suturecut needle driver instrument broke off and fell into the patient. All of the fragments were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal pin is broken at the proximal clevis. Engineering concluded that damage was likely due tip overloading with excessive force applied. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2520131 | 2495000 | 2391733 | 2955842201200122 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2012-02-20 | 2012-03-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2012-02-20 | Device Not Retu | No | Reuse | 2012
40027202 | 2012-02-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the active blade of the harmonic ace curved shears insert accessory, which was installed on the harmonic ace curved shears instrument, broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 675 x. 085 piece of the curved blade broken off. The curved blade fractured near the clamp arm pin. It was determined that the blade may have been making contact with the clamp arm pin during energy activation, thus cracking the blade and leading to breakage. The broken piece was not returned. No other damage was found. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2519999 | 2494866 | 2391599 | 2955842201200118 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-17 | 2012-03-16 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-02-17 | Device Not Retu | Yes | Initial | 2012
IS3000 A60P6 | 2012-02-01 | Other |  | It was reported that after removing the cannula accessories after a da vinci si gynecological procedure was completed, the physician assistant found that the patient's port incisions exhibited mild burns. No other information was provided. | Isi has made several attempts to obtain additional information concerning the reported event, however, no other information has been provided. A follow up mdr will be submitted if additional information is received concerning this report. While no additional information was provided, intuitive surgical field service engineering (fse) investigated a similar event reported by the site that occurred during a second da vinci si procedure performed on that same day. The fse investigation found that the issue experienced by the customer was due to a defective grounding pad. (refer to medwatch report 2955842-2012-00098). Intuitive surgical investigation has shown that it is possible for the electrosurgical unit (esu) current to pass through the patient side manipulator arm to ground, through the cannula accessory to the patient, it the patient is not properly grounded. As of (b)(4) 2012, the site has continued to use the da vinci si system and there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2520000 | 2494871 | 2372570 | 2955842201200119 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-02-15 | 2012-03-16 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2012-02-15 | No | No | Reuse | 2012
42020505 | 2012-02-17 | Malfunction |  | It was reported that during a da vinci si pancreatectomy procedure, the plastic insulation on the tip of the fenestrated bipolar forceps instrument broke off. The site was unable to confirm if the broken instrument piece fell into the patient. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On march 13, 2012, the initial reporter at university hospital indicated that he is unable to confirm if the broken instrument piece remains in the patient. No patient harm, adverse outcome or injury occurred and to the best of his knowledge, the patient has not returned to the hospital due to any post operative complications. No other information was provided. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2520002 | 2494872 | 2391605 | 2955842201200117 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-02-17 | 2012-03-16 | No | Yes | Health Professional | 42020505 | No | No | 2012-02-17 | No | No | Unkown | 2012
420171 | 2012-02-16 | Other | Other | Event Description: --/-- It was reported that during a da vinci s coronary artery bypass graft (cabg) procedure, arcing from the micro bipolar forceps instrument was observed when the surgeon was not applying cautery energy. The site contacted isi technical support engineering (tse) for troubleshooting assistance, however, the site declined support, as the surgeon did not have time to trouble shoot the issue with the tse. The planned surgical procedure was completed and no patient harm adverse outcome or injury was reported. | Manufacturer Narrative: --/-- On (b)(6) 2012, intuitive surgical received uf/importer report (b)(4) from the fda. The event details are provided below: event description: patient for a cabg with a da vinci robot. Surgeon attempted to use the da vinci bipolar forceps, but they did not work. Circulator checked the electrosurgical unit (esu and noted that ground pad (return electrode for esu was not plugged in. Ground pad connected. Surgeon stated there was no noise and the foot pedal was not depressed, but as soon as the surgeon touched the diaphragm with the bipolar forceps, the machine buzzed and the diaphragm had a small burn. Circulator checked the esu and noted that the forceps cable was plugged into two of the three monopolar 2 sockets. Bipolar was plugged into bipolar sockets, machine functioned normally throughout the rest of the case. Patient had two (2) ground pads on, unsure which was the actual lot number involved in event. Based on the additional information provided in the site's u/f report (b)(4) indicating that the patient's diaphram was burned, on (b)(6) 2012, isi contacted risk manager, (b)(6) to advise that during isi's initial investigation into the reported event, it was reported by the site's da vinci coordinator, (b)(6) initial reporter) that no patient harm, adverse outcome or injury occurred. (b)(6) indicated that contrary to the initial information provided, a burn to the patient's diaphram occurred, however, no repair of the affected area was required, the planned surgical procedure was completed, and there was no report that the patient experienced any post-operative complications. She does not know if the patient has had to return to the hospital due to any post-operative complications related to the reported event. As indicated in isi's initial medwatch report submitted to the fda on (b)(4) 2012, isi's investigation into the reported event found that the issue experienced by the site was due to improper connection of the micro bipolar forceps instrument to the electrical surgical unit (esu). The instruments and accessories user manual specifically states: general precautions and warnings o please refer to the individual esu manufacturer's operator's manual for operating instructions. Set the esu to bipolar output. Set the power as low a possible to achieve adequate hemostasis - do not use bipolar instruments with a monopolar source output as this may cause damage to the instrument and harm to the patient or medical personnel system functional testing performed by the fse found that the system functioned within specification. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by field service engineering (fse) found that the issue experienced by the site was due to improper connection of the micro bipolar forceps instrument to the electrical surgical unit (esu). The hospital's biomedical department indicated to the fse that during the surgical procedure, the micro bipolar forceps instrument was found to be incorrectly plugged into the monopolar connection output on the electrical surgical unit (esu). The site immediately resolved the issue by connecting the instrument to the bipolar connection site output on the esu as indicated in our ifu. The instruments and accessories user manual specifically states: general precautions and warnings o please refer to the individual esu manufacturer's operator's manual for operating instructions. Set the esu to bipolar output. Set the power as low a possible to achieve adequate hemostatsis o do not use bipolar instruments with a monpolar source output as this may cause damage to the instrument and harm to the patient or medical personnel system functional testing performed by the fse found that the system functioned within specification. On (b)(4) 2012, isi followed up with the initial reporter at university of medical center and he confirmed that no patient harm, adverse outcome or injury occurred. The patient tolerated the planned surgical procedure well and has not returned to the hospital due to any post operative complications. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | MICRO BIPOLAR FORCEPS INSTRUMENT | Unkown | 2520005 | 2494890 | 2391623 | 2955842201200120 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2012-03-08 | 2012-03-16 | Yes | Yes | Health Professional | 420171 | No | No | 2012-03-08 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-02-17 | Other | Other | It was reported that in the process of docking the patient cart to the patient, the surgical staff experienced system error code 31030. A technical support engineer (tse) attempted to troubleshoot the issue via telephone and had the site restart the system, however, the same error code recurred upon power up. The tse reviewed the system error logs and observed system error code 23034, indicating a stuck cannula port clutch switch on one of the system arms. The tse was able to confirm with the surgical staff that one of the system arms could not be clutched. The surgical staff powered down the system again, actuated the clutch switch of the affected instrument arm multiple times, however, the issue persisted upon power up. The patient was under anesthesia and the port incisions had been made when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes 31030 and 23034 experienced by the site were associated with a patient side manipulator (psm). The psm is an instrument arm located on the surgical cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 31030 appears when the da vinci safety system determines that a subcomponent of the system did not successfully complete its startup process. System error code 23034 appears when the safety system determines that after a specified amount of time, a valid event has not been seen for one of the remote compute engine switches. Upon determining these conditions, the system records the fault and transitions to a safe state. The psm was returned for evaluation. Engineering was unable to replicate the reported issue, however, the cannula port clutch switch was replaced per the errors observed in the system logs. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2520050 | 2494911 | 2391644 | 2955842201200121 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-17 | 2012-03-16 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-02-17 | Device Not Retu | No | Reuse | 2012
M0043804300306 | 2012-03-14 | Malfunction |  | Patient was having a lap assisted robotic hysterectomy/bso (bilateral salpingo oophorectomy)done, when about half way through the procedure the robot alarmed and the arms and instruments froze which allowed the surgeon no control to continue with the procedure. We called technical support from intuitive and were instructed to power down the robot and reboot it. When the robot was restarted and calibrated it regained function and we were able to proceed with the case. This process of waiting and trying to fix the problem took somewhere between 15 and 25 minutes. There was no patient harm. |  | DA VINCI PS2000 | SYSTEM ROBOTIC SURGICAL | 2526035 | 2500225 | 2396958 | 2500225 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-03-16 | 2012-03-16 | No | Yes | Invalid Data | M0043804300306 | No | No Answer Provided |  |  |  |  | 2012
420205 | 2012-02-13 | Injury | Required Intervention | It was reported that while performing right side dissection of the broad ligament with a fenestrated bipolar forceps instrument during a da vinci s sacrocolpopexy w/ hysterectomy procedure, the surgeon inadvertently contacted the patient's ureter with the instrument while it was energized. As a result, the patient's ureter was injured. A urologist was called into the operating room, who re-implanted the patient's ureter to repair the injury. The procedure was converted to traditional open surgical techniques to complete the planned procedure. No additional patient harm was reported. | Intuitive surgical concluded that the damage to the patient's ureter was caused by user error and is unrelated to the performance of the maryland bipolar forceps instrument. Per the information provided, it was determined that the instrument worked as intended. No information has been received which suggests that the instrument malfunctioned in a way that may have caused or contributed to the patient's injury. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2516317 | 2491746 | 2388479 | 2955842201200116 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-13 | 2012-03-14 | Yes | No | Health Professional | 420205 | No | No | 2012-02-13 | No | No | Unkown | 2012
420184 | 2012-03-07 | Malfunction |  | During a robotic surgery, the tip of the robotic spatula instrument broke off while being removed from the robotic arm. Surgeon explored the wound thoroughly and no foreign objects were found in the patient. X-rays were not taken because missing pieces are not radioopaque. The broken pieces were not found. Surgeon aware. Intuitive surgical rep was present and witnessed the incident. This particular instrument was purchased in 2008 and is a limited life instrument that is resterilized in our facility. This particular instrument can be used for 10 times before replacement is required. This device had been used 7 times. |  |  | INSTRUMENT ROBOTIC SURGICAL CA | 2520584 | 2495307 | 2392040 | 2495307 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-03-13 | 2012-03-14 | No | Yes | Physician | 420184 | Yes | No Answer Provided |  |  |  |  | 2012
42017207 | 2012-01-21 | Malfunction |  | Event Description: --/-- It was reported that while performing a da vinci s myomectomy procedure, the surgeon tried to cauterize a vessel with the maryland bipolar forceps instrument. During this time the forceps emitted sparks from the tip. The instrument was removed and replaced with another maryland bipolar forceps instrument. The planned surgical procedure was completed with the new maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. The delay in reporting is due to an administrative oversight where the aware date was inadvertently not updated by a customer service representative. | Manufacturer Narrative: --/-- (b)(4). --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. The engineering investigation found charring and localized melting on both yaw pulleys at the base grips, between the grips. The charring indicates an arcing event occurred. It was determined that the charring was likely due to a breach in the instrument insulation, or carbonized tissue which created a conductive path, or high generator settings. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 2514819 | 2490353 | 2387106 | 2955842201200115 | 1 | NAY | Manufacturer | ForeignDistributor | Other | InitialFollowup | 2012-02-10 | 2012-03-13 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-02-10 | Device Not Retu | No |  | 2012
|  | Injury | Required Intervention | The following patient story concerning a da vinci gall bladder procedure was found posted on the (b)(6) by intuitive surgical on (b)(6) 2012: (b)(6), my son was almost killed after a surgeon used the da vinci machine to remove his gall bladder. No other details were provided. Isi was unable to contact the author of the post concerning the reported event, as the author has blocked their (b)(6) account. | On march 4, 2012, the author of the patient story originally posted on (b)(6) 2012, posted the following additional information: (b)(6): my son had surgery on his gall bladder, the da vinci machine was recommended we were told it would be less invasive and shorter recovery time that was a big mistake, he almost died. They cut a small hole in his intestine, on the third day, he looked like a (b)(6) victim. He was sent home i then took him to another hospital. He has had 3 more surgeries and needs more. Intuitive surgical has made several attempts to contact the initial reporter concerning the reported event. However, no response has been received. A follow-up mdr will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2512971 | 2488683 | 2385436 | 2955842201200114 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2012-02-11 | 2012-03-12 | Yes | No | Health Professional |  | No | Yes | 2012-02-11 | No | No | Unkown | 2012
IS3000 A60P6 | 2012-02-10 | Other |  | It was reported that during a da vinci si prostatectomy procedure, the site experienced system error code 23025, prompting the site to disable the right master tool manipulator(mtmr). With the assistance of isi technical support engineering (tse) the site performed a power cycle of the system and exercized the affected axis on the right mtm, however, the error code 23025 reoccurred along with system error code 23008. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the system error codes experienced by the customer were associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 23025 occurs when the remote arm controller hardware detects an encoder quadrature fault. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the faults and transitions to a safe state. The mtmr was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported issue and determined that the axis-4 motor/encoder had malfunctioned, thus generating the system error codes experienced by the customer. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2509275 | 2485603 | 2382356 | 2955842201200111 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-02-10 | 2012-03-09 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-02-10 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-02-08 | Other | Other | Event Description: --/-- It was reported that while beginning a da vinci si prostatectomy procedure, the surgeon experienced double vision intermittently through the high resolution stero viewer (hrsv). Prior to contacting isi for technical support, the site recalibrated the endoscope, replaced the endoscope, the camera head and restarted the system, however, the issue persisted. Unable to resolve the issue and prior to contacting isi for technical support, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the site was associated with the high resolution stereo viewer (hrsv). The fse found that an internal digital video interface (dvi) cable within the hrsv was loose, which caused the double vision experienced by the surgeon. The dvi cable transmits the video signal from the display electronics drivers to the video monitors located inside of the hrsv. The system was repaired by reseating the affected dvi cable. As of march 9, 2012, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2012, the da vinci coordinator, (b)(6) followed up with intuitive surgical and indicated that the surgeon made the decision to abort the planned surgical procedure and rescheduled it for the next day. Anesthesia had been administered to the patient and the port incisions had been created. No patient harm, adverse outcome or injury occurred. The rescheduled da vinci si prostatectomy procedure was performed on (b)(6) 2012 and the surgeon has reported that the patient is recovering well. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2509280 | 2485628 | 2382407 | 2955842201200112 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-02-08 | 2012-03-09 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-02-08 | Device Not Retu | No |  | 2012
42018304 | 2012-02-10 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, upon inspection of the permanent cautery hook (pch) instrument prior to use, the scrub technician found that a notch from the instrument joint that meets the instrument shaft was missing. The scrub technician informed the surgeon and the bedside surgeon assistant that the pch instrument had a missing piece, however, the surgeon made the decision to use the instrument. Upon insertion of the instrument into the cannula, a fragment from the instrument shaft was observed to be dangling from the instrument and a piece broke off and fell into the patient. The instrument fragment was retrieved, however, the surgeon continued to use the instrument in its damaged condition, until the surgical team insisted that the instrument be removed. It was indeterminable if any other material from the instrument shaft fell into the patient. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the proximal clevis was broken at the main tube connection. An area approximately 0. 36 x 0. 13 had broken off the clevis. Engineering concluded that the damage was likely due to excess force applied to the cautery hook and clevis. The instruments and accessories user manual specifically states: - proper care and handling is essential for satisfactory performance of surgical instruments and accessories. Examine the instrument or accessory - including all of its components - thoroughly before and after each use. If any abnormality is found, do not use it. Use the device for its intended purpose only. - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2509301 | 2485633 | 2382386 | 2955842201200113 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-02-10 | 2012-03-09 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2012-02-10 | Device Not Retu | No | Reuse | 2012
42018406 | 2011-12-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure, sparks were observed coming from the distal end of the permanent cautery spatula instrument. The instrument was immediately removed and inspection of the instrument by the surgical staff found that the ceramic covering was broken. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the ceramic covering was broken, however, the instrument does not exhibit any evidence that an arcing event occurred. The ceramic sleeve has a. 090 x. 085 piece broken off at the distal end. The broken piece was not returned. Engineering concluded that the damage is likely due to mishandling/misuse. The broken instrument piece was not returned. The instrument passed electrical continuity testing and no other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2507955 | 2484425 | 2327090 | 2955842201200106 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2012-02-07 | 2012-03-08 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2012-02-07 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-01-30 | Malfunction |  | It was reported that during a da vinci si prostatectomy arcing from the monopolar curved scissors (mcs) instrument with the tip cover accessory installed was observed. Examination of the tip cover accessory by the surgical staff found that the tip cover had a hole. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The hospital indicated that the affected tip cover will not be returned for evaluation. | Since the tip cover accessory will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2507906 | 2484426 | 2381233 | 2955842201200108 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-07 | 2012-03-08 | No | Yes | Health Professional | 40018010 |  | Yes | 2012-02-07 | No | Yes | Unkown | 2012
40018010 | 2012-01-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy, arcing from the monopolar curved scissors (mcs) instrument with the tip cover accessory installed was observed. Examination of the tip cover accessory by the surgical staff found that the tip cover had a hole. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The hospital indicated that the affected tip cover will not be returned for evaluation. | Since the tip cover accessory will not be returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2507905 | 2484467 | 2381260 | 2955842201200107 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-07 | 2012-03-08 | No | Yes | Health Professional | 40018010 | No | Yes | 2012-02-07 | No | Yes | Unkown | 2012
42018306 | 2012-02-09 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure, plastic fell off of the permanent cautery hook instrument tip and was retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor cap was missing from the distal end. As a result, the yaw pulley is not seated properly in the distal clevis. The boss feature on the yaw pulley hub that mates with the cap is broken off. This breakage likely allowed the cap to dislodge. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2508020 | 2484531 | 2381304 | 2955842201200110 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-02-10 | 2012-03-08 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-02-10 | Device Not Retu | No | Reuse | 2012
400272 | 2012-01-12 | Malfunction |  | Event Description: --/-- On february 9, 2012, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda, which contained the following information: event desc: a piece of the jaw of the harmonic curved shears broke off in the patient during surgery. The piece was removed by the surgeon, intact, with no harm to the patient. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the curved blade was broken off. Evidence of contact was found at the pin on the clamp arm. No other damage was found. There are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. --/-- Additional Manufacturer Narrative: --/-- The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the device is returned or if additional information is received. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | Unkown | 2508022 | 2484534 | 2329837 | 2955842201200109 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-02-09 | 2012-03-08 | No | Yes | Health Professional | 400272 | Device Returned To Manufacturer | No | 2012-04-02 | Device Not Retu | Yes |  | 2012
42000306 | 2012-02-03 | Malfunction |  | It was reported that during a da vinci si surgical procedure, while using the 8mm small clip applier instrument, one of the jaws broke and fell into the patient. The surgeon immediately retrieved the instrument fragment and the planned procedure was completed. No patient harm, injury or adverse event was reported. | The instrument has not been returned for evaluation. Multiple attempts have been made to obtain further information however as of the time of this report the site has not responded. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM SMALL CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 2501401 | 2477069 | 2373862 | 2955842201200096 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-03 | 2012-03-05 | No | Yes | Health Professional | 42000306 | No | No | 2012-02-03 | No | No | Unkown | 2012
IS2000 A51P8 | 2012-02-07 | Other | Other | It was reported that during set up for a da vinci s gynecological procedure the site received a system fault. With the assistance of an isi technical support engineer the site cleaned and reseated the fiber cable and restarted the system, however the system fault continued to occur. The patient was under anesthesia with no port incisions made, when the surgeon decided to abort the planned surgical procedure. | The investigation conducted by field service engineering concluded that system fault experienced by the customer was associated with the red fiber cable and the surgeon side cart (ssc) rear i/o panel. The fiber cable provides communication from the ssc to the patient side cart. The i/o panel provides the connection point for the fiber cable into the ssc. The system was repaired by cleaning the red fiber cable and replacing the ssc i/o panel. As of march 2, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2501430 | 2477075 | 2373868 | 2955842201200104 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-02-07 | 2012-03-05 | Yes | Yes | Service Personnel | IS2000 A51P8 | Yes | No | 2012-02-07 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-01-31 | Other | Other | It was reported that post a successful da vinci si partial nephrectomy procedure, the patient returned to the hospital for a fasciotomy to repair compartment syndrome. Per the robotics coordinator present for the case, the da vinci si system did not malfunction in any way nor contribute to the patient's outcome, however the patient was obese and the procedure took longer than normal. The robotics coordinator stated that the patient's positioning and duration of the procedure likely caused the compartment syndrome. | The robotics coordinator stated that the system in no way contributed to the patient injury. Based on the information provided, it was determined that the da vinci si surgical system did not malfunction and did not contributed to the patient's outcome. The hospital has continued to use the da vinci si surgical system to perform surgery on patients. As of march 2, 2012 no adverse events have been experienced with the site's system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2501443 | 2477080 | 2372966 | 2955842201200105 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-09 | 2012-03-05 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No | 2012-02-09 | Device Not Retu | No | Reuse | 2012
42017207 | 2012-02-02 | Malfunction |  | It was reported that during a da vinci si thoracic sympathectomy procedure, while using the (b)(4) bipolar forceps instrument, the instrument began to smoke and an arc was visible. The instrument was removed and replaced with an instrument of the same type and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found both yaw pulleys exhibited charring and localized melting at the grip base. Engineering concluded that the charring likely occurred due to a breach in insulation, carbonized tissue creating a conductive path, and or a high generator setting. Electrical continuity passed, no other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2501442 | 2477081 | 2373874 | 2955842201200103 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-02-02 | 2012-03-05 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-02-02 | Device Not Retu | No | Reuse | 2012
| 2012-02-08 | Malfunction |  | A surgical procedure note indicates that the robotic precise bipolar forceps tip malfunctioned and wouldn't align right. The instrument was cleaned and given to rn to forward to risk management. |  | DA VINCI PRECISE BIPOLAR FORCEPS | INSTRUMENT ROBOTIC SURGICAL BI | 2507246 | 2483858 | 2380645 | 2483858 | 1 | NAY | User Facility |  |  | Initial | 2012-03-05 | 2012-03-05 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 | 2012-02-01 | Injury | Required Intervention | It was reported that after removing the cannula accessories after a da vinci si sacrocolpopexy procedure was completed, charring was observed at two of the port incision sites. The surgeon excised the burnt skin at one of the port locations prior to closing the incision and the other incision was closed primarily. No additional harm was reported. | Intuitive surgical investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. Intuitive surgical field service engineering contacted the hospital's robotics coordinator, who indicated that the grounding pad coming from the generator that goes under the patient was found to be faulty and was replaced. The robotics coordinator confirmed that the damage to the patient's tissue was caused by the faulty grounding pad and there were no issues with the da vinci si surgical system. As of (b)(6), 2012, the site has continued to use the da vinci si system and there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2500014 | 2475784 | 2372570 | 2955842201200098 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-02-01 | 2012-03-02 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-02-01 | No | No | Reuse | 2012
IS3000 A60P6 | 2012-02-01 | Other | Other | It was reported that during a da vinci si liver resection procedure the patient began bleeding heavily. The initial reporter indicated that the patient bleeding was unrelated to the use of the da vinci si surgical system. The bleeding resulted in confusion in the operating room, causing a surgical staff member to trip on a system cable. This action generated non-recoverable system error code 25100. An isi technical support engineer was contacted via telephone for troubleshooting assistance, however, the surgeon was unable to take the time to walk through how to recover from this error due to the patient bleeding. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was due to a loose system cable connecting the patient cart to the vision cart. The cable passes video, audio, and data during system operation. It was determined that the cable became loose when a surgical staff member accidentally tripped on it. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25100 appears when the da vinci safety system determines that valid data was not received, which is caused by a communication problem on the fiber optic cable. Upon determining this condition, the system records the fault and transitions to a safe state. As of february 29, 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2500368 | 2476173 | 2372966 | 2955842201200099 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-01 | 2012-03-02 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-02-01 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2012-02-03 | Other | Other | It was reported that during a da vinci s myomectomy procedure, the customer was unable to remove instruments from two of the patient side manipulator arms. After several unsuccessful attempts to remove the instruments, and prior to calling intuitive surgical for technical support, the surgeon made the decision to convert the procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer found that the issue experienced by the site was due to improper removal of the instruments from the patient side manipulators (psm). The surgical staff attempted to release one of the instruments from the sterile adapter installed on the first psm prior to clasping the instrument release levers, which jammed the instrument into the sterile adapter, preventing removal of the instrument. The customer then used the instrument in the first psm to straighten the wrist of the instrument installed on the 2nd psm, which broke the instrument, preventing removal of the instrument. The patient side manipulator is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The fse removed the instruments from the affected psms and functional testing of the system using training instruments and accessories found that the system functioned within specification. The fse also retrained the site's biomedical department on proper removal of the instruments from the psm arms. On (b)(4), 2012, isi followed up with the initial reporter at (b)(6) medical center and she indicated that there was no harm to the patient, the patient tolerated the open procedure well and to the best of her knowledge the patient has not returned to the hospital due to any post-operative complications. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. As of (b)(4), 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2500471 | 2476189 | 2372982 | 2955842201200100 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-02-03 | 2012-03-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-02-03 | Device Not Retu | No | Reuse | 2012
42018406 | 2012-02-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the spatula on the permanent cautery spatula instrument broke and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis ear is broken and the conductor wire cap is missing. The ceramic sleeve is broken at the spatula base. The cables and wires do not exhibit any damage. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2500442 | 2476191 | 2372957 | 2955842201200101 | 1 | NAY | Manufacturer | Company RepresentativeOther | Nurse | InitialFollowup | 2012-02-02 | 2012-03-02 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2012-03-05 | Device Not Retu | No |  | 2012
40020504 | 2012-02-02 | Injury | Required Intervention | Event Description: --/-- It was reported that during a da vinci hysterectomy procedure, after the bipolar cord was connected to the electrosurgical unit, energy was released from the bipolar instrument. As a result, an injury to the patient's bowel occurred. The bowel damage was secured by suture and the planned surgical procedure was completed. No additional harm was reported. | Manufacturer Narrative: --/-- On (b)(6) 2012, the hospital provided intuitive surgical a copy of medwatch uf/importer report (b)(4), which contained the following information: clinical engineering analysis: when connecting the bipolar cord to the esu, apparently the nurse inadvertently plugged the cord into the monopolar jack instead of the bipolar jack. Since the tips of the bipolar forceps were closed, this triggered the monopolar cut mode of the esu and it began delivering energy. Staff noticed the esu tone and smoke visible on the video display and immediately shut off the power to the esu. The monopolar jack on the esu has three holes. Connecting the two outer holes will activate the cut mode, as this is how the normal pencil switch works. The disposable bipolar cord has individual banana plugs on the machine end, which makes it possible to plug into the wrong jack. If the machine end had a single molded connector, the spacing of bipolar pins would prevent plugging into a monopolar jack. Per the additional information provided above, intuitive surgical concluded that the issue experienced by the hospital was due to improper connection of the bipolar instrument to the electrosurgical unit (esu). The instruments and accessories user manual specifically states: caution: please refer to the individual esu manufacturer's operator's manuals for operating instructions. Set the esu to bipolar output. Set the power as low as possible to achieve adequate hemostatis. Warning: do not use bipolar instruments with a monopolar source output as this may cause damage to the instrument and harm to the patient or medical personnel. --/-- Additional Manufacturer Narrative: --/-- The hospital has been contacted multiple times for further details and information concerning this incident. To date, no response has been received. As a result, the root cause of the reported event cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Unkown | 2500555 | 2476271 | 2373037 | 2955842201200102 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2012-02-03 | 2012-03-02 | Yes | Yes | Health Professional | 40020504 | No | No | 2012-04-04 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2012-01-17 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the patient was not securely velcroed to the operating table, causing the patient to slide. The system camera arm that was docked to the patient was able to hold the patient on the table and prevent them from fully sliding off. Due to the sliding, tearing occurred at the trocar port incision location. The tear was repaired at the end of the surgical procedure. No additional patient harm was reported. After the event occurred, the surgeon indicated that the camera arm was difficult to move and felt like it was pulsating during manual movement. | Investigation was conducted by field service engineering, who inspected the system. A visual examination was performed on the endoscopic camera manipulator (ecm) and trocar mound and no signs of fatigue or stress were found. The ecm is the camera arm located on the patient cart which provides the sterile interface for the 3d endoscope. Additionally, no issues were observed when manipulating the ecm. System verification tests were then performed with no issues detected. It was determined that the damage to the patient's skin was caused by the surgical staff not properly securing the patient to the operating room table. As of march 1, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2498778 | 2474756 | 2371542 | 2955842201200097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Health Professional | Initial | 2012-01-31 | 2012-03-01 | Yes | No | Health Professional | IS3000 A60P6 | Yes | Yes | 2012-01-31 | Device Not Retu | No | Reuse | 2012
420194 | 2011-12-29 | Malfunction |  | On (b)(6) 2012, medwatch uf/importer report (b)(4) was received: event desc: piece of grasper teeth broke off into patient and could not be retrieved. (device was on 6th use. ) what was the original intended procedure? hysterectomy device usage problem: other. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2496656 | 2472622 | 2369435 | 2955842201200093 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2012-01-31 | 2012-02-29 | No | Yes | Health Professional | 420194 | No | No | 2012-01-31 | No | No | Unkown | 2012
IS3000 A60P6 |  | Death | Death | It was reported that approximately 1 day post-op, a successful da vinci si lysis of adhesions procedure, the patient exhibited signs of cellulitis at one of the trocar sites and experienced decreased urine output. A ct scan performed showed the patient had normal gastrointestinal and genitourinary tract, however, within a few hours, the patient's health began to deteriorate and was taken to the icu. Several specialists consulted suspected that the patient had necrotizing fasciitis and the patient was then taken to the operating room. The general surgeon consulted found that the patient's small bowel had a perforation. The patient had developed septic shock, acidosis and expired. It is the surgeon's belief that the infection developed by the patient was very atypical presentation for peritonitis and perforation and that maybe the patient's immunosuppressive status caused the event, as damage to the patient's bowel was not thermal damage and was not caused by the trocar as she performed an open entry. The delay in reporting is due to an oversight in initial capturing of the complaint by isi customer service. | On february 23, 2012, isi contacted the risk manager at (b)(6) hospital and she stated that damage to the patient's bowel was possibly introduced during the da vinci si lysis of adhesions procedure. However, there was no malfunction of the da vinci si system, instruments and/or accessories that caused or contributed to the patient's injury and subsequent demise. The risk manager was unable to provide additional information stating hippa regulations prohibited her from providing any other information. Isi advised the risk manager that hippa regulations do not apply to investigations related to adverse events that are being investigated and reported to the fda. However, she advised that she would be open to discussing the reported event with the fda. As of february 29, 2012, no adverse events have been experienced with the site's system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2497302 | 2473426 | 2370219 | 2955842201200095 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2011-12-27 | 2012-02-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-12-29 | No | No | Reuse | 2012
400272 | 2011-12-01 | Malfunction |  | On (b)(6) 2012, isi received uf/importer report (b)(4) (b)(6). The event details are provided below: during a robotic assisted laparoscopic myomectomy, the harmonic scalpel jaw broke inside the patient's abdomen. Post operative x-ray was done and showed no residual. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the insert accessory is returned or if additional information is received. On (b)(4) 2011, isi contacted the risk manager at (b)(6) and she reported that the jaw of the harmonic ace insert accessory broke off and fell into the patient's abdomen. The broken piece was retrieved and an x-ray performed of the patient's abdomen revealed that there were no remaining pieces from the insert accessory. There was no harm, adverse outcome or injury to the patient. In addition, the patient has not returned to the hospital due to experiencing any post operative complications related to the reported event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2497312 | 2473427 | 2370220 | 2955842201200094 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-31 | 2012-02-29 | No | Yes | Health Professional | 400272 | No | No | 2012-01-31 | No | Yes | Unkown | 2012
40018010 | 2012-01-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument was failing and arcing and a hole in the tip cover were observed. The tip cover was immediately removed and replaced and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2491884 | 2467834 | 2364664 | 2955842201200082 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-30 | 2012-02-24 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-30 | No | Yes | Unkown | 2012
40018010 | 2012-01-31 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument was failing and arcing and a hole in the tip cover were observed. The tip cover was immediately removed and replaced and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2491879 | 2467851 | 2364664 | 2955842201200080 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-31 | 2012-02-24 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-31 | No | Yes | Unkown | 2012
IS1200 A43P9 | 2012-01-31 | Other | Other | It was reported that prior to starting a da vinci prostatectomy procedure, the site experienced a blue tint in the right eye of the surgeon's console viewer while performing a white balance of the camera. The site contacted a technical support engineer for troubleshooting assistance and was instructed to swap both ends of the camera cable, however, the issue recurred. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the site was related to the right camera control unit (ccu). The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the right ccu. As of february 23, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2491909 | 2467854 | 2364667 | 2955842201200083 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-01-31 | 2012-02-24 | Yes | Yes | Health Professional | IS1200 A43P9 | No | No | 2012-01-31 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-01-25 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument was failing and arcing and a hole in the tip cover were observed. The tip cover was immediately removed and replaced and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2491910 | 2467855 | 2359937 | 2955842201200081 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-25 | 2012-02-24 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-25 | No | Yes | Unkown | 2012
IS1200 A43P9 | 2012-01-27 | Other | Other | Event Description: --/-- It was reported that while attempting to home the system prior to starting a da vinci prostatectomy procedure, the surgical staff noticed that an instrument arm would not move and system error code 20008 occurred. The site restarted the system multiple times, however, the issue persisted. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm was reported. | Manufacturer Narrative: --/-- The patient side manipulator was returned and evaluated. The error code experienced by the customer was confirmed in the system logs, however, engineering was unable to replicate the reported failure during internal testing. Based on the system error logs, an axis potentiometer and motor were replaced as a precaution. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the system issue experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the surgical cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 20008 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. As of february 23, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | Reuse | 2491914 | 2467857 | 2364670 | 2955842201200084 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2012-01-27 | 2012-02-24 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2012-04-16 | Device Not Retu | No |  | 2012
IS1200 A43P9 | 2012-01-16 | Death | Death | It was reported that after an uncomplicated da vinci hysterectomy procedure was successfully completed, the patient expired later that night. | On february 13, 2012, the hospital's operating room director indicated that an autopsy was performed, however, the patient's cause of death was inconclusive. The operating room director also mentioned that toxicology testing was conducted and the results would be provided to intuitive surgical when completed. No other information was provided. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2491894 | 2467890 | 2364706 | 2955842201200085 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-26 | 2012-02-24 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2012-01-26 | No | No | Reuse | 2012
IS3000 60P6 | 2012-01-04 | Other | Other | It was reported that a nurse got her hand smashed between the da vinci si surgical system and the wall while she was driving the patient side cart. The nurse obtained an x-ray of her hand to confirm that there was no fracture. No further harm was reported. | The investigation conducted by the field service engineer found the patient side cart throttle control was functioning properly in both forward and reverse directions. System verification was performed and found to be within specification. While going over the event with the nurse and or manager it was determined that the customer reported issue was associated with user error. Onsite training was performed to ensure no further events occur. As of february 22, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2492315 | 2468240 | 2365036 | 2955842201200079 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-27 | 2012-02-24 | Yes | No | Health Professional | IS3000 60P6 | Yes | Yes | 2012-01-27 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-01-25 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found a hole approximately 0. 028 inches in diameter with a burn mark on the edge of the hole's wall. The orientation of the hole confirmed energy leakage likely occurred. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2492330 | 2468242 | 2365038 | 2955842201200078 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-25 | 2012-02-24 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-01-25 | Device Not Retu | Yes | Initial | 2012
40018010 |  | Malfunction |  | On (b)(6) 2012, user facility medwatch report (b)(4) was received: during surgery using the laparoscopic da vinci, the surgeon noticed arcing from the accessory tip cover that was covering the scissors. Upon closer inspection, small holes were discovered in the cover allowing current to escape. The surgeon voiced concern about an increased risk of an inadvertent burn to surrounding tissues. The same problem has been reported twice within the last week. No patient injury was sustained on either event. Additionally it was noted on both used tips that the ink that is used labeling is flaking off making it illegible. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however, upon follow up with the robotic lab technician at (b)(6) hospital it was discovered that tears were seen in the tip cover accessories and that the same surgeon reported both events. The site was provided excerpts from the ifu regarding tip cover arcing prevention tips. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2492324 | 2468249 | 2365065 | 2955842201200087 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2012-01-31 | 2012-02-24 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-31 | No | Yes | Unkown | 2012
40018010 | 2012-01-11 | Malfunction |  | On (b)(6) 2012, user facility medwatch report (b)(4) was received: during surgery using the laparoscopic da vinci, the surgeon noticed arcing from the accessory tip cover that was covering the scissors. Upon closer inspection, small holes were discovered in the cover allowing current to escape. The surgeon voiced concern about an increased risk of an inadvertent burn to surrounding tissues. The same problem has been reported twice within the last week. No patient injury was sustained on either event. Additionally it was noted on both used tips that the ink that is used labeling is flaking off making it illegible. | The tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found two holes near the tip of the tip cover. The holes are approximately 0. 3 inches in diameter and are connected with a crater like shape. The shape and location of the holes indicates an arcing event likely occurred. Upon follow up with the robotic lab technician at (b)(6) hospital, it was discovered that tears were seen in the tip cover accessories and that the same surgeon reported both events. The site was provided experts from the ifu regarding tip cover arcing prevention tips. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2492335 | 2468265 | 2365058 | 2955842201200086 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2012-01-31 | 2012-02-24 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-01-31 | Device Not Retu | Yes | Initial | 2012
42018406 | 2011-12-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the tip on the permanent cautery spatula instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece of the spatula broken off. The broken piece was not returned, however, it measures approximately. 140 x. 085. A piece broke at the forming line where the metal transitions to the flattened spatula shape. A side view of distal end shows a slight bend/curvature directly below the fracture surface, indicating a load was applied in a direction perpendicular to the spatula face. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2492354 | 2468270 | 2365083 | 2955842201200089 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-27 | 2012-02-24 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2012-01-27 | Device Not Retu | No | Initial | 2012
40018010 | 2012-01-27 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while dissecting the right side pelvic lymph nodes, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed, causing a small hole in the patient's right iliac vein. The surgeon repaired the affected area intra-operatively with a suture stitch and was able to complete the planned surgical procedure with a replacement monopolar curved scissors instrument and mcs tip cover accessory. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(6) 2012, isi followed up with (b)(6) and the risk manager reported that the patient's hospitalization was not prolonged and to the best of her knowledge, the patient has not returned to the hospital due to any post-operative complications. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2492355 | 2468273 | 2365086 | 2955842201200088 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-27 | 2012-02-24 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-27 | No | No | Unkown | 2012
42018305 | 2012-01-26 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the permanent cautery hook instrument malfunctioned. After the instrument was removed from the patient, the surgical staff noticed that the collar area was cracked with material missing. The surgeon was able to retrieve a portion of the missing material. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. The engineering investigation found the proximal clevis to be dislodged and cracked at the base, with a piece missing. The missing piece is approximately 0. 2 x 0. 09 in size. A broken piece was also returned with the instrument, and measures approximately 0. 07 x 0. 1 in size. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2492341 | 2468274 | 2365087 | 2955842201200090 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-26 | 2012-02-24 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2012-01-26 | Device Not Retu | No | Reuse | 2012
42018304 | 2012-01-25 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the half of the tip on the permanent cautery hook instrument broke off and fell into the patient's abdomen. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2492389 | 2468283 | 2365096 | 2955842201200091 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-27 | 2012-02-24 | No | Yes | Health Professional | 42018304 | No | Yes | 2012-01-27 | No | No | Unkown | 2012
42017910 | 2012-01-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing/smoking from the monopolar curved scissors (mcs) instrument with the tip cover accessory installed was observed. Examination of the instrument intra-operatively by the surgical staff found that the tip of the mcs instrument was bent. The instrument was immediately removed and the planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension broken at main tube connection. Char marks and melted material can be found. Engineering concluded that breakage was likely due to excess force applied to the tube extension. The instruments and accessories user manual specifically states: proper care and handling is essential for satisfactory performance of surgical instruments and accessories. Examine the instrument or accessory - including all of its components - thoroughly before and after each use. If any abnormality is found, do not use it. Use the device for its intended purpose only. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not used a damaged cannula or reducer. Damaged cannulae and reducers may abrade the instrument shaft and generate particles that may fall inside the patient. The hospital reported that the tip cover accessory was discarded, therefore, evaluation of the tip cover accessory could not be performed. On (b)(4) 2012, the initial reporter indicated that there was no injury to the patient and to the best of his knowledge the patient has not returned to the hospital due to any post-operative complications as a result of the reported event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2492370 | 2468305 | 2365098 | 2955842201200092 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-25 | 2012-02-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-01-25 | Device Not Retu | No | Reuse | 2012
| 2012-02-24 | Malfunction |  | Surgeon informed staff that the maryland bipolar forceps tip was broken. It was discarded from the field. New bipolar forceps were added to the field and the procedure continued without furhter event. Bipolar instrument lifespan 2 of the projected 10 uses. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL MA | 2496190 | 2472121 | 2368949 | 2472121 | 1 | NAY | User Facility |  |  | Initial | 2012-02-24 | 2012-02-24 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-02-24 | Malfunction |  | Rn draping the robot discovered holes in the 4 arm camera drape. It was discarded from the field and a new drape opened. Case proceeded without incident. |  | DA VINCI S | ACCESSORY ROBOTIC SURGICAL 4 A | 2496192 | 2472123 | 2368951 | 2472123 | 1 | NAY | User Facility |  |  | Initial | 2012-02-24 | 2012-02-24 | No | Yes | Nurse |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42018405 | 2012-01-11 | Malfunction |  | Event Description: --/-- It was reported that (b)(6) days post a successful da vinci s hysterectomy procedure, while the patient was undergoing a non related secondary procedure, a small piece of plastic was found in the surgical wound. The site reported that the fragment looked related to the da vinci permanent cautery spatula instrument. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear was broken off at its base. Clevis ear on the opposite side of the conductor wire is broken off, resulting in partial dislodgment of yaw pulley. The broken piece was not returned, but is roughly. 285 x. 235 in size. Additionally engineering observed a missing conductor cap and broken ceramic sleeve. Cap is missing from the distal end. Conductor cap likely detached as a result of clevis ear breakage. Yaw pulley boss feature that mates with cap is broken. Ceramic sleeve is broken and missing pieces, including the entire smaller diameter cylindrical piece at the distal end. This damage is likely the result of mishandling. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULAINSTRUMENT | Reuse | 2490508 | 2466079 | 2362892 | 2955842201200076 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2012-01-25 | 2012-02-23 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2012-03-09 | Device Not Retu | No |  | 2012
IS2000A51P8 | 2012-01-24 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the surgical staff experienced scope alignment issues, focus issues, a recognition issue on one of the instrument arms, and poor image color quality. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issues were associated with the site's endoscopes and camera heads. The fse also determined that the instrument recognition issue was caused by a faulty instrument cannula mount. The fse repaired the system by replacing the camera heads and affected instrument cannula mount. A replacement endoscope was also sent to the site. The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The camera head produces three dimensional images to the da vinci system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. Each instrument arm has a cannula mount installed. The cannula mount is a clamp which holds the cannula to the instrument arm. The cannula mount sensor is inside the cannula mount and detects which type of cannula is use so that surgical motion can be controlled accurately. As of february 22, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2490803 | 2466367 | 2363183 | 2955842201200077 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-24 | 2012-02-23 | Yes | Yes | Health Professional | IS2000A51P8 | Device Returned To Manufacturer | No | 2012-01-24 | Device Not Retu | No | Reuse | 2012
400180 | 2012-02-23 | Malfunction |  | Resident noted tear in clear silicone of accessory tip cover. Did not extend into dark gray area. Valleylab esu set at 35 cut/ 50 coag on dessicate. No arcing noted; no unintended tissue burning noted. Tip cover in use for 1 hour 45 min prior to tear being noticed. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2495492 | 2471538 | 2368351 | 2471538 | 1 | NAY | User Facility |  |  | Initial | 2012-02-23 | 2012-02-23 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2012
| 2012-02-23 | Malfunction |  | Surgeon refused pk dissecting forceps and reported that tip was broken. Discarded from field. New pk dissecting forceps were added to the field and case proceeded without event. No documented inury to patient. Broken tip survived only 5 of projected 10 lives. |  | DA VINCI S | INSTRUMENT ROBOTIC SURGICAL PK | 2496145 | 2472118 | 2368946 | 2472118 | 1 | NAY | User Facility |  |  | Initial | 2012-02-23 | 2012-02-23 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
400180 | 2012-02-21 | Malfunction |  | After 2 hours and 45 minutes of use, it was noted that there was a tear in the clear silicone of the accessory tip cover. Valleylab esu set at 35 cut/ 50 coag, set on dessicate. No arcing noted. Tip changed out as soon as tear was noticed. No known patient harm. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2496146 | 2472119 | 2368947 | 2472119 | 1 | NAY | User Facility |  |  | Initial | 2012-02-23 | 2012-02-23 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2012
DAVINCI SL | 2012-01-27 | Injury | DisabilityHospitalizationRequired Intervention | The patient safety concern regarding the davinci robot is a possible product design issue. The patient suffered a compartment syndrome of the right lower extremity (rle) after the robotic-assisted laparoscopy (with procedures). It is felt that the robotic arm, while docked on the right side of the patient, either bumped and/or rested on the patient's rle. This was a 7 hour procedure. The patient required emergency, 4-compartment fasciotomy followed by another surgery for skin graft (plus wound vac) and wound closure. |  | DAVINCI SL | SYSTEMSURGICALCOMPUTER CONTROL | 2495265 | 2471339 | 2368152 | 2471339 | 1 | NAY | User Facility |  |  | Initial | 2012-02-22 | 2012-02-22 | Yes | Yes | Invalid Data | DAVINCI SL | Yes | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2012-01-23 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure the surgical staff experienced system error code 212. The system was restarted multiple times, however, system error code 212 continued to recur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the system records the fault and transitions to a safe state. The mtmr was returned for failure analysis investigation. Engineering evaluation determined that the harness cable had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2487389 | 2463123 | 2359936 | 2955842201200073 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-23 | 2012-02-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-01-23 | Device Not Retu | No | Reuse | 2012
40018010 | 2012-01-24 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument was failing and arcing and a hole in the tip cover were observed. The tip cover was immediately removed and replaced and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the tip cover is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2487374 | 2463144 | 2359937 | 2955842201200075 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-24 | 2012-02-21 | No | Yes | Health Professional | 40018010 | No | No | 2012-01-31 | No | Yes | Unkown | 2012
42017910 | 2012-01-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a piece of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension dislodged from the main tube and the entire tube extension wall circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No instrument pieces were found to be missing, therefore, the customer reported issue cannot be confirmed. It was determined that the instrument damage may have been caused by excessive side loading or other type of mishandling. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2487383 | 2463153 | 2359966 | 2955842201200074 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-01-23 | 2012-02-21 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-01-23 | Device Not Retu | No | Reuse | 2012
| 2012-02-10 | Malfunction |  | Curved scissor designed to have 10 uses; reportedly dull after 5 uses. Surgoen refuses scissor due to dullness. New curved scissor opened to the field and procedure continues without incident. No documented injury sustained in surgeon's post-operative note. |  | DA VINCI S | INSTRUMENT ROBOTIC SURGICAL MO | 2495254 | 2471332 | 2368145 | 2471332 | 1 | NAY | User Facility |  |  | Initial | 2012-02-21 | 2012-02-21 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-02-10 | Malfunction |  | Maryland bipolar forceps jaw would not close. Surgeon refused forceps. Forceps had 6 previous uses and did not survive recomended 10 uses. A second new maryland bipolar forceps added to the field and surgery proceeded with out incident - no complication documented in surgeon's postoperative notes. |  | DA VINCI S | INSTRUMENT ROBOTIC SURGICAL MA | 2495255 | 2471333 | 2368146 | 2471333 | 1 | NAY | User Facility |  |  | Initial | 2012-02-21 | 2012-02-21 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-02-08 | Malfunction |  | Rn scrub nurse noticed holes in the sterile arm drape. Discarded it and a second drape was opened. Rn noticed holes in second sterile drape also. It was discarded and a third drape opened. Third drape was satifactory. Case proceded without incident. |  | DA VINCI S | ACCESSORY ROBOTIC SURGICAL 4 A | 2495257 | 2471335 | 2368148 | 2471335 | 1 | NAY | User Facility |  |  | Initial | 2012-02-21 | 2012-02-21 | No | Yes | Nurse |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
| 2012-02-08 | Malfunction |  | Rn scrub nurse noticed holes in the sterile arm drape. Discarded it and a second drape was opened. Rn noticed holes in second sterile drape also. It was discarded and a third drape opened. Third drape was satifactory. Case proceded without incident. |  | DA VINCI S | ACCESSORY ROBOTIC SURGICAL 4 A | 2495257 | 2471335 | 2368148 | 2471335 | 1 | NAY | User Facility |  |  | Initial | 2012-02-21 | 2012-02-21 | No | Yes | Nurse |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2012
42017909 | 2012-01-18 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a section of the left blade had broken off. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2483729 | 2459898 | 2356714 | 2955842201200065 | 1 | NAY | Manufacturer | ForeignDistributor | Other | InitialFollowup | 2012-01-18 | 2012-02-17 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2012-02-21 | No | No |  | 2012
40027202 | 2012-01-18 | Malfunction |  | It was reported that during a da vinci si liver resection procedure, the jaw of the harmonic ace curved shears insert accessory (installed on the harmonic ace curved shears instrument) broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the insert accessory is returned or if additional information is received. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2483712 | 2459902 | 2356718 | 2955842201200066 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-18 | 2012-02-17 | No | Yes | Health Professional | 40027202 | No | No | 2012-01-18 | No | Yes | Unkown | 2012
IS2000 A51P8 | 2012-01-17 | Death | Death | It was reported that 4 hours into a da vinci s prostatectomy procedure, it was noted by the anesthesiologist that the patient's health was declining. Several attempts by the surgical team were made to reverse the patient's condition, however, the attempts were unsuccessful and the patient expired. No other additional information was provided. | On (b)(4) 2010, isi contacted the risk manager at (b)(6) hospital and she indicated that due to hippa regulations, she was unable to provide any additional information concerning the reported event. However, she stated that the patient's demise was unrelated to the da vinci s surgical system and there was no malfunction of the system, instruments and/or accessories that caused or contributed to the patient's demise. As of february 17, 2012, no adverse events have been experienced with the site's system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2485080 | 2461093 | 2357909 | 2955842201200067 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-01-19 | 2012-02-17 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-01-19 | No | No | Reuse | 2012
40018010 | 2012-01-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory and/or the mcs instrument are returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2485162 | 2461191 | 2358007 | 2955842201200068 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-19 | 2012-02-17 | No | Yes | Health Professional | 40018010 | No | Yes | 2012-01-19 | No | Yes | Unkown | 2012
40016907 | 2012-01-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the harmonic curved shears insert accessory broke and a piece fell into the patient. The piece was retrieved and removed from the patient. No patient harm, adverse outcome, or injury was reported. | The insert was returned and evaluated and engineering found the clamp arm detached from the distal end of the instrument. Shaft features that mate with clamp arm pins are bent backwards, indicating possible hyperextension of the clamp arm. Curved blade has a small gouge below midpoint. Evidence is not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. -do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL | 2485188 | 2461198 | 2358014 | 2955842201200071 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2012-01-19 | 2012-02-17 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2012-01-19 | Device Not Retu | No | Unkown | 2012
420208 | 2012-01-19 | Malfunction |  | It was reported that during a da vinci si lung biopsy procedure, while using the thoracic grasper instrument, the instrument shaft rubbed against the cannula accessory causing shavings from the main tube to fall into the patient. Irrigation was performed to remove the instrument fragments and the planned surgical procedure was completed. No patient injury was reported. | The instrument has not been returned for evaluation. Per the isi representative present for the procedure, no post operative complications have been reported. Attempts have been made to obtain further information about this event however, as of the time of this report the site has not responded. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. If the instrument is returned or if further information is received, a follow-up mdr will be submitted. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2485191 | 2461203 | 2358019 | 2955842201200069 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2012-01-19 | 2012-02-17 | No | Yes | Health Professional | 420208 | No | No | 2012-01-19 | No | No | Unkown | 2012
42017910 | 2012-01-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the site observed arcing from the monopolar curved scissors (mcs) instrument with tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint as microscopic inspection of the instrument's clevis and tube extension did not reveal any evidence of arcing. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2485165 | 2461207 | 2358023 | 2955842201200070 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-19 | 2012-02-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-01-19 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2012-01-21 | Death | Death | It was reported that two days after a successful da vinci si tongue base resection procedure and while still admitted, the patient experienced bleeding from the operative area and expired. It was reported that the patient did experience bleeding from the lingual artery during the procedure; however, the bleeding was controlled with cautery. The site has been contacted for additional information without success. | Based on the information provided, it was undetermined as to how this event occurred. The site has been contacted for additional information and once this information has been received, a follow up report will be sent to the fda. At this time, the site has not reported a system, instrument or accessory malfunction. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2485276 | 2461298 | 2358114 | 2955842201200072 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-23 | 2012-02-17 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-01-23 | No | No | Reuse | 2012
| 2011-10-21 | Malfunction |  | Pt underwent a robotic-assisted radical hysterectomy with bilateral pelvic lymph node dissection. During the lymph node dissection, there was a noted cautery injury to the right external iliac artery. This was noted to be a superficial burn (defect) measuring about 3 to 4mm. An intraoperative consultation was obtained from vascular surgery. The defect was repaired with two (2) 5-0 prolene through and through sutures incorporating the thermal injured area. Upon inspection of the cautery shears, it was noted that the "tip cover accessory" had a small hole at the tip which allowed current to arc and cause the thermal injury. There was no measureable blood loss associated with this portion of the procedure. |  | DAVINCI ROBOTIC SURGICAL SYSTEM TIP COVER ACCESSORY | TIP COVER | 2488287 | 2462463 | 2359296 | MW5024250 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-11-03 | 2012-02-17 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2012
8 MM | 2012-02-16 | Malfunction |  | Patient for a cabg with da vinci robot. Surgeon attempted to use the da vinci bipolar forceps, but they did not work. Circulator checked the electrosurgical unit (esu) and noted that ground pad (return electrode for esu) was not plugged in. Ground pad connected. Surgeon stated there was no noise and the foot pedal was not depressed, but as soon as the surgeon touched the diaphragm with the bipolar forceps, the machine buzzed and the diaphragm had a small burn. Circulator checked the esu and noted that the forceps cable was plugged into two of the three monopolar 2 sockets. Bipolar was plugged into bipolar sockets, machine functioned normally throughout the rest of the case. Patient had two (2) ground pads on, unsure which was the actual lot number involved in event. ======================manufacturer response for system,surgical computer controlled instrument, da vinci (per site reporter). ======================intuitive representative came to the facility and tested the machine. No issues identified. |  | MICROBIPOAR | INSTRUMENT ROBOTIC SURGICAL BI | 2491716 | 2467656 | 2364472 | 2467656 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-02-17 | 2012-02-17 | No | Yes | Physician | 8 MM | Yes | No Answer Provided |  |  |  |  | 2012
8 MM | 2012-02-16 | Malfunction |  | Patient for a cabg with da vinci robot. Surgeon attempted to use the da vinci bipolar forceps, but they did not work. Circulator checked the electrosurgical unit (esu) and noted that ground pad (return electrode for esu) was not plugged in. Ground pad connected. Surgeon stated there was no noise and the foot pedal was not depressed, but as soon as the surgeon touched the diaphragm with the bipolar forceps, the machine buzzed and the diaphragm had a small burn. Circulator checked the esu and noted that the forceps cable was plugged into two of the three monopolar 2 sockets. Bipolar was plugged into bipolar sockets, machine functioned normally throughout the rest of the case. Patient had two (2) ground pads on, unsure which was the actual lot number involved in event. ======================manufacturer response for system,surgical computer controlled instrument, da vinci (per site reporter). ======================intuitive representative came to the facility and tested the machine. No issues identified. |  | MICROBIPOAR | INSTRUMENT ROBOTIC SURGICAL BI | 2491716 | 2467656 | 2364472 | 2467656 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-02-17 | 2012-02-17 | No | Yes | Physician | 8 MM | Yes | No Answer Provided |  |  |  |  | 2012
400180 | 2012-01-16 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, a hole was observed in the tip cover accessory and then damage was found to the patient's proximal tissue. No repair was required as the surgeon felt the burn was superficial and the planned surgical procedure was completed with no adverse outcome reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2478485 | 2454427 | 2351243 | 2955842201200062 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-16 | 2012-02-16 | Yes | Yes | Health Professional | 400180 | No | No | 2012-01-16 | No | Yes | Unkown | 2012
IS3000 A60P5 | 2012-01-17 | Other | Other | It was reported that during a da vinci si ovarian cystectomy procedure the customer experienced a non-recoverable system error code 25553. With the assistance of an isi technical support engineer (tse) the customer power cycled the system and performed an emergency power off of the patient side cart, however, the system error code reoccurred. The surgeon made the decision to complete the planned surgical procedure using traditional open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the customer was associated with a remote arm controller (rac) of the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25553 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The rac was returned to isi for failure analysis investigation. Engineering investigation revealed that the rac was contaminated with fluid, thus generating the system error code experienced by the customer. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2481889 | 2458104 | 2354920 | 2955842201200063 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-17 | 2012-02-16 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2012-01-17 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2010-10-22 | Other | Hospitalization | Event Description: --/-- On (b)(6) 2012, a pt who underwent a da vinci s prostatectomy procedure on (b)(6) 2010, provided intuitive surgical information regarding their postsurgical experience. The pt indicated that their procedure was completed in 5. 5 hours and the day after surgery, their jp drain filled with urine and leaking outside of the catheter occurred. The pt spent 6 recovery days in the hospital. Their leaking catheter was in place for 5 weeks, with the jp drain in place for 5+ weeks. The pt reported that the procedure left them incontinent and impotent. The pt also indicated that a second surgery has been planned to install an ius (internal urethral support). The pt alleges that their surgery outcome may have been different if the procedure had been performed by a "qualified" surgeon. | Manufacturer Narrative: --/-- On (b)(6) 2012, the hospital's director of risk management informed intuitive surgical that the patient's postsurgical complications were in no way related to the use of the da vinci s surgical system. The director also indicated that the hospital was unable to release any more information without the patient's consent. --/-- Additional Manufacturer Narrative: --/-- Postoperative incontinence and impotence is considered a known and assessed normal risk of prostatectomy procedures, and can develop regardless of the modality used to perform the surgery. Per the information provided, no system malfunctions were alleged. Our records indicate that the surgeon who performed the pt's procedure was trained to use the da vinci s surgical system while conducting their residency. A request for additional information has been made to the hospital's risk management department. To date, no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | Reuse | 2524014 | 2467769 | 2364582 | 2955842201200064 | 1 | NAY | Manufacturer | ConsumerOther | UNKNOWN | FollowupInitial | 2012-01-17 | 2012-02-16 | Yes | No | Health Professional | IS2000 A51P8 | No |  | 2012-03-09 | Device Not Retu | No |  | 2012
65504009 | 2012-01-17 | Injury | Required Intervention | In middle of partial cystectomy, davinci si robot failed. After trying to recover fault several times with intuitive, fault became non-recoverable and surgeon decided to not proceed with the surgery. When service was performed on (b)(4) 2012, problem was caused by fluid during the procedure pouring down the column and into the components. Unit was repaired by replacing with part #655040-09 / serial #(b)(4). Fault code during procedure was 25553. |  | INTUITIVE SURGICAL DAVINCI SI ROBOT | DAVINCI | 2487992 | 2462528 | 2359344 | MW5024262 | 1 | NAY | Voluntary |  | Nurse | Initial | 2012-02-13 | 2012-02-14 | Yes | No | Health Professional | 65504009 |  | Yes |  |  |  |  | 2012
40018010 | 2012-01-12 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, while performing the para-aortic lymph node dissection using the monopolar curved scissors instrument with a tip cover accessory installed, electricity leaked and burned the patient's atherosclerotic aorta. A vascular surgeon was called and the case was converted to traditional laparoscopic techniques to repair the patient's aorta. The patient lost 2 liters of blood but was reported as recovering well as of (b)(6) 2012. | The tip cover accessory was discarded by the site, however the isi representative onsite for the procedure examined the tip cover accessory after the surgery and stated that it exhibited a small hole. The surgeon changed the tip cover accessory twice during the procedure to ensure adequate insulation, however the arcing event occurred. If further information is provided a follow-up mdr will be submitted. As of (b)(6) 2012 the site has not reported any similar occurrences. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2469859 | 2446094 | 2342930 | 2955842201200056 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-12 | 2012-02-10 | Yes | Yes | Health Professional | 40018010 | No | No | 2012-01-12 | No | Yes | Unkown | 2012
IS2000 A51P8 | 2012-01-11 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon experienced limited movement of a patient side manipulator (psm) while articulating the master tool manipulators from the surgeon side cart. Prior to contacting isi technical support, the site undocked the psm and re-docked in an attempt to realign the set-up joint of the psm, however, the issue persisted. Another port incision was created to obtain a better angle for the psm and the endoscopic camera manipulator arm was adjusted to allow viewing of the surgical area from a different angle. Unable to resolve the issue and prior to contacting isi technical support, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) was unable to replicate the issue experienced by the site. The fse concluded the site's da vinci s system functioned properly, however, the limited motion of the patient side manipulator (psm) experienced by the surgeon was likely due to poor port placement and or/or system setup. Recreation of the system set up with the site revealed that the pitch of the affected psm was collapsed, forcing the psm axis 90 degrees to align with the movement of instrument tip when the surgeon articulated the master tool manipulators. The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). No parts were replaced to resolve the issue, however, training was performed to avoid future occurrences. As of february 10,2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2470757 | 2446761 | 2343597 | 2955842201200057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-11 | 2012-02-10 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2012-01-11 | Device Not Retu | No | Reuse | 2012
42017910 | 2012-01-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the monopolar curved scissors (mcs) instrument sparked and burnt through the mcs tip cover accessory installed on the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument did not exhibit any evidence indicative of arcing. The instrument metal components located at the distal end of the instrument does not have any char marks and does not have any burrs and/or damage on the components that would cause an arcing event. The mcs tip cover accessory used in the conjuction with the instrument was not returned. The instrument passed electrical continuity testing. No physical damage to the instrument was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2470759 | 2446785 | 2343621 | 2955842201200058 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-11 | 2012-02-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2012-01-11 | Device Not Retu | No | Reuse | 2012
42009308 | 2012-01-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure a piece broke off of the prograsp forceps instrument and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a segment at the proximal end of one of the grips is broken off. The curved section of the grip adjacent to the pin that mates with the force amp pulley broke off. The broken piece was not returned, however, it measures approximately. 125 x. 040. The cross section of the grip at the fracture is roughly triangular in shape and inspection of fractured surface does not reveal any material flaws. No other damage was found. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2470809 | 2446825 | 2343661 | 2955842201200060 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2012-01-16 | 2012-02-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-01-16 | Device Not Retu | No | Reuse | 2012
42018406 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the permanent cautery spatula instrument sparked. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed an arc track at the yaw pulley. The yaw pulley had a char mark and melted material, indicating that an arcing event occurred. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2470810 | 2446827 | 2343663 | 2955842201200061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-01-16 | 2012-02-10 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-01-16 | Device Not Retu | No | Reuse | 2012
42022703 | 2012-01-04 | Malfunction |  | It was reported that during a da vinci s hysterectomy one of the tips on the pk dissecting forceps instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient injury, harm or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip is broken off at the yaw pulley grip base. The broken piece was not returned. There appears to be a slight twist in the grip at the fracture surface, indicating that twisting motion/mechanical overload may have occurred, thus causing the grip to break. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2468938 | 2444962 | 2341797 | 2955842201200055 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-01-10 | 2012-02-09 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-01-10 | Device Not Retu | No | Reuse | 2012
| 2012-01-27 | Malfunction |  | During the surgical procedure, a small tear was noted on the distal end of the silicon tip cover while being used inside the patient. The tip was inspected prior to incision and no tear was detected at that time. Once the tip tore past 1mm, it was removed from the patient and replaced with a new tip. No injury to the patient. Cautery settings: 35 cut/35 fulgurate coag. New style tip cover. Traditional approach robotic partial nephrectomy. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2473741 | 2449503 | 2346339 | 2449503 | 1 | NAY | User Facility |  |  | Initial | 2012-02-09 | 2012-02-09 | No | Yes | Invalid Data |  |  | No Answer Provided |  |  |  |  | 2012
SI | 2012-02-01 | Malfunction |  | Robot was in use. Cable accidentally knocked out creating a recoverable error. Robot would not move. |  | DAVINCI | SURGICAL SYSTEM | 2480369 | 2452316 | 2349152 | MW5024234 | 1 | NAY | Voluntary |  | Other | Initial | 2012-02-09 | 2012-02-09 | No | Yes | Health Professional | SI |  | Yes |  |  |  |  | 2012
40015604 | 2012-01-06 | Malfunction |  | It was reported that during a da vinci si surgical procedure the 5mm cautery hook a-sure accessory broke and fell into the patient. The accessory fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The accessory was returned and evaluated. Per the customer reported complaint engineering found the metal hook to be detached from the plastic sleeve. The washer that goes in the sleeve is missing. The sleeve can be screwed onto a monopolar cautery instrument indicating the threading is within specification. Engineering concluded the hook was likely pulled away from the sleeve causing the breakage. | 5MM CAUTERY HOOK ASUREACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2465300 | 2441381 | 2338217 | 2955842201200054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-10 | 2012-02-07 | No | Yes | Health Professional | 40015604 | Device Returned To Manufacturer | Yes | 2012-01-10 | Device Not Retu | Yes | Initial | 2012
420179 | 2012-01-25 | Malfunction |  | During surgery, the instrument began to lose conductivity. It would not cut or coagulate. The ground dispersive pad was changed, and the monopolar cable was changed, both without fixing the problem. Then the surgeon noticed that the instrument began to break apart. The surgery was paused and the instrument was replaced. The patient was inspected for signs of injury, and there were no signs of injury. The surgery was completed with another instrument. |  | ENDOWRIST HOT SHEARS | INSTRUMENT ROBOTIC SURGICAL MO | 2491245 | 2466842 | 2363658 | 2466842 | 1 | NAY | User Facility |  |  | Initial | 2012-02-07 | 2012-02-07 | No | Yes | Physician | 420179 |  | No Answer Provided |  |  |  |  | 2012
IS3000 A60P3 | 2012-01-05 | Other | Other | It was reported that while setting up for a da vinci si hysterectomy procedure, the site was unable align the endoscope. With the assistance of isi technical support engineering(tse)the site checked the image through the camera head and found that the image was out of focus. The tse instructed the site to power off the system, reseat the camera cable into the double camera control unit (doco) and tighten the camera cable connection. After restarting the system, the issue persisted. The site swapped to a backup camera head and camera cable, which resolved the issue. Approximately 1 hour later the site contacted isi technical support to report video loss in the left eye of the surgeon side cart. However, prior to contacting isi technical support for further troubleshooting assistance, the surgeon made the decision to complete the procedure using traditional open surgical techniques as observation of the patient's uterus revealed adhesions that would have prevented the surgeon from completing the procedure using the da vinci si system. No patient harm or injury was reported. | The investigation conducted by the field service engineer concluded that the vision issues experienced by the site were related to the site's camera and the double camera controller (doco). The fse confirmed that the site's first camera would not focus and determined that the replacement camera failure was a result of a defective doco, which most likely failed due to the site disconnecting the original camera and camera cable before shutting down the system. The system was repaired by replacing the affected camera and doco. The camera produces three dimensional images to the da vinci si system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. The da vinci si user manual recommends that all system instruments, accessories and auxilliary devices be removed prior to shutting down the system. On (b)(6) 2012, the site reported that the patient's status was normal, and she was released from the hospital on (b)(6) 2012 and had not returned to the hospital due to any post-operative complications. The site has continued to perform procedures and as of (b)(6) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2462223 | 2438882 | 2335718 | 2955842201200053 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-01-05 | 2012-02-03 | Yes | No | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2012-01-05 | Device Not Retu | No | Reuse | 2012
40018010 | 2011-10-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the 8 mm monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover accessory was removed and replacement before the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2460428 | 2437327 | 2334183 | 2955842201200052 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-03 | 2012-02-02 | No | Yes | Health Professional | 40018010 | No | Yes | 2011-10-05 | No | Yes | Initial | 2012
| 2012-01-27 | Injury | Required InterventionOther | While dissecting the right side pelvic lymph nodes, with the arm of the da vinci, an arc occurred at the site of the tip cover ("tip cover accessory") causing a small hole in the patient's vessel. The surgeon repaired the injury with a stitch. Electrical surgical unit was set at 40 and then turned down to 35. The esu appears to be working fine - per the clinical engineer. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2464252 | 2440628 | 2337464 | 2440628 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-02-02 | 2012-02-02 | Yes |  | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
IS3000 A60P6 | 2012-01-02 | Other | Other | It was reported that prior to starting a da vinci si hysterectomy procedure, the site had difficulties calibrating the endoscope for use. The site attempted to troubleshoot the problem by using a back up camera cable and camera head; however, the issue persisted. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the calibration issues experienced by the site was associated with the camera cables. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cables. The camera cables were returned to intuitive surgical for failure analysis investigations. During internal evaluation, engineering was able to replicate the reported failure modes. Conductor damaged was found on the camera cables thus causing the customer reported issue. As of january 31, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2458810 | 2435666 | 2332522 | 2955842201200051 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-02 | 2012-02-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2012-01-02 | Device Not Retu | No | Reuse | 2012
42000606 | 2012-01-03 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the surgeon noticed an internal collision while using the large needle driver instrument. Upon further inspection of the instrument, it was noticed that the main tube exhibited scratches and fragments had fallen into the patient. The instrument fragments were retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument main tube insulation to have scratches and gouge marks consistent with internal collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings "handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. " do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2456752 | 2433638 | 2330494 | 2955842201200046 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-01-06 | 2012-01-31 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2012-01-06 | Device Not Retu | No | Reuse | 2012
40018010 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon observed arcing coming from monopolar curved scissors (mcs) instrument with the tip cover accessory installed. The tip cover accessory was removed and inspected, and a hole was noticed during inspection. A replacement tip cover accessory was installed on the mcs instrument and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover cover has a hole burned through the silicone. Slight burn marks were found on the edge of the hole wall and the size of the hole is approximately 0. 07 x 0. 08. Engineering concluded that the orientation of the hole opening allowed energy to escape the tip cover. There were no mechanical tears in the silicone. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2456963 | 2433843 | 2330739 | 2955842201200047 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-01-05 | 2012-01-31 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-01-05 | Device Not Retu | Yes | Unkown | 2012
40018010 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon observed arcing coming from monopolar curved scissors (mcs) instrument with the tip cover accessory installed. The tip cover accessory was removed and inspected, and a hole was noticed during inspection. A replacement tip cover accessory was installed on the mcs instrument and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover had a couple of holes burned though the silicone. The silicone exhibits localized melting around the holes, indicating an arcing event occurred. The hole sizes are approximately. 025 to. 035 with varying shapes (round and oblong), and are located from. 315 to. 370 above the silicone to sleeve interface. Engineering concluded that multiple holes indicate multiple arcing events occured. The tip cover accessory also exhibits mechanical tearing at the edge of each hole. No other damage found. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2456986 | 2433883 | 2330739 | 2955842201200048 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-01-05 | 2012-01-31 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2012-01-05 | Device Not Retu | Yes | Unkown | 2012
IS3000 A60P6 | 2012-01-04 | Other | Other | It was reported that while preparing to start a da vinci si splenectomy procedure, the site experienced system error code 48100. With the assistance of isi technical support engineering (tse)the site powered down the system, turned off the breaker on vision side cart for approximately 30 seconds and unplugged their s-video cable, however, the system error code reoccurred. The patient was under anesthesia and the port incisions had not yet been made when the surgeon made the decision to complete the planned procedure using traditional open surgical techniques. No patient injury, harm or adverse outcome was reported. | The investigation conducted by field service engineering determined that the system error code 48100 experienced by the site was associated with the personality module, auxiliary/video (pmav) box. The pmav is located on the vision side cart (vsc) and allows system video to be output to external video sources and/or video streams to be brought in to the system for display on the surgeon's viewer. The system alarm (system generated fault code) functioned as designed. System error code 48100 appears when the da vinci si system detects a communication signal was outside of the expected range of values. Upon determining this condition, the safety system puts da vinci s in a non-recoverable safe state. The system was repaired by replacing the affected pmav. As of january 31, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2456976 | 2433887 | 2330743 | 2955842201200049 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-04 | 2012-01-31 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2012-01-04 | Device Not Retu | No | Reuse | 2012
40016907 | 2011-12-19 | Malfunction |  | It was reported that during a da vinci si thyroidectomy procedure, the harmonic curved shears (hcs) inset installed on the harmonic curved shears instrument did not work. The hcs instrument was removed and inspection of the insert found that the shaft of the insert broke off and fell into the patient. The insert fragment was retrieved and the planned procedure was completed with a replacement insert. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint engineering observed that the insert was returned with a piece of the curved blade detached from the distal end. The size of the broken piece is approximately. 430 x. 090. The blade fractured near where it transitions from straight to curved. The blade does not exhibit any scratch marks. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2457006 | 2433909 | 2330765 | 2955842201200050 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2012-01-05 | 2012-01-31 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2012-01-05 | Device Not Retu | Yes | Unkown | 2012
42018406 | 2011-12-12 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s procedure a piece of the working element on the permanent cautery spatula instrument broke off and fell inside the patient's abdomen. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a piece of the spatula tip broken off. The broken spatula piece was also returned and measures approximately. 150 x. 090. The tip fractured at the forming line where it transitions to a flattened spatula shape. Engineering also observed damage to the ceramic sleeve, where a. 090 x. 070 piece of the distal end of the sleeve is broken off. The sleeve has a couple of scratches below the fracture. It was determined that the damage to the spatula and sleeve may be due to overloading at the tip or other type of mishandling/misuse. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2453392 | 2430218 | 2327090 | 2955842201200040 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | FollowupInitial | 2011-12-29 | 2012-01-27 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-01-30 | Device Not Retu | No |  | 2012
42018406 | 2011-12-13 | Malfunction |  | Event Description: --/-- It was reported that after a da vinci s procedure was performed, the sterile processing department observed that a piece of the protective cautery sheath was missing from the permanent cautery spatula instrument. It is unknown if the missing piece had fallen into the patient during the surgical procedure. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the ceramic sleeve had a piece broken off at the distal end. The broken piece is approximately. 070 x. 090 in size. The sleeve also exhibits scratch marks below the broken section. It was determined that the damage is likely due to mishandling/misuse. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2453414 | 2430254 | 2327090 | 2955842201200039 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | InitialFollowup | 2011-12-29 | 2012-01-27 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-02-21 | Device Not Retu | No |  | 2012
IS3000 A60P5 | 2012-01-03 | Other | Other | It was reported that during the beginning of a da vinci si prostatectomy procedure the surgeon dropped the camera. As a result, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by an intuitive surgical field service engineer concluded that the camera issue experienced by the customer was caused by surgeon dropping the camera. The system was repaired by replacing the affected camera. The camera was returned and evaluated and the light guide cable was found to be defective. It was determined that the camera damage was caused by shock. The da vinci si surgical system user manual explicitly states that, environmental or equipment failures may cause the da vinci si system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2453397 | 2430258 | 2327094 | 2955842201200041 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-01-03 | 2012-01-27 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2012-01-03 | Device Not Retu | No | Reuse | 2012
42018405 | 2011-12-27 | Malfunction |  | It was reported that during a da vinci si procedure the tip of the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the proximal clevis ear broken off at the yaw pulley. It was determined that the damage was most likely due to excess force applied to the instrument. Engineering also observed the ceramic sleeve cracked and broken at the spatula base, which was likely caused by mishandling. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2453717 | 2430478 | 2327314 | 2955842201200042 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-28 | 2012-01-27 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2011-12-28 | Device Not Retu | No | Reuse | 2012
42014204 | 2011-12-28 | Malfunction |  | It was reported that during a da vinci s transoral otolaryngology surgery (tors) procedure, black fiber from the distal part of the 5 mm monopolar cautery instrument shaft started peeling off near the tip. The instrument was then used for another procedure and it was reported that the damage became worse with burning. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a large area of the main tube insulation was damaged. It was determined that the damage was likely due to mishandling from contact with the cannula accessory or other instruments. The damage points are aligned in the same orientation and direction. The instruments and accessories user manual specifically states: - proper care and handling is essential for satisfactory performance of surgical instruments and accessories. Examine the instrument or accessory - including all of its components - thoroughly before and after each use. If any abnormality is found, do not use it. Use the device for its intended purpose only. - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not used a damaged cannula or reducer. Damaged cannulae and reducers may abrade the instrument shaft and generate particles that may fall inside the patient. | 5 MM MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2453741 | 2430487 | 2327328 | 2955842201200044 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-30 | 2012-01-27 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | Yes | 2011-12-30 | Device Not Retu | No | Reuse | 2012
42017910 | 2011-12-08 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure a nurse observed that the plastic part of the tip of the monopolar curved scissors (mcs) instrument and the tip cover accessory installed on the instrument were burnt out. It was determined that electric energy leaked through the instrument and tip cover. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring at the proximal end of the tube extension, near the main tube interface. The charred area indicates an arcing event occurred. Engineering also observed a broken tube extension, where the tube extension is dislodged from the main tube. The entire tube extension wall is circumferentially broken at the base of the axial keys and as a result, the tube extension can be rotated 360 degrees. It was determined that the damage may be due to excessive side loading or other type of mishandling, which likely contributed to the arcing event. No other damage was found. The hospital reported that the tip cover accessory was discarded, therefore, evaluation of the tip cover accessory could not be performed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2453756 | 2430497 | 2327338 | 2955842201200043 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-29 | 2012-01-27 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2011-12-29 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2011-12-28 | Other | Other | It was reported that prior to starting a da vinci s hysterectomy procedure, while docking the patient side cart (psc) to the patient, a member of the surgical staff stepped on a fiber cable, causing the system to fault. When the site restarted the system to clear the fault, the error message patient side cart not connected occurred. The site then reseated the fiber cable and restarted the system, however, the error message reoccurred. With the assistance of an isi technical support engineer the site powered down the system, reset the console and the psc breakers and performed an emergency power off. The error message continued to occur after restarting the system. While attempting to provide additional troubleshooting techniques, the surgeon made the decision to convert the procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system fault and error message experienced by the site were associated with a defective fiber cable. A review of the site's system log by the fse found multiple occurrences of system error code 23. The fse reseated the fiber cable on both ends and performed an emergency power off of the patient side cart, however, the error message patient side cart not detected appeared. The fse determined that the fiber cable had internal damage, which caused the system to fault. The fiber cable connects the patient side cart to the surgeon console and passes the video, audio and data during system operation. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The system was repaired by replacing the fiber cable. The fiber cable was returned to isi for evaluation. Engineering was able to replicate the issue experienced by the site. The system error message patient side cart not detected was generated during functional testing of the fiber cable on an in-house da vinci s system. Engineering observation also found that the fiber cable exhibited damage due to cuts and was contaminated with a yellow discoloration. On (b)(6) 2012, the circulating nurse reported that the patient is recovering well and has not reported experiencing any post-operative complications. As of january 27, 2012 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2453836 | 2430586 | 2327422 | 2955842201200045 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-28 | 2012-01-27 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-12-28 | Device Not Retu | No | Reuse | 2012
SS 3000 | 2011-11-25 | No Answer Provided |  | During a transoral robotic-assisted surgical procedure, the robotic monopolar cautery was being used and suddenly a small flame appeared at the surgical site. Md immediately stopped surgery, pa withdrew instrument, and sterile saline was poured into the field extinguishing the flame. It appeared the disposable cautery spatula tip was what had caught on fire as it was burned and a piece fell off - it was removed from mouth. Moistened raytec were in place surrounding the mouth and retractors. Charge nurse and mda (md anesthesiologist) notified. Mouth and throat were examined by surgeon and mda. No apparent harm to patient. Cautery tip was removed and replaced with a new one. Patient was being ventilated with 100% oxygen prior to incident due to low spo2 when n2o was used and after incident water was poured into the mouth and throat and a small cuff leak was noticed. Pt stable and case continued without incident after all parties concurred it was safe to continue. Family notified by md after procedure. |  | DA VINCI | SYSTEMSURGICALCOMPUTER CONTROL | 2453510 | 2430296 | 2327117 | 2430296 | 1 | NAY | User Facility |  |  | Initial | 2012-01-26 | 2012-01-26 | No | No | Physician | SS 3000 | Yes | No Answer Provided |  |  |  |  | 2012
400272 | 2012-01-12 | Malfunction |  | A piece of the jaw of the harmonic curved shears broke off in the patient during surgery. The piece was removed by the surgeon, intact, with no harm to the patient. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL HA | 2455933 | 2432981 | 2329837 | 2432981 | 1 | NAY | User Facility |  |  | Initial | 2012-01-26 | 2012-01-26 | No | Yes | Physician | 400272 | Yes | No Answer Provided |  |  |  |  | 2012
420209 | 2012-01-23 | Malfunction |  | While using the davinci driver, the instrument failed. The surgeon immediately removed it from the patient. Upon removal, it appeared the internal wire at the tip of the instrument snapped. |  | DA VINCI NEEDLE DRIVER | INSTRUMENT ROBOTIC SURGICAL NE | 2455931 | 2432979 | 2329835 | 2432979 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-01-25 | 2012-01-25 | No | Yes | Physician | 420209 |  | No Answer Provided |  |  |  |  | 2012
| 2012-01-19 | Malfunction |  | Patient with prostate cancer undergoing prostectomy - during procedure the harmonic curved shears insert was not cutting and with cauterizing it would spray the blood without cutting or cauterizing. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL HA | 2455932 | 2432980 | 2329836 | 2432980 | 1 | NAY | User Facility |  |  | Initial | 2012-01-25 | 2012-01-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
| 2012-01-12 | Injury | Required Intervention | During a robotic assisted modified radical hysterectomy and bilateral para - aortic lymphadenectomy, the insulating sheath on the monopolar scissors malfunctioned and leaked electricity on to the atherosclerotic aorta. This created an aortic tear which necessitated converting to a laparotomy and a vascular surgeon to repair the aortic injury. |  | DA VINCI | TIP COVER ACCESSORY MONOPOLAR | 2563484 | 2534528 | 2431281 | 2534528 | 1 | NAY | User Facility | Unknown | Other | Initial | 2012-01-25 | 2012-01-25 | No | Yes | Health Professional |  | Device Returned To Manufacturer | No |  | No Answer Provi |  | Invalid Data | 2012
40027202 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the harmonic ace instrument, the harmonic ace insert accessory broke and fell into the patient. The accessory fragment was retrieved and the harmonic ace insert accessory was replaced. The case continued as planned however, the replacement harmonic ace insert accessory broke and fell into the patient. The accessory fragment was retrieved and the procedure was completed with a replacement harmonic ace insert accessory. No patient harm, injury or adverse outcome was reported. | The harmonic ace insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(6) 2012 there have been no reported recurrences of the issue at the hospital. The following medwatch report was sent to the fda for the second harmonic ace insert accessory: mfr report 2955842-2012-00031. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2445273 | 2424066 | 2320908 | 2955842201200032 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-22 | 2012-01-20 | No | Yes | Health Professional | 40027202 | No | No | 2011-12-22 | No | Yes | Unkown | 2012
40027202 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the harmonic ace instrument, the harmonic ace insert accessory broke and fell into the patient. The accessory fragment was retrieved and the harmonic ace insert was replaced. The case continued as planned however, the replacement harmonic ace insert accessory broke and fell into the patient. The accessory fragment was retrieved and the procedure was completed with a replacement harmonic ace insert accessory. No patient harm, injury or adverse outcome was reported. | The harmonic ace insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(4) 2012, there have been no reported recurrences of the issue at the hospital. The following medwatch report was sent to the fda for the second harmonic ace insert accessory: mfr report 2955842-2012-00032. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2445301 | 2424073 | 2320908 | 2955842201200031 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-22 | 2012-01-20 | No | Yes | Health Professional | 40027202 | No | No | 2011-12-22 | No | Yes | Unkown | 2012
IS3000 A60P6 | 2011-12-22 | Other | Other | It was reported that during a da vinci si hysterectomy procedure the surgical staff experienced system error code 23008. The system was restarted multiple times, however, system error code 23008 continued to recur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault codes) functioned as designed. System error code 23008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the faults and transitions to a safe state. The mtmr was returned for failure analysis investigation and engineering was able to replicate the reported failure. During evaluation, a broken screw was found on an axis gear clamp, thus generating the system error code experienced by the customer. As of january 19, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2445316 | 2424089 | 2320924 | 2955842201200030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-22 | 2012-01-20 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-12-22 | Device Not Retu | No | Reuse | 2012
42018405 | 2011-12-14 | Malfunction |  | It was reported that during a da vinci si lung wedge resection procedure while using the permanent cautery spatula instrument, a small piece of the instrument's insulation broke off into the patient. The surgeon attempted to remove the instrument fragment for approximately 30 minutes but was unsuccessful. No patient injury, harm or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the distal clevis to be broken. The clevis ear on the side opposite the conductor wire was found to be broken with a piece approximately. 265 inches by. 200 inches in size missing. Engineering concluded that the damage was likely caused by overloading at the instrument tip. Per the director of perioperative services at (b)(6) the surgeon felt there was more risk in searching for the instrument fragment during the procedure due to the patient's additional time under anesthesia. The patient was made aware of the retained fragment during a post op visit and there are currently no plans to retrieve the fragment. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2445333 | 2424116 | 2320952 | 2955842201200033 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-22 | 2012-01-20 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2011-12-22 | Device Not Retu | No | Reuse | 2012
IS200051P8 | 2011-12-24 | Other | Other | It was reported that while preparing for a da vinci s prostatectomy procedure, upon system start up, the site received a system error 253. With the assistance of an isi technical support engineer, the surgeon checked the system cabling, reset the system power breaker and restarted the system, however the system error 253 continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 253 experienced by the customer was associated with the multiple servo driver (msd) pca board. The system contains five multiple servo driver (msd) pca boards, which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected msd pca board. The msd board was returned to isi for failure analysis investigation. Engineering confirmed the customer reported system error code 253 and found the board output to be out of the acceptable range as of january 19, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2445335 | 2424118 | 2320954 | 2955842201200034 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-12-24 | 2012-01-20 | Yes | Yes | Health Professional | IS200051P8 | Device Returned To Manufacturer | No | 2011-12-24 | Device Not Retu | No | Reuse | 2012
40027202 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the harmonic ace instrument, the harmonic ace insert broke and a piece from the insert fell into the patient. The accessory fragment was retrieved and the harmonic ace insert was replaced. The case continued as planned, however, after some time of use the replacement harmonic ace insert broke and a piece fell into the patient. Unable to locate the accessory fragment, the surgeon made the decision to complete the procedure with a replacement harmonic ace insert. Immediately following completion of the surgical procedure, x-rays of the patient's abdomen revealed the location of the broken piece and it was removed laparoscopically. No patient harm, injury or adverse outcome was reported. | The harmonic ace insert has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert is returned (post engineering evaluation). Per the information provided by the customer, the broken insert piece was retrieved. As of (b)(6) 2011 there have been no reported recurrences of the issue at the hospital. The following medwatch report was sent to the fda regarding the 2nd harmonic ace curved shears insert reported: mfr report 2955842-2012-00037. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2445380 | 2424165 | 2321000 | 2955842201200036 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-22 | 2012-01-20 | No | Yes | Health Professional | 40027202 | No | No | 2011-12-22 | No | Yes | Unkown | 2012
42017910 |  | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the surgical staff observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the clevis and tube extension did not exhibit any damage due to arcing. In addition, arcing from the instrument and intuitive surgical in-house tip cover accessory installed was not observed. --/-- Additional Manufacturer Narrative: --/-- The instrument and mcs tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Several attempts to obtain additional information from the site have been made concerning the reported event; however, no other information has been provided. A follow-up mdr will be submitted if the instrument and/or tip cover accessory is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2445392 | 2424173 | 2321008 | 2955842201200035 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2011-12-30 | 2012-01-20 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-12-30 | Device Not Retu | No |  | 2012
40027202 | 2011-12-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the harmonic ace instrument, the harmonic ace insert broke and a piece from the insert fell into the patient. The accessory fragment was retrieved and the harmonic ace insert was replaced. The case continued as planned, however, after some time of use the replacement harmonic ace insert broke and a piece fell into the patient. Unable to locate the accessory fragment, the surgeon made the decision to complete the procedure with a replacement harmonic ace insert. Immediately following completion of the surgical procedure, x-rays of the patient's abdomen revealed the location of the broken piece and it was removed laporascopically. No patient harm, injury or adverse outcome was reported. | The harmonic ace insert has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the customer, the broken insert piece was retrieved. As of (b)(6) 2011, there have been no reported recurrences of the issue at the hospital. The following medwatch report was sent to the fda regarding the 1st harmonic ace curved shears insert reported: mfr report 2955842-2012-00036. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2445526 | 2424294 | 2321000 | 2955842201200037 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-22 | 2012-01-20 | No | Yes | Health Professional | 40027202 | No | No | 2011-12-22 | No | Yes | Unkown | 2012
42019408 | 2011-12-23 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed a piece from the mega needle driver instrument fall into the patient. The piece was retrieved and the procedure was successfully completed with no harm to the patient. | Manufacturer Narrative: --/-- The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one carbide insert to be missing from one of the grip tips. The adhesive residues were found on the grip surface, indicating the insert was previously installed. No further damage was observed. | MEGA NEEDLE DRIVER INSTRUMENT | Reuse | 2445558 | 2424368 | 2321204 | 2955842201200038 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2011-12-23 | 2012-01-20 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-02-02 | Device Not Retu | No |  | 2012
IS1200 A43P1 | 2011-12-14 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the surgical staff experienced several recurring system faults while homing the system and observed slow movement of a patient side manipulator (psm). With the assistance of an isi technical support engineer (tse), the site replaced the sterile adapter and instrument; however, the issue recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. The delay in reporting is due to an oversight in initial review of the complaint by the responsible complaint handler. | After completion of the surgical procedure, the site contacted isi technical support engineering (tse) for further troubleshooting assistance. With the assistance of the tse, the site inspected the control cable of the affected patient side manipulator (psm). Upon inspection of the cables the site found that the sterile drape on the psm was pinched under the cables. The site cleared the drape and attempted to home the system, however, the system would not home. Review of the site's system log by the tse found system error code 21003. The site was instructed by the tse to perform a power cycle. After completion of the power cycle, the system homed with no other issues noted. The investigation conducted by field service engineering (fse) confirmed the system error code 21003 experienced by the customer and determined that the error code was generated due to drape interference on the patient side manipulator(psm), thus causing the reoccurring system faults and slow movement of the psm during homing. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. Inspection and functional testing by the fse of the patient side manipulators found that there were no other foreign materials present and the system performed to specification. On january 19, 2011, follow-up with the initial reporter indicated that the patient tolerated the open procedure well and to the best of her knowledge the patient is doing well and has not returned to the hospital due to any post-operative complications. No other information was provided. As of january 19, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2443826 | 2422475 | 2319310 | 2955842201200029 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-14 | 2012-01-19 | Yes | Yes | Health Professional | IS1200 A43P1 | Yes | No | 2011-12-14 | Device Not Retu | No | Reuse | 2012
42011006 | 2011-11-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the precise bipolar forceps instrument sparked. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a broken conductor wire and charred yaw pulley. One conductor wire is broken at the yaw pulley exit and the wire is detached from its connection at the grip. One yaw pulley (located on the side with the broken wire) has a char mark next to the drive cable, which indicates that an arcing event occurred. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2441632 | 2420454 | 2317290 | 2955842201200028 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-19 | 2012-01-18 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2011-12-19 | Device Not Retu | No | Reuse | 2012
400180 VER10 | 2012-01-11 | Malfunction |  | During surgery using the laparoscopic davinci, the surgeon noticed arcing from the accessory tip cover that was covering the scissors. Upon closer inspection, small holes were discovered in the cover allowing current to escape. The surgeon voiced concern about an increased risk of an inadvertent burn to surrounding tissues. The same problem has been reported twice within the last week. No patient injury was sustained on either event. Additionally it was noted on both used tips that the ink that is used labeling is flaking off making it illegible. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2448272 | 2426441 | 2323277 | 2426441 | 1 | NAY | User Facility |  |  | Initial | 2012-01-18 | 2012-01-18 | No | Yes | Physician | 400180 VER10 | Yes | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2011-12-19 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the customer experienced a patient side cart not connected system error message. With the assistance of an isi technical support engineer the site cleaned and reconnected the fiber cable, however, the system error message continued to occur. The surgeon made the decision to convert the procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site were associated with a defective fiber cable. The fiber cable connects the patient side cart to the surgeon console and passes the video, audio and data during system operation. The system was repaired by replacing the fiber cable. As of january 17, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2439926 | 2418612 | 2315448 | 2955842201200027 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-12-19 | 2012-01-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-12-19 | Device Not Retu | No | Reuse | 2012
| 2012-01-17 | Malfunction |  | Patient undergoing robotic surgery. Small tear noted at distal end of tip cover accessory. No bovie arcing occurred. No known patient harm. Tip cover replaced. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2444191 | 2422781 | 2319616 | 2422781 | 1 | NAY | User Facility |  |  | Initial | 2012-01-17 | 2012-01-17 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2012
| 2011-12-29 | Malfunction |  |  |  | MEGA NEEDLE DRIVER | INSTRUMENT ROBOTIC SURGICAL NE | 2444193 | 2422782 | 2319617 | 2422782 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2012-01-16 | 2012-01-16 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2012
400180 | 2011-12-02 | No Answer Provided |  |  |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2444194 | 2422783 | 2319618 | 2422783 | 1 | NAY | User Facility |  |  | Initial | 2012-01-16 | 2012-01-16 | No |  | Physician | 400180 | No Answer Provided | No Answer Provided |  |  |  |  | 2012
IS2000 A51P8 | 2011-12-16 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon was unable to position the instrument in their desired location. The surgical staff redocked the affected instrument arm to change the reach, however, this did not resolve the issue. No system error messages were generated and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure prior to calling intuitive surgical technical support for assistance. No patient harm was reported. | A field service engineer (fse) went onsite and evaluated the system. System tests were performed and the fse was unable to replicate the customer reported failure. The system performed as intended and no issues were found. While onsite, the initial reporter informed the fse that the instrument positioning issue experienced by the surgeon was likely caused by port placement. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2436842 | 2415662 | 2312498 | 2955842201200021 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-16 | 2012-01-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2011-12-16 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2011-12-19 | Other | Other | It was reported that while docking an instrument arm to the patient for a da vinci si prostatectomy procedure, non-recoverable system error code 31030 occurred. The system was restarted multiple times, however, the error code continued to appear upon start up. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 31030 was associated with a setup joint (suj) arm. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from it's own joint positions to the remote arm controller. The system was repaired by replacing the affected suj. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 31030 appears when the da vinci safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining this condition, the system records the fault and transitions to a safe state. As of (b)(4), 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2436836 | 2415679 | 2312515 | 2955842201200023 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-19 | 2012-01-13 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-12-19 | Device Not Retu | No | Reuse | 2012
IS1200 A43P9 | 2011-12-18 | Other | Other | It was reported that during a da vinci prostatectomy procedure the surgical staff experienced system error code 20013, which continued to appear after the fault override button was pressed. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 20013 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 20013 system error code appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the system records the fault and transitions to a safe state. The mtml was returned for failure analysis investigation. Engineering evaluation determined that the potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of january 12, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2436859 | 2415681 | 2312517 | 2955842201200022 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-12-18 | 2012-01-13 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-12-18 | Device Not Retu | No | Reuse | 2012
42018305 | 2011-12-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, while using the permanent cautery hook instrument, the yellow plastic piece on the side of the hook broke and fell off into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient injury, harm or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the yaw pulley ear, proximal clevis and grip cable to be broken. Engineering concluded the damage was likely due to excessive force applied to the hook. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2437260 | 2416045 | 2312881 | 2955842201200024 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-16 | 2012-01-13 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2011-12-16 | Device Not Retu | No | Reuse | 2012
IS3000 A60P6 | 2011-12-14 | Other | Other | It was reported that while preparing for a da vinci si surgical procedure the site experienced a system error code 252 when starting the system. With the assistance of an isi technical support engineer the site restarted the system and checked all connections, however upon start up the system error code 252 recurred. The patient was under anesthesia when the site decided to abort the planned procedure. No patient harm, injury or adverse event was reported. | The investigation conducted by the field service engineering concluded that the system error code 252 was associated with the dual camera controller unit (doco). After further inspection it was found that the doco was turned off by the site. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by turning on the doco. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of january 12, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINSI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2437268 | 2416047 | 2312883 | 2955842201200025 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-12-14 | 2012-01-13 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-12-14 | Device Not Retu | No | Reuse | 2012
42009308 | 2011-12-06 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip of the prograsp forceps instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a loose pin on the jaw. Evidence of crimping was observed. No damage was found on clevis nor the cable. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2437273 | 2416050 | 2312886 | 2955842201200026 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-15 | 2012-01-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2011-12-15 | Device Not Retu | No | Reuse | 2012
42020505 | 2011-12-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a piece of broken cable from the fenestrated bipolar forceps instrument fell into the patient. The broken cable was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch up cable broken at the distal clevis hub. The cable segment that contains the crimp is missing from the hub. As a result of the breakage, the pitch down cable is loose from loss of tension. The other cables at the wrist are not damaged. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2434458 | 2413504 | 2310340 | 2955842201200017 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-13 | 2012-01-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2011-12-13 | Device Not Retu | No | Reuse | 2012
42019408 | 2011-12-14 | Malfunction |  | Event Description: --/-- It was reported that while the surgeon was suturing the vaginal cuff during a da vinci si hysterectomy procedure, the mega needle driver instrument insert separated from the tip and fell into the patient. The insert was retrieved and the planned surgical was completed with an instrument of the same type. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the carbide insert broken off the grip tip. The detached insert was also returned and found to be intact, with no signs of cracking. Slight adhesive residue was found on the grip tip surface where the carbide insert was installed. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | Reuse | 2434519 | 2413558 | 2310394 | 2955842201200016 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-12-14 | 2012-01-12 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-03-23 | Device Not Retu | No |  | 2012
IS3000 A60P6 | 2011-12-13 | Death | Death | It was reported that during a da vinci si hysterectomy procedure, the patient expired from a pulmonary embolism. | On (b)(6) 2012, the physician who performed the surgical procedure provided the following information: the patient expired after the da vinci si surgical system portion of the procedure was completed - no system issues occurred during use - the cause of death is still undetermined. Additional information was requested. To date, no further information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2434704 | 2413676 | 2310512 | 2955842201200018 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-13 | 2012-01-12 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-12-13 | No | No | Reuse | 2012
IS2000 51P8 | 2011-12-07 | Death | Death | On (b)(6) 2011, a patient underwent a da vinci s hysterectomy procedure. It was reported that the procedure took 5 hours and 25 minutes to complete, with the patient in steep trendelenburg position during the case. 6-8 hours post-op, the patient suffered sudden bilateral pneumothorax (tension) which rendered her comatose and in a persistent vegetative state. Life support was removed on (b)(6) 2011, and the patient expired. | On (b)(6) 2011, the hospital or manager indicated that no malfunctions occurred with the da vinci s surgical system during the procedure and that the patient's family refused an autopsy. The or manager also informed intuitive surgical that the hospital is unwilling to provide any further information regarding this event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2434751 | 2413729 | 2310565 | 2955842201200019 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-13 | 2012-01-12 | Yes | No | Health Professional | IS2000 51P8 | No | No | 2011-12-13 | No | No | Reuse | 2012
IS3000 A60P6 | 2011-12-13 | Other | Other | It was reported that during set up of the da vinci si system, the surgeon experienced flickering in both eyes of the high resolution stereo viewer (hrsv). With the assistance of an isi technical support engineer, the site reset factory setting and re-seated the camera cable at both ends; however, the issue persisted. The patient was under anesthesia when the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected camera cable. The da vinci si surgical system user manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have back up equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of january 11, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2435187 | 2414136 | 2310972 | 2955842201200020 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-13 | 2012-01-12 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-12-13 | Device Not Retu | No | Reuse | 2012
400180 | 2012-01-11 | Malfunction |  | At the conclusion of the 260 minute robotic surgery, a hole was noted to have developed in the tip cover accessory with charring on the distal portion of the tip. No stray bovie arcing noted to have occurred. No patient harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2438573 | 2417367 | 2314203 | 2417367 | 1 | NAY | User Facility |  |  | Initial | 2012-01-11 | 2012-01-11 | No | Yes | Physician | 400180 |  | No Answer Provided |  |  |  |  | 2012
400180 | 2011-12-05 | Malfunction |  | The patient was undergoing a robotic laparoscopic total hysterectomy when a small insulation failure occurred with the tip cover accessory of the r1 endoshears resulting in a small spot serosal touch injury on the sigmoid colon. This was irrigated and reinforced with 2 sutures. No permanent harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2434290 | 2413283 | 2310119 | 2413283 | 1 | NAY | User Facility |  |  | Initial | 2012-01-10 | 2012-01-10 | No | Yes | Physician | 400180 |  | No Answer Provided |  |  |  |  | 2012
42018304 | 2011-12-07 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, some of the rubber from the permanent cautery hook instrument was observed floating around in the patient's abdomen. All of the rubber that could be retrieved was and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis ear was broken off of the instrument. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2425992 | 2404953 | 2301789 | 2955842201200014 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-14 | 2012-01-09 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2011-12-14 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2011-12-13 | Other | Other | It was reported that during a da vinci s lung biopsy and exploratory surgical procedure, the surgeon experienced a double image issue at the console. An intuitive surgical technical support engineer (tse) attempted to troubleshoot the issue via telephone and requested that the endoscope be replaced, however, the issue was not resolved after replacement. The surgical staff was also having an issue with one of the system arms, where the jaws of the instrument installed on the arm were not opening and closing correctly. The site reseated the instrument sterile adapter, however, the issue remained. The tse requested that the surgeon re-drape the system arm, however, the surgeon declined. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the customer was associated with the camera. The camera produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. A camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera. The patient side manipulator (psm) arm that the surgical staff was experiencing an instrument issue with was also evaluated by the fse. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The fse was unable to duplicate the reported issue. The fse then verified normal functionality of the psm with the customer. The camera was returned to the original equipment manufacturer (oem) for evaluation. The oem observed that the camera stage assembly had been damaged from mechanical shock, thus resulting in the vision issue experienced by the customer. As of january 5, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2426001 | 2404959 | 2301795 | 2955842201200015 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-13 | 2012-01-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-12-13 | Device Not Retu | No | Reuse | 2012
IS1200 A43P9 | 2011-12-07 | Other | Other | It was reported that during a da vinci nephrectomy procedure, the site experienced a system error codes 20002 and 20008. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer confirmed that system error codes 20002 and 20008 were associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error codes 20002 and 20008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and/or the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering discovered that a flex circuit assembly had malfunctioned, thus generating the system error codes experienced by the customer. As of january 5, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2424257 | 2403089 | 2299925 | 2955842201200013 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-07 | 2012-01-08 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-12-07 | Device Not Retu | No | Reuse | 2012
IS2000 A51P8 | 2011-12-08 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site experienced a system error codes 23008. The patient was under anesthesia and port incisions had been made when the surgeon decided to abort the planned procedure. No patient injury, harm or adverse outcome was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the right master tool manipulator. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The mtmr was returned to isi for evaluation. Engineering discovered that a flex circuit assembly had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2424259 | 2403091 | 2299927 | 2955842201200012 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-08 | 2012-01-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-12-08 | Device Not Retu | No | Reuse | 2012
IS3000 A60P5 | 2011-12-01 | Death | Death | It was reported that several days after a successful da vinci si tongue base resection (tors) procedure and after being discharged home, the patient experienced bleeding from the operative area and expired. No additional information has been provided; however, the site has been contacted and additional information is expected. | Based on the information provided, it was undetermined as to how this event occurred. The site has been contacted for additional information and once this information has been received, a follow up report will be sent to the fda. At this time, the site has not reported a system, instrument or accessory malfunction. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2422052 | 2400903 | 2297739 | 2955842201200008 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-06 | 2012-01-06 | Yes | No | Health Professional | IS3000 A60P5 | No | No | 2011-12-06 | No | No | Reuse | 2012
IS3000 A60P6 | 2011-09-15 | Other | Other | It was reported that prior to starting a da vinci si surgical procedure, the site experience no image in the right eye of the high resolution video (hrsv) monitor. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No patient harm or injury was reported. No additional information has been provided. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the site was associated the high resolution video (hrsv) monitor. The hrsv monitor provides the video image for the surgeon side console (ssc). The system was repaired by replacing the affected high resolution video (hrsv) monitor. As of december 4, 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2422095 | 2400953 | 2297789 | 2955842201200009 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-12-06 | 2012-01-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-12-06 | Device Not Retu | No | Reuse | 2012
42030901 | 2011-12-09 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, one of the tips on the medga suturecut needle driver instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragment was retrieved from the patient and the procedure was successfully completed without incident. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2423686 | 2402476 | 2299312 | 2955842201200011 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-09 | 2012-01-06 | No | Yes | Health Professional | 42030901 | No | No | 2011-12-09 | No | No | Unkown | 2012
40018010 | 2011-12-07 | Malfunction |  | It was reported that approximately 5 minutes into a da vinci s prostatectomy procedure, while the surgeon was dissecting tissue using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, the surgical staff smelled burning plastic. Intraoperative observation of the mcs instrument and tip cover found that both the instrument and accessory had evidence of burned material. Both the mcs instrument and tip cover were removed and the planned surgical procedure was successfully completed with a replacement mcs instrument and mcs tip cover accessory. No patient harm, injury or adverse outcome was reported. | The tip cover accessory was returned and evaluated. Engineering confirmed the customer reported failure mode and found that the tip cover has a. 375 x. 115 hole burned through the silicone and ultem sleeve. Charring on the sleeve indicates that an arcing event occurred. The damage appears to have started near the ultem to sleeve interface. The mcs instrument used during the procedure was also returned and evaluated. Engineering found that the tube extension and proximal clevis on the instrument also have charring in the same location as the burned tip cover. No other damage was found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2423730 | 2402556 | 2299392 | 2955842201200010 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-09 | 2012-01-06 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-12-09 | Device Not Retu | Yes | Unkown | 2012
42018305 | 2011-12-07 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, while using the permanent cautery hook instrument, the instrument broke on one side and a fragment fell into the patient. The instrument fragment was immediately removed and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. --/-- Additional Event Description: --/-- On (b)(4) 2012 maude event report (b)(4) was received from the fda: scheduled form a robotic abdominal perineal resection on 2011 was started at 0933, the robotic portion was started approx at 1010, the intuitive instruments were inserted into the robotic arms, the procedure was started, dr started using the intuitive permanent cautery hook, and the site of the instrument broke off into the patient. It is presumed that this piece of instrument is left in the patient. The piece of instrument is not able to be xrayed. Dr assessed the patient with the da vinci camera internally, she instilled 600ml normal saline thru blake drain into the patient's pelvis, the patient was turned from right to left, and into reverse trendelenburg to see if the piece would leave the patient thru irrigation fluid. All lap sponges and raytexs were felt for a piece of equipment. The specimen - rectum and sigmoid were palpitated x2 and fluid from rectum and sigmoid were checked for any instrument piece. The intuitive rep was present throughout the case. Dr came into the room, he was notified, suggested a cat scan of the patient, dr notified of a cat scan. Patient with no apparent injury taken to pacuvia bed. Charge nurse notified. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering found one distal clevis ear, on the conductor wire side, to be completley broken off of it's base resulting in dislodgement of the yaw pulley. The broken piece is approximately. 285 inches by. 220 inches in size and was not returned with the instrument. --/-- Additional Manufacturer Narrative: --/-- Based on the information from maude event report (b)(4), the fragment was not retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 2420697 | 2399142 | 2295978 | 2955842201200007 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | FollowupInitial | 2012-01-13 | 2012-01-05 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2012-01-13 | Device Not Retu | No |  | 2012
40018010 | 2011-12-07 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, when the surgical staff removed the monopolar curved scissors (mcs) instrument, the tip cover accessory was not installed. The patient's abdomen was searched and the tip cover accessory was found and removed laparoscopically. A replacement tip cover accessory was placed on the mcs instrument and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. Multiple attempts have been made to obtain additional information from the site, however as of the date of this report no additional information has been provided, therefore the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2420698 | 2399143 | 2295979 | 2955842201200006 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-07 | 2012-01-05 | No | Yes | Health Professional | 40018010 | No | No | 2011-12-07 | No | Yes | Unkown | 2012
IS3000 A60P6 | 2011-12-06 | Other | Other | It was reported that approximately 2 hours during a da vinci si hysterectomy procedure the surgeon experienced difficulty articulating an instrument installed on a specific patient side manipulator (psm) arm. Prior to contacting technical support to assist with troubleshooting the issue, the surgeon decided to convert to traditional open surgical techniques in order to control unexpected patient bleeding. The planned procedure was completed and no adverse outcome was reported. | The instrument has not been returned for evaluation. Multiple attempts to obtain additional information about the event, the patient's post operative status, and device return have been made, however as of the date of this report the site has not responded. The root cause of the customer reported failure mode cannot be determined. If the additional information is received, a follow-up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2420715 | 2399149 | 2295985 | 2955842201200004 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-12-06 | 2012-01-05 | Yes | Yes | Health Professional | IS3000 A60P6 | No | Yes | 2011-12-06 | No | No | Reuse | 2012
IS2000 51P8 | 2011-12-06 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the site experienced a flickering yellow image in the right eye of the high resolution stereo viewer (hrsv). Prior to contacting technical support to assist with troubleshooting the issue, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineering concluded that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2420718 | 2399155 | 2295991 | 2955842201200003 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-12-06 | 2012-01-05 | Yes | Yes | Health Professional | IS2000 51P8 | Yes | No | 2011-12-06 | Device Not Retu | Yes | Reuse | 2012
IS2000 A51P8 | 2011-11-16 | Injury | Other | It was reported that 9 days post a successful da vinci s hysterectomy procedure, the patient returned to the hospital with complications and was found to have a damaged ureter. | Based on the information provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the patient's injury. Multiple attempts have been made to obtain additional information about the event and the patient's post operative recovery, however as of the date of this report the site has not responded. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2420739 | 2399186 | 2296022 | 2955842201200005 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2011-12-06 | 2012-01-05 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2011-12-06 | Device Not Retu | No | Reuse | 2012
42017207 | 2011-11-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure one of the jaws from the maryland bipolar forceps instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip detached from the distal end. The grip had fractured within the yaw pulley, at the cross section of the conductor wire hole. Installation of the broken grip back into the yaw pulley shows a slight offset at the tips and separation between the yaw pulley and pulley cover. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2419075 | 2397676 | 2294509 | 2955842201200002 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-12-06 | 2012-01-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2011-12-06 | Device Not Retu | No | Reuse | 2012
420142 | 2011-11-25 | Malfunction |  | Event Description: --/-- It was reported that while the surgeon was maintaining hemostasis with the 5 mm monopolar cautery instrument during a da vinci si tongue base resection procedure, the instrument caught on fire. The instrument and 5 mm cautery spatula tip accessory that was installed on the instrument were immediately removed and the planned surgical procedure was completed with a replacement instrument and tip accessory. No patient harm was reported. | Manufacturer Narrative: --/-- On (b)(6) 2012, the hospital provided a video of the surgical procedure to intuitive surgical. During analysis of the video, it was determined that the polypropylene sleeve of the spatula tip accessory had ignited. After further communications with the hospital, intuitive surgical was informed that the patient was administered anesthesia via an endotracheal tube with a cuff installed and a mixture of oxygen, nitrous oxide, and sevoflurane were given to the patient. Cautery for the instrument was supplied by a valleylab esu. During the procedure the patient became de-saturated (blood level went down) and the oxygen level was increased to 100%. The anesthesiologist announced the oxygen level increase to the surgeon, however, the announcement was not heard by surgeon. The flame occurred when the oxygen level was at 100%. Based on the additional information provided, it was concluded that the use of oxygen at 100% was directly related to the product malfunction experienced by the customer. It was determined that the high oxygen enriched environment, followed by the spark source (esu cautery), allowed the spatula tip accessory to ignite. On (b)(6) 2012, intuitive surgical informed the hospital that the or safety precautions provided by valleylab specifically states, avoid oxygen enriched environments. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2011, a hospital representative indicated that the affected instrument and tip accessory will not be returning for evaluation as they were discarded, thus, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 5 MM MONOPOLAR CAUTERY INSTRUMENT | Unkown | 2417229 | 2395627 | 2292463 | 2955842201200001 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-12-07 | 2012-01-03 | No | Yes | Health Professional | 420142 | No | No | 2012-02-24 | Device Not Retu | Yes |  | 2012
IS2000 A51P8 | 2011-12-05 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, system error code 212 occurred. The surgical staff power cycled the system several times, however, system error code 212 continued to recur after power up. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the system records the fault and transitions to a safe state. The mtml was returned for failure analysis investigation. Engineering evaluation determined that the harness cable had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2415525 | 2393590 | 2290425 | 2955842201100422 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-05 | 2012-01-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-12-05 | Device Not Retu | No | Reuse | 2012
42018306 | 2011-10-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the permanent cautery hook (pch) instrument sparked and then smoke was observed coming from the joint at the tip of the instrument. The planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2415553 | 2393595 | 2242634 | 2955842201100423 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-12-01 | 2012-01-02 | No | Yes | Health Professional | 42018306 | No | No | 2011-12-01 | No | No | Unkown | 2012
IS3000 A60P5 | 2011-12-01 | Other | Other | It was reported that at the beginning of a da vinci si revascularization procedure, the monopolar spatula instrument would not energize when the surgeon applied cautery energy. An isi representative present during the surgical procedure contacted isi's technical support engineering (tse) for troubleshooting assistance. With the assistance of the tse, the site swapped the cautery cable and reseated the energy activation cable, however, the instrument would not cauterize. At this time the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. | Upon follow-up with the site by an isi field service engineer (fse), the or manager indicated that no field investigation by the fse was required, as functional testing of their da vinci system, instruments and accessories found that the they functioned properly and that the surgeon's inability to apply cautery energy was likely due to poor patient grounding and/or improper system setup. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2415554 | 2393596 | 2290431 | 2955842201100424 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-01 | 2012-01-02 | Yes | No | Health Professional | IS3000 A60P5 | No | No | 2011-12-01 | No | No | Reuse | 2012
| 2012-06-25 | Injury | Required Intervention | During robot assisted surgery the monopolar curved scissors came apart, the pieces had to be retrieved. |  | INTUITIVE DAVINCI SURGICAL SYSEMS | MONOPOLAR CURVED SCISSORS | 2674776 | 2643177 | 2539793 | MW5026011 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2012-06-27 | 2012-06-27 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2012
| 2012-05-09 | Injury |  | Monopolar insulator for scissors tip has hole. |  | DAVINCI | MONOPOLAR INSULATOR FOR SCISSO | 2606617 | 2575632 | 2472345 | MW5025371 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2012-05-10 | 2012-05-10 | Yes | No | Health Professional |  | No | Yes |  |  |  |  | 2012
42027804 | 2011-11-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si gastric bypass procedure, the instrument arm that the graptor instrument was installed on banged against another instrument arm. As a result, the instrument housing came off and pieces of the instrument fell into the patient. The fragmented pieces were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the part is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that two of the four housing snap tabs and three adjacent housing pins were broken. As a result of the damage, the housing can be removed. The housing is located on the back end of the instrument and does not come into patient contact during the surgical procedure. The front end of the instrument was found to be intact. Since there is no patient contact with the housing, the initial reporter was contacted for additional information regarding the reported product problem. On (b)(4) 2012, the initial reporter indicated that the information provided when the complaint was initially reported was incorrect and that no instrument pieces fell into the patient during the surgical procedure. | GRAPTOR RETRACTOR INSTRUMENT | Reuse | 2415143 | 2393447 | 2290262 | 2955842201100419 | 1 | GEI | Manufacturer | User facility | Other | FollowupInitial | 2011-11-25 | 2011-12-30 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No |  | 2011
40018010 | 2011-11-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that a flash arced from the monopolar curved scissors (mcs) tip cover accessory (installed on the mcs instrument) and burnt the tip cover and the patient's ovarian tube. The initial reporter indicated that this was not an issue since the patient's ovarian tubes and ovaries were being removed as part of the planned surgical procedure. No further issues were reported and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the interface between the silicone and ultem sleeve. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole size is approximately. 045 in diameter. A small tear was also observed on the distal end of the silicone tip. The mcs instrument used with the affected tip cover accessory when the malfunction occurred was also returned and evaluated. Engineering observed a char mark on the instrument proximal clevis, adjacent to one of the window features that exposes the clevis seal. The tip cover hole position aligns with the clevis char mark, which leads engineering to believe that the char mark is likely where energy escaped through the tip cover. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2415144 | 2393448 | 2290263 | 2955842201100418 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-30 | 2011-12-30 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-11-30 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-12-01 | Other | Other | It was reported that while preparing to start a da vinci s prostatectomy procedure, the site experienced color bars in the right and left eye of the surgeon side console and found the led light on the digitizer to be off. With the assistance of an isi technical support engineer, the site checked the camera controller unit (ccu), reseated the power cable and reseated the system cable, however the digitizer would not restart. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for the following day. No patient injury, harm or adverse outcome was reported. | An investigation conducted by field service engineering confirmed the customer reported problem was associated with the digitizer. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images. The system was repaired by replacing the digitizer. As of december 30, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2415357 | 2393492 | 2290327 | 2955842201100420 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-12-01 | 2011-12-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-12-01 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-18 | Other | Other | It was reported that while preparing for a da vinci s pharyngectomy procedure, the customer experienced a system error code 23007. The patient was under anesthesia and the port incisions had been made when the surgeon decided abort planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error code 23007 experienced by the customer was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23007 appears when the da vinci s safety system determines, on startup, that one or more robotic joints on the manipulator did not make the prescribed test motion to within the specified tolerance. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The ecm was returned for evaluation. Engineering was able to replicate the customer reported system error code and found the axis 3 insertion drive to have loose bearings which casued the ecm arm to require significant force to move. All 5 bearings in axis 3 were replaced. As of december 30, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2415359 | 2393493 | 2290328 | 2955842201100421 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-01-18 | 2011-12-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-18 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-12-01 | Other | Other | It was reported that while preparing to start a da vinci s prostatectomy procedure, the site experienced the left eye in the surgeon side console to not be working. The isi representative onsite for the procedure found that the camera controller unit (ccu) power supply was off. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for the following day. No patient injury, harm or adverse outcome was reported. | An investigation conducted by field service engineering confirmed the customer reported problem was associated with the camera control unit (ccu) power supply. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu power supply. As of december 30, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2415388 | 2393496 | 2290331 | 2955842201100417 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-12-01 | 2011-12-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-12-01 | Device Not Retu | No | Reuse | 2011
420184 | 2011-12-12 | Injury | Required Intervention | A piece of the working element on a permanent cautery spatula being used in the o. R. Broke off inside the abdomen during a procedure. Piece was visualized and removed. Case continued with a new device. |  | DA VINCI S SURG | PERMANENT CAUTERY SPATULA | 2419807 | 2396742 | 2293578 | MW5023644 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-12-16 | 2011-12-29 | No | Yes | Health Professional | 420184 | Yes | No |  |  |  |  | 2011
420230 | 2011-12-08 | Malfunction |  | Da vinci hemo-lok clip applier failed to open. Rep (b)(6) attempted to troubleshoot with no luck. Instrument was new, believed to be manufacturer flaw. |  | LARGE HEMOLOCK CLIP | DA VINCI SURGICAL | 2419789 | 2396766 | 2293599 | MW5023643 | 1 | FZP | Voluntary |  | RISK MANAGER | Initial | 2011-12-16 | 2011-12-29 | No | Yes | Health Professional | 420230 | Yes | No |  |  |  |  | 2011
420184 | 2011-12-13 | Malfunction |  | Sterile processing dept. Found a piece of the protective cautery sheath was missing after the procedure. |  | DA VINCI SURGICAL | PERM CAUTERY SPATULA | 2419744 | 2396771 | 2293604 | MW5023645 | 1 | HHT | Voluntary |  | RISK MANAGER | Initial | 2011-12-16 | 2011-12-29 | No | Yes | Health Professional | 420184 | Yes | No |  |  |  |  | 2011
420184 | 2011-12-22 | Injury | Disability | Cautery spatula was arcing during use in an intra-abdominal robotic procedure. Reason for use: laparoscopic robotic hysterectomy. |  | CAUTERY SPATULA REF 420184 | CAUTERY SPATULA | 2419694 | 2396775 | 2293608 | MW5023650 | 1 | NAY | Voluntary |  | Nurse | Initial | 2011-12-28 | 2011-12-28 | No | Yes | Health Professional | 420184 | Yes | Yes |  |  |  |  | 2011
IS2000 A51P8 | 2011-11-22 | Other | Other | It was reported that during a da vinci s nephrectomy procedure, vision was lost on the touchscreen and slave monitors when a surgical staff member tripped over the power cord connected to the vision cart. With the assistance of an intuitive surgical technical support engineer, the site attempted to troubleshoot the issue by verifying that the power cords were fully seated and then power cycling the vision cart multiple times. These actions were unable to resolve the issue and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the vision issue experienced by the customer was associated with the video digitizer. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images (such as the user interface icons and text or an auxilliary image from an ultrasound or other source). The system was repaired by replacing the video digitizer. The digitizer was returned and evaluated. Failure analysis investigation was performed and engineering was unable to replicate the reported failure mode. As of december 20, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2407566 | 2386433 | 2283289 | 2955842201100414 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-11-22 | 2011-12-22 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-11-22 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-07-25 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing was observed at the tip of the monopolar curved scissors instrument. When the instrument was removed from the patient, a small hole was noticed on the tip cover accessory that was installed on the mcs instrument. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole on the tip cover wall, approximately 0. 04 in diameter. A slight burn mark was found on the edge of the hole wall. Orientation of the hole opening confirms energy leakage from the scissor blades. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2407611 | 2386458 | 2283313 | 2955842201100413 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-11-29 | 2011-12-22 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-11-29 | Device Not Retu | Yes | Initial | 2011
IS3000 A60P6 | 2011-11-29 | Other |  | It was reported that while docking the patient cart to the patient for a da vinci si prostatectomy procedure the surgical staff experienced system error code 25326. The error was cleared by power cycling the system, however, error code 23025 then occurred. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the system error codes experienced by the customer were associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 25326 is generated when a remote arm controller detects an issue which is isolated to its associated arm and triggers a fault reaction only on that arm network. System error code 23025 occurs when the remote arm controller hardware detects an encoder quadrature fault. Upon determining these conditions, the system records the faults and transitions to a safe state. The psm was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported issue and determined that the axis-1 motor/encoder had malfunctioned, thus generating the system error codes experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2407675 | 2386550 | 2283405 | 2955842201100415 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-29 | 2011-12-22 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-11-29 | Device Not Retu | No | Reuse | 2011
42009308 | 2011-11-28 | Other |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure that utilized the prograsp forceps instrument, the patient was burned. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. As the instrument is non cautery capable the customer reported complaint could not be confirmed. No cautery damage was evident on the instrument and no damage was found. Multiple attempts have been made to gather additional information from the account, however as of the time of this report the site has not responded. If additional information is received a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for failure analysis investigation. The root cause of the customer reported failure mode cannot be determined. Multiple attemtps have been made to gain further information about this event, however, as of the date of this report the site has not responded. If additional information is received a follow-up mdr will be submitted. | PROGRASP FORCEPS INSTRUMENT | Reuse | 2408138 | 2386942 | 2283798 | 2955842201100416 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2011-12-22 | 2011-12-22 | Yes | No | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2011-12-22 | Device Not Retu | No |  | 2011
42018406 | 2011-11-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the insulation near the tip of the permanent cautery spatula instrument broke and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found a piece of the ceramic sleeve to be broken off. The distal end of the sleeve has a. 105 inch by. 080 inch piece missing and there is a small gap/separation between the spatula shank and the intact section of the distal end of the sleeve. Engineering concluded that the damage was likely due to mishandling and or misuse. No other damage was found. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2404505 | 2383353 | 2280209 | 2955842201100406 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-11-21 | 2011-12-20 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-11-21 | Device Not Retu | No | Reuse | 2011
400180 | 2011-12-05 | Malfunction |  | It was reporting that toward the end of a da vinci si hysterectomy procedure arcing was observed coming from the tip of the monopolar curved scissors instrument with tip cover accessory installed. The surgical staff observed bowel harm at the superficial serosa to sigmoid area which was reinforced with suture. The procedure was successfully completed and no additional patient harm warm was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2404545 | 2383356 | 2280212 | 2955842201100408 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-05 | 2011-12-20 | No | Yes | Health Professional | 400180 | No | No | 2011-12-05 | No | Yes | Unkown | 2011
420142 | 2011-12-01 | Other | Other | It was reported that during a da vinci si tongue base resection the surgical staff observed two intraoral lesions after the use of the 5mm monopolar cautery instrument. The procedure was successfully completed and no additional patient injury was reported. User facility medwatch report (b)(4) was received on (b)(4) 2011, regarding this event. As of the date of this response, no additional information has been provided from the site, despite making attempts. | Multiple attempts have been made, however at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | 5MM MONOPOLAR CAUTERY INSTRUME | ENDOSCOPIC ELECTROSURGICAL INS | 2404566 | 2383357 | 2259178 | 2955842201100409 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-12-02 | 2011-12-20 | Yes | No | Health Professional | 420142 | No | No | 2011-12-02 | No | No | Unkown | 2011
40018010 | 2011-11-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The instrument and tip cover accessory were removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The hospital also reported that during visual inspection of the tip cover accessory, a small hole was observed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole approximately 0. 03 in diameter on tip cover wall. Two cut openings approximately 0. 18 and 0. 07 in length were also found on the tip cover wall. Slight burn mark was found along the edge of the hole wall. No other damage was found. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2404570 | 2383361 | 2280217 | 2955842201100411 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-11-21 | 2011-12-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-11-21 | No | Yes | Initial | 2011
40018010 | 2011-11-15 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The instrument and tip cover accessory were removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The hospital also reported that during visual inspection of the tip cover accessory, a small hole was observed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed two holes approximately 0. 04 and 0. 025 in diameter on the tip cover wall. A slight burn marks were found on the edge of the hole walls. No other damage was found. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2404549 | 2383362 | 2280217 | 2955842201100412 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-11-21 | 2011-12-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-11-21 | No | Yes | Initial | 2011
IS1200 A43P9 | 2011-11-22 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site experienced a system error code 20008. With the assistance of an isi technical support engineer (tse), the site undocked the system from the patient and attempted to exercise the patient side cart (psc) arm axis, apply emergency power off (epo) multiple times as well as reseat system cables; however, the system error code 20008 continued to occur. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering found that system error code 20008 as well as system error code 20105 and 20205. It was determined that these system error codes experienced by the customer were associated with a patient side manipulator (psm) arm and a remote interface adapter printed circuit assembly (ria pca) boards. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. An ria pca is assigned to each system arm and performs numerous functions related to the function, control and hardware fault monitoring for its assigned arm. The system was repaired by replacing the affected psm and ria pca. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The 20008, system error code appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20105 appears when an auxiliary voltage error is detected during self-test. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2404547 | 2383369 | 2280225 | 2955842201100407 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-22 | 2011-12-20 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-11-22 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-11-18 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The instrument and tip cover accessory were removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The hospital also reported that during visual inspection of the tip cover accessory, a small hole was observed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed two holes approximately 0. 015 in diameter were found on the tip cover wall. A slight burn mark was found on the edge of the hole wall. No other damage found. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2404571 | 2383384 | 2280217 | 2955842201100410 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-11-21 | 2011-12-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-11-21 | No | Yes | Initial | 2011
400180 | 2011-11-11 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed smoke on the video image and then blood coming from the patient's vessel. The surgeon placed a temporary clamp on the vessel and a vascular surgeon was called in to permanently repair the vessel using a vascular clamp. The procedure was converted to traditional open surgical techniques for the repair. Upon examination of the monopolar curved scissors instrument with the tip cover accessory installed the surgical staff observed a small hole in the tip cover accessory. No adverse patient outcome was reported. | The tip cover accessory was not returned for evaluation. Therefore the root cause of the customer reported failure mode cannot be determined. The monopolar curved scissors instrument was returned for evaluation, however, engineering was unable to find any evidence of arching, char marks or other related physical damage. The instrument was found to be fully functional with no trouble found. The surgeon confirmed that the injury to the patient's vessel was a thermal injury. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2402635 | 2381732 | 2278588 | 2955842201100403 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-11-16 | 2011-12-19 | Yes | Yes | Health Professional | 400180 | No | No | 2011-11-16 | No | Yes | Initial | 2011
IS2000 A51P8 | 2011-11-29 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the site experienced multiple occurrences of recoverable system error code 23008. The site disabled the affected patient side manipulator (psm) arm and contacted isi technical support. The technical support engineer attempted to troubleshoot with the customer, however, the surgeon decided to convert the procedure to traditional open surgical techniques. The planned procedure was completed and no patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23008 was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a non-recoverable safe state. The psm was returned to isi for evaluation and engineering confirmed the customer reported problem and found the psm faulted immediately. The axis 7 motor encoder and potentiometers were replaced. As of december 20, 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2402636 | 2381733 | 2278589 | 2955842201100405 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-29 | 2011-12-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-11-29 | Device Not Retu | No | Reuse | 2011
42020505 | 2011-10-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the fenestrated bipolar forceps instrument broke and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one grip to be broken off. The broken piece was not returned with the instrument. The grip broke approximately. 090 inches above the top surface of the yaw pulley. The fractured surface lines up closely to the bottom most tooth of the grip and the grip is severely bent towards the other grip below the fractured surface. Engineering concluded the damage was likely due to mishandling and or misuse. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2402637 | 2381734 | 2278590 | 2955842201100404 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-22 | 2011-12-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2011-11-22 | Device Not Retu | No | Reuse | 2011
| 2011-12-07 | Other | Other | (b)(6) scheduled for a robotic abdominal perineal resection on (b)(6), 2011 was started at 0933, the robotic portion was started approx at 1010, the intuitive instruments were inserted into the robotic arms, the procedure was started, dr. (b)(6) started using the intuitive permanent cautery hook, and the side of the instrument broke off into the pt. It is presumed that this piece of instrument is left in the pt. The piece of instrument is not able to be xrayed. Dr. (b)(6) assessed the pt with the davinci camera internally, she instilled 600ml normal saline thru a blake drain into the pts pelvis, the pt was turned from right to left, and into reverse trendelenburg to see if the piece would leave the pt thru irrigation fluid. All lap sponges, and raytexs were felt for a piece of equipment. The specimen - rectum and sigmoid were palpated x2 and fluid from rectum and sigmoid were checked for any instrument piece. The intuitive rep, jt cullison was present throughout the case. Dr. (b)(6) came into the room, he was notified, suggested a cat scan of the pt, dr. (b)(6) notified of a cat scan. Pt with no apparent injury taken to pacuvia bed. (b)(6) -charge nurse, (b)(6), supervisor, and (b)(6) notified. |  | INTUITIVE CAUTERY HOOK | INTUITIVE CAUTERY HOOK | 2411740 | 2388722 | 2285558 | MW5023552 | 1 | NAY | Voluntary |  | Other | Initial | 2011-12-19 | 2011-12-19 | No | Yes | Health Professional |  | No | Yes |  |  |  |  | 2011
42030901 | 2011-11-16 | Malfunction |  | It was reported that while the surgeon was suturing the vaginal cuff during a da vinci si hysterectomy procedure the cables on the mega suture cut needle driver instrument snapped and fragments fell into the patient's omentum. All of the fragments were retrieved laparoscopically and nothing was left in the patient. The planned surgical procedure was completed using a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable broken at the distal idlers. The broken cable segment sticks out at the instrument wrist. The idler pulley spins freely and does not exhibit damage. The other cables at the wrist are not damaged. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2399940 | 2379339 | 2276195 | 2955842201100400 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-11-16 | 2011-12-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2011-11-16 | Device Not Retu | No | Reuse | 2011
400180 | 2011-11-18 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory which was installed on the mcs instrument. As a result, the patient's uterus was burnt. Since the patient's uterus was being removed as part of the planned surgical procedure, no patient harm was reported. The hospital indicated that the affected tip cover was lost, therefore, will not be returning for evaluation. | Since the tip cover accessory will not be returning to for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2399977 | 2379351 | 2276207 | 2955842201100401 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-11-18 | 2011-12-16 | Yes | Yes | Health Professional | 400180 | No | No | 2011-11-18 | No | Yes | Unkown | 2011
42017909 | 2011-11-18 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si salpingectomy procedure while using the 5mm needle driver instrument, the instrument broke off. The site did not observe pieces of the instrument within the patient. The procedure was completed with no patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip is broken off at the base and broken piece was not returned. The grip fractured where the arm meets the hub. There is a 90 degree angle near the fracture surface, which is a possible stress concentration area. Other grip is not damaged. No other damage found. | 5MM NEEDLE DRIVER INSTRUMENT | Reuse | 2400908 | 2380107 | 2276963 | 2955842201100402 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-11-18 | 2011-12-16 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2012-01-05 | Device Not Retu | No |  | 2011
40018010 | 2011-11-03 | Malfunction |  | Event Description: --/-- On (b)(4), 2011, medwatch uf/importer report (b)(4) was received from the food and drug administration with the following information: patient underwent robotic radical prostatectomy and bilateral pelvic lymph node dissection via da vinci robotic surgery. Tip cover accessory had developed a hole earlier and was replaced with new tip cover. This cover also developed a hole allowing the bovie to arc into the patient's body and required replacement. No apparent patient harm. What was the original intended procedure? robotic surgery. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working) device usage problem: device malfunction - that is, the device did not do what it was supposed to do. The hospital had initially reported this product problem to intuitive surgical on (b)(6), 2011. | Manufacturer Narrative: --/-- An unfocused photo of the tip cover accessory was provided to intuitive surgical on (b)(4) 2011. Engineering evaluation of the photo observed possible damage near the silicone midpoint. --/-- Additional Manufacturer Narrative: --/-- Despite several requests, the tip cover accessory has not be returned for evaluation. As a result, the root cause of the reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the tip cover accessory is returned or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | Unkown | 2385425 | 2364781 | 2261617 | 2955842201100396 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | FollowupInitial | 2011-11-09 | 2011-12-09 | No | Yes | Health Professional | 40018010 | No | No | 2011-12-12 | Device Not Retu | Yes |  | 2011
40018010 | 2011-11-03 | Malfunction |  | Event Description: --/-- On (b)(4), 2011, medwatch uf/importer report (b)(4) was received from the food and drug administration with the following information: da vinci robotic surgery tip cover accessory developed a hole in the tip cover allowing the arc from the bovie to escape into the patient's body. No apparent injury. Defective tip removed and replaced with new one. What was the original intended procedure? robotic radical prostatectomy and bilateral pelvic lymph node dissection device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working) device usage problem: device malfunction - that is, the device did not do what it was supposed to do the hospital had initially reported this product problem to intuitive surgical on (b)(6), 2011. | Manufacturer Narrative: --/-- Despite several requests, the tip cover accessory has not be returned for evaluation. As a result, the root cause of the reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the tip cover accessory is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- A photo of the tip cover accessory was provided to intuitive surgical on (b)(4) 2011. Engineering evaluation of the photo observed a hole burnt through the silicone, below the distal tip. --/-- Additional Manufacturer Narrative: --/-- | 8 MM MCS TIP COVER ACCESSORY | Unkown | 2385437 | 2364789 | 2261617 | 2955842201100397 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | FollowupFollowup | 2011-11-09 | 2011-12-09 | No | Yes | Health Professional | 40018010 | No | No | 2011-12-12 | Device Not Retu | Yes |  | 2011
42018304 | 2011-11-11 | Malfunction |  | It was reported that during a da vinci si low anterior resection procedure, the permanent cautery hook instrument broke at the tip and 2 pieces fell into the patient. The instrument fragments were retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one distal clevis ear, the clevis ear on the conductor wire side, to be completely broken off at its base. The broken piece is approximately. 285 by. 220 inches in size and was not returned with the instrument. Additionally, the instrument's conductor cap was found to be missing and the yaw pulley boss feature is bent. No other damage was found. | PERMANENT CAUTERY HOOKINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2385900 | 2365160 | 2262016 | 2955842201100398 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-11-11 | 2011-12-09 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2011-11-11 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-11-04 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the surgeon was cutting around the prostate tissue to perform nerve sparing, one of the tips of the monopolar curved scissors instrument broke. The instrument fragment was removed from the patient and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found both scissor blades to be intact and no components at the distal tip appear to be missing. The tube extension is dislodged from the main tube and the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2385901 | 2365161 | 2262017 | 2955842201100399 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-11-11 | 2011-12-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-11-11 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-11-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced non-recoverable system error code 23008. With the assistance of a field service engineer (fse),the surgical team exercised an axis on the right master tool manipulator (mtmr) and performed an emergency power off of the system several times, however, the system error continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineering concluded that system error code 23008 was associated with the right master tool manipulator, as review of the system log found multiple instances of system error code 23008. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The mtm was returned to isi for failure analysis investigation. Engineering evaluation confirmed the customer reported failure mode. Sensor testing of the mtm revealed that the pot to encoder reading was out of specification, thus generating the system error code experienced by the customer. As of (b)(6) 2011, the patient was reported to be doing well and was released from the hospital (b)(6) after the surgical procedure. No post - operative complications have been reported. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2382992 | 2362411 | 2259267 | 2955842201100395 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-11-07 | 2011-12-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-11-07 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-11-08 | Other | Other | It was reported that during a da vinci si procedure a surgical staff member inadvertently kicked the video cable, which resulted in non-recoverable system error code 54. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was due to a loose video cable connection on the surgeon console. It was determined that the cable became loose when it was inadvertently kicked by the surgical staff member. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 54 appears when the da vinci safety system determines that the strength of the signal received at the end of the blue fiber has fallen below the low warning limit. Upon determining this condition, the system records the fault and transitions to a safe state. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2380591 | 2359972 | 2256827 | 2955842201100389 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-08 | 2011-12-06 | Yes | Yes | Health Professional | IS3000 A60P6 | No | Yes | 2011-11-08 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-11-03 | Injury | Other | It was reported that during the beginning of a da vinci si prostatectomy and lymphadenectomy procedure, the surgeon grasped the patient's colon with the maryland bipolar forceps instrument and energy began to emit without the foot pedal having been activated. The surgeon immediately removed the maryland bipolar forceps instrument and found the electrical surgical unit (esu) setting was placed on auto. The esu setting was changed and the planned surgical procedure was completed. The patient's colon was repaired and no further adverse event was reported. | The investigation concluded that the patient injury was related to a setup issue with the electrical surgical unit (esu) which is not supplied by intuitive surgical. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2380916 | 2360194 | 2257049 | 2955842201100390 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-11-07 | 2011-12-06 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-11-07 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-11-09 | Other | Other | It was reported that during a da vinci si pediatric colectomy procedure the surgeon experienced a yellow tint in the right eye of the high resolution stereo viewer (hrsv). Prior to contacting isi technical support to assist in troubleshooting the issue, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected camera cable. The da vinci si surgical system user manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have back up equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of december 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2380915 | 2360195 | 2257050 | 2955842201100393 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-09 | 2011-12-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-11-09 | Device Not Retu | No | Reuse | 2011
IS2000 | 2011-11-09 | Other | Other | It was reported that while preparing to start da vinci s hysterectomy procedure, the site experienced a non-recoverable system error 23034. With the assistance of an isi technical support engineer the site restarted the system and exercised the clutch button the endoscopic camera manipulator (ecm) arm, however upon restarting the system the 23034 system error recurred. The patient was under anesthesia and the port incisions were made when the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23034 was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23034 appears when the da vinci s safety system determines that after a specified amount of time, a valid event has not been seen for one of the remote compute engine switches. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2380923 | 2360199 | 2257054 | 2955842201100392 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-11-09 | 2011-12-06 | Yes | Yes | Health Professional | IS2000 | Device Returned To Manufacturer | No | 2011-11-09 | Device Not Retu | No | Reuse | 2011
400180 | 2011-11-09 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the surgeon observed a spark coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2380921 | 2360207 | 2257063 | 2955842201100391 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-09 | 2011-12-06 | No | Yes | Health Professional | 400180 | No | Yes | 2011-11-09 | No | Yes | Unkown | 2011
42034403 | 2011-10-27 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure, burning and charring was observed at the distal tip of the curved bipolar dissector instrument. The instrument was removed and replaced with another instrument to successfully complete the procedure. It was noted that the site was using a high esu setting above recommended levels. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | CURVED BIPOLAR DISSECTORINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2381306 | 2360532 | 2257388 | 2955842201100394 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2011-11-07 | 2011-12-06 | No | Yes | Health Professional | 42034403 | No | Yes | 2011-11-07 | No | No | Unkown | 2011
420142 | 2011-12-01 | Malfunction |  | Surgeon stated that two small intraoral lesions resulted from using this 5mm robotic monopolar cautery instrument. Cautery tip was on the instrument. Company called to inquire if additional cover is needed on the instrument. The answer was "no. ". |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL MO | 2382921 | 2362322 | 2259178 | 2362322 | 1 | NAY | User Facility |  |  | Initial | 2011-12-05 | 2011-12-05 | No | Yes | Physician | 420142 | No | No Answer Provided |  |  |  |  | 2011
42000306 | 2011-10-10 | Malfunction |  | It was reported that during a da vinci si coronary artery bypass graft (cabg) procedure, the tip of the small clip applier instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument has one grip broken off near its midpoint and a. 080 x. 045 piece of the grip is missing. The intact grip is bent slightly near the midpoint. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage was found. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2377345 | 2356577 | 2253433 | 2955842201100386 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-11-03 | 2011-12-02 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2011-11-03 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-11-02 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced system error code 600 while changing surgeon's at the surgeon side console (ssc). With the assistance of an isi technical support engineer, the site applied emergency power off (epo), cycled all switches on several subcomponents, however, the error persisted. At this time, the surgeon made the decision to convert the planned surgical procedure to a traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site was associated with a uninterruptible power source (ups). The ups is a subsystem that can sustain power to the system electronics for a finite period of time after the loss of ac line power. The system was repaired by replacing the affected ups. System error codes 600 appears when the davinci safety system determines that the ready switch is not operating correctly. Upon determining these conditions, the safety system put davinci in a recoverable safe state. As of (b)(6), 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2377521 | 2356702 | 2253559 | 2955842201100388 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-11-02 | 2011-12-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2011-11-02 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-11-01 | Other | Other | It was reported that during a da vinci si hysterectomy procedure there was no video displayed on the right side monitor of the surgeon console. The issue was resolved by replacing the camera cable, however, the customer then experienced no video on the left side monitor. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the right side vision issue was associated with the camera cable and the left side vision issue was associated with the double camera controller (doco). The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The dr refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected camera cable and doco. The camera cable and doco were returned to intuitive surgical for failure analysis investigation. During internal evaluation, engineering was able to replicate the reported failure modes. Open pins were found on the camera cable and engineering determined that the left ccu had malfunctioned on the doco. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2375466 | 2355182 | 2252038 | 2955842201100385 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-01 | 2011-12-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-11-01 | Device Not Retu | No | Reuse | 2011
| 2011-10-20 | Injury | OtherHospitalizationLife ThreateningDisabilityRequired Intervention | Event Description: --/-- Rptr was having a simple robotic procedure. Dr inserted trocar and it punctured the mesenteric wall, posterior and anterior, perforated straight through the aorta and small intestines. She lost 3500 ml blood, was placed in icu for 4 days for this supposed 1 day procedure. She woke up on a ventilator, had a ng tube and was hospitalized 6 days. She had a laparotomy for the extreme hemorrhage with 17 pints of blood transfused. All drs' reports indicate a trocar injury. The physician didn't know she had punctured the aorta. Rptr now has heart problems and kidney problems. Procedure performed at winthrop hosp in (b)(6). --/-- Additional Event Description: --/-- Add'l info received from reporter (b)(6) 2011: i went for a robotic/laparoscopic hysterectomy. During the procedure, my aorta was punctured through and through, small intestine punctured (required dissection, puncture to mesenteric and peritoneal lacerations. ) i was transfused with 17 pints of blood, and was in intensive care unit hospitalized for 6 days - this was a life threatening event. | Manufacturer Narrative: --/-- | DAVINCI ROBOTICS |  | 2378448 | 2356629 | 2253486 | MW5023228 | 1 | GCJ | Voluntary |  | Patient | InitialFollowup | 2011-12-01 | 2011-12-01 | Yes | No | Health Professional |  | No | No |  |  |  |  | 2011
42027202 | 2011-11-03 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the bottom part of the harmonic ace insert accessory fell inside the patient. The accessory fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragment was retrieved from the patient and the procedure was successfully completed without incident. | HARMONIC ACE INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2372061 | 2351927 | 2248783 | 2955842201100383 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-03 | 2011-11-29 | No | Yes | Health Professional | 42027202 | No | Yes | 2011-11-03 | No | Yes | Unkown | 2011
40027202 | 2011-10-12 | Malfunction |  | Event Description: --/-- On (b)(4), 2011, medwatch uf/importer report (b)(4) was received: during a laparoscopic myomectomy, surgeon noted the robotic instrument was not working. The system was rebooted and reset. The instrument worked intermittently. However, the jaw fractured while grasping sheath. Tissue fragment retrieved, perfect insert match. Post-op x-rays taken per protocol - clear. No patient harm. | Manufacturer Narrative: --/-- The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Multiple attempts to obtain additional information about this event from the site have been made, however as of the time of this report no response has been received. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragment was retrieved from the patient and the procedure was successfully completed without incident. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2012, the risk manager at (b)(6), reported that the surgeon confirmed the harmonic ace insert accessory broke and fell into the patient. The accessory fragment was immediately retrieved and an x-ray of the surgical area was performed to confirm that there were no foreign bodies retained in the patient. No patient complications occurred and the patient was discharged from the hospital as planned. Additionally the patient has not returned due to any post-operative complications. --/-- Additional Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a fractured jaw. The insert accessory was found with the curved blade broken at approximately 0. 4 from the tip. Slight burn mark was found near the base where potential contact with the clamp arm pin was located. No damage was found on the clamp arm. The associated instrument was also returned and evaluated. The instrument was placed on a da vinci si surgical system for testing. The instrument was recognized by the system. The instrument moved intuitively with full range of motion in all directions. | HARMONIC ACE INSERT ACCESSORY | Unkown | 2372062 | 2351928 | 2248784 | 2955842201100382 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | FollowupFollowup | 2012-01-27 | 2011-11-29 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-03-05 | Device Not Retu | Yes |  | 2011
42017910 | 2011-11-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, material from the monopolar curved scissors instrument was shredding off the main tube and falling into the patient. The instrument fragments were retrieved and irrigated out of the patient. The planned surgical procedure was completed and no patient injury, harm or adverse outcome was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragments were retrieved from the patient and the procedure was successfully completed without incident. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2372030 | 2351929 | 2248785 | 2955842201100381 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-11-01 | 2011-11-29 | No | Yes | Health Professional | 42017910 | No | No | 2011-11-01 | No | No | Unkown | 2011
400180 | 2011-10-31 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, while dissecting tissue, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The arc went to the patient's uterus, which was removed as part of the planned procedure. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the additional information is received, a follow-up mdr will be submitted. On (b)(6) 2011, the site reported that the patient had been discharged as planned and no adverse effects from the incident were reported. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2372067 | 2351944 | 2248800 | 2955842201100384 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-11-03 | 2011-11-29 | Yes | Yes | Health Professional | 400180 | No | Yes | 2011-11-03 | No | Yes | Unkown | 2011
400180 | 2011-11-28 | Malfunction |  | Patient underwent a laparoscopic robotic assisted left partial nephrectomy with intra-operative sonography. Operative length 210 minutes. Tear in da vinci tip cover accessory noted at conclusion of procedure. No known patient harm. Operative report indicates no accidental punctures or lacerations experienced by patient. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2374826 | 2354308 | 2251164 | 2354308 | 1 | NAY | User Facility |  |  | Initial | 2011-11-29 | 2011-11-29 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
MEGA SUTURECUT NEEDLE DRIVER | 2011-11-22 | Malfunction |  | Da vinci mega suture cut needle driver broke again during robotic surgery. |  | DA VINCI MEGA SUTURECUT NEEDLE DRIVER | INSTRUMENT ROBOTIC SURGICAL NE | 2371546 | 2351461 | 2248317 | 2351461 | 1 | NAY | User Facility |  |  | Initial | 2011-11-28 | 2011-11-28 | No | Yes | Physician | MEGA SUTURECUT NEEDLE DRIVER | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-10 | Malfunction |  | Patient underwent a robotic-assisted laparoscopic radical prostatectomy. Procedure length 360+ minutes tip cover accessory reportedly developed hole during procedure and was replaced. No known patient harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371548 | 2351463 | 2248319 | 2351463 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-11 | Malfunction |  | Patient underwent a laparoscopic robotic assisted left partial nephrectomy. Tip cover accessory was reportedly checked during a break in the case - tip found to be splitting near monopolar scissors tips. Operative time about 204 minutes. No patient harm reported. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371550 | 2351484 | 2248340 | 2351484 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-18 | Malfunction |  | Patient undewent a laparoscopic robotic assisted left partial nephrectomy. Silicon tip reportedly was used on patient for 93 minutes. Tip was inspected, removed and replaced. Tip had shown a tear in the tip of the scissor during use on the patient. No reported harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371551 | 2351485 | 2248341 | 2351485 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-14 | Malfunction |  | Patient underwent a robotically assisted laparoscopic radical prostatectomy. Tip cover accessory developed both a hole and split, requiring device removal and replacement. No apparent patient harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371553 | 2351486 | 2248342 | 2351486 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Invalid Data | 400180 |  | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-17 | Malfunction |  | Patient underwent a robotic-assisted laparoscopic right partial nephrectomy. This second (replacement) robotic tip cover accessory reportedly failed after about 1 hour of use and was replaced again after it developed a tear. No patient harm reported. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371556 | 2351489 | 2248345 | 2351489 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-17 | Malfunction |  | Patient underwent robotic-assisted laparoscopic right partial nephrectomy. Tip cover accessory split at side after about 1 hour of use. Tip removed and replaced. No known patient harm. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371558 | 2351492 | 2248348 | 2351492 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-11 | Malfunction |  | Patient underwent a robotic assisted prostatectomy. Operative time 140 minutes. Upon inspection after use of robotic instrument, the silicon tip appeared to have a split near the working end of the tip that was not there at the beginning of use. The instrument and tip were used for approximately 90 minutes. No patient harm reported. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2371559 | 2351494 | 2248350 | 2351494 | 1 | NAY | User Facility |  |  | Initial | 2011-11-25 | 2011-11-25 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-10-24 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the site smelled something burning and observed the monopolar curved scissors tip cover accessory to be burned. The site replaced the tip cover accessory to successfully complete the procedure with no harm to the patient. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed arcing damage near the location of the distal clevis on the instrument. The damage area is approximately 0. 05 in circumference. No other damage found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2367977 | 2348330 | 2245186 | 2955842201100378 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-10-24 | 2011-11-23 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-10-24 | Device Not Retu | Yes | Initial | 2011
IS1200 A43P9 | 2011-10-25 | Injury |  | It was reported that during a da vinci hysterectomy procedure the surgeon experienced insufflation loss at the camera trocar. The site replaced the camera trocar; however, the insufflation loss persisted. At this time the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No system malfunction occurred. | Based on the investigation performed by an isi field service engineer, it was determined that the loss of insufflation was due to the site vaginally inserting an occulator device that is a non isi supplied product. While onsite, the fse tested the da vinci surgical system and determined that the da vinci surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the converted procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT ACCESSOR | 2367952 | 2348345 | 2245201 | 2955842201100379 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-10-25 | 2011-11-23 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2011-10-05 | No | Yes | Unkown | 2011
IS3000 A60P6 | 2011-10-01 | Other | Other | It was reported that during a da vinci si pyeloplasty procedure, the surgeon observed energy coming from both the monopolar and bipolar instruments when the monopolar cautery was intended to be activated. The patient experienced a small bowel injury which was successfully resected. The patient was hospitalized for 2. 5 weeks before being discharged home. No further information has been provided. | The investigation being conducted by the isi field service engineer (fse) is ongoing; however, the results of this investigation will be provided in a follow-up mdr upon completion. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2367955 | 2348350 | 2245206 | 2955842201100380 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-10-27 | 2011-11-23 | Yes | No | Health Professional | IS3000 A60P6 | Yes | No | 2011-10-27 | Device Not Retu | No | Reuse | 2011
400180 | 2011-11-22 | Malfunction |  | Patient underwent a laparoscopic robotic assisted right partial nephrectomy using robotic surgery. No complications noted from this device failure. No identified patient harm. Surgical time 3 hours. Tip 1 (previous medsun report of this date) failed at 65 minutes. Tip was replaced with tip 2. Tip 2 failed at 80 minutes. Silicone noted to be tearing at wrist end of cover. No arcing noted. Tip removed and replaced immediately. Robotic scissor 420179 version 10. Valley lab force fx esu settings of 35 cut / 35 coag. Davinci robot arms ccf#102404. Davinci robot #sh0199 (tower#). |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2369499 | 2349314 | 2246170 | 2349314 | 1 | NAY | User Facility |  |  | Initial | 2011-11-23 | 2011-11-23 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-22 | Malfunction |  | Patient underwent a laparoscopic robotic assisted right partial nephrectomy using robotic surgery. No complications noted from this device failure. No identified patient harm. Surgical time 3 hours. Two tip failures tip 1 failed at 65 minutes (this report) and was replaced. Tip cover noted to be splitting while watching procedure on video screen. No electrosurgical arcing noted. Used with valleylab forcefx esu with settings 35 cut/35 coag. Robotic arms tip used for approximately one hour and 5 minutes. Scissor instrument version 10. Tip 2 (separate medsun report) also failed at 80 minutes. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2369500 | 2349315 | 2246171 | 2349315 | 1 | NAY | User Facility |  |  | Initial | 2011-11-23 | 2011-11-23 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
42018306 | 2011-10-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed a spark and smoke from the joint of the permanent cautery hook instrument after approximately 30 minutes of use. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, injury or adverse outcome was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and or if additional information is received. | PERMANENT CAUTERY HOOKINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2365382 | 2345778 | 2242634 | 2955842201100377 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-28 | 2011-11-22 | No | Yes | Health Professional | 42018306 | No | Yes | 2011-10-28 | No | No | Unkown | 2011
42022703 | 2011-09-13 | Malfunction |  | On (b)(4), 2011, a copy of uf medwatch report (b)(4) was received from the food and drug administration. The report contained the following information: event desc: during a robotic-assisted laparotomy, a screw broke off the da vinci s pk (plasmakinetic) dissecting forceps and fell into the abdominal cavity. (the device continued to work after this event. ) the screw was located and retrieved from the abdominal cavity. No harm to patient. Vendor was in the or at the time of the malfunction. What was the original intended procedure? robotic removal of tvt tape; removal of sacrocolpopexy mesh; and cystoscopy. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The instrument has not been received for evaluation, therefore, the root cause of the customer reported failure cannot be determined. Multiple attempts have been made to obtain further information, including who the vendor referenced as being in the or at the time of the event was, however as of the date of this report no response has been provided by the hospital. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2365383 | 2345779 | 2242635 | 2955842201100376 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-10-25 | 2011-11-22 | No | Yes | Health Professional | 42022703 | No | No | 2011-10-25 | No | No | Unkown | 2011
420194 | 2011-11-15 | Malfunction |  | Surgical staff reports while doing surgery on patient with davinci robot the cable in the mega needle driver broke making the device useless. Needle driver switched out with alternate one and case proceeded with no further issues. No injury to patient. Device is available for evaluation purposes. |  | DA VINCI MEGA NEEDLE DRIVER | INSTRUMENT ROBOTIC SURGICAL NE | 2367313 | 2347799 | 2244655 | 2347799 | 1 | NAY | User Facility |  |  | Initial | 2011-11-22 | 2011-11-22 | No | Yes | Physician | 420194 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-15 | Malfunction |  | Patient undergoing a robotic-assisted ureteral reimplant with boari flap surgery. Da vinci tip cover accessory had been used for 2+ hours in this procedure. Tear of tip cover discovered upon inspection. |  | MICROTEK DA VINCI | INSTRUMENT ROBOTIC SURGICAL TI | 2368260 | 2348525 | 2245381 | 2348525 | 1 | NAY | User Facility |  |  | Initial | 2011-11-22 | 2011-11-22 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
42017207 | 2011-09-15 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the jaws of the maryland bipolar forceps instrument became tangled with the prograsp forceps instrument. A jaw of the maryland bipolar forceps instrument broke as a result and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361153 | 2342238 | 2239094 | 2955842201100363 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-10-18 | 2011-11-18 | No | Yes | Health Professional | 42017207 | No | No | 2011-10-19 | No | No | Unkown | 2011
40018010 | 2011-10-21 | Injury | Life Threatening | It was reported that while using the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed to dissect the aortic lymph nodes near the iliac artery during a da vinci si hysterectomy procedure, a small hole was burnt into the patient's iliac artery. Another surgeon was called in, who repaired the injury using a prolene suture. The instrument and tip cover accessory were removed and replaced and the planned surgical procedure was completed. No additional patient harm, adverse outcome or injury was reported. The hospital also reported that during visual inspection of the tip cover accessory, a small hole was observed. | The tip cover accessory was not returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2361154 | 2342239 | 2239095 | 2955842201100361 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2011-10-21 | 2011-11-18 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-10-21 | No | Yes | Unkown | 2011
42017207 | 2011-10-17 | Malfunction |  | It was reported that approximately one hour into a da vinci s hysterectomy procedure, while using the maryland bipolar forceps instrument, a piece broke off and fell into the patient. The instrument fragment was immediately retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient injury, harm or adverse event was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361155 | 2342240 | 2239096 | 2955842201100362 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-19 | 2011-11-18 | No | Yes | Health Professional | 42017207 | No | Yes | 2011-10-19 | No | No | Unkown | 2011
42018406 | 2011-10-25 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the insulation at the tip of the permanent cautery spatula instrument broke and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found a the ceramic sleeve to be broken off and completely missing. One yaw pulley cable was found to be derailed at the pulley. No other damage was found. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361161 | 2342242 | 2239098 | 2955842201100364 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-28 | 2011-11-18 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-10-28 | Device Not Retu | No | Reuse | 2011
42018406 | 2011-10-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the permanent cautery spatula instrument broke and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found a piece of the ceramic sleeve to be broken off. The distal end of the sleeve has a. 095 inch by. 095 inch piece missing and the sleeve exhibits scratch marks and a cracked section where it meets the yaw pulley. Engineering concluded that the damage was likely due to mishandling and or misuse. No other damage was found. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361163 | 2342243 | 2239098 | 2955842201100365 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-10-25 | 2011-11-18 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-10-25 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-10-12 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure while using the monopolar curved scissors instrument, the surgical staff smelled something burning and saw smoke in the surgical field. The monopolar curved scissors instrument was removed and found to have a burn mark on the shaft. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the distal end of the main tube exhibits a hole burned through approximately. 530 inches above the interface with the tube extension. The instrument was disassembled to find char marks on the hypotubes and localized melting on the tube extension in the same location as the main tube hole, indicating an arcing event occurred. The main tube has a hairline crack in the axial direction located directly below the hole that appears to have started on the outer tube surface. Engineering concluded that the crack contributed to the arcing damage. No other damage was found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361164 | 2342244 | 2239120 | 2955842201100369 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-10-24 | 2011-11-18 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-10-24 | Device Not Retu | No | Initial | 2011
42022703 | 2011-09-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the pk dissecting forceps instrument broke and cable fragments fell into the patient. The surgeon retrieved as many cable fragments as they could find. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Engineering also observed that both conductor wires are broken near the distal clevis wire holes. The wires appear to have been cut with a sharp object. The wires stick out at the wrist and electrical continuity fails. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361165 | 2342245 | 2239121 | 2955842201100368 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-20 | 2011-11-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2011-10-20 | Device Not Retu | No | Reuse | 2011
42020504 | 2011-10-05 | Malfunction |  | It was reported that near the end of a da vinci s hysterectomy procedure, a small black piece of the fenestrated bipolar forceps instrument broke off onto the patient's bowel. The fragment was immediately retrieved. The planned surgical procedure was completed with an instrument of the same type and no patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the yaw pulley to be missing a. 15 inch by. 08 inch piece opposite the conductor wire, allowing the grip to move within the pulley and have a larger range of motion in yaw. The conductor wire was found to be broken at the yaw pulley exit and detached from the grip connection. Electrical continuity failed. No other damage was found. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361186 | 2342246 | 2239122 | 2955842201100367 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-05 | 2011-11-18 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2011
42022703 | 2011-09-02 | Malfunction |  | On (b)(6) 2011, it was initially reported that during a da vinci si hysterectomy procedure the pk dissecting forceps instrument stopped cauterizing. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The instrument was returned and during failure analysis evaluation, engineering observe a piece missing from the instrument. The hospital was informed of the missing piece and on (b)(6), 2011, the hospital indicated that a video recording of the surgical procedure was reviewed and it was noted that a small piece of the instrument broke off and fell into the patient's uterus while the surgeon was performing retraction and coagulation. The broken piece appeared to have been suctioned out. The hospital also mentioned that no postsurgical complications have been experienced by the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire was broken at the yaw pulley exit. On the opposite side of the conductor wire, a small piece of the yaw pulley had broken off, likely causing the conductor wire to come loose and break. The size of the broken piece is approximately. 08 x. 17. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361187 | 2342247 | 2239123 | 2955842201100366 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-25 | 2011-11-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2011-10-25 | Device Not Retu | No | Reuse | 2011
42009308 | 2011-10-26 | Malfunction | Other | It was reported that during a da vinci s prostatectomy procedure, when the surgeon grasped and retracted the patient's bladder using the prograsp forceps instrument, the bladder slipped out of the instrument grips, causing a tear to the orifice of the patient's bladder. Removal and inspection of the prograsp forceps instrument by the surgical team found that its wire was broken. The planned surgical procedure was completed with an instrument of the same type. | The instrument was returned and evaluated. Engineering found that the drive cables at the wrist were not broken, however, one grip cable exhibited damage. One grip open cable derailed from the force amp pulley, resulting in non-intuitive motion in yaw. The same cable was also fraying at the force amp pulley and the frayed strands stick out slightly at the force amp pulley. Engineering concluded that derailment most likely contributed to the fraying of the cable. No other damage found. Multiple requests for additional information have been made, however, as of the date of this report the site has not responded. A follow-up mdr will be submitted if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361891 | 2342749 | 2239605 | 2955842201100372 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-10-26 | 2011-11-18 | Yes | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2011-10-26 | Device Not Retu | No | Reuse | 2011
42018406 | 2011-10-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure, after approximately 1 hour of use, the tip on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed the ceramic sleeve on the instrument to be cracked and broken at the base of the spatula. The spatula did not exhibit any burn marks. The broke pieces were not returned with the instrument. Damage to the clevis was not observed. No other damage was found. Multiple requests for additional information have been made, however, as of the date of this report the site has not responded. A follow-up mdr will be submitted if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2361893 | 2342751 | 2239607 | 2955842201100371 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-21 | 2011-11-18 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-10-21 | Device Not Retu | No | Reuse | 2011
42000606 | 2011-10-21 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy with hysterectomy procedure, a fragment from the large needle driver instrument fell into the patient and was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument was returned with the carbide insert broken off of one grip. The brazing surface on the grip and on the detached insert have incomplete surface area coverage of filler material. The grip has a couple of scratch marks at the tip. No additional damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2361979 | 2342844 | 2239700 | 2955842201100373 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-24 | 2011-11-18 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2011-10-24 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-10-24 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure the surgical staff observed the system message preventative maintenance message. In an attempt to clear the message the site applied emergency power off multiple times and reseated cables from the back end of the system; however, the message persisted. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the system message experienced by the site was associated with the system video processor printed circuit assembly (svp pca) board. The s5vp pca is the video processor for the system which creates the display for the high resolution stereo viewer. The system was repaired by replacing the affected svp pca board the svp pca was returned to isi for evaluation, however, engineering was unable to replicate the customer reported failure mode. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2361990 | 2342869 | 2239725 | 2955842201100374 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-24 | 2011-11-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-10-24 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-10-21 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed an arc from the monopolar curved scissors instrument. It was also discovered that the patient had sustained an injury to their artery which was successfully repaired with the assistance of a cardiovascular surgeon. The original procedure was reported as being completed. No further informatin has been provided. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2362001 | 2342877 | 2239733 | 2955842201100375 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-21 | 2011-11-18 | Yes | No | Health Professional | 42017910 | No | Yes | 2011-10-21 | No | No | Unkown | 2011
951021 | 2011-11-01 | No Answer Provided |  | Procedure: prostatectomy, radical, with robotic assistance. During placement of the da vinci instrumentation, one of the bifurcated light guide cables became loose. This was discovered, but not before six small burns occurred in the first layer of the blue drapes. There was no smoke or flame observed. There were cloth towels placed beneath the drape and a towel prevented the patient's skin (face/chest) from being burned. There was no injury to the patient or staff. The md was notified and the procedure continued. Device thought to be ignition source: bifurcated light guide cable. Device thought to be the fuel source: da vinci drape. The patient was intubated (endotracheal tube) and receiving oxygen via a ventilator. It is believed by surgery staff that this occurrence resulted from human error. The light cable connections were not double checked to assure that they were secure. One person set up the procedure and another person came in and did the case. Surgery leadership were asked to review information regarding the "preventing surgical fires" initiative on the fda website. |  | DA VINCI | ACCESSORY ROBOTIC SURGICAL BIF | 2362397 | 2343127 | 2239983 | 2343127 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-11-10 | 2011-11-17 | No | No | Invalid Data | 951021 | Yes | No Answer Provided |  |  |  |  | 2011
951021 | 2011-11-01 | No Answer Provided |  | Procedure: prostatectomy, radical, with robotic assistance. During placement of the da vinci instrumentation, one of the bifurcated light guide cables became loose. This was discovered, but not before six small burns occurred in the first layer of the blue drapes. There was no smoke or flame observed. There were cloth towels placed beneath the drape and a towel prevented the patient's skin (face/chest) from being burned. There was no injury to the patient or staff. The md was notified and the procedure continued. Device thought to be ignition source: bifurcated light guide cable. Device thought to be the fuel source: da vinci drape. The patient was intubated (endotracheal tube) and receiving oxygen via a ventilator. It is believed by surgery staff that this occurrence resulted from human error. The light cable connections were not double checked to assure that they were secure. One person set up the procedure and another person came in and did the case. Surgery leadership were asked to review information regarding the "preventing surgical fires" initiative on the fda website. |  | DA VINCI | ACCESSORY ROBOTIC SURGICAL BIF | 2362397 | 2343127 | 2239983 | 2343127 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-11-10 | 2011-11-17 | No | No | Invalid Data | 951021 | Yes | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2011-10-17 | Injury | Required Intervention | It was reported that during a da vinci hysterectomy procedure the surgeon accidentally cut the patient's hypogastric nerve. The procedure was converted to traditional open surgical techniques to complete the planned procedure and repair the patient's hypogastric nerve. No system malfunction occurred. | Based on the information provided by hospital, it was determined that the da vinci surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the patient's injury. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2358145 | 2339742 | 2236618 | 2955842201100360 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-10-17 | 2011-11-16 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2011-10-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-10-10 | Other | Other | It was reported that during a da vinci si prostatectomy procedure,the surgical staff experienced system error code 500. With the assistance of an isi technical support engineer (tse), the surgical staff attempted to resolve the issue by cycling power and swapping the blue fiber cables on the vision sice card (vsc) and surgeon side console (ssc) without success. The system was powered down again, the fiber cables and the console connector were cleaned with compressed air before cycling the breakers to the vsc again; however, system error code 500 re-occurred. The patient was under anesthesia when the surgeon made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the site was associated with a remote arm controller (rac). The rac consists of five printed circuit assembly boards which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The rac was returned for evaluation. Engineering was able to replicate the reported failure and determined that a remote controller module had malfunctioned, thus generating the system error code experienced by the customer. As of november 09, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2350845 | 2332398 | 2229274 | 2955842201100359 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-10 | 2011-11-10 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-10-11 | Device Not Retu | No | Reuse | 2011
400180 | 2011-10-27 | Malfunction |  | Da vinci robotic tip accessory developed damage on the inside of the tip cover. Hole did not go through the silicon full thickness. The cautery tip was inspected at the first opportunity after about 1 hour of use and an inconsistency in the silicon tip was noted. No patient harm, but there was the potential for harm of an unintended cautery burn should the erosion have progressed through the tip. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2358677 | 2340111 | 2236987 | 2340111 | 1 | NAY | User Facility |  |  | Initial | 2011-11-10 | 2011-11-10 | No | Yes | Invalid Data | 400180 |  | No Answer Provided |  |  |  |  | 2011
400180 | 2011-10-25 | Malfunction |  | Tip cover accessory was being used on a robotic instrument. Tip cover found to have a hole in it after 3 hours of use. The hole posed the situation of causing a cauterization zone where it was not intended, which would have the potential to injure the patient. No harm known to have occurred in this event. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2358678 | 2340112 | 2236988 | 2340112 | 1 | NAY | User Facility |  |  | Initial | 2011-11-10 | 2011-11-10 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-07 | No Answer Provided |  | Tip cover accessory covering the monopolar curved scissors (420179) sustained a small hole in the product. No known patient injury. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2358681 | 2340115 | 2236991 | 2340115 | 1 | NAY | User Facility |  |  | Initial | 2011-11-10 | 2011-11-10 | No |  | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-10-25 | Malfunction |  | Hole discovered in silicone tip cover accessory during procedure after 2 hours of use, posing a potential patient harm situation. Tip cover acts as insulation on the metal of the instrument to protect from unwanted cauterization of the patient's tissues. No known patient injury. Hole found at the conclusion of the procedure. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2358682 | 2340116 | 2236992 | 2340116 | 1 | NAY | User Facility |  |  | Initial | 2011-11-10 | 2011-11-10 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
420179 | 2011-10-10 | Malfunction |  | On (b)(6) 2011, medwatch uf/importer report (b)(4) was received stating the following: event desc: while the davinci monopolar curved scissors instrument was being used through a port inside the patient, the distal end of the scissor partially broke away from the handle. The wire inside of the scissor remained intact within the instrument. All pieces were removed from the laparoscopic port. | The initial reporter has been contacted on several occasions to request the return of the instrument and additional information related to this event; however, as of the date of this report, the instrument has not been returned nor has additional information been provided. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2348823 | 2330519 | 2227395 | 2955842201100358 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2011-10-10 | 2011-11-09 | No | Yes | Health Professional | 420179 | No | No | 2011-10-10 | No | No | Unkown | 2011
VER10 | 2011-11-07 | Malfunction |  | Towards the end of case, it was noted that a davinci instrument had separated at a joint, instrument removed from field and retained for examination. (b)(4). |  | DAVINCI SURGICAL SYSTEM | MONOPOLAR CURVED SCISSORS | 2357188 | 2337034 | 2233910 | MW5023057 | 1 | NAY | Voluntary |  | Other | Initial | 2011-11-08 | 2011-11-09 | No | Yes | Health Professional | VER10 | Yes | No |  |  |  |  | 2011
42019408 | 2011-09-26 | Malfunction |  | Event Description: --/-- On (b)(6) 2011, intuitive surgical received a copy of uf medwatch report (b)(4) from (b)(6). The report contained the following information: instrument examined per protocol prior to use. All functions within normal limits. The mega needle holder was inside the cannula, getting ready to pick up needle for first stitch when it was noted that the jaw liner from one side had broken off and fell inside the patient. Instrument removed, broken metal piece retrieved and removed. Instrument taken off field. X-ray performed post-op per protocol. No retained foreign body noted. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the carbide insert broken off one of the grips. Slight adhesive residue was found on the insert mating surface with the grip, indicating the insert was installed prior to use. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | Reuse | 2346313 | 2328052 | 2224919 | 2955842201100357 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | InitialFollowup | 2011-10-12 | 2011-11-08 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-02-27 | Device Not Retu | No |  | 2011
40018010 | 2011-10-07 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the surgeon observed an arc from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument and tip cover accessory were removed and inspection of the tip cover revealed a hole. The procedure continued with a replacement mcs instrument and tip cover accessory, however, the surgeon observed arcing once again. The replacement mcs instrument and tip cover accessory were removed and inspection of the replacement tip cover revealed a hole. The procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. User facility medwatch report (b)(4) was received on (b)(6) 2011 regarding this event. | Manufacturer Narrative: --/-- The site provided a photographic image of the tip cover accessory used during the procedure. Visual analysis revealed multiple holes burned through the silicone indicating multiple arcing events occurred. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory will not been returned for evaluation as the site wants retain them. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The patient was released from the hospital on october 10, 2011 with no post operative complications. The following medwatch report was sent to the fda regarding this event: mfr report 2955842-2011-00356. | 8MM MCS TIP COVER ACCESSORY | Unkown | 2346265 | 2328071 | 2224927 | 2955842201100355 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2011-12-14 | 2011-11-08 | No | Yes | Health Professional | 40018010 | Yes | No | 2011-12-14 | Device Not Retu | Yes |  | 2011
40018010 | 2011-10-07 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the surgeon observed an arc from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument and tip cover accessory were removed and inspection of the tip cover revealed a hole. The procedure continued with a replacement mcs instrument and tip cover accessory, however, the surgeon observed arcing once again. The replacement mcs instrument and tip cover accessory were removed and inspection of the replacement tip cover revealed a hole. The procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. User facility medwatch report (b)(4) was received on (b)(6) 2011 regarding this event. | Manufacturer Narrative: --/-- The tip cover accessory will not been returned for evaluation as the site wants retain them. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The patient was released from the hospital on (b)(6) 2011 with no post operative complications. The following medwatch report was sent to the fda regarding this event: mfr report 2955842-2011-00355. --/-- Additional Manufacturer Narrative: --/-- The site provided a photographic image of the tip cover accessory used during the procedure. Visual analysis found an oblong hole burned through the silicone indicating an arcing event occurred. A tear was also observed at the tip of the tip cover accessory. | 8MM MCS TIP COVER ACCESSORY | Unkown | 2346291 | 2328072 | 2224927 | 2955842201100356 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2011-12-14 | 2011-11-08 | No | Yes | Health Professional | 40018010 | Yes | No | 2011-12-14 | Device Not Retu | Yes |  | 2011
5 METAL TROCAR KIT | 2011-11-07 | Malfunction |  | The #5 metal trocar from si robot kit became bent during the procedure (robotic assisted right lobectomy). The si robot representative was in the room and is aware of the issue. |  | DA VINCI SI | INSTRUMENT ROBOTIC SURGICAL TR | 2357821 | 2339452 | 2236328 | 2339452 | 1 | NAY | User Facility |  |  | Initial | 2011-11-08 | 2011-11-08 | No | Yes | Physician | 5 METAL TROCAR KIT |  | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-03 | Malfunction |  | Patient underwent robotic radical prostatectomy and bilateral pelvic lymph node dissection via da vinci robotic surgery. Tip cover accessory had developed a hole earlier and was replaced with new tip cover. This cover also developed a hole allowing the bovie to arc into the patient's body and required replacement. No apparent patient harm. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2344884 | 2327020 | 2223896 | 2327020 | 1 | NAY | User Facility |  |  | Initial | 2011-11-07 | 2011-11-07 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-11-03 | Malfunction |  | Da vinci robotic surgery tip cover accessory developed a hole in the tip cover allowing the arc from the bovie to escape into the patient's body. No apparent injury. Defective tip removed and replaced with new one. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2344906 | 2327021 | 2223897 | 2327021 | 1 | NAY | User Facility |  |  | Initial | 2011-11-07 | 2011-11-07 | No | Yes | Invalid Data | 400180 |  | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-10-05 | Injury | Life Threatening | It was reported that 30 minutes into a da vinci si endometriosis resection procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. As a result, the patient's bowel was burnt. The injured bowel was repaired and a sigmoidoscopy was performed to ensure that no other bowel damaged occurred. The planned surgical procedure was completed and no additional patient harm, adverse outcome or injury was reported. The hospital also reported that when examining the tip cover accessory, a small hole was observed. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2342747 | 2324918 | 2221794 | 2955842201100345 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-10-05 | 2011-11-04 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-10-05 | No | Yes | Unkown | 2011
400180 | 2011-09-30 | Injury |  | It was reported that while taking down the patient's uterus during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. Additionally, the user facility indicated that no postsurgical complications have been experienced by the patient. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2342725 | 2324938 | 2221815 | 2955842201100346 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-10-05 | 2011-11-04 | No | Yes | Health Professional | 400180 | No | No | 2011-10-05 | No | Yes | Unkown | 2011
IS1200 A43P9 | 2011-10-14 | Other | Other | It was reported that the surgeon had just sat down at the surgeon console to begin a da vinci prostatectomy procedure when system error code 20008 occurred. To clear the error, the surgical staff powered down the system, back drove the left master tool manipulator gimbal, and restarted the system. After restarting the system, the same error code recurred. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code 20008 was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtml gimbal. The system alarm (system generated fault code) functioned as designed. System error code 20008 occurs when the da vinci safety system determines the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the system records the fault and transitions to a safe state. The mtml gimbal was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported failure during internal testing and determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2342727 | 2324939 | 2221814 | 2955842201100347 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-10-14 | 2011-11-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-10-14 | Device Not Retu | No | Reuse | 2011
400156 | 2011-09-29 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the 5 mm monopolar cautery hook tip accessory installed on the 5 mm monopolar cautery instrument broke off and fell into the patient. The broken piece was successfully retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The hospital indicated that the tip accessory will not be returned for evaluation as it was discarded, however, the instrument used in conjunction with the tip accessory was returned and evaluated. Per the customer reported complaint, engineering observed a-sure rectangular tube remaining inside the monopolar adapter at the top of the instrument snake wrist. Microscopic inspection found no evidence of spot weld at the end of the a-sure tube, which may have allowed the tip to dislodge. | 5 MM MONOPOLAR CAUTERY HOOK TIP ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2342757 | 2324946 | 2221822 | 2955842201100348 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-06 | 2011-11-04 | No | Yes | Health Professional | 400156 | No | No | 2011-10-06 | No | Yes | Unkown | 2011
420179 | 2011-10-06 | Injury | Required Intervention | It was reported that during a da vinci si partial nephrectomy procedure, the patient side assistant pushed the monopolar curved scissors instrument through the patient's spleen during insertion. After the partial nephrectomy was completed, the surgical staff decided that the patient's spleen should be removed. The splenectomy was successfully completed using the system. It is believed that an instrument was holding down the patient's spleen and when the instrument was removed, the spleen moved into the path of the new instrument insertion. | No information has been received which suggests that the instrument malfunctioned in a way that may have caused or contributed to the spleen damage experienced by the patient. Intuitive surgical concluded that the patient's surgical complication was unrelated to the performance of the instrument. The da vinci si surgical system user manual specifically states: caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the patient. Caution: removal of instruments during a procedure should be done very carefully and only with the knowledge and full view of the surgeon console operator. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2343252 | 2325446 | 2222322 | 2955842201100349 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-10-06 | 2011-11-04 | Yes | No | Health Professional | 420179 | No | No |  | No | No | Reuse | 2011
IS2000 A51P7 | 2011-10-05 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure the surgeon was unable to take control of a specific patient side manipulator (psm) arm. Prior to contacting isi for technical support troubleshooting assistance, the surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by isi field service engineering was unable to replicate the issue experienced by the customer. Review of the system log found a single event of system error code 22008 which occurred at the time the surgeon was unable to take control of the psm. A 100 - system error code 22008 occurs when an inserted tool, instrument or accessory, is not accepted by the da vinci s surgical system. Based on the information provided by the site and additional onsite troubleshooting, it was determined that the cannula installed was likely not detected by the system. This system error can be resoved with simple troublehooting, however, the surgeon decided to convert the procedure prior to contacting isi technical support. The surgeon and surgical staff were advised that in the future if system errors occur, isi technical support should be contacted for assistance. On (b)(6) 2011, the site reported that the patient was discharged from the hospital on (b)(6) 2011 and was recuperating at home in stable condition and did not experience any post-operative complications. As of (b)(6) 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2343367 | 2325538 | 2222414 | 2955842201100350 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-10-05 | 2011-11-04 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2011-10-05 | Device Not Retu | No | Reuse | 2011
A43P9 IS1200 | 2011-10-07 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site was unable to focus the image through the high resolution stereo viewer (hrsv). Trouble shooting with the assistance of an isi technical support engineer (tse) revealed that the camera head was defective. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. There was no patient harm, adverse outcome or injury reported. | The investigation conducted by the isi field service engineer concluded that the focusing issue experienced by the customer was associated with the system camera head. With the assistance of a fse, the site replaced the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The site performed and successfully completed the procedure the same day. The camera head was evaluated by the original equipment manufacturer (oem). The stage was found to be defective, thus preventing functionality of the camera head focus feature. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2343469 | 2325643 | 2222519 | 2955842201100351 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-10-07 | 2011-11-04 | Yes | Yes | Health Professional | A43P9 IS1200 | Device Returned To Manufacturer | Yes | 2011-10-06 | Device Not Retu | No | Reuse | 2011
420003 | 2011-10-06 | Other |  | It was reported that during a da vinci si single vessel arrested heart revascularization procedure, the clips that were being used with the small clip applier instrument were not staying closed. Before calling isi for technical support, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | An investigation performed by an isi representative, found that the site was using clips that are not compatible with the small clip applier instrument. Since this event, the site has been retrained on the proper clips to use and as of (b)(4) 2011, there has not been a reoccurrence of this issue at this site. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2343505 | 2325708 | 2222584 | 2955842201100354 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-06 | 2011-11-04 | Yes | No | Health Professional | 420003 | No | No | 2011-10-06 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-10-06 | Other | Other | It was reported that while preparing to start a da vinci s hysterectomy procedure, the site experienced a system message that the patient side cart (psc) was not connected. The psc was not powered on and the battery indicator leds were not illuminated. With the assistance of a isi field service engineer, the system emergency power off (epo) was applied and the power cord was swapped; however, the issue persisted. The power cord bracket was disconnected and reseated to the psc with no change. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the customer reported complaint was related to the patient side power distribution (ppd) pcs board. The patient side power distribution board is a printed circuit (pca) board located on the patient side cart that divides the electrical power feed into subsidiary circuits while providing a protective circuit breaker for each circuit. The system was repaired by replacing the ppd. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2343523 | 2325731 | 2222607 | 2955842201100352 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-06 | 2011-11-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-10-06 | Device Not Retu | No | Reuse | 2011
400180 | 2011-10-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the inner plastic part of the 8mm cannula seal accessory was found to be missing during removal of the obturator. It is not clear if the piece fell into the patient; however, laproscopic abdominal exploration was performed. No pieces were located. As of (b)(6) 2011, there has not been a report of patient harm, adverse outcome or injury as a result of this event. | The cannula seal accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned and/or if additional information is received. | 8MM CANNULA SEAL ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2343524 | 2325732 | 2222608 | 2955842201100353 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-10-07 | 2011-11-04 | No | Yes | Health Professional | 400180 | No | No | 2011-10-07 | No | Yes | Unkown | 2011
400180 | 2011-11-01 | Malfunction |  | Tip cover accessory on the da vinci monopolar curved scissors #63063 developed a small tear on the side. Tip cover removed and replaced. No patient harm or injury. The tip functions to prevent arcing of the electrical cautery. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2343197 | 2325414 | 2222290 | 2325414 | 1 | NAY | User Facility |  |  | Initial | 2011-11-03 | 2011-11-03 | No | Yes | Physician | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-10-03 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgical staff experienced reocurring system error code 212. With the assistance of an isi technical support engineer (tse), the site exercised all axes on the right master tool manipulator; however, the system error continued. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2339911 | 2320888 | 2217764 | 2955842201100344 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-03 | 2011-11-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-10-03 | Device Not Retu | No | Reuse | 2011
DA VINCI | 2011-06-20 | Malfunction | Other | Pt undergoing robot assisted laparoscopic prostatectomy, left nerve sparing with bilateral lymph node dissection. During dissection, a nonrecoverable fault was experienced on the robot. Technical support notified. Remote diagnostic and advised action on site were unsuccessful. A determination was made to proceed with open radial prostatectomy. |  | DA VINCI | DA VINCI | 2347248 | 2328028 | 2224884 | MW5022981 | 1 | NAY | Voluntary |  | Other | Initial | 2011-11-02 | 2011-11-02 | No | Yes | Health Professional | DA VINCI | Yes | Yes |  |  |  |  | 2011
42011704 | 2011-10-05 | Malfunction |  | It was reported that during a da vinci s surgical procedure while using the 5mm needle driver instrument, the instrument broke and a fragment fell into the patient. The fragment was removed and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, injury or adverse outcome was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2338009 | 2318154 | 2215029 | 2955842201100342 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-10-05 | 2011-11-01 | No | Yes | Health Professional | 42011704 | No | Yes | 2011-10-05 | No | No | Unkown | 2011
IS2000 A51P8 | 2010-11-03 | Injury | Other | It was reported by the patient that three weeks post, a successful da vinci s hysterectomy procedure, the patient was readmitted into the hospital and found to have a severed ureter. | The patient stated that on (b)(6) 2010, she had a nephrostomy tube placed, then on (b)(6) 2011, she underwent surgery to reattach her ureter and bladder. The patient also stated that she experienced hematomas at the cannula port sites. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. Based on the information provided, including review of the system error logs, it was determined that the da vinci s surgical system did not malfunction in a way that would have led or contributed to the patient's injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2338016 | 2318160 | 2215035 | 2955842201100343 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-10-03 | 2011-11-01 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-10-03 | Device Not Retu | No | Reuse | 2011
| 2011-10-31 | Malfunction |  | Patient undergoing robotic laparoscopic cystectomy of the bladder. This was a long surgery lasting 7 hours. Silicon tip cover accessory used on robotic instrument split from the tip of the cover backwards for 1/8 inch exposing the instrument. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2340136 | 2321283 | 2218159 | 2321283 | 1 | NAY | User Facility |  |  | Initial | 2011-10-31 | 2011-10-31 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2011
400180 | 2011-09-29 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, after approximately 20 minutes of use, arcing was observed from the tip cover accessory installed on the monopolar curved scissors(mcs) instrument. The mcs instrument was removed and the tip cover accessory was replaced. The planned surgical procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was discarded by the site and as of the date of this report the monopolar curved scissors instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Several attempts to obtain additional information from the site have been made concerning the reported event; however, no other information has been provided. If additional information is received, a follow up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2334616 | 2313976 | 2210852 | 2955842201100341 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-29 | 2011-10-28 | No | Yes | Health Professional | 400180 | No | No | 2011-09-29 | No | Yes | Unkown | 2011
IS2000 A51P8 | 2011-09-26 | Other | Other | It was reported that during a da vinci s splenectomy procedure, the surgical staff experienced a broken cannula mount. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by engineering concluded that the customer reported problem was associated with the endoscopic camera manipulator (ecm) cannula mount. The ecm cannula mount is designed for the camera arm and attaches to a disposable cannula that serves as a port of entry for the endoscope. The system was repaired by replacing the ecm cannula mount. The ecm cannula mount was returned for evaluation. Engineering discovered the ecm cannula mount to be broken in two pieces. The cannula mount base fractured at the 90 degree corner feature directly behind the back plate. The cross sectional area of base is also the smallest where it fractured. Temperature sensitive sticker has not changed color, indicating part has not exposed to high temps. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2331029 | 2309749 | 2206644 | 2955842201100340 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-09-26 | 2011-10-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-09-26 | No | No | Reuse | 2011
42014304 | 2011-09-22 | Malfunction |  | It was reported that during a da vinci si thyroid lobectomy procedure the 5mm maryland dissector instrument became jittery when articulating. The instrument was removed and found to have tube insulation peeling away with fragments having fell into the patient. The instrument fragments were retrieved and the procedure was completed with a replacement instrument of the same type. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the shaft to have a large section with damage to the black insulation, approximately 3. 75 inches above the instrument's snake wrist. The insulation has a 1. 7 inch by. 170 inch section missing, exposing the metal shaft below. The edges of the damaged section are rough. Engineering concluded that damage was likely due to mishandling and or misuse. No other damage was found. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2326357 | 2304980 | 2201876 | 2955842201100339 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-27 | 2011-10-24 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | Yes | 2011-09-27 | Device Not Retu | No | Reuse | 2011
IS2000 51P8 | 2011-09-27 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced a left eye video loss system error. With the assistance of an isi technical support engineer, the site checked their video cable to ensure it was properly seated. The video flickered when the cable was moved. The patient was under anesthesia and the port incisions were made when the surgical staff made the decision to abort the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the vision cable not being properly seated. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by reseating the camera cable. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2326388 | 2305024 | 2201920 | 2955842201100338 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-27 | 2011-10-24 | Yes | Yes | Health Professional | IS2000 51P8 | Yes | Yes | 2011-09-27 | Device Not Retu | No | Reuse | 2011
42019405 | 2011-09-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the grip pad on the jaw of the mega needle driver instrument fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. The broken piece will not be returned with the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one carbide insert broken off. The missing piece is approximately. 140 x. 090. The exposed part of the brazing surface on the grip has a large void and the filler material appears present only at the edges. The outer surfaces of both grips also exhibit light scratching. No other damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2324542 | 2303558 | 2200454 | 2955842201100333 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-21 | 2011-10-21 | No | Yes | Health Professional | 42019405 | Device Returned To Manufacturer | No | 2011-09-21 | Device Not Retu | No | Reuse | 2011
42009308 | 2011-09-02 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a cable from the prograsp forceps instrument fell into the patient. The broken cable piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable was broken at the distal clevis hub and the cable segment that contains the crimp was missing. No other damage was found. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2324552 | 2303560 | 2200456 | 2955842201100334 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-27 | 2011-10-21 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2011-09-27 | Device Not Retu | No | Reuse | 2011
42018304 | 2011-09-26 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the physician was using the permanent cautery hook instrument to hold the uterus out of the way while cauterizing around the cervix. During use, the instrument broke and pieces fell into the patient. A grasper was used to retrieve the broken pieces. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal clevis was dislodged from the main tube and a piece approximately. 45 x. 18 in size was found missing from the clevis. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2324546 | 2303582 | 2200478 | 2955842201100332 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-09-26 | 2011-10-21 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2011-09-26 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-09-22 | Other | Other | It was reported that during the beginning of a da vinci radical cystectomy procedure, the surgical staff experienced issues with an instrument arm not recognizing the instrument sterile adapter. The issue was resolved, however, the surgical staff had difficulty exchanging instruments on the same instrument arm during the middle of the procedure. The sterile adapter was reseated and the system began alarming. The system was rebooted, however, the affected instrument arm failed to home after restart. Multiple system reboots were attempted, however, the homing issue could not be resolved. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient post-surgical complications were reported. | An intuitive surgical technical support engineer spoke to the initial reporter and it was determined that the homing issue experienced by the customer was caused by the patient side manipulator (psm) insertion axis becoming stuck due to the cabling. The initial reporter reset the cabling and was able to home the system without any issues. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. An intuitive surgical field service engineer went onsite and replaced the psm. The psm was returned to intuitive surgical for failure analysis investigation. During internal testing, engineering was unable to replicate the customer reported homing and recognition/engagement failures. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2324647 | 2303614 | 2200510 | 2955842201100335 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-09-22 | 2011-10-21 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-09-22 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-09-22 | Other | Other | It was reported that during a da vinci si beating heart revascularization procedure, the surgeon was unable to activate the monopolar cautery function using the surgeon console foot pedal. The system was power cycled and the connection was switched to another input, however, the surgical staff was unable to resolve the issue. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the technical support engineer concluded that the monopolar cautery activation issue experienced by the customer was associated with the energy activation cable that connects the electrosurgical unit to the surgeon console. An intuitive surgical clinical sales representative was onsite and repaired the system by replacing the affected energy activation cable. As of october 20, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2325189 | 2304041 | 2200937 | 2955842201100336 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-09-22 | 2011-10-21 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-09-22 | Device Not Retu | No | Reuse | 2011
400180 | 2011-09-12 | Injury | Other | It was reported that post a successful da vinci si prostatectomy procedure, the patient returned to the hospital experiencing complications. It was discovered that the patient had a perforation in his right colon, approximately 6 cm from the caecum. It was noted that the monopolar curved scissors instrument used during the procedure was inspected prior to use. Additionally, no complications were observed by the surgeon or operating room staff with the monopolar curved scissors instrument or the tip cover accessory during the procedure. | Multiple attempts to obtain additional information from the surgeon have been made, however, at the time of this report limited information has been received. The monopolar curved scissors instrument used during the procedure was returned as the surgeon requested verification that unintended electrical arcing from the instrument did not cause the patient injury. The instrument was evaluated and engineering was unable to find any evidence of arcing or melted materials. The instrument was placed on an in house test system for performance testing and no trouble was found. The tip cover accessory used during the procedure was discarded by the hospital and will not be returned. The root cause of the customer reported event cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2323378 | 2302813 | 2199709 | 2955842201100330 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-09-21 | 2011-10-20 | Yes | No | Health Professional | 400180 | No | No | 2011-09-21 | No | Yes | Unkown | 2011
IS2000 51P8 | 2011-09-20 | Other | Other | It was reported that prior to starting a da vinci s nephrectomy procedure, the surgical staff reported issues turning the patient side cart (psc) to the left and forward. It was determined that a field service engineer (fse) would arrive on site to troubleshoot the issue. The patient had been under anesthesia and port incisions had been made when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the linkage between the patient side cart and the wheel assembly. The system was repaired by replacing the damaged linkage. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2323428 | 2302853 | 2199749 | 2955842201100331 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-20 | 2011-10-20 | Yes | Yes | Health Professional | IS2000 51P8 | Yes | Yes | 2011-09-20 | Device Not Retu | No | Reuse | 2011
IS3000 A60P4 | 2011-09-19 | Other |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff experienced system error code 25521. The system was rebooted multiple times, however, system error code 25521 continued to recur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 25521 was associated with the left master tool manipulator (mtm). The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25521 appears when the system detects that it has lost communication of sensor information between two printed circuit assemblies in the arm control system. In the event of these communication issues, the system records the fault and transitions to a safe state. The mtml was returned to intuitive surgical for failure analysis investigation. Engineering was unable to replicate the reported failure during internal testing, however, multiple printed circuit assembly boards were replaced and various upgrades were performed as a precaution. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2319250 | 2299568 | 2196464 | 2955842201100329 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-09-19 | 2011-10-18 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2011-09-19 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-09-19 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, the site experienced a system error code 297. With the assistance of an isi technical support engineer the site restarted the system, however, the system error 297 reoccurred upon start up when a specific set up joint (suj) was moved. The patient was under anesthesia and port incisions had been placed when the surgical staff made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 297 was associated with a patient side manipulator (psm) setup-joint (suj). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instrument. The setup-joint (suj) is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 297 occurs when reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The suj was returned to isi for failure analysis investigation. Engineering duplicated the customer reported system error and found the suj cable harness to be defective. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2318025 | 2298872 | 2195768 | 2955842201100327 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-19 | 2011-10-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-09-19 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-09-20 | Other | Other | It was reported that during a da vinci si pyeloplasty procedure, the site experienced system error 23008 multiple times. With the assistance of an isi technical support engineer (tse), the site exercised axis 4 on the endoscopic camera manipulator (ecm) arm, powered down the system and restarted, cycled the power breaker on the patient side cart and re-exercised axis 4 however, system error 23008 re-appeared. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci si safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The ecm was returned to isi for evaluation. Engineering was able to reproduce the customer reported problem and found the motor and encoder to be defective. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2318040 | 2298874 | 2195770 | 2955842201100328 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-20 | 2011-10-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-09-20 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-09-14 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, non-recoverable system error code 25580 occurred. The surgical staff attempted to resolve the issue by pressing the emergency power off button on the patient cart and then restarting the system, however, system error code 25580 recurred during startup. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 25580 was associated with a remote arm controller (rac). The rac consists of five printed circuit assembly boards which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25580 is reported when the da vinci si safety system determines a hardware voltage level is out of range, in this case on the rac. Upon determining this condition, the system records the fault and transitions to a safe state. The rac was returned for evaluation. Engineering was able to replicate the reported failure and determined that a remote drive module had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2315499 | 2296439 | 2193335 | 2955842201100326 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-14 | 2011-10-14 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-09-14 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-08-16 | Malfunction |  | It was reported that while the surgeon was dissecting tissue during a da vinci si hysterectomy procedure, arcing was observed coming from the 8 mm monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover accessory was removed and inspected, and a hole was noticed during inspection. A replacement tip cover accessory was installed on the mcs instrument and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. The tip cover accessory was visually inspected under magnification and no holes were observed. The tip cover accessory does not exhibit any signs of arcing or mechanical tearing of the silicone. The tip cover accessory was also installed on an in-house mcs instrument without difficulty. No physical damage was found and the customer reported failure cannot be confirmed. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2315551 | 2296477 | 2193373 | 2955842201100325 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-14 | 2011-10-14 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-09-14 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-08-16 | Malfunction |  | It was reported that while the surgeon was dissecting tissue during a da vinci si hysterectomy procedure, arcing was observed coming from the 8 mm monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover accessory was removed and inspected, and a hole was noticed during inspection. A replacement tip cover accessory was installed on the mcs instrument and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is oblong, approximately. 050 x. 030 in size, and is located. 190 above the silicone to sleeve interface. The tip cover accessory also has a mechanical tear in the silicone, near the hole. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2315572 | 2296505 | 2193373 | 2955842201100324 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-14 | 2011-10-14 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-09-14 | Device Not Retu | Yes | Initial | 2011
400180 | 2011-10-14 | Malfunction |  | Patient underwent da vinci robotic laparoscopic retropubic radical prostatectomy surgery. During the procedure the accessory tip cover developed a hole. The tip was immediately removed off the surgical field. The patient suffered no apparent injury. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2317783 | 2298658 | 2195554 | 2298658 | 1 | NAY | User Facility |  | Invalid Data | Initial | 2011-10-14 | 2011-10-14 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
400180 | 2011-10-14 | Malfunction |  | Patient underwent da vinci robotic laparoscopic retropubic radical prostatectomy surgery. During the procedure the accessory tip cover developed a hole. The tip was immediately removed. The patient suffered no apparent injury. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2317784 | 2298660 | 2195556 | 2298660 | 1 | NAY | User Facility |  |  | Initial | 2011-10-14 | 2011-10-14 | No | Yes | Invalid Data | 400180 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-09-13 | Other | Other | It was reported that during set up for a da vinci s sacrocolpopexy, the surgical staff experienced system error code 23013. The site contacted an isi technical support engineer to assist in troubleshooting this issue; however, it was determined that this system error code would be likely to re-occur during the case. The patient was under anesthesia and the port incisions had been placed when the surgeon made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering found that issue experienced by the customer to be associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23013 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2312144 | 2293507 | 2190403 | 2955842201100323 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-13 | 2011-10-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-09-13 | Device Not Retu | No | Reuse | 2011
400180 | 2011-10-07 | Malfunction |  | A recurrent problem has been occurring with the silicone tip covering the scissors and bovie of the da vinci robotic surgery instrument whereby holes develop in the tip cover at the flexible joint of the instrument. This requires replacement of the tip cover one or more times during surgery. The bovie can fire through this hole, potentially burning blood vessel or bowel. |  | MICROTEK DA VINCI TIP COVER ACCESSORY | INSTRUMENT ROBOTIC SURGICAL TI | 2317780 | 2298655 | 2195551 | 2298655 | 1 | NAY | User Facility |  |  | Initial | 2011-10-13 | 2011-10-13 | No | Yes | Invalid Data | 400180 |  | No Answer Provided |  |  |  |  | 2011
420227 VER03 | 2011-09-13 | Malfunction |  | During a robotic-assisted laparotomy, a screw broke off the da vinci s pk (plasmakinetic) dissecting forceps and fell into the abdominal cavity. (the device continued to work after this event). The screw was located and retrieved from the abdominal cavity. No harm to patient. Vendor was in the or at the time of the malfunction. |  | DA VINCI S | INSTRUMENT ROBOTIC SURGICAL PK | 2315743 | 2296624 | 2193520 | 2296624 | 1 | NAY | User Facility |  |  | Initial | 2011-09-29 | 2011-10-12 | No | Yes | Physician | 420227 VER03 | Yes | No Answer Provided |  |  |  |  | 2011
| 2011-10-11 | Malfunction |  | During a gyn robotic case, the mega suturecut needle driver broke, and the cable snapped. The instrument was cleaned and risk management notified. |  | MEGA SUTURECUT NEEDLE DRIVER | GUIDE NEEDLE SURGICAL | 2335701 | 2315393 | 2212269 | 2315393 | 1 | GDF | User Facility |  |  | Initial | 2011-10-11 | 2011-10-11 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2011
42019408 | 2011-09-09 | Malfunction |  | It was reported that while suturing during a da vinci si myomectomy procedure, one side of the flat plate at the tip of the mega needle driver instrument fell into the patient. The procedure had been in progress for 3. 5 hours when the malfunction occurred. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the carbide insert was detached from one grip. The brazing surface on the affected grip exhibits some voids in the filler material. A couple of scratch marks were also observed on the outer surface of the grip. The detached carbide insert was returned with the instrument. No other damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2300562 | 2282715 | 2179611 | 2955842201100320 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-09 | 2011-10-07 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2011-09-09 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-09-08 | Other | Other | It was reported that while performing lymph node dissection during a da vinci hysterectomy procedure, the surgeon master controllers pulled forward and then moved in the superior direction. As a result, non-intuitive movement occurred with the instrument arms assigned to the master controllers and the system immediately faulted. System error code 20013 was displayed and the surgeon pressed the fault override button to continue with the procedure. The same issue occurred two more times and after the third occurrence, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the non-intuitive movement experienced by the customer, which resulted in system error code 20013 being displayed, was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the system records the faults and transitions to a safe state. The psm was returned to intuitive surgical for failure analysis investigation. Engineering was unable to replicate the customer reported failure mode during internal testing and the psm passed all tests. As a precaution, the axis-2 potentiometer and motors were replaced. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2300557 | 2282729 | 2179625 | 2955842201100319 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-09-08 | 2011-10-07 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-09-08 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-09-09 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the site observed sparking from the monopolar curved scissors (mcs) instrument with tip cover accessory attached. No additional information was provided. No patient harm was reported. An attempt to gather additional information has been unsuccessful. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint. Instrument was returned with a tip cover installed, however, no damage could be identified on the tip cover accessory or the instrument using magnification. There were no signs of arcing or localized melting observed on silicone. The tip cover was removed to inspect the distal end of instrument for char marks. No char marks were observed at the wrist components or on the main tube and tube extension interface. The instrument was driven on an in-house is3000 system and moved intuitively with a full range of motion in all directions. An attempt to create arcing through the tip cover using a generator and monopolar cord were unsuccessful. Scissors opened and closed properly. Latex cut test and electrical continuity passed. Instrument is fully functional with no physical damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2301653 | 2283678 | 2180574 | 2955842201100321 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-09 | 2011-10-07 | No | Yes | Health Professional | 42017910 | No | No | 2011-09-21 | Device Not Retu | Yes | Initial | 2011
42000606 | 2011-09-08 | Malfunction |  | It was reported that 3 hours into a da vinci s prostatectomy procedure, and while suturing, one side of the flat plate located on the tip of the large needle driver instrument, fell into the patient and was retrieved. The large needle driver instrument was replaced with another instrument of the same type and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument was returned with the carbide insert broken off of one grip. The grip surface where the carbide insert was installed has a smooth surface with no adhesive residue. Indicating a clean break off the grip. No additional damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2301670 | 2283691 | 2180587 | 2955842201100322 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-09-08 | 2011-10-07 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2011-09-08 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-09-06 | Injury | Life Threatening | It was reported that during a da vinci surgical procedure, the tip cover accessory installed on the monopolar curved scissors (mcs) instrument became perforated. As a result, arcing occurred and the patient's vein was burnt. The same incident then occurred two additional times with two other separate tip cover accessories before the case was completed. This medwatch mfr report is for the third occurrence of arcing. Medwatch mfr report 2955842-2011-00316 and 2955842-2011-00317 were submitted for the first and second occurrences of arcing. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 010 x. 040 oval shaped hole in the silicone. The hole is located. 100 above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. A mechanical tear in the middle section of the silicone was also observed. Multiple requests for additional information have been made to the customer. To date, no additional information has been received. General precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2298852 | 2280951 | 2177847 | 2955842201100318 | 1 | NAY | Manufacturer | ForeignCompany Representative | Health Professional | Initial | 2011-09-06 | 2011-10-06 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-09-06 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-09-06 | Injury | Life Threatening | It was reported that during a da vinci surgical procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument became perforated. As a result, arcing occurred and the patient's vein was burnt. The same incident then occurred two additional times with two other separate tip cover accessories before the case was completed. This medwatch mfr report is for the first occurrence of arcing. Medwatch mfr report 2955842-2011-00317 and 2955842-2011-00318 were submitted for the second and third occurrences of arcing. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed four holes burnt through the silicone. The silicone exhibits localized melting around the holes, which is indicative of arcing. The holes range from. 015 to. 040 in diameter and are located. 060 to. 520 above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. A tear was also observed on the distal end of the tip cover accessory. Multiple requests for additional information have been made to the customer. To date, no additional information has been received. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2298869 | 2280953 | 2177849 | 2955842201100316 | 1 | NAY | Manufacturer | ForeignCompany Representative | Health Professional | Initial | 2011-09-06 | 2011-10-06 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-09-06 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-09-06 | Injury | Life Threatening | It was reported that during a da vinci surgical procedure the tip cover accessory installed on the monopolar curved scissors (mcs) instrument became perforated. As a result, arcing occurred and the patient's vein was burnt. The same incident then occurred two additional times with two other separate tip cover accessories before the case was completed. This medwatch mfr report is for the second occurrence of arcing. Medwatch mfr report 2955842-2011-00316 and 2955842-2011-00318 were submitted for the first and third occurrences of arcing. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 025 x. 035 oval shaped hole in the silicone. The hole is located. 280 above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. A mechanical tear at the distal end of the silicone was also observed. Multiple requests for additional information have been made to the customer. To date, no additional information has been received. General precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2298864 | 2280964 | 2177860 | 2955842201100317 | 1 | NAY | Manufacturer | ForeignCompany Representative | Health Professional | Initial | 2011-09-06 | 2011-10-06 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-09-06 | Device Not Retu | Yes | Initial | 2011
REF 420309 VER 01 | 2011-10-04 | Malfunction |  | Megas suture cut needle driver tips snapped during procedure. |  | DAVINCI | INSTRUMENT ROBOTIC SURGICAL SU | 2309986 | 2291668 | 2188564 | 2291668 | 1 | NAY | User Facility |  |  | Initial | 2011-10-06 | 2011-10-06 | No | Yes | Physician | REF 420309 VER 01 | Yes | No Answer Provided |  |  |  |  | 2011
42018406 | 2011-09-02 | Injury | Life Threatening | It was reported that while the surgeon was cauterizing tissue during hilar dissection for a da vinci s lung biopsy procedure the permanent cautery hook instrument arced and burnt the patient's azygos vein. The burn required no treatment and the planned surgical procedure was completed. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2290060 | 2270470 | 2167366 | 2955842201100314 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-02 | 2011-09-30 | Yes | Yes | Health Professional | 42018406 | No | No | 2011-09-02 | No | No | Unkown | 2011
400180 | 2011-09-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was removed and the tip cover accessory was replaced. The planned surgical procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. | Multiple attempts have been made to obtain additional information from the site about this event, however, at the time of this report no information has been provided. Additionally, the tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2290869 | 2271105 | 2168001 | 2955842201100315 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-09-09 | 2011-09-30 | No | Yes | Health Professional | 400180 | No | No | 2011-09-09 | No | Yes | Unkown | 2011
420179 | 2011-05-25 | Malfunction |  | Surgeon observed sparks coming out of the instrument. Defective instrument was removed and exchanged for a new one. |  | HOT SHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2293769 | 2274773 | 2171669 | 2274773 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-09-30 | 2011-09-30 | No | Yes | Physician | 420179 |  | No Answer Provided |  |  |  |  | 2011
42027202 | 2011-08-16 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the tip of the harmonic ace curved shears insert broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the insert accessory is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 670 x. 080 piece of the curved blade detached from the distal end. The blade had fractured within the shaft, near the clamp arm pin. The insert was disassembled to check for signs of contact between the blade and clamp arm. A dark mark was found on the blade, directly above the fracture surface on the side that the clamp arm pin is located on. The clamp arm pin also exhibits a dark mark at the center. It was determined that the breakage may have been due to contact between the pin and clamp arm during harmonic cautery activation. No other damage was found. There are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. | HARMONIC ACE CURVED SHEARS INSERT | Unkown | 2288387 | 2269018 | 2165913 | 2955842201100312 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-08-31 | 2011-09-29 | No | Yes | Health Professional | 42027202 | Device Returned To Manufacturer | No | 2011-10-10 | Device Not Retu | Yes |  | 2011
IS3000 A60P6 | 2011-07-18 | Death | Death | It was reported that 56 hours after a da vinci si tors (transoral robotic surgery) procedure was successfully performed, the patient experienced heavy bleeding and cardiopulmonary arrest ensued. Cpr was performed in excess of 20 minutes before the patient's pulse was restored and as a result, irreversible anoxic brain injury occurred. On (b)(6) 2011, the patient expired. At the time of the report, the cause of the patient bleeding was undetermined. It was also reported that during the surgical procedure, there were no system failures, no injury to the patient, and no complications experienced. | On (b)(6) 2011, the surgeon who performed the tors procedure informed intuitive surgical that they were advised not to provide any additional information. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2288388 | 2269019 | 2165914 | 2955842201100313 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-30 | 2011-09-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-08-30 | No | No | Reuse | 2011
42027202 | 2011-08-24 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the tip of the harmonic ace curved shears insert broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the insert accessory is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 670 x. 080 piece of the curved blade detached from the distal end. The blade had fractured within the shaft, near the clamp arm pin. The insert was disassembled to check for signs of contact between the blade and clamp arm. A dark mark was found on the blade, directly above the fracture surface on the side that the clamp arm pin is located on. It was determined that the breakage may have been due to contact between the pin and clamp arm during harmonic cautery activation. No other damage was found. There are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. | HARMONIC ACE CURVED SHEARS INSERT | Unkown | 2288358 | 2269026 | 2165922 | 2955842201100311 | 1 | NAY | Manufacturer | Company Representative | Other | FollowupInitial | 2011-08-31 | 2011-09-29 | No | Yes | Health Professional | 42027202 | Device Returned To Manufacturer | No | 2011-10-10 | Device Not Retu | Yes |  | 2011
| 2011-08-17 | Malfunction |  | Scrub tech was attaching harmonic device to robotic arm and they would not connect. A second device was used without incident. It is unknown if the second device was the same or different lot number. |  | HARMONIC ACE CURVED SHEARS INSERT | INSTRUMENT ROBOTIC SURGICAL HA | 2318908 | 2299330 | 2196226 | 2299330 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-09-29 | 2011-09-29 | No | Yes | Other |  | No | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-08-29 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon experienced difficulty moving the patient side manipulatory axis. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found the issue experienced by the site to be associated with the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the psm. The psm was returned and evaluated. Engineering observed that the assembly brake was damaged, thus causing the customer reported complaint. As of (b)(4), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2287427 | 2268374 | 2165269 | 2955842201100310 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-08-29 | 2011-09-28 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-09-29 | No | No | Reuse | 2011
420179 | 2011-09-26 | Malfunction |  | While the davinci monopolar curved scissor instrument was being used through a port inside the patient, the distal end of the scissor partially broke away from the handle. The wire inside of the scissor remained intact within the instrument. All pieces were removed from the laparoscopic port. |  | HOT SHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2293767 | 2274772 | 2171668 | 2274772 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-09-28 | 2011-09-28 | No | Yes | Physician | 420179 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-08-29 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgical staff experienced system error code 23013. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer determined that system error code 23013 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 23013 appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. The psm was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported failure mode during testing and determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2285731 | 2266489 | 2163385 | 2955842201100308 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-08-29 | 2011-09-27 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-08-29 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-09-02 | Other | Other | It was reported that the camera cannula mount broke when the surgeon was docking the system to the patient for a da vinci s procedure. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The camera cannula mount is designed for the camera arm and attaches to a disposable cannula that serves as a port of entry for the endoscope. An intuitive surgical technical support engineer spoke to the customer and informed them to order a replacement camera cannula mount. On (b)(4) 2011, intuitive surgical confirmed with the customer that a camera cannula mount was purchased and the affected camera cannula mount was replaced. As of (b)(4), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2285732 | 2266490 | 2163386 | 2955842201100309 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-09-02 | 2011-09-27 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2011-09-02 | No | No | Reuse | 2011
400180 | 2011-08-31 | Other | Other | It was reported that during the beginning of a da vinci si hysterectomy procedure, while the surgeon was working on the round ligament, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs was immediately removed and the tip cover accessory was replaced. The procedure continued however, shortly after the replacement tip cover accessory was in use, arcing was observed coming from the mcs instrument. The arc hit the back side of the patient's uterus causing a superficial burn. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, injury or adverse outcome was reported. | The tip cover accessories have not been returned for evaluation as they were discarded by the hospital. The monopolar curved scissors (mcs) instrument used during the procedure was returned and evaluated. Engineering was unable to find any evidence of arcing and no burrs or damage in the area where arcing may have occurred. A new tip cover accessory was placed on the instrument in attempt to re-create the customer reported failure, however after multiple attempts arcing was not observed. Electrical continuity passed and a cut test was successful. No physical damage was found. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch report was sent to the fda in relation to this event: mfr report 2955842-2011-00305. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2283538 | 2264081 | 2160981 | 2955842201100306 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-31 | 2011-09-26 | Yes | Yes | Health Professional | 400180 | No | No | 2011-08-31 | No | Yes | Initial | 2011
400180 | 2011-08-31 | Other | Other | It was reported that during the beginning of a da vinci si hysterectomy procedure, while the surgeon was working on the round ligament, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs was immediately removed and the tip cover accessory was replaced. The procedure continued however, shortly after the replacement tip cover accessory was in use, arcing was observed coming from the mcs instrument. The arc hit the back side of the patient's uterus causing a superficial burn. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, injury or adverse outcome was reported. | The tip cover accessories have not been returned for evaluation as they were discarded by the hospital. The monopolar curved scissors (mcs) instrument used during the procedure was returned and evaluated. Engineering was unable to find any evidence of arcing and no burrs or damage in the area where arcing may have occurred. A new tip cover accessory was placed on the instrument in attempt to re-create the customer reported failure, however after multiple attempts arcing was not observed. Electrical continuity passed and a cut test was successful. No physical damage was found. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch report was sent to the fda in relation to this event: mfr report 2955842-2011-00306. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2283537 | 2264085 | 2160981 | 2955842201100305 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-31 | 2011-09-26 | Yes | Yes | Health Professional | 400180 | No | No | 2011-08-31 | No | Yes | Initial | 2011
42018405 | 2011-08-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the permanent cautery spatula instrument broke and fell into the patient. The fragments were retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow up mdr will be submitted. | PERMANENT CAUTERY SPATUAL INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2283549 | 2264092 | 2160988 | 2955842201100307 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-29 | 2011-09-26 | No | Yes | Health Professional | 42018405 | No | Yes | 2011-08-29 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-08-23 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced multiple occurrences of system error code 23013. Prior to calling into intuitive surgical for technical support, the surgeon made the decision to convert the procedure to traditional open surgical techniques. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23013 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The psm was returned to isi for evaluation. Engineering confirmed the occurrence of system error code, thus confirming the reported failure mode experienced by the customer. The psm's potentiometer and motor were replaced. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2280825 | 2260958 | 2157854 | 2955842201100302 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2011-08-23 | 2011-09-23 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-08-23 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-07-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while suturing the vaginal cuff, the metal grip insert came off of the mega needle driver instrument. The instrument fragment was immediately observed and retrieved with no impact to the patient. The planned surgical procedure was completed with additional issues or patient injury. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the carbide insert to have fallen off of one of the grip tips. The entire piece of the insert was detached. The grip tip surface was found to be clean with no adhesive residue. No damage was found to the instrument's clevis or cables. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2280827 | 2260987 | 2157883 | 2955842201100303 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-08-23 | 2011-09-23 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2011-08-23 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-08-01 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the customer experienced a multiple recoverable system error 23025. Before contacting intuitive surgical's technical support, the surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. It was reported that the patient had experienced post operative bleeding and was taken to the operating room. The site has been contacted to request additional information regarding this event; however, no additional information has been provided. | The investigation conducted by field service engineering concluded that the system error codes experienced by the customer were associated with the master tool manipulator (mtm). The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtm. The mtm was returned to isi for failure analysis investigation. Engineering was able to replicate the system error and replaced the motor to repair the returned mtm. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2282476 | 2262699 | 2159595 | 2955842201100304 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-08-26 | 2011-09-23 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2011-08-26 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-08-19 | Other | Other | It was reported that while preparing for a da vinci prostatectomy procedure, the surgical staff experienced difficulty with the patient side manipulator (psm) arm not recognizing the instrument cannula accessory. The site attempted to troubleshoot the problem without success. At this time, the patient had been under anesthesia and the port holes had been created when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. | The investigation conducted by the field service engineer concluded that the issue experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned for failure analysis investigation. Engineering evaluation observed that the axis 2 rolling loop was deformed causing interference and binding of the cable for insertion. The rolling loop was replaced to resolve the issue with the psm. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2273862 | 2252309 | 2149225 | 2955842201100301 | 1 | NAY | Manufacturer | ForeignCompany Representative | Health Professional | Initial | 2011-08-19 | 2011-09-17 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-08-19 | Device Not Retu | No | Reuse | 2011
42017909 | 2011-08-09 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a piece of the monopolar curved scissors instrument broke off and fell inside of the patient. The piece was retrieved and the planned surgical procedure was completed using a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received of if the instrument is returned, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2273199 | 2251758 | 2148654 | 2955842201100300 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-18 | 2011-09-16 | No | Yes | Health Professional | 42017909 | No | Yes | 2011-08-18 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-07-07 | Death | Death | It was reported that approximately 45 minutes into a da vinci s ovarian cystectomy procedure, the surgeon encountered excessive adhesions within the patient's anatomy which made it difficult to perform the procedure laparoscopically. At this time, the surgeon made the decision to convert to traditional open techniques to successfully complete the procedure with no issues noted. Approximately two to three weeks post procedure, the patient returned to the hospital and was diagnosed with ileus. It is unknown how the patient was treated; however, the patient expired after being diagnosed. The site's risk management department has been contacted several times to obtain additional information; however, our attempts have been unsuccessful. The delay in reporting is due to an isi employee's misunderstanding of complaint reporting criteria. The site advised that there was not correlation between our system and the event and our isi representative understood the site's feedback as not requiring the initiation of a complaint record. Since this misunderstanding was discovered, retraining has been performed with the isi representative to mitigate future occurrences. | Based on our investigation, there were no complaints against the da vinci system, instruments or accessories and the site has continued to use the system to perform surgery on patients without incident. The site has been contacted for additional information related to this event. If additional information becomes available, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2271626 | 2250270 | 2147186 | 2955842201100299 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-07-07 | 2011-09-15 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-07-07 | No | No | Reuse | 2011
420172 | 2011-09-06 | Malfunction |  | During a surgical procedure, the surgical resident was disconnecting the new davinci endowrist bipolar maryland instrument. During this time, the connecting bipolar prongs at the top of the instrument had slipped out of the instrument. It was not determined whether or not the wires were internally disconnected. Witnessed by manufacturer rep. |  | MARYLAND BIPLOAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2284397 | 2265098 | 2193252 | 2265098 | 1 | NAY | User Facility |  |  | Initial | 2011-09-14 | 2011-09-14 | No | Yes | Physician | 420172 | No Answer Provided | No Answer Provided |  |  |  |  | 2011
| 2011-08-02 | Malfunction |  | During end of the procedure, the robot fault alarmed. The over ride button was pressed and the robot fault alarmed again. Manufacturer technical support was notified. Technician said the left master arm was broken beyond repair. A cable was snapped. Converted to laparoscopic prostatectomy. |  | DAVINCI S SYSTEM | DAVINCI SURGICAL SYSTEM | 2284398 | 2265099 | 2161995 | 2265099 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-09-13 | 2011-09-13 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-08-11 | Other | Other | It was reported that while preparing to start a da vinci s prostatectomy procedure and while docking the system to the patient, the site experienced system error code 1. With the assistance of an isi technical support engineer (tse), the site applied emergency power off (epo) to the patient side cart (pst) and the surgeon side cart (ssc); however, the error persisted. The patient was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the customer reported complaint was related to the patient side power distribution (ppd) pcs board. The patient side power distribution board is a printed circuit (pca) board located on the patient side cart that divides the electrical power feed into subsidiary circuits while providing a protective circuit breaker for each circuit. The system was repaired by replacing the ppd. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2265768 | 2244226 | 2141142 | 2955842201100297 | 1 | DHB | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-11 | 2011-09-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-08-11 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-08-15 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure the customer was unable to calibrate the endoscope. The isi representative onsite exchanged the endoscope, camera, and camera cable, however the endoscope would not white balance. With the assistance of an technical support engineer the site restarted the system however the vision issue continued to occur. The patient was under anesthesia when the surgical staff made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) and review of the system error logs determined that the vision issue experienced by the customer was associated with the double camera controller (doco) and the pmva pca board. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The pmva is a printed circuit assembly (pca) which receives the stereo video images from the camera controller and outputs the audio and video for the touchscreen monitor. The system was repaired by replacing the affected doco and pmva pca board. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2266035 | 2244442 | 2141358 | 2955842201100298 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-15 | 2011-09-12 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-08-15 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-08-11 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced system error code 20008 and 21008. They site attempted to troubleshoot the issue by restarting the system and applying emergency power off (epo), however, the system error code persisted. At this time, the surgeon made the decision to convert the planned surgical procedure to a traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. System error codes 20008 and 21008 appears when the davinci safety system, the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. The psm was returned for failure analysis investigation. Engineering evaluation observed that the axis 7 cables had become derailed from its pulleys. The psm was repaired by replacing the cables. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2263099 | 2241809 | 2138748 | 2955842201100293 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-08-11 | 2011-09-09 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-08-11 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-08-05 | Death | Death | It was reported that after a da vinci si bilateral salpingo-oophorectomy (bso) was completed, the patient returned to the hospital the next day or two days post procedure complaining of pain. A general surgeon saw the patient and made the decision to open the patient for exploratory surgery which resulted in a diagnosis of sepsis. It is unknown how the patient was treated. The patient expired on (b)(6) 2011. Since receiving this initial report, the site has been contacted to request additional information without success. | The site's risk management department was contacted to obtain additional information related to this event and risk management advised that isi product did not cause or contribute to this event and that the no additional information would be provided included any medical records or copies of an autopsy report. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2263142 | 2241885 | 2138804 | 2955842201100291 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-09 | 2011-09-09 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-08-09 | No | No | Reuse | 2011
42017206 | 2011-07-19 | Malfunction |  | It was reported that while performing hemostatis on the patient's liver during a da vinci si liver resection procedure, the surgeon was grasping tissue for removal when a 1cm flame was observed coming from the maryland bipolar forceps instrument. The first instrument was removed and replaced with another maryland bipolar forceps instrument, however, a 1cm flame was also observed with this second instrument after 30 minutes of use. The second maryland bipolar forceps instrument was removed and replaced with a third instrument of the same type to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. This report addresses the second instrument used. Medwatch mfr report 2955842-2011-00294 was submitted for the first maryland bipolar forceps instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on both yaw pulleys at the base grips. It was determined that the charring was likely due to a breach in the instrument insulation or carbonized tissue which created a conductive path. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2263496 | 2242222 | 2139153 | 2955842201100295 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-08-03 | 2011-09-09 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2011-08-03 | Device Not Retu | No | Reuse | 2011
42017206 | 2011-07-19 | Malfunction |  | It was reported that while performing hemostatis on the patient's liver during a da vinci si liver resection procedure, the surgeon was grasping tissue for removal when a 1cm flame was observed coming from the maryland bipolar forceps instrument. The first instrument was removed and replaced with another maryland bipolar forceps instrument, however, a 1cm flame was also observed with this second instrument after 30 minutes of use. The second maryland bipolar forceps instrument was removed and replaced with a third instrument of the same type to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. This report addresses the first instrument used. Medwatch mfr report 2955842-2011-00295 was submitted for the second maryland bipolar forceps instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on both yaw pulleys at the base grips. It was determined that the charring was likely due to a breach in the instrument insulation or carbonized tissue which created a conductive path. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2263495 | 2242237 | 2139153 | 2955842201100294 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-08-03 | 2011-09-09 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2011-08-03 | Device Not Retu | No | Reuse | 2011
420184 | 2011-08-08 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure, arcing was observed to have come from the permanent cautery spatula instrument. No patient harm was reported. The site reported that the procedure was successfully completed. Since this report was originally received, several attempts have been made to gather additional information without success. | Manufacturer Narrative: --/-- Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal clevis to exhibits charring and localized melting on the outer surface of one ear. The yaw pulley also exhibits charring near where the cable crimp is installed. The damage is indicative of an arcing event. Additional observations are a damaged conductor wire and broken ceramic sleeve. The conductor wire has a section with missing insulation at the yaw pulley. The bare wire is exposed, creating a path for arcing. The distal end of ceramic sleeve is broken off and the sleeve broke where it changes diameter. Based on these finding it is likely that the damage to the wire and sleeve are due to mishandling/misuse. No other damage found. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 2258612 | 2237765 | 2134681 | 2955842201100292 | 1 | NAY | Manufacturer | User facility | Health Professional | InitialFollowup | 2011-08-08 | 2011-09-08 | No | Yes | Health Professional | 420184 | Device Returned To Manufacturer | Yes | 2011-08-08 | Device Not Retu | No |  | 2011
IS1000 | 2011-02-03 | Malfunction |  | Robot failure and malfunction during robotic assisted radical prostatectomy. Instruments attached to robot arm, surgeon at console with no difficulties noted. Robot docked for surgery and dissection begun by surgeon. Dissection accomplished through perivesical space when system malfunctioned for approximately 30 minutes. Twice, robot failed during additional dissection requiring redocking of robot, turning off unit, and restarting. System finally did not reboot and surgeon had to switch to open procedure. ======================manufacturer response for davinci robot, intuitive davinci (per site reporter)======================robot noted to have defective parts. Parts replaced include:-installed: part 650340-13-03 (no old part removed). -removed: 370967-01 (assemply, uip ii membrane & backer, english). -removed: 380099-01 (assembly, uip, removable). New part installed. Following this action, manufacturer noted no other issues. |  | DAVINCI SURGICAL SYSTEM | DAVINCI SURGICAL SYSTEM | 2272470 | 2250774 | 2147690 | 2250774 | 1 | NAY | User Facility |  |  | Initial | 2011-09-07 | 2011-09-07 | No | Yes | Physician | IS1000 | Yes | No Answer Provided |  |  |  |  | 2011
42019408 | 2011-08-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while starting the cuff closure and using the mega needle driver instrument, a piece of the instrument jaw broke off and fell into the patient. The instrument fragment was retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the carbide insert to have fallen off of one of the grip tips. The entire piece of the insert was detached. The grip tip surface was found to be clean with no adhesive residue. No damage was found to the instrument's clevis or cables. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2253873 | 2233157 | 2130073 | 2955842201100287 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2011-08-05 | 2011-09-02 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-08-05 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-08-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, after the colpotomy and while using the monopolar curved scissors (mcs) instrument, the surgeon attempted to remove the patient's uterus transvaginally. The instruments were then exchanged to sew the vaginal cuff when the mcs instrument was found to be broken at the instrument wrist. The mcs instrument could not be removed through the cannula due to the breakage, and therefore, the cannula and instrument had to be removed from the patient. It was noted that fragments fell into the patient and were removed. No patient harm, injury or adverse outcome were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument's entire tube extension to be dislodged from the main tube. The tube extension wall was found to be circumferentially broken at the base of the axial keys, however, no pieces are missing. Engineering concluded that the damage was likely due to excessive side loading or other type of mishandling. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2253875 | 2233159 | 2130075 | 2955842201100288 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-09 | 2011-09-02 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2011-08-09 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-07-12 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found various holes burned through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The ultem sleeve was found to be melted as one hole is located at the interface between the silicone and ultem sleeve. The hole sizes range from. 140 inches to under. 010 inches and are located from 0 inches to. 300 inches above the silicone to sleeve interface. The tip cover accessory was also found to have mechanical tears in the vicinity of the holes, including one tear that connects two holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2253883 | 2233190 | 2130106 | 2955842201100289 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-10 | 2011-09-02 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-08-10 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-07-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is a oblong and measures approximately. 015 inches by. 030 inches and is located. 460 inches above the silicone to sleeve interface. The tip cover accessory was found to have mechanical tears in the vicinity of the holes, and an axial tear at the tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2253907 | 2233191 | 2130107 | 2955842201100290 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-10 | 2011-09-02 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-08-10 | Device Not Retu | Yes | Initial | 2011
IS3000 A60P6 | 2011-08-02 | Other | Other | It was reported that prior to starting a da vinci si surgical procedure, the site experience a system error code stating that the patient side cart (psc) was not connected. With the assistance of an isi technical support engineer, the site applied emergency power off (epo) to the psc and swapped the blue fiber cable from the surgeon side cart (ssc) and the psc with no change. The site then epo'd the system again and cycled the breaker on the vision side cart with no change. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No patient harm or injury was reported. | The investigation conducted by the field service engineer concluded that system power issue was associated the set up joint (suj) wiring harness. The suj wire harness assembly is a set of electrical wires inside the setup joint which transmits power, control and sensor signals between circuit boards in the slave arms and the base of the patient side cart. The system was repaired by replacing the affected ipd pca board. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2252908 | 2232012 | 2121552 | 2955842201100285 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-08-02 | 2011-09-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-08-02 | Device Not Retu | No | Reuse | 2011
42034402 | 2011-08-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, a spark was observed coming from the curved bipolar dissector instrument. No patient harm, adverse outcome or injury was reported. Attempts to gain additional information from this site have been unsuccessful to date. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed burn marks at both grip tips and at the bases of the grip jaws. Melted material was also observed at the yaw pulley. The damage indicates that an arcing event occurred. Intuitive surgical internal laboratory testing has determined that the observed arc track failure mode is most likely to occur when a bipolar instrument is energized with the grips in a wide-open position with no tissue between the grips. Internal arcing of bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive surgical has concluded that the likelihood of injury due to this failure is remote. Users are instructed to retain back-up instruments in case an instrument fails to perform and are also instructed in proper technique to minimize the likelihood of internal arcing. | CURVED BIPOLAR DISSECTORINSTRUMENT | Reuse | 2252900 | 2232039 | 2128955 | 2955842201100286 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2011-08-02 | 2011-09-01 | No | Yes | Health Professional | 42034402 | Device Returned To Manufacturer | No | 2011-09-01 | Device Not Retu | No |  | 2011
IS2000 A51P8 | 2011-08-01 | Other | Other | It was reported that prior to starting a da vinci s mitral valve repair procedure, the surgical staff was unable to start the surgeon side console (ssc). With the assistance of an isi field service engineer (fse) the site attempted to emergency power off the system without success. The patient was under anesthesia when the surgeon made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering found that issue experienced by the customer was associated with the primary power supply(pps). The pps is a +48v power supply with battery backup and is mounted outside of the system control electronics chassis. The system was repaired by replacing the pps. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2250605 | 2230101 | 2127010 | 2955842201100284 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-01 | 2011-08-31 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-08-01 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-07-27 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the camera clutch pedal (located on the surgeon console) stopped working. The surgical staff power cycled the system and non-recoverable system error code 3 occurred when the system was powered back on. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 3 was associated with the user interface & status (uis) board. The uis performs multiple functions, one of which is to read and report the status of the 5 console foot switches. The system was repaired by replacing the affected uis. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 3 appears when the da vinci s safety system determines that a redundent switch was missing its ground sense, or the contacts did not report as expected at startup. Upon determining these conditions, the system records the fault and transitions to a safe state. The uis was returned to intuitive surgical for failure analysis investigation and engineering was able to replicate the reported failure mode during internal testing and evaluation. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245494 | 2224628 | 2121543 | 2955842201100275 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-27 | 2011-08-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-07-27 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-08-02 | Other | Other | It was reported that during a da vinci si prostatectomy procedure the surgical staff experienced system error code 297 and the leds on the illuminator and doco (double camera controller) turned orange. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the site reboot the illuminator and reseat the connection to the doco multiple times, however, the issue could not be resolved. The entire system was also rebooted with the same result. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was performed by a field service engineer (fse), who reviewed the system logs and observed error codes 297 and 48238. System error code 297 appears when the system detects that an electronic component was reporting an incorrect configuration. System error code 48238 indicates an illuminator communication issue. Upon determining these conditions, the system records the faults and transitions to a safe state. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The fse concluded that the issue experienced by the customer was associated with the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245497 | 2224632 | 2121547 | 2955842201100276 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-08-02 | 2011-08-26 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-08-02 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-08-05 | Other | Other | It was reported that during a da vinci partial nephrectomy procedure there was no video displayed on the left side monitor of the surgeon console. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. The hrsv was returned to intuitive surgical for evaluation. Engineering discovered that the left side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245498 | 2224634 | 2121549 | 2955842201100277 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-08-05 | 2011-08-26 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-08-05 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-08-04 | Injury | Life Threatening | It was reported that during a da vinci si hysterectomy procedure, an instrument would not advance on its own so the instrument arm clutch button was pressed to manually advance the instrument. When the button was pressed, the instrument advanced into the patient's perineum and nicked a superficial vein. The surgical staff removed all of the instruments and converted to open surgical techniques to address the patient injury and complete the planned procedure. The hospital confirmed that the patient was not adversely affected by the incident. | Investigation was conducted by a field service engineer, who reviewed the system logs and did not find any errors associated with the affected patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. As a precaution, the affected psm was replaced. The psm was returned to intuitive surgical for failure analysis investigation. Engineering observed that the linear bearings on link-5 are not aligning, thus causing excessive friction and binding on the axis-3 insertion drive. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245500 | 2224637 | 2121552 | 2955842201100278 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-08-04 | 2011-08-26 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-08-04 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-05-02 | Other | Hospitalization | It was reported that after a da vinci s low anterior resection procedure was successfully completed on (b)(6) 2011, the patient experienced ischemia reperfusion in his left leg on the same day. Cardiovascular acute incompetence then occurred as a result of the ischemia reperfusion. Due to the post surgical complications, the patient's hospitalization was extended. The patient recovered and was released on (b)(6) 2011. | On (b)(6) 2011, the console surgeon who performed the procedure indicated that the cause of the patient's ischemia reperfusion was due to the patient's body positioning during the case. The surgeon also reported that there were no malfunctions of the da vinci s surgical system during the procedure and that the patient's post surgical complications are unrelated to the use of the system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended. No information has been received which suggests that the system malfunctioned in a way that may have caused or contributed to the patient's post surgical complications. Intuitive surgical concluded that the patient's post surgical complications are unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245557 | 2224655 | 2121570 | 2955842201100279 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-07-28 | 2011-08-26 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-07-28 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-07-27 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the customer experienced a dark image in both eyes of the surgeon side console, high resolution stereo viewer. The surgeon made the decision to convert the procedure to traditional open surgical techniques. The planned procedure was completed and no patient harm was reported. | The investigation conducted by field service engineering (fse) found the site had set the camera controller unit (ccu) gain button to high which caused the dark image as reported by the customer. Adjustment of the ccu settings resolved the vision issue. As of (b)(6), 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245629 | 2224701 | 2121616 | 2955842201100280 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-07-27 | 2011-08-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-07-27 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-08-02 | Other | Other | It was reported that while preparing for a da vinci s prostatectomy procedure, the customer experienced a green tint in the left eye of the surgeon side console high resolution stereo viewer (hrsv). With the assistance of an isi technical support engineer (tse) the site replaced the camera cable and camera head, however the green tint continued to occur. The tse had the site restart the camera controller unit and attempt to black and white balance the image, however it was discovered that the camera cable was disconnected from the ccu during the initial camera cable exchange. The patient was under anesthesia for approximately 2. 5 hours when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer confirmed the customer reported problem was associated with the digitizer. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images. The system was repaired by replacing the affected digitizer. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245598 | 2224710 | 2127010 | 2955842201100281 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-08-02 | 2011-08-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-08-02 | Device Not Retu | No | Reuse | 2011
400180 | 2011-07-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2245636 | 2224727 | 2121642 | 2955842201100282 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-08-01 | 2011-08-26 | No | Yes | Health Professional | 400180 | No | Yes | 2011-08-01 | No | Yes | Unkown | 2011
IS3000 A60P6 | 2011-07-18 | Other | Other | It was reported that after da vinci si hysterectomy procedure the patient sustained a richter hernia. A portion of the small bowel herniated through the fascial defect left by the 8 mm cannula post procedure. The surgeon identified the problem and on (b)(6) conducted surgery to reset the herniated portion of the bowel. No additional patient harm, adverse outcome or injury was reported. | No malfunctions of the da vinci si surgical system, instrument or accessories were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients without issue. As of (b)(6), 2011, no similar instances have been reported to isi. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2245949 | 2224986 | 2121902 | 2955842201100283 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-07-28 | 2011-08-26 | Yes | No | Health Professional | IS3000 A60P6 | No | Yes | 2011-07-28 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-07-25 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, while changing instruments, the customer experienced system error code 23007. With the assistance of an isi technical support engineer (tse) the site restarted the system multiple times and exercised the right master tool manipulator (mtmr), however, the system error code reoccurred. The surgeon converted to traditional laparoscopic surgical techniques to complete the planned procedure, however to provide the patient with water tight anastomosus, the surgeon decided to convert to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23007 experienced by the customer was associated with the a gimal on the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23007 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The mtmr gimbal was returned to isi for failure analysis investigation. Engineering found the axis 7 bevel gear clamp to have a broken screw, which was stuck between the axis 7 gear causing it to stick. The axis 7 gear clamp was replaced. As of (b)(6), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2243257 | 2222297 | 2119213 | 2955842201100274 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-07-25 | 2011-08-24 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-07-25 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-07-25 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, while dissecting the patient's kidney, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The tip cover accessory was removed and a hole was found. The tip cover accessory was replaced to successfully complete the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2241767 | 2220922 | 2117838 | 2955842201100272 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-07-25 | 2011-08-23 | No | Yes | Health Professional | 40018010 | No | No | 2011-07-25 | No | Yes | Unkown | 2011
IS3000 A60P5 | 2011-07-24 | Other | Other | It was reported that at the end of a da vinci si gall bladder procedure, the surgeon inadvertently punctured the patient's liver. Repair was performed using cauterization and a coagulating gel. No additional patient harm was reported. | The investigation conducted by an isi representative concluded that system was working as intended with no reported or observed malfunctions. The surgeon is aware of the user related actions that contributed to this event and on (b)(4) 2011, additional training was conducted with the surgeon. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2241930 | 2221088 | 2118004 | 2955842201100273 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-07-24 | 2011-08-23 | Yes | No | Health Professional | IS3000 A60P5 | Yes | No | 2011-07-24 | Device Not Retu | No | Reuse | 2011
|  | Injury | Required Intervention | Sterile processing dept found cable sticking out of robotic cautery spatula when cleaning equipment after the case. |  | INTUITIVE SURGICAL | CAUTERY SPATULA | 2247626 | 2224423 | 2121338 | MW5021850 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial |  | 2011-08-23 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
400180 | 2011-07-15 | Other | Required InterventionHospitalization | It was reported that during a da vinci si surgical procedure and while dissecting tissue arcing was observed to have come from the monopolar curved scissors instrument with a tip cover accessory installed. The patient sustained an injury to their left iliac vein. The physician stopped the bleeding initially with a prograsp instrument and sutured the deficit using a prolene suture and vascular pledgets. The patient's hospitalization was extended one day as a precaution. No additional patient injury was reported. The delay in reporting is due to an administrative oversight where the complaint handling representative misfiled the report for reporting the day after it was due. | The tip cover accessory was not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2239095 | 2217861 | 2130106 | 2955842201100271 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-19 | 2011-08-19 | Yes | Yes | Health Professional | 400180 | No | Yes | 2011-07-19 | No | Yes | Unkown | 2011
42018305 | 2011-07-20 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, while using the permanent cautery hook instrument, a piece of the plastic guard on the distal end of the instrument broke off and fell into the patient. A portion of the fragment was retrieved; however, one piece was not located. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the hook broken off of the instrument. The hook had fractured within the yaw pulley, at the cross section of the hole feature. The detached hook was not returned. Engineering also observed that one side of the ceramic sleeve is broken off and missing. The size of the missing piece is approximately. 155 x. 075. It was determined that the damage to the hook and sleeve may be due to overloading at the tip or other mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMENANT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2239115 | 2217949 | 2114865 | 2955842201100270 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-07-20 | 2011-08-19 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2011-07-20 | Device Not Retu | No | Reuse | 2011
REF420227 | 2011-08-10 | Malfunction |  | Cable broke. |  | PK DISSECTING FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2244758 | 2223953 | 2120868 | 2223953 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-08-19 | 2011-08-19 | No | Yes | Invalid Data | REF420227 | Yes | No Answer Provided |  |  |  |  | 2011
REF420309 | 2011-08-11 | Malfunction |  | The cable broke on the davinci suturecut instrument while inside the patient. No visible parts lost. |  | MEGA SUTURECUT NEEDLE DRIVER | ENDOWRIST NEEDLE DRIVER INSTRU | 2244761 | 2223954 | 2120869 | 2223954 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-08-19 | 2011-08-19 | No | Yes | Physician | REF420309 | Yes | No Answer Provided |  |  |  |  | 2011
| 2011-06-17 | Injury | Required Intervention | Cautery spatula used during a lung procedure was missing part of the protective sheath on the end of the instrument. |  | DA VINCI S SURG | CAUTERY SPATULA | 2240828 | 2216943 | 2113859 | MW5021776 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-08-04 | 2011-08-16 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
| 2011-05-20 | Injury | Required Intervention | Part of protective sheath on the tip of the robotic spatula was missing. |  | INTUITIVE SURGICAL DAVINCI | CAUTERY SPATULA | 2240760 | 2216950 | 2113866 | MW5021774 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-08-04 | 2011-08-16 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
| 2011-07-28 | Injury | Required Intervention | The seal on the housing of the davinci robotic large needle driver broke. Three of the four clips that hold the housing together are broken. |  | INTUITIVE SURGI | ROBOTIC LG NEED | 2240791 | 2216951 | 2113867 | MW5021775 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-08-03 | 2011-08-16 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
42018305 | 2011-06-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, while using the permanent cautery hook instrument, a 'plastic' piece broke off and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument distal clevis ear to be broken off at the grip tip. No evidence of arc trac was found on the clevis or the cable. Engineering concluded that the instrument likely broke due to overloading at the hook causing the clevis ear to break. | PERMENANT CAUTERY HOOK INSTRUMENT | Reuse | 2233287 | 2212308 | 2109223 | 2955842201100269 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2011-08-17 | 2011-08-15 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2011-08-17 | Device Not Retu | No |  | 2011
42019408 | 2011-07-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the mega needle driver instrument, a piece fell off and into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury, or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the carbide insert to be broken off of one grip, the insert was returned separated from the instrument. No damage was found to the instrument clevis or cable. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2231447 | 2209896 | 2106812 | 2955842201100265 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-14 | 2011-08-12 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-07-14 | Device Not Retu | No | Reuse | 2011
IS2000 51P8 | 2011-07-13 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced system error 23013. The site restarted the system 4 times, however the system error continued to occur. With the assistance of an isi technical support engineer the site emergency powered off the patient side cart and burnished the endoscopic camera manipulator (ecm) arm, however the system continued to fault. The surgeon decided to convert to traditional surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error code 23013 experienced by the customer was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23013 system error code appears when the da vinci s safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The ecm was returned to isi for failure analysis investigation. Engineering was able to replicate the system error experienced by the customer and found the potentiometer to encoder were not matching on axis 2. The potentiometer and motor encoder assembly were replaced to repair the ecm. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2231431 | 2209897 | 2106813 | 2955842201100266 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-13 | 2011-08-12 | Yes | Yes | Health Professional | IS2000 51P8 | Device Returned To Manufacturer | Yes | 2011-07-13 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-07-19 | Other | Other | It was reported that during a da vinci si prostatectomy procedure the site experienced a non recoverable system error 5. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field engineering concluded that the system error code 5 experienced by the customer was associated with the generic power distribution (gpd) printed circuit assembly board (pca). The gpd pca converts the 48 volt supply from the power supply to other voltage levels and distributes it as appropriate to the other boards in the surgeon side cart. The system was repaired by replacing the addected gpd pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 5 appears when the da vinci si safety system determines one or more fans are not moving as desired. Upon determining this condition the safety system puts da vinci in a recoverable safe state. As of (b)(6) 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2231432 | 2209898 | 2106814 | 2955842201100267 | 1 | NAY | Manufacturer | ForeignDistributor | Nurse | Initial | 2011-07-19 | 2011-08-12 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-07-19 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-07-15 | Other | Other | It was reported that while preparing for a da vinci prostatectomy procedure, the customer experienced system error code 209. The patient was under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 209 experienced by the customer was associated with a motor channel on a motor servo driver (msd) board. The msd is a printed circuit assembly which powers, controls and monitors the motors on an instrument arm, with one channel for each motor. The system was repaired by replacing the affected multiple servo drive (msd) board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 209 appears when the da vinci safety system determines that the a fault occurred on one or more motor channels on a msd board. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2231635 | 2210056 | 2106972 | 2955842201100264 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-07-15 | 2011-08-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-07-15 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-07-13 | No Answer Provided | Other | It was reported that during a da vinci s sacrocolpopexy procedure, the site experienced difficulty in moving a patient side manipulatory (psm) arm. With the assistance of an isi technical support engineer (tse), the site reseated the sterile adaptor; however, the system did not put the instrument automatically in following mode. The surgical staff decided not to continue with troubleshooting the reported problem and converted to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by tse could not reproduce the customer reported issue. The tse contacted the surgical staff after the procedure to discuss the event. The staff recycled power to the unit and the system issue was resolved. Additional testing performed by the staff did not reveal any system errors or problems with the system. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2231983 | 2210818 | 2107751 | 2955842201100268 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-07-13 | 2011-08-12 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-07-13 | Device Not Retu | No | Reuse | 2011
42017909 | 2010-04-28 | Malfunction |  | Event Description: --/-- On july 11, 2011, a copy of cmdsnet sentinel report (b)(4) was received from (b)(6). The following event description was provided in the report: robotic monopolar curved scissors broke inside the patient. Piece retrieved inside the patient. Instrument kept for the company to assess if there was any problem with them. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Several requests for additional information have been made to the hospital, along with multiple requests for the instrument's return. To date, the instrument has not been returned and no additional information has been provided. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The affected instrument was received for evaluation on (b)(4) 2010. At the time of receipt, the reported malfunction was a snapped blade and the reporter indicated that no instrument pieces fell into the patient. During evaluation, engineering observed one scissor blade broken off of the instrument. The blade had fractured at the cross section of the cable crimp and the fracture plane is straight. The broken blade was returned with the instrument. The intact blade does not exhibit any damage at the tip. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2222328 | 2200255 | 2097171 | 2955842201100262 | 1 | NAY | Manufacturer | Foreign | Other | FollowupInitial | 2011-07-11 | 2011-08-11 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2011-07-11 | Device Not Retu | No |  | 2011
IS3000 A60P6 | 2011-07-11 | No Answer Provided | Other | It was reported that during a da vinci si cystectomy procedure, the site experienced trouble engaging the sterile adapter on the system's patient side manipulator (psm). With the assistance of an isi technical support engineer (tse), the site attempted to engage the sterile adapter several times before the sterile adapter was successfully engaged. The site continued on with the procedure; however, the issue reoccurred. With the assistance of a tse, the site performed an emergency stop and fault override without success. The surgeon decided not to continue with troubleshooting the reported problem and converted to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering could not reproduce the sterile adapter issue. The system was evaluated for any potential issues which could have caused or contributed to this issue, however, the system was found to be operating within specification with no malfunctions observed. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2222477 | 2200416 | 2097332 | 2955842201100263 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-07-11 | 2011-08-11 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-07-11 | Device Not Retu | No | Reuse | 2011
| 2011-07-18 | Injury | Required Intervention | Sterile processing dept inspected da vinci cautery spatula after it was used. Found one of the cables on the articulating tip was no longer seated properly and was sticking out. |  | DA VINCI S SURG | CAUTERY SPATULA | 2235570 | 2211486 | 2091998 | MW5021705 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-07-31 | 2011-08-10 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
| 2011-08-03 | Malfunction |  | Failure of the third operating arm of the da vinci si robot. The intuitive rep was present and on the phone with tech support, but was unable to recover the fault the robot was indicating. |  | DAVINCI | DA VINCI SURGICAL SYSTEM | 2239636 | 2213255 | 2110171 | MW5021721 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-08-10 | 2011-08-10 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
42011704 | 2011-07-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the surgeon was having difficulty getting a good angle with the 5 mm needle driver instrument (used with the omniguide laser) while trying to manipulate the patient's large fibroid when the plastic coating of the instrument became compromised and pieces of the coating fell into the patient. Some of the pieces were retrieved, however, it cannot be confirmed if all pieces were retrieved. The instrument was removed and replaced with an instrument of the same type, however, the same issue recurred with the replacement instrument. Some pieces of the second instrument were retrieved, however, it cannot be confirmed if all pieces were retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00 was submitted for the second 5 mm needle driver instrument used during the procedure. --/-- Additional Event Description: --/-- It was reported that during a da vinci si myomectomy procedure the surgeon was having difficulty getting a good angle with the 5 mm needle driver instrument (used with the omniguide laser) while trying to manipulate the patient's large fibroid when the plastic coating of the instrument became compromised and pieces of the coating fell into the patient. Some of the pieces were retrieved, however, it cannot be confirmed if all pieces were retrieved. The instrument was removed and replaced with an instrument of the same type, however, the same issue recurred with the replacement instrument. Some pieces of the second instrument were retrieved, however, it cannot be confirmed if all pieces were retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00261 was submitted for the second 5 mm needle driver instrument used during the procedure. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various scratch marks on the distal end of the main tube with light material removed and a rough surface finish. --/-- Additional Manufacturer Narrative: --/-- | 5 MM NEEDLE DRIVER INSTRUMENT | Reuse | 2218870 | 2195565 | 2092501 | 2955842201100260 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2011-07-13 | 2011-08-08 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2011-07-13 | Device Not Retu | No |  | 2011
42011704 | 2011-07-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure the surgeon was having difficulty getting a good angle with the 5 mm needle driver instrument (used with the omniguide laser) while trying to manipulate the patient's large fibroid when the plastic coating of the instrument became compromised and pieces of the coating fell into the patient. Some of the pieces were retrieved, however, it cannot be confirmed if all pieces were retrieved. The instrument was removed and replaced with an instrument of the same type, however, the same issue recurred with the replacement instrument. Some pieces of the second instrument were retrieved, however, it cannot be confirmed if all pieces were retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00260 was submitted for the initial 5 mm needle driver instrument used during the procedure. --/-- Additional Event Description: --/-- It was reported that during a da vinci si myomectomy procedure the surgeon was having difficulty getting a good angle with the 5 mm needle driver instrument (used with the omniguide laser) while trying to manipulate the patient's large fibroid when the plastic coating of the instrument became compromised and pieces of the coating fell into the patient. Some of the pieces were retrieved, however, it cannot be confirmed if all pieces were retrieved. The instrument was removed and replaced with an instrument of the same type, however, the same issue recurred with the replacement instrument. Some pieces of the second instrument were retrieved, however, it cannot be confirmed if all pieces were retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00 was submitted for the initial 5 mm needle driver instrument used during the procedure. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed various scratch marks on the distal end of the main tube with light material removed and a rough surface finish. | 5 MM NEEDLE DRIVER INSTRUMENT | Reuse | 2218881 | 2195573 | 2092501 | 2955842201100261 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2011-07-13 | 2011-08-08 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2011-07-13 | Device Not Retu | No |  | 2011
PK DISSECTING | 2011-06-22 | Malfunction |  | (b)(4) dissecting forcep spring wire separated causing malfunction of device. |  | DAVINCI PK DISSECTING FORCEP | PK DISSECTING FORCEP | 2233588 | 2209606 | 2106522 | MW5021698 | 1 | NAY | Voluntary |  | Other | Initial | 2011-08-08 | 2011-08-08 | No | Yes | Health Professional | PK DISSECTING | Yes | Yes |  |  |  |  | 2011
IS2000 A51P8 | 2011-07-08 | Other | Other | It was reported that while preparing to begin a da vinci s prostatectomy procedure, the endoscopic camera manipulator (ecm) cannula mount broke at the base connection point into the ecm. The patient side cart was docked to the patient and the camera was being inserted into the ecm when the cannula mount broke. The site did not have a spare ecm cannula mount. The patient was under anesthesia and the port incisions were made when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm was reported. | Isi field service engineering provided the site additional training on how to replace the ecm cannula mount and the part number to obtain spares for their inventory. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The cannula mount accessory is attached to the ecm and designed to hold the cannula securely in place, outside of the patient cavity. The system was repaired by replacing the affected ecm cannula mount accessory. The da vinci s surgical system user's manual explicitly states that: environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have back up equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2011 there have been no reported recurrences od the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2216972 | 2193027 | 2089963 | 2955842201100253 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-07-08 | 2011-08-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-07-08 | Device Not Retu | No | Reuse | 2011
42011704 | 2011-06-29 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, while using the 5mm needle driver instrument, the instrument's arm broke off and fell into the patient's abdominal cavity. The fragment was immediately removed and the instrument was replaced with another of the same type. The procedure continued with the replacement 5mm needle driver instrument, however, the replacement instrument's arm broke in a similar fashion as the initial instrument. The second fragment was removed and the planned surgical procedure was completed with a third replacement 5mm needle driver instrument. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda and is related to this event: 29558842-2011-00xxx. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip to be broken off and missing. The grip fractured at the junction between the grip arm and grip hub, where there is a 90 degree angle. The fractured surface exhibits a reddish / orange substance, possibly bio debris. The remaining grip does not exhibit damage. Per the case notes, the instrument fragment was successfully removed from the patient. As of august 4, 2011 there have been no reported recurrences of the issue at this hospital. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2216980 | 2193028 | 2089987 | 2955842201100258 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-06 | 2011-08-05 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | Yes | 2011-07-06 | Device Not Retu | No | Initial | 2011
IS2000 A51P8 | 2011-07-08 | Other | Other | It was reported that while preparing to start a da vinci s prostatectomy procedure, the site experienced loss of video in the left eye of the high resolution stereo viewer. With the assistance of an isi technical support engineer (tse), the site restarted the camera controller unit (ccu) and the vision was restored to normal. Approximately 4 hours later, the site called back stating that the right eye of the high resolution stereo viewer appeared distorted. The tse discovered that the site had reversed the ccu cable connections. The site unplugged and reseated the ccu connection and the image appeared normal. The tse then continued to assist the site with the vision set up and found the hrsv left eye become black. The patient was under anesthesia for approximately 2 hours and the port incisions were made when the surgeon decided to abort the planned procedure. No patient harm was reported. | The investigation conducted by an isi field service engineer found that the console vision cable was defective and difficult to insert into the vision side cart. The console vision cable connects the vision cart to the surgeon side console (ssc) and passes the video signals between the ssc and the vision cart. The system was repaired by replacing the console vision cable. As of (b)(4) 2011 there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2216951 | 2193047 | 2089983 | 2955842201100254 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-08 | 2011-08-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2011-07-08 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-07-12 | No Answer Provided | Other | It was reported that at the beginning of a da vinci si hysterectomy procedure, the site experienced recurring system error codes 23013. With the assistance of an isi technical support engineer, the site exercised the right master tool manipulator (mtmr) and restarted the system, however, the system error 23013 returned. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23013 experienced by the customer was associated with the right master tool manipulator (mtmr) gimball. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr gimball. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23013 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The mtmr gimball was returned to isi for evaluation. Engineering found a broken gear clamp screw, thus generating the system error code experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2216977 | 2193049 | 2089985 | 2955842201100255 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-12 | 2011-08-05 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-07-12 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-07-08 | Malfunction |  | It was reported that during a da vinci s pulmonary wedge resection procedure, while using the permanent cautery spatula instrument, the site noticed a small black object inside the patient. The surgeon removed the object and believed it to be part of the permanent cautery spatula instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument to be returned with a dark red and brown object approximately. 095 inches by. 095 inches in size and separated from the instrument. Visual inspection of the object under magnification concluded that the object is not part of the instrument, however, it is a charred piece of bio debris. Engineering also observed the spatula exhibits melting and thermal deformation at the distal end and the surface has a dark discoloration. The ceramic sleeve was found to be broken at the distal end, exposing a section of the spatula shank. Broken pieces were not returned with the instrument. The proximal clevis was found to be cracked between the ears, with the cracking observed consistent with overloading at the tip. Engineering concluded the damage found was likely due to mishandling and or misuse of the instrument. Per the case notes all fragments were retrieved from the patient during the procedure. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2216978 | 2193050 | 2089986 | 2955842201100256 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-07-08 | 2011-08-05 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-07-08 | Device Not Retu | No | Reuse | 2011
42011704 | 2011-06-29 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, while using the 5mm needle driver instrument, the instrument's arm broke off and fell into the patient's abdominal cavity. The fragment was immediately removed and the instrument was replaced with another of the same type. The procedure continued with the replacement 5mm needle driver instrument, however, the replacement instrument's arm broke in a similar fashion as the initial instrument. The second fragment was removed and the planned surgical procedure was completed with a third replacement 5mm needle driver instrument. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda and is related to this event: 29558842-2011-00xxx. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one grip to be broken off and missing. The grip fractured at the junction between the grip arm and grip hub, where there is a 90 degree angle. The fractured surface exhibits a reddish / orange substance, possibly bio debris. The remaining grip does not exhibit damage. Per the case notes the instrument fragment was successfully removed from the patient. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2216979 | 2193051 | 2089987 | 2955842201100257 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-06 | 2011-08-05 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | Yes | 2011-07-06 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-18 | Other | Other | On (b)(6) 2011 (b)(4) reported that post a da vinci s hysterectomy procedure, which took place on (b)(6) 2011, the patient returned to the hospital with complications and was found to have a bowel perforation. | On (b)(4) 2011, (b)(4) hospital (b)(4), stated that the patient injury was not believed to be a result of a da vinci s system malfunction. Intuitive surgical has made multiple attempts to obtain detailed information concerning the patient, their diagnosis, pre-existing medical conditions and / or any other related additional information. However, on (b)(4) 2011, (b)(4), stated that she is unable to provide any additional information about this case. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2216992 | 2193069 | 2090005 | 2955842201100259 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-07-07 | 2011-08-05 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-07-07 | Device Not Retu | No | Reuse | 2011
420227 | 2011-08-03 | Malfunction |  | Davinci forcep cable broke in jaw of item. |  | PK DISSECTING FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2221910 | 2199856 | 2096772 | 2199856 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-08-05 | 2011-08-05 | No | Yes | Physician | 420227 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-07-05 | Other | Other | It was reported that during a da vinci s hysterectomy procedure the site experienced a backwards image on the surgeon side console monitor. With the assistance of an isi technical support engineer (tse) it was discovered that the site did not have the endoscope turned on. The site powered on the endoscope and experienced a pink and green image. The tse had the site perform scope alignment and the image appeared normal. The site continued with the procedure. Approximately 90 minutes later the site called back stating they were experiencing a green image in the right eye of the high resolution stereo viewer. The tse had the site replace the camera cable, however the green image remained. The tse had the site replace the camera head and the image appeared normal. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. Uf medwatch report (b)(4) was received (b)(6) 2011 regarding this event. | The investigation conducted by isi field service engineering concluded that the issue experienced by the customer was associated with the camera head. The camera produces a three dimensional image to the vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The da vinci s surgical system user's manual explicitly states that: environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have back up equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. On (b)(6) 2011 the isi clinical sales representative and field service engineer performed in-service training with the site's surgical team. This training included troubleshooting techniques and component exchanges. As of (b)(6) 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURIGCAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2215280 | 2190788 | 2087724 | 2955842201100252 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-07-05 | 2011-08-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-07-05 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-07-04 | Other | Other | It was reported that prior to performing any surgical tasks for a da vinci prostatectomy procedure, no image was displayed in the surgeon console viewer or the assistant monitor. A technical support engineer attempted to troubleshoot the issue via telephone, however, the issue could not be resolved. The patient had been under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | Investigation was conducted by a field service engineer (fse), who went onsite and evaluated the system. The fse determined that the vision cable had been improperly connected, thus resulting in the vision issue experienced by the customer. As of august 2, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2214077 | 2189468 | 2086404 | 2955842201100251 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-07-04 | 2011-08-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2011-07-04 | Device Not Retu | No | Reuse | 2011
| 2011-07-19 | Malfunction |  | Robotic cautery spatula was used to divide up a uterine fibroid so that it could be removed vaginally. During division of specimen, there was arcing from the spatula over the protective sheath to the plastic portion of the instruments where the cables for articulation are located. The plastic on the end of the robotic cautery spatula was damaged, appears charred and melted. No harm to pt as fibroid had already been dissected away from pt. |  | DA VINCI S SURG | ROBOTIC CAUTERY SPATULA | 2223713 | 2195062 | 2091998 | MW5021613 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-07-26 | 2011-08-03 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2011
| 2011-07-13 | Malfunction |  | During robotic portion of case, md noted piece of 5mm cannula seal broke off during insertion of robotic instrument. |  | 5MM CANNULA SEAL | ENDOWRIST 5MM ACCESSORY INSTRU | 2221911 | 2199857 | 2096773 | 2199857 | 1 | NAY | User Facility |  | Nurse | Initial | 2011-08-02 | 2011-08-02 | No | Yes | Physician |  |  | Yes |  |  |  |  | 2011
400309 | 2011-07-28 | Malfunction |  | Davinci 8mm mega suturecut needle driver endowrist instrument. Tip broken during use. This is a reposable instrument with 8 out of 10 uses left. No harm to patient. |  | MEGA SUTURECUT NEEDLE DRIVER | ENDOWRIST NEEDLE DRIVER INSTRU | 2221929 | 2199866 | 2096759 | 2199866 | 1 | NAY | User Facility |  |  | Initial | 2011-08-01 | 2011-08-01 | No | Yes | Physician | 400309 | Yes | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-05-10 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the site observed arcing coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The surgical staff observed burns that were local into fatty tissue in the bladder region. The tip cover accessory was changed to successfully complete the surgery. The site reported that the patient was not injured beyond the local charring at the level of the distal tip and that the burns did not require any added intervention. The patient was reported as recovering normally and no follow up visit was required as result of burns. | The tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found the tip cover to have multiple holes burned through the silicone. The silicone exhibits localized melting around the hole which is indicative that an arcing event occurred. The hole sizes are roughly. 015 and. 050, and are located. 220 and. 330 above the silicone to sleeve interface. The tip cover also has various mechanical tears and scratches in the general vicinity of the holes. The site also returned a video of the procedure where various collisions between silicone part of tip cover and other instruments can be observed prior to arcing. Evidence not conclusive, however, the instrument collisions with the tip cover accessory have likely contributed to the arcing event. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings to handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The monopolar curved scissors instrument instructions for use. Inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2209971 | 2184632 | 2081568 | 2955842201100250 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-30 | 2011-07-30 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-06-30 | Device Not Retu | Yes | Initial | 2011
IS3000 A60P6 | 2011-06-30 | Other | Other | It was reported that during a da vinci si prostatectomy procedure the vision cart power turned off and the cart could not be restarted. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the power issue experienced by the customer was related to the vision cart's power supply. The system was repaired by replacing the affected power supply. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2209402 | 2184097 | 2081032 | 2955842201100249 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-06-30 | 2011-07-29 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2011-06-30 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-06-30 | Other | Other | It was reported that during a da vinci s sternotomy lima-lad procedure, the instruments installed on one of the instrument arms did not open/close all the way and wrist movement was sluggish. As a result, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by a field service engineer (fse), who went onsite and evaluated the affected instrument arm. Cable tension was measured and found to be within specifications and the sine cycle test passed with no issues. The fse performed a system verification test and drove the system with the sterile adapters used during the procedure. No issues were found and the fse was unable to replicate the instrument issues experienced by the customer. The hospital has continued to perform procedures using the system and as of july 28, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2209371 | 2184116 | 2081052 | 2955842201100248 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-30 | 2011-07-29 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2011-06-30 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-06-30 | Malfunction |  | It was reported that approximately one hour into a da vinci si hysterectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The planned procedure was completed, and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover to have various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The hole sizes are. 015 inches and. 020 inches and are located from. 190 inches to. 295 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears in the vicinity of the holes, including at the edge of the holes. The plastic sleeve has two cracks in the axial direction at the proximal end. No other damage was found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2208026 | 2182455 | 2079391 | 2955842201100243 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-30 | 2011-07-28 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-06-30 | Device Not Retu | Yes | Initial | 2011
400169 | 2011-05-26 | Malfunction |  | (b)(4). Event desc: patient undergoing prostatectomy via robotic assistance. Surgeon was using the harmonic curved shears insert for the davinci when, during the procedure, he noticed the tip of the harmonic had come off. Harmonic tip noted to be off at approximately 3 hours into the surgery. It was all retrieved and surgery was completed. Patient sent to recovery. Device usage problem: device malfunction - that is, the device did not do what it was supposed to do. | Multiple attempts with the site have been made; however, at the time of this report, the accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned and/or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2208126 | 2182600 | 2034583 | 2955842201100247 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-06-28 | 2011-07-28 | No | Yes | Health Professional | 400169 | No | No | 2011-06-28 | No | Yes | Unkown | 2011
42019408 | 2011-06-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was suturing with the mega needle driver instrument, the metal plate below the jaw fell off into the patient. The piece was retrieved and the planned surgical procedure was completed with no patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument to be returned with the carbide insert detached from one grip. A small piece of the corner of the insert remains attached to the grip. The brazing surface of grip has some voids with missing filler material and the outer surface of damaged grip exhibits scratch marks. No other damage found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2208194 | 2182675 | 2079611 | 2955842201100246 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-28 | 2011-07-28 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-06-28 | Device Not Retu | No | Reuse | 2011
420117 | 2011-07-13 | Malfunction |  | The surgeon was using the omni guide laser in conjunction with the 5mm needle driver. When dedocking the robot, the needle driver was removed, it was identified that the coating on the instrument shaft had melted and pieces had fallen into the patient which were later recovered. |  | NEEDLE DRIVER | ENDOWRIST 5MM INSTRUMENT | 2221912 | 2199858 | 2096774 | 2199858 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-07-28 | 2011-07-28 | No | Yes | Physician | 420117 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
42011704 | 2011-06-21 | Malfunction |  | It was reported that while the surgeon was sewing the ureter during a da vinci s pyeloplasty procedure, one end of the 5mm needle instrument broke off and fell into the patient. The broken piece was successfully retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation despite several requests. As a result, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned, or if additional information is received. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2204956 | 2179459 | 2076395 | 2955842201100241 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-23 | 2011-07-27 | No | Yes | Health Professional | 42011704 | No | No | 2011-06-23 | No | No | Unkown | 2011
42022703 | 2011-06-20 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while removing the pk dissecting forceps instrument from use, the surgical staff observed the tip of the instrument to be broken. Pieces of the instrument were visualized inside of the patient and removed by the patient side assistant. The planned surgical procedure was completed with a replacement instrument of the same type and no patient injury, harm or adverse outcome was reported. The delay in reporting is due to a miscommunication between the site and the isi representative present during the procedure. The site advised they would report the event; however, they did not report the complete details of the event to our customer service department until july 22, 2011, which was beyond the original 30 days for mdr reporting. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire to be broken at the yaw pulley exit. The wire is detached from its connection at the grip. Electrical continuity failed. The yaw pulley shows no signs of arcing; however, the yaw pulley is missing a. 150 x. 060 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2205315 | 2179817 | 2076753 | 2955842201100244 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-20 | 2011-07-27 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2011-06-20 | Device Not Retu | No | Reuse | 2011
420194 | 2011-07-26 | Malfunction |  | Operating staff report that while surgeon was utilizing an 8mm mega needle driver in a robotic assisted surgery one of the cables broke. The cable evidently moves the tip of the device when handpiece is triggered (moved). You can visualize the cable sticking out the tip of the device. The device is evidently meant to have "10 lives. " after each case the device is sent to our sterile processing department and re-used until the 10 uses have expired. The device in question was on its 10th use. No injury to patient; surgeon confident no pieces of the cable entered the surgical site. Device sequestered and removed from field. New device of same type/model obtained and used to complete the case with no further issues. |  | MEGA NEEDLE DRIVER FOR DAVINCI SYSTEM | ENDOWRIST INSTUMENT NEEDLE DRI | 2213343 | 2188545 | 2085481 | 2188545 | 1 | NAY | User Facility |  |  | Initial | 2011-07-27 | 2011-07-27 | No | Yes | Physician | 420194 | Yes | No Answer Provided |  |  |  |  | 2011
42000606 | 2011-06-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while using the large needle driver instrument, a piece of the tip broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned, and/or if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2202723 | 2176755 | 2073691 | 2955842201100235 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-22 | 2011-07-26 | No | Yes | Health Professional | 42000606 | No | Yes | 2011-06-22 | No | No | Unkown | 2011
42000606 | 2011-06-22 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, after the surgeon drove a suture through the vaginal cuff and tied the first knot, a metal fragment was noted on the patient's sigmoid in the pelvis. The fragment was immediately removed and the instruments used were inspected. The site found the large needle driver instrument to be missing a piece of the grip layer on the right side. The large needle driver instrument was replaced with another instrument of the same type and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument was returned with the carbide insert broken off of one grip. No damage was found on the clevis or cable. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2204502 | 2179005 | 2075941 | 2955842201100238 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-22 | 2011-07-26 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2011-06-23 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-06-28 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the site experienced system error code 23002. The site was able over ride the fault and continue with the case, however the fault continued to occur. With the assistance of an isi technical support engineer the site restarted the system and repositioned the master tool manipulator (mtm) , however, the system error code 23002 continued to occur. The surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23002 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23002 system error code appears when the da vinci s safety system determines that the angular position of one of the joints, as measured by the primary sensor, is outside the range of acceptable values. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The mtmr was returned to isi for failure analysis investigation and the fault experienced by the customer was duplicated. The flex circuit assembly was replaced to address the system error code 23002. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2204507 | 2179010 | 2075946 | 2955842201100242 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-06-28 | 2011-07-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-06-28 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-06-27 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the permanent cautery hook instrument, the surgical staff noticed the yellow plastic tip fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument grip to be missing the conductor wire cap at the yaw pulley. No damage was found to the pulley or cable, however, one of the grip cables was derailed from the pulley. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2204615 | 2179118 | 2076054 | 2955842201100239 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-27 | 2011-07-26 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | Yes | 2011-06-27 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-06-23 | Other | Other | It was reported that post a da vinci si hysterectomy procedure, after undocking the system from the patient, the surgical staff found burn marks on the patient's skin surrounding the port incisions where the cannula accessories were located. The patient required no additional medical treatment and no additional patient harm, adverse outcome or injury was reported. | Isi field service engineering contacted the site's biomedical engineering group and recommended that the electrical surgical unit (esu) used for the procedure be tested for rf leakage as well as the cables connecting from the esu to the instruments be exchanged. No system errors were found in the system error log. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(6) 2011, the site has continued to use the da vinci si system and there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2204657 | 2179160 | 2076096 | 2955842201100237 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-23 | 2011-07-26 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-06-23 | Device Not Retu | No | Reuse | 2011
42018406 | 2011-06-09 | Malfunction |  | It was reported that while the surgeon was performing a lobectomy dissection during a da vinci s lung biopsy procedure, a piece of the permanent cautery spatula instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument ceramic sleeve was damaged and pieces were missing. As a result, the base of the spatula is exposed. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2204667 | 2179170 | 2076106 | 2955842201100236 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-23 | 2011-07-26 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2011-06-23 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-06-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was suturing the vaginal cuff and while using the mega needle driver instrument, the insert from the mega needle driver instrument came off and fell into the patient. The fragment was immediately retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument was returned with a broken carbide insert at one of the grips. The insert was broken near the midpoint of the grip. No damage was found on the other grip or at the instrument clevis. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2204762 | 2179265 | 2076201 | 2955842201100240 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-30 | 2011-07-26 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-06-30 | Device Not Retu | No | Reuse | 2011
IS2000 A51P7 | 2011-06-20 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the surgical staff experienced four occurrences of system error code 45049. For each occurrence, a system restart was required to clear the error. After the fourth occurrence, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. The system had been in use for approximately 6 hours when the conversion occurred. No patient harm was reported. | Investigation was performed and it was concluded that system error code 45049 was associated with an update made to the domain name system (dns) configuration on the internet used by da vinci s onsite to communicate with the onsite server. Onsite for the da vinci s surgical system is an optional upgrade that enables field service engineers to obtain system information remotely for the purpose of remotely diagnosing system issues. To communicate with the onsite server infrastructure, the da vinci s onsite hardware is connected to the hospital internet protocol (ip) network based on the network parameters provided by the hospital information technology department. The dns is used to translate server names to ip addresses. The da vinci s surgical system talks to the hospital dns server and performs a dns lookup to obtain the ip address of the onsite server. When the update was made to the dns configuration on the internet, the response received from the hospital dns server caused a memory error to occur. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 45049 appears when the da vinci s safety system detects a communication timeout with the software running the da vinci onsite application. When the memory error occurred the software was unable to respond and the communication timeout was detected. Upon determining these conditions, the system records the fault and transitions to a safe state. The dns update was reverted back to its prior configuration on (b)(6) 2011, and as of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2201483 | 2175210 | 2072146 | 2955842201100234 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-20 | 2011-07-25 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2011-06-20 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-06-26 | Death | Death | It was reported that after a da vinci sacrocolpopexy with hysterectomy procedure was completed on (b)(6) 2011, the surgeon returned to the hospital later on the same day to remove the pt's y-mesh using open surgical techniques. The reason for removing the y-mesh is unk. On (b)(6) 2011, the pt experienced severe pain and it was determined by the medical staff that the pt had developed sepsis due to a perforated bowel. Subsequently, the pt expired on (b)(6) 2011. | A video of the da vinci si procedure was reviewed by a hospital physician. It was noted that the surgeon did not close the pt's peritoneum. The physician informed intuitive surgical that no issues were observed during the case. The physician also indicated that there were no system, instrument or accessory issues experienced by the surgical staff during the procedure. To the best of the physician's knowledge, there were no pt complications during the surgical procedure. The hospital is currently performing an investigation of the event. To date, it is unk how the damage to the pt's bowel was introduced and the exact cause of death is also unk. A follow-up medwatch report will be submitted if additional info is received. | DE VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2219454 | 2192397 | 2089333 | 2955842201100245 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-06-25 | 2011-07-25 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-06-25 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-06-20 | Other | Other | It was reported that during the beginning of a da vinci si prostatectomy procedure, the site experienced error code 23017 and the endoscopic camera manipulator's(ecm) led was illuminated yellow. With the assistance of a isi technical support engineer (tse), the site attempted to troubleshoot the issue by applying emergency power off (epo); however, non recoverable system error code 25588 occurred. At this time, the decision was made to perform the surgery using traditional open techniques. There was no patient injury reported. | The investigation conducted by field service engineering found system error code 23017 and 25588 in the system error logs. It was determined that these errors were associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. System error code 25588 appears when the da vinci s safety system determines that a flat flex cable connecting to the embedded serializer in master platform (espm) board was not fully inserted. Upon determining this condition, the safety systems put da vinci in a non-recoverable safe state. The ecm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and replaced the motor encoder and fully inserted the flat flex cable into the espm to repair the ecm. As of july 18, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2200117 | 2173978 | 2070914 | 2955842201100231 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-20 | 2011-07-24 | Yes | No | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-06-20 | Device Not Retu | No | Reuse | 2011
400180 |  | Other |  | It was reported that during a da vinci si surgical procedure, arcing was observed to have come from the monopolar curved scissors instrument with a tip cover accessory installed. No patient injury was reported. | The tip cover accessory was not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2200144 | 2173988 | 2070924 | 2955842201100233 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-19 | 2011-07-24 | No | Yes | Health Professional | 400180 | No | No | 2011-06-19 | No | Yes | Unkown | 2011
400180 |  | Other |  | It was reported that during a da vinci si surgical procedure, arcing was observed to have come from the monopolar curved scissors instrument with a tip cover accessory installed. No patient injury was reported. | The tip cover accessory was not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2200139 | 2174002 | 2070936 | 2955842201100232 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-19 | 2011-07-24 | No | Yes | Health Professional | 400180 | No | No | 2011-06-19 | No | Yes | Unkown | 2011
IS2000 A51P7 | 2011-06-15 | Other | Other | It was reported that during a da vinci s arrested heart revascularization procedure one of the instrument arms was unresponsive when an instrument was installed. While discussing the issue with an intuitive surgical technical support engineer via telephone, the system alarmed and the patient cart switched to battery. While troubleshooting these issues, the surgeon experienced cautery issues and decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | An intuitive surgical field service engineer (fse) went onsite and evaluated the system. Prior to arrival, the patient cart (pc) had been plugged into an electrical outlet, powered on for approx 2 hours, and did not lose power during this time period. The system was moved into the operating room (or) and plugged into the same electrical outlets that were used during the surgical procedure. The fse confirmed that the battery was fully charged, reseated the power cables and confirmed that the lockdown screws holding the power cables were tight. Proper power was verified. The power cables were also visually inspected with no issues found. The system was driven multiple times and the fse could not duplicate the customer complaint of power loss. The pc remained powered up for 2 more hours with no power failure. The or staff had mentioned to the fse that the pc went into battery mode when moving the pc to the patient. Per this information, the fse determined that the power cable may have been pulled from the wall during the move, resulting in the pc switching to battery mode. Regarding the instrument recognition issue reported by the customer, the fse evaluated all of the instrument arms and no problems were found. As a precaution, the fse reviewed proper instrument sterile adapter installation with the or staff. Per the cautery issues experienced by the surgeon, the surgical staff indicated that those issues were related to their electrosurgical unit. As of (b)(6) 2011, the site has continued to perform procedures and there have been no reported recurrences of the issues at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2193118 | 2166712 | 2063648 | 2955842201100226 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-15 | 2011-07-18 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2011-06-15 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-06-15 | Other | Other | It was reported that while beginning a da vinci s prostatectomy procedure, the surgical staff lost the image in the surgeon side console, high resolution stereo viewer (hrsv). With the assistance of an isi technical support engineer the site changed the camera controller unit settings, reseated the vision cable and restarted the system. The system powered on and the hrsv image returned, however, the surgeon decided to abort the planned procedure. The patient was under anesthesia and the port incisions had been made when the surgeon decided to reschedule the procedure for a later date. No patient harm, injury or adverse outcome was reported. | The investigation conducted by isi field service engineering found that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv provides the video image for the surgeon side console (ssc). The system was repaired by replacing the affected hrsv. The hrsv was returned to isi for evaluation. Engineering was unable to replicate the customer report vision issue, however as a precaution the left and right monitors were replaced. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2193461 | 2167339 | 2064275 | 2955842201100227 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-15 | 2011-07-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-06-15 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-27 | Other | Other | On (b)(6) 2011, medwatch uf/importer report (b)(4) was received: while separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. | Multiple attempts with the site have been made, however at the time of this report, the monopolar cautery scissors instrument and tip cover accessories have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument or accessories are returned and / or if additional information is received. The following medwatch mfr report was also sent to the fda pertaining to this event: 2955842-2011-00229. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2193439 | 2167357 | 2064293 | 2955842201100228 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2011-06-15 | 2011-07-18 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-06-15 | No | Yes | Unkown | 2011
40018010 | 2011-05-27 | Other | Other | On (b)(4) 2011, medwatch uf/importer report (b)(4) was received: while separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. | Multiple attempts with the site have been made, however at the time of this report, the monopolar cautery scissors instrument and tip cover accessories have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument or accessories are returned and / or if additional information is received. The following medwatch mfr report was also sent to the fda pertaining to this event: 2955842-2011-00228. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2193464 | 2167359 | 2064293 | 2955842201100229 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2011-06-15 | 2011-07-18 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-06-15 | No | Yes | Unkown | 2011
42019408 | 2011-05-31 | Other | Other | It was reported that 17 days post a successful da vinci si hysterectomy procedure, the patient returned to the hospital with severe pelvic symphsis pain. An x-ray found a 5-6 mm metallic object in the patient's pelvis. The patient underwent a laparoscopic procedure on (b)(6) 2011 to retrieve the metallic object using fluoroscopic guidance. The object was believed to be a 'thin rough cover plate' from the mega needle driver instrument. The patient was found to have a ruptured ovarian cyst that was believed to be the cause of her pain. | The instrument will not be returning for evaluation as the site has discarded it. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2193441 | 2167360 | 2064296 | 2955842201100230 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-15 | 2011-07-18 | Yes | Yes | Health Professional | 42019408 | No | Yes | 2011-06-15 | No | No | Reuse | 2011
40027202 | 2011-07-12 | Malfunction |  | The harmonic ace curved shears would not connect correctly to the housing unit, making it so the jaws would not open or close. ======================manufacturer response for endowrist 8mm disposable instrument accessory, harmonic ace curved shears insert (per site reporter)======================manufacturer provided rga# for product return evaluation. |  | HARMONIC ACE CURVED SHEARS INSERT | ENDOWRIST 8MM DISPOSABLE INSTR | 2202948 | 2177058 | 2073994 | 2177058 | 1 | NAY | User Facility |  |  | Initial | 2011-07-15 | 2011-07-15 | No | Yes | Physician | 40027202 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420184 | 2011-04-27 | Malfunction |  | On (b)(6), 2011, medwatch uf / importer report (b)(4) was received: patient underwent gynecological surgery utilizing da vinci robotics. An intuitive cautery spatula was utilized during the surgery. While moving the camera and cautery spatula within the abdominal cavity, the two objects came in contact and a small plastic part was dislodged from the tip of the cautery spatula. It was not possible to identify the piece and the surgeon elected to leave it in the abdominal cavity. This was the instrument's final use. | Multiple attempts with the site have been made, however at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2182864 | 2156617 | 2007609 | 2955842201100217 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 420184 | No | Yes | 2011-06-14 | No | No | Unkown | 2011
420183 | 2011-05-09 | Malfunction |  | On (b)(6), 2011, medwatch uf / importer report (b)(4) was received: patient undergoing robotic myomectomy. The da vinci permanent cautery hook broke into three pieces in the intra-abdominal cavity. A scope was utilized and the cautery hook pieces were recovered. This was the 7th time this instrument had been used and expected life is ten uses. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2182871 | 2156619 | 2007043 | 2955842201100219 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 420183 | No | Yes | 2011-06-14 | No | No | Reuse | 2011
40018010 |  | Malfunction |  | It was reported that after a da vinci s urology procedure, burn marks were present on the monopolar curved scissors (mcs) instrument tip cover accessory installed on the monopolar curved scissors instrument. No arcing was observed during the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burned through the silicone. The silicone exhibits localized melting around the hole which is indicative of an arcing event. The hole is oblong and approximately. 050 inches by. 025 inches in size and is located. 345 inches above the silicone to sleeve interface. If additional information becomes available, a follow-up mdr will be provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2182888 | 2156623 | 2053578 | 2955842201100220 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-10 | 2011-07-12 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-06-10 | Device Not Retu | Yes | Initial | 2011
40018010 |  | Malfunction |  | It was reported that after a da vinci s urology procedure, burn marks were present on the monopolar curved scissors (mcs) instrument tip cover accessory installed on the monopolar curved scissors instrument. No arcing was observed during the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole which is indicative of an arcing event. The hole sizes range from. 055 inches to under. 010 inches and are located from. 175 inches to. 210 inches above the silicone to sleeve interface. The tip cover also has mechanical tears at the edges of the two holes. If additional information becomes available, a follow-up mdr will be provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2182889 | 2156624 | 2053579 | 2955842201100221 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-10 | 2011-07-12 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-06-10 | Device Not Retu | Yes | Initial | 2011
IS3000 A60P6 | 2011-06-13 | Other | Other | It was reported that during a da vinci si gastric bypass procedure, the site experienced an unknown error. With the assistance of an isi technical support engineer the site restarted the system multiple times, however, upon start up system error code 31030 appeared. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that system error code 31030 experienced by the customer was associated with the surgeon side console touchpad. The surgeon side console touchpad is located in the middle of the surgeon console armrest and provides the means for selecting various system functions such as system status including instrument arms, camera arm, and energy controls. The system was repaired by replacing the affected touchpad. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the system records the fault and transitions to a safe state. The touchpad was returned for failure analysis evaluation. Engineering was able to replicate the reported failure and found a defective fuse holder. The fuse holder was replaced to address the system error code experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2182892 | 2156626 | 2053581 | 2955842201100223 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-13 | 2011-07-12 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-06-13 | Device Not Retu | No | Reuse | 2011
420093 | 2011-02-24 | Malfunction |  | On (b)(6), 2011, medwatch uf / importer report (b)(4) was received: patient underwent a da vinci assisted laparoscopic bilateral pelvic lymph node dissection, extraction of bladder stones time four and radical prostatectomy, non-nerve-sparing. During the procedure, a prograsp forcep broke within a trocar. It was not possible to remove the device until the instrument within the trocar had been straightened with one of the other robotic arms. It was noted that there was a small foreign body loose within the abdomen after the forcep broke. Before the end of the procedure, the foreign body was located and removed. This was the 9th use with expected use of 10 times. | Multiple attempts with the site have been made, however at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2182869 | 2156633 | 2007042 | 2955842201100218 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 420093 | No | Yes | 2011-06-14 | No | No | Reuse | 2011
42018305 | 2011-05-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while using the permanent cautery hook instrument, the surgeon found the instrument to be broken. A piece was found in the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the ceramic sleeve to be completely removed from the base of the cautery hook, exposing the base of the instrument. No damage was found to the hook or the clevis and no broken cables or wires were found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL ACC | 2182874 | 2156637 | 2053593 | 2955842201100222 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2011-06-14 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-06-08 | Malfunction |  | It was reported that during a da vinci s low anterial resection, the harmonic curved shears insert instrument fell into the patient and was immediately retrieved. The insert was replaced; however, the replacement harmonic curved shears insert also broke with pieces falling into the patient. The insert was removed and replaced to successfully complete the procedure. No patient harm, adverse outcome, or injury was reported. Please see mdr 2955842-2011-00225 for second insert related to this event. | The insert was returned and evaluated and engineering found the clamp arm detached from the distal end of the instrument. One side of shaft feature that mates with clamp arm pins is bent. The distal end of shaft exhibits various scratch marks on the black paint. Evidence is not conclusive, but damage is likely due to mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL | 2184166 | 2157957 | 2054958 | 2955842201100224 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2011-06-14 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-06-08 | Malfunction |  | It was reported that during a da vinci s low anterial resection, the harmonic curved shears insert instrument fell into the patient and was immediately retrieved. The insert was replaced; however, the replacement harmonic curved shears insert also broke with pieces falling into the patient. The insert was removed and replaced to successfully complete the procedure. No patient harm, adverse outcome, or injury was reported. Please see mdr 2955842-2011-00224 for first insert related to this event. | The insert was returned and evaluated and engineering found the clamp arm completely broken off from the grip of the insert. The clamp arm based connection point was found with bent joints, indicating excess force applied to clamp arm. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL | 2184179 | 2158003 | 2054958 | 2955842201100225 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-14 | 2011-07-12 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2011-06-14 | Device Not Retu | No | Reuse | 2011
420117 | 2011-06-29 | Malfunction |  | Davinci robotic 5mm needle driver tip broke in the patient's abdominal cavity during the procedure. The surgeon asked for another and the second tip also broke. All pieces of broken tips visualized and retrieved by md. ======================manufacturer response for endowrist needle holder 5mm instrument, needle driver for davinci s (per site reporter)======================nothing yet. |  | NEEDLE DRIVER FOR DAVINCI S | ENDOWRIST NEEDLE HOLDER 5MM IN | 2198405 | 2172282 | 2069218 | 2172282 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-07-11 | 2011-07-11 | No | Yes | Physician | 420117 | No | No Answer Provided |  |  |  |  | 2011
420205 | 2011-05-02 | Injury | Required InterventionLife Threatening | It was reported that during a da vinci s thymic tumor resection procedure the fenestrated bipolar forceps instrument attached to the patient's right internal thoracic vein. As a result, a hole approximately 2mm in diameter was created and bleeding (approximately 20cc) occurred. The bleeding was controlled by angiopressure and applying styptic. The planned surgical procedure was completed. | The hospital indicated that the affected instrument would not be returned for evaluation. As a result, it cannot be determined if the root cause of the customer reported failure was due to a device malfunction or another factor(s). A follow-up mdr will be submitted if the instrument is returned or if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2180568 | 2154185 | 2051141 | 2955842201100214 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-07 | 2011-07-08 | Yes | Yes | Health Professional | 420205 | No | No | 2011-06-07 | No | No | Unkown | 2011
42011704 | 2011-05-11 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure, one of the jaws of the 5mm needle driver instrument broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse events, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip to be broken off. The grip fractured at junction between the grip hub and grip arm, where there is a 90 degree angle. The fracture surface does not show any obvious material defect. The visual inspection of intact grip under 40x magnification shows a faint hairline crack at the same general location as the fractured grip. No other damage. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2180954 | 2154461 | 2051417 | 2955842201100215 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-06-08 | 2011-07-08 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2011-06-08 | Device Not Retu | No | Reuse | 2011
42022703 | 2011-05-10 | Other | Required Intervention | It was reported that on (b)(6) 2011, the da vinci si surgical system was used to remove an instrument cable fragment from a patient. The cable fragment is believed to have broken off of a intuitive surgical endowrist instrument used during the patient's da vinci si gynecological procedure previously performed on (b)(6) 2011. | On (b)(6) 2011, the fragments removed from the patient were evaluated by isi representatives. It was determined that the fragments were a cable and crimp from the pk dissecting forceps instrument used during the patient's initial da vinci si surgical procedure performed on (b)(6) 2011. The affected instrument was received for evaluation on (b)(6) 2011, prior to the patient's second da vinci si surgical procedure. At the time of receipt, the reported malfunction was a broken cable with no indication of missing components. During evaluation, failure analysis confirmed that the pitch cable was broken at the distal clevis hub, but did not identify the missing section of cable and crimp. As per standard analysis procedures, the instrument was discarded after failure analysis investigation was performed and before the patient became aware that the fragment had fallen into her body. Had the customer reported observation included a reference that "a broken cable had fallen into the patient", instrument analysis would have been expanded to include a detailed inspection of the missing cable location. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2231407 | 2169691 | 2066618 | 2955842201100216 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-07 | 2011-07-07 | Yes | Yes | Health Professional | 42022703 | Yes | Yes | 2011-06-07 | Yes | No | Reuse | 2011
42020706 | 2011-06-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si cystocele repair procedure, a fragment from the tenaculum forceps instrument was observed to be in the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned for evaluation. Per the customer reported complaint the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | Reuse | 2176434 | 2149060 | 2046015 | 2955842201100212 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-06-06 | 2011-07-05 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2011-06-06 | Device Not Retu | No |  | 2011
IS2000 A51P8 | 2011-06-06 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgical staff experienced a black screen on the assistant monitor and within the high resolution stereo viewer (hrsv) of the surgeon side console (ssc). With the assistance of an isi technical support engineer, the site found that the patient side cart (psc) was unplugged. The site plugged in the psc and restarted the system, however, the ssc would not restart. The patient was under anesthesia when the surgeon made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering found that issue experienced by the customer was associated with the system control processor (scp) board. The system control processor (scp) is a pca board that resides in the surgeon side console (ssc) that provides real time processing and control, system supervisor and fault monitoring, and external system communication control. The system was repaired by replacing the defective scp. As of july 5, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2177128 | 2149970 | 2046926 | 2955842201100213 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-06-06 | 2011-07-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-06-06 | Device Not Retu | No | Reuse | 2011
400180 | 2011-05-31 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the site observed an arc coming from the tip of monopolar curved scissors instrument with the tip cover accessory installed. No patient harm or adverse outcome was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2174181 | 2147098 | 2044053 | 2955842201100204 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-31 | 2011-07-01 | No | Yes | Health Professional | 400180 | No | Yes | 2011-05-31 | No | Yes | Initial | 2011
IS3000 A60P6 | 2011-05-31 | Injury | Required Intervention | Event Description: --/-- It was reported from a facility in europe that during a da vinci si hysterectomy procedure and while performing para-aortic lymph node dissection, the surgeon positioned the right master tool manipulator (mtmr) toward the uppermost boundary of his available mtm workspace. The surgeon removed his fingers from the mtmr while his head was in the surgeon console viewer (contrary to a warning in the labeling), and the instrument associated with the mtmr dropped down in an uncontrolled manner, and cut the patient's abdominal aorta. The planned surgical procedure was converted to traditional open surgery and the patient's aorta was repaired by placement of a stent. The planned hysterectomy procedure was also completed. As of the date of this report, the patient is reported to be recovering with no complications and was discharged from the hospital a week after the operation. | Manufacturer Narrative: --/-- The master tool manipulator refers to the master controller which provides the means for the surgeon to control the instruments inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The instruments follow the motions of the mtms to conduct the procedure, which is referred to as following mode. As a safety feature, when the surgeon removes his head from the viewer portion of the surgeon console, the instruments freeze in position. An investigation was conducted by a local isi field service engineer (fse) following the event. The fse was unable to reproduce the reported event; however, the fse did observe that the mtmr spring guide component had failed. The system was repaired by replacing the affected mtmr, and the original mtmr was returned for analysis. The spring guide component that was found to be damaged contains two springs, which are part of the gravity compensation strategy for the shoulder and elbow joints of the mtm, which are joints that support the mtm handles that the surgeon grasps with his or her fingers. The spring guide secures the springs in place such that motion of the joint compresses the springs, partially counteracting the force of gravity. The spring counterbalance is supplemented by a motorized gravity compensation algorithm. For the comfort and precision of the surgeon, the spring and motor compensation is to reduce the apparent weight of the robotic arm as perceived by the surgeon. It is believed that in instances where the spring guide component has failed, the gravity compensation system can be compromised, resulting in a maximum imbalance of 0. 6 lbs. Thus, if contrary to the below warning included in the product labeling, the surgeon releases his grip on an mtm while his or her head is in the viewer, this component failure may cause the mtm and its associated instrument, to move in an uncontrolled manner. If the surgeon complies with the label warning below, no uncontrolled motion should occur. The da vinci si user manual warning states: warning: once in following, the surgeon console operator must not remove his or her hands from the masters until removing his or her head from the surgeon console viewer - thereby taking the system out of following mode. Failure to do so may result in uncontrolled movement of the masters, resulting in serious harm to the patient. A corrective action is underway to install metal reinforcing caps for the retention component in the mtms on all is3000 systems to prevent the failure of the spring guide component. Additional information concerning the corrective action is provided with the fedex copy of this report. (b)(4). --/-- Additional Manufacturer Narrative: --/-- The master tool manipulator (mtmr) was returned to isi for evaluation. The link 1 cover was removed to verify the spring guide failure. When the cover was removed to inspect the damaged spring guide it was noted that the counterbalance spring was pushing against the inside of the exterior cover. When the cover was replaced to take force measurements, it was necessary to compress the spring slightly to tighten the cover mounting screw. The joint 2 (shoulder) spring guide end flange was observed to have cracked away from the rest of the spring guide tube. The end flange remains attached to the tube by a small thickness of material at the end of a visible tear in the plastic. The fracture initiates at a high stress location near a mounting screw head. The spring extension resulting from the crack was measured with the cover removed to be approximately 1. 0 inch +/-. 1 inch. Visual inspection of the mtmr did not reveal any other anomalies, aside from the cracked spring guide, that would account for loss in gravity compensation. The root cause of the spring guide failure was a stress concentration in the spring guide where the flange intersects the thin wall of the tube in a sharp corner, combined with cyclic loading of the spring guide due to mtm motion. The stress concentration resulted in a crack that initiated at the high stress location and then slowly propagated due to cyclic loading as the force from the spring varied during use of the mtm. Stress was induced in the fracture location by tensile loading on the guide from spring compression. Stress near the end-flange on the spring guide restraining the joint 2 counterbalance spring was increased by the clamp force from the screw head that mounts the guide to the mtm. The reduction in gravity compensation on this mtmr is due to cracking near the joint 2 (shoulder) spring guide end flange. The spring tube contains stress concentrations that are inherent to the part geometry and cyclic loading can cause cracking and gradual tearing of the plastic part. The separation of the end flange from the tube and resulting reduction in gravity compensation is likely to have occurred over the course of several months. Intuitive surgical initiated a corrective action, 2955842-070111-002 c, issued 7/1/2011, in relation to this spring counterbalance subsystem failure of the master tool manipulator that occurred at (b)(4) hospital. Correction of affected systems has been performed using reinforcing caps for the retention component to correct the problem with the mtm subsystem. Affected mtm spring guides were inspected by field service engineers and those spring guides having no cracks or gaps less than or equal to 0. 156 inches were retrofitted with reinforcing caps in the field. Mtm's containing spring guides with gaps larger than 0. 156 inches were returned through the rma process and spring guides were replaced prior to reinforcing cap installation. All affected systems have been retrofitted with the reinforcing caps. In addition, reemphasis of warnings listed in the approved user manual labeling provided with each system warning: once in following, the surgeon console operator must not remove his or her hands from the masters until removing his or her head from the surgeon console viewer - thereby taking the system out of following mode. Failure to do so may result in uncontrolled movement of the masters, resulting in serious harm to the patient. Has been provided to all affected sites. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2174889 | 2147492 | 2044447 | 2955842201100205 | 1 | NAY | Manufacturer | User facilityOther | Physician | InitialFollowup | 2011-07-19 | 2011-07-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-31 | Device Not Retu | No |  | 2011
42019408 | 2011-06-06 | Malfunction |  | It was reported that approximately 2 hours into a da vinci si hysterectomy procedure, while using the mega needle driver instrument, the tip of the instrument became dislodged and fell into the patient. The fragment was removed and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient injury, harm or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one of the carbide inserts to be missing from the grip tip. No adhesive residue was found on the connecting surface. No other damage was found at the instrument grip or clevis. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2175309 | 2147698 | 2044654 | 2955842201100206 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-06-06 | 2011-07-01 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2011-06-06 | Device Not Retu | No | Reuse | 2011
42014204 | 2011-06-01 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while using the 5mm monopolar cautery instrument, the hook broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient injury, harm or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the hook to be detached and missing. The distal end exhibits melting and thermal damage on one side of the monopolar adapter. The threading is deformed due to thermal damage, preventing installation of a disposable cautery tip. The inside pocket of the adapter is charred and one of the electrical contacts within the adapter is bent downwards, preventing the disposable cautery tip from fully seating in the adapter. No other damage was found. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2175291 | 2147699 | 2044655 | 2955842201100207 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-02 | 2011-07-01 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | Yes | 2011-06-02 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-02 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with an mcs tip cover accessory installed. The planned surgical procedure was completed with a replacement mcs instrument and tip cover. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole to be burned through the silicone and the silicone exhibits localized melting around the hole, indicative of arcing. The hole is approximately. 012 inches in diameter and is located. 250 inches above the silicone to sleeve interface. The silicone has a mechanical tear at the edge of the hole. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2175292 | 2147700 | 2044656 | 2955842201100208 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-03 | 2011-07-01 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-06-03 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-04-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with an mcs tip cover accessory installed. The planned surgical procedure was completed with a replacement mcs instrument and tip cover. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole to be burned through the silicone and the silicone exhibits localized melting around the hole, indicative of arcing. The hole is approximately. 020 inches in diameter and is located. 275 inches above the silicone to sleeve interface. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2175312 | 2147702 | 2044658 | 2955842201100209 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-03 | 2011-07-01 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-06-03 | Device Not Retu | Yes | Initial | 2011
IS3000 60 P5 | 2011-06-05 | Other | Other | It was reported that during a da vinci si surgical procedure, the site could not get the right and left eye of the endoscope to focus together. The endoscope was replaced and the site continued with the procedure. Shortly after the endoscope was replaced, the surgeon stated that the color seemed to be off. The surgical staff was in the process of troubleshooting when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted found no system errors in the error log. The isi technical support engineer (tse) had the customer install the camera head used during the procedure and perform white balance and 3d auto focus, and found the image to be within specification. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2175344 | 2147746 | 2044701 | 2955842201100210 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-05 | 2011-07-01 | Yes | Yes | Health Professional | IS3000 60 P5 | Yes | Yes | 2011-06-05 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-06-06 | Other | Other | It was reported that while starting a da vinci prostatectomy procedure, the site loss vision in the right eye of the camera controller unit (ccu). The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer found that the right camera controller unit (ccu) power supply was defective. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu power supply. As of (b)(4) 2011 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2175345 | 2147748 | 2044703 | 2955842201100211 | 1 | NAY | Manufacturer | DistributorCompany Representative | Other | Initial | 2011-06-06 | 2011-07-01 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-06-06 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-05-30 | Other | Other | It was reported that while starting a da vinci si hysterectomy procedure the site experienced a system error code 25326. With the assistance of an isi technical support engineer the site restarted the system. The system restarted without error and the site decided to continue with the planned surgical procedure. During anastomosis the surgeon had difficulty engaging an instrument installed on a patient side manipulator (psm) arm. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the system error code 25326 experienced by the customer was associated with a remote arm controller (rac)pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The fse was unable to replicate the instrument engagement issue reported by the customer. Multiple insertions of the instrument used in the procedure were made by the fe without any issue. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25326 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put davinci in a recoverable safe state. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2169222 | 2141421 | 2038397 | 2955842201100203 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-30 | 2011-06-27 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-30 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-25 | Malfunction |  | On (b)(6) 2011, a copy of medwatch uf/importer report (b)(4) was received from (b)(6) hospital with the following information: during robotic hysterectomy, guard on tip of davinci mega needle driver fell off inside patient. Guard was able to be retrieved. | The hospital indicated that the instrument would not be returned for evaluation. An intuitive surgical representative was able to obtain a photograph of the affected instrument, which was then evaluated by engineering. Per the photograph, engineering observed that the carbide insert had detached from one of the instrument grips. Since the instrument was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2168329 | 2140068 | 2037043 | 2955842201100197 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-02 | 2011-06-25 | No | Yes | Health Professional | 42019408 | No | Yes | 2011-06-02 | No | No | Unkown | 2011
42022703 | 2011-06-01 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while using the pk dissecting forceps instrument, a plastic piece fell into the patient. The site was unable to locate the fragment or confirm if the fragment was still retained in the patient. The planned surgical procedure was completed and no patient injury, harm or adverse outcome was reported. | Multiple attempts have been made, however as of the date of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2167913 | 2139660 | 2036636 | 2955842201100192 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-01 | 2011-06-24 | No | Yes | Health Professional | 42022703 | No | No | 2011-06-01 | No | No | Unkown | 2011
40027202 | 2011-05-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory was returned and evaluated. Per the customer reported complaint engineering observed the teflon pad is detached from the clamp arm. The middle section of the teflon pad is still attached to the clamp arm, but a u shaped piece of the clamp arm is detached. An isi performed a site visit on june 14, 2011 to observed cases, surgeon and surgical staff technique, and provided feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2167914 | 2139661 | 2036637 | 2955842201100193 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2011-05-26 | Device Not Retu | Yes | Initial | 2011
40027202 | 2011-05-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed the insert to be returned with a broken piece approximately. 670 inches by. 085 inches in size. The blade fractured below the clamp arm, within the shaft, and the clamp arm is intact. An isi product manager performed a site visit on (b)(6) 2011 to observe cases, surgeon and surgical staff technique, and provide feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2167915 | 2139662 | 2036638 | 2955842201100194 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2011-05-26 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-06-08 | Other | Other | It was reported that during the beginning of a da vinci s prostatectomy procedure, system error code 23008 occurred. In an attempt to clear the error, the surgical staff power cycled the system and back drove the left master tool manipulator. The system was then restarted, however, the same error code recurred after restart. The surgeon made the decision to abort the planned procedure and reschedule it for a later date. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. The mtml was returned for evaluation. Engineering was able to replicate the reported failure mode and determined that the main cable harness had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2167946 | 2139693 | 2036669 | 2955842201100195 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-08 | 2011-06-24 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-06-08 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-06-06 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the permanent cautery spatula instrument broke and the surgical staff had difficulty removing the instrument from the cannula due to the angle of the instrument wrist. As a result, the system arm was undocked from the patient with the instrument still installed and inserted in the cannula. While undocking, a piece of the instrument fell into the patient and was retrieved. The site was able to remove the broken instrument, the system arm was redocked, and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed the yaw pulley and spatula assembly detached from the distal end of the instrument. The conductor wire has been cut at the distal end. One of the distal clevis ears is broken off at its base. The broken ear piece was not returned, with the missing piece approximately. 280 x. 235 in size. The conductor wire cap is also missing. The yaw pulley has the extruded boss feature at the center of the hub and one side of the cable crimp pocket broken off. It was determined that the initial failure was likely a broke clevis ear and the spatula/yaw pulley assembly was likely detached to remove the instrument from the cannula. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2167947 | 2139709 | 2036664 | 2955842201100196 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-06-06 | 2011-06-24 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-06-06 | Device Not Retu | No | Reuse | 2011
42011704 | 2011-05-24 | Malfunction |  | It was reported that during a da vinci s nissen paraesophageal repair procedure the tip of the 5mm needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one needle driver grip broken off where the grip arm meets the grip hub. The grip has a 90 degree angle at the fracture location. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2167951 | 2139712 | 2036687 | 2955842201100198 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-02 | 2011-06-24 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2011-06-02 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-05-25 | Malfunction |  | It was reported that while coagulating some tissue during a da vinci si hysterectomy procedure, the monopolar curved scissors instrument sparked. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a slight char mark on the tube extension connection point with the main tube. The connection is broken throughout the circumference. Some material is melted at the char mark, indicating an arcing event occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2167952 | 2139713 | 2036688 | 2955842201100199 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-25 | 2011-06-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-05-25 | Device Not Retu | No | Reuse | 2011
IS2000 51P8 | 2011-05-27 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff was unable to insert an instrument on the patient side manipulator (psm). The site attempted to troubleshoot by rebooting the system; however, the issue persisted. The patient had been under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the psm. The psm was returned and evaluated. Engineering was not able to replicate the reported failure mode; however, suspect components were replaced to repair the psm. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2168049 | 2139779 | 2036755 | 2955842201100201 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-27 | 2011-06-24 | Yes | Yes | Health Professional | IS2000 51P8 | Device Returned To Manufacturer | Yes | 2011-05-27 | Device Not Retu | No | Reuse | 2011
42000606 | 2011-05-24 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a fragment from the large needle driver instrument fell into the patient and was retrieved. The planned procedure was successfully completed with no harm to the patient. | The instrument was returned and evaluated. Engineering observed the instrument to be returned with the carbide insert detached from one grip. The brazing surface on the grip exhibits various voids in the filler material. The outer surface of grips do not exhibit damage. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2168035 | 2139794 | 2036769 | 2955842201100200 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-25 | 2011-06-24 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2011-05-25 | Device Not Retu | No | Reuse | 2011
40017908 | 2011-05-20 | Malfunction |  | It was reported that during a da vinci nephrectomy procedure, the surgical staff observed the tip of the monopolar curved scissors instrument break resulting in a piece falling into the patient. The piece was removed and the procedure was successfully completed with no harm to the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the left blade to be broken off of the instrument. The broken segment was not returned. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2168060 | 2139807 | 2036782 | 2955842201100202 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2011-05-26 | Device Not Retu | No | Reuse | 2011
42017207 | 405928 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument broke and a fragment fell into the patient. While trying to locate the broken piece, there was a delay in the procedure. The broken piece was successfully retrieved and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip broken off at the base of the clevis. As a result, there is separation of the yaw pulley and pulley cover at the glue joint. The broken grip was returned with the instrument. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2164546 | 2136853 | 2033829 | 2955842201100191 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-05-23 | 2011-06-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 405928 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-12 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was suturing and using the mega needle driver instrument, the serrated teeth plate fell into the patient. The fragment was retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | Multiple attempts with the site have been made, however, as the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2161916 | 2134479 | 2031455 | 2955842201100190 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-12 | 2011-06-20 | No | Yes | Health Professional | 42019408 | No | No | 2011-05-12 | No | No | Unkown | 2011
P5300 | 2011-05-24 | Injury | Required Intervention | While docking robot, a non-recoverable fault alarm sounded. Troubleshooting measures did not work. Manufacturer contacted. This problem resulted in converting from laparoscopic to open surgery. Another case scheduled for that day had to be cancelled. Eval by manufacturer found 16 faults and no power to arm 2. The set up joint was replaced and equipment was tested. |  | DAVINA ROBOT | ROBOT | 2168624 | 2138245 | 2035221 | MW5021068 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-06-14 | 2011-06-20 | No | Yes | Health Professional | P5300 | Yes | Yes |  |  |  |  | 2011
400272 | 2011-04-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, when the surgeon began to amputate the uterus from the cervix, the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory has not been returned for evaluation, however, an isi product manager performed a site visit on (b)(6) 2011 to observe cases, surgeon and surgical staff technique, and provide feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. The root cause of the customer reported failure mode cannot be determined due to the harmonic ace insert accessory not being returned for evaluation. If additional information is received, a follow-up mdr will be submitted. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2159836 | 2132658 | 2029632 | 2955842201100179 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-18 | 2011-06-17 | No | Yes | Health Professional | 400272 | No | Yes | 2011-05-18 | No | Yes | Unkown | 2011
40018302 | 2011-05-13 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, arcing was observed coming from the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a char mark on the yaw pulley and yaw pulley exit. The conductor wire has surface damage but is not broken. The ceramic sleeve also exhibits damage. No other damage found. The charring indicates an arcing event occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2159842 | 2132662 | 2029636 | 2955842201100180 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-19 | 2011-06-17 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | No | 2011-05-19 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy, the plate inside tip of the mega needle driver instrument came off, fell into the patient and was retrieved. The planned surgical procedure was completed with no patient harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering noticed the carbide insert detached from one of the instrument grips and no adhesive residue was observed on the grip tip surface. The detached insert was returned with the instrument. No damage was found on the insert on the other grip tip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2160398 | 2132954 | 2029928 | 2955842201100181 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-23 | 2011-06-17 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-05-23 | Device Not Retu | No | Initial | 2011
IS2000 | 2011-05-20 | Other | Life ThreateningOther | It was reported that while isolating the pulmonary vessel during a da vinci s thoracic sympathectomy procedure, a tear on the underside of the patient's pulmonary vessel occurred when the surgeon pulled on a calcified lymph node. Due to the location of the tear, the surgeon was unable to angle the system camera to view the affected area, therefore, unable to repair the artery using the system. The surgeon made the decision to convert to traditional open surgical techniques to repair the tear and complete the planned procedure. The da vinci procedure had been in progress for approx 2 hours when the event occurred. No patient is reportedly doing well and no post surgical complications have been reported. | An intuitive surgical field service engineer (fse) went onsite and evaluated the system. Visual and mechanical inspections were performed with no issues found. System verification testing was performed and passed. Based on the information provided, it was determined that the da vinci s surgical system worked as intended. No information has been received which suggests that the da vinci s surgical system malfunctioned in a way that may have caused or contributed to the patient's injury. Intuitive surgical concluded that the patient injury was unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160402 | 2132958 | 2029932 | 2955842201100183 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-20 | 2011-06-17 | Yes | No | Health Professional | IS2000 | Yes | No | 2011-05-20 | Device Not Retu | Yes | Reuse | 2011
IS3000 A60P6 | 2011-05-20 | Other | Other | It was reported that during a da vinci si nephrectomy procedure, the site experienced a loss of vision in the right eye image of the high stereo resolution viewer (hrsv). With the assistance of an isi technical support engineer the surgical staff reseated the vision cable on both ends and restarted the system, however, the vision issue was not resolved. The surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160410 | 2132966 | 2029940 | 2955842201100182 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-20 | 2011-06-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-20 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-05-24 | Other | Other | It was reported that during a da vinci si surgical procedure the site experienced system error code 16 and lost power on a patient side manipulator (psm) arm. With the assistance of an isi technical support engineer the site restarted the system and tried to manually reposition the psm, however the psm did not regain power. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system error code 16 was associated with a patient side manipulator (psm) setup-joint axis. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The setup-joint (suj) is a non-motorized support arm which holds a patient side manipulator (psm) during surgery. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 16 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the psm. Upon determining this condition, the system records the fault and transitions to a safe state. The suj was returned for failure analysis evaluation. Engineering was able to replicate the reported failure and replaced all internal electrical cables to address the system error codes experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160412 | 2132968 | 2029942 | 2955842201100184 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-24 | 2011-06-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-24 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-05-24 | Other | Other | It was reported that during a da vinci s hysterectomy the surgeon experienced a depth perception issue between the left and right eye of the surgeon console. The system was restarted, however, this action was unable to resolve the vision issue. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The camera head was returned and evaluated by the original equipment manufacturer (oem). The oem observed that the camera stage was damaged, thus causing the vision issue experienced by the customer. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160415 | 2132971 | 2029945 | 2955842201100185 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-05-24 | 2011-06-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-05-24 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-05-20 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s nissen fundoplication procedure, the plastic sheeting of the permanent cautery hook instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | Unkown | 2160493 | 2133036 | 2030012 | 2955842201100186 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-05-20 | 2011-06-17 | No | Yes | Health Professional | 42018305 | Yes | Yes | 2011-05-20 | Device Not Retu | No |  | 2011
42017910 | 2011-05-19 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the site observed the tip of the monopolar curved scissors (mcs) instrument to have broken. The piece was immediately retrieved and the instrument was replaced before the planned surgery was successfully completed without incident. The patient was reported as doing well with no post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension to be broken and missing a. 280 x. 210 piece at the prox clevis interface. The scissor blades are intact and the clevis is dislodged from tube extension. Evidence not conclusive, but tube extension likely broke due to excessive side loading. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2160495 | 2133038 | 2030014 | 2955842201100187 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-05-19 | 2011-06-17 | No | Yes | Health Professional | 42017910 | No | No | 2011-04-13 | No | Yes | Initial | 2011
IS2000 51P8 | 2011-05-18 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff was unable to focus the camera. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the system; however they did not have a back up camera cable or camera head. The patient had been under anesthesia for 2 hours when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the camera cable, camera head and video digitizer. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images (such as the user interface icons and text or an auxilliary image from an ultrasound or other source). The system was repaired by replacing the camera cable, camera head and video digitizer. The components were returned and evaluated. Engineering not able to replicate the reported failure mode with the camera cable and digitizer; however, did reproduce the reported failure mode on the returned camera head. The camera head showed evidence of mechanical shock resulting in damaged stage assembly thus causing the reported complaint. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160496 | 2133039 | 2030015 | 2955842201100188 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-18 | 2011-06-17 | Yes | Yes | Health Professional | IS2000 51P8 | Device Returned To Manufacturer | Yes | 2011-05-18 | Device Not Retu | No | Reuse | 2011
IS1200 43P9 | 2011-05-20 | Malfunction |  | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff experienced the left eye of the surgeon side consoles high resolution stereo viewer (hrsv) to be black and the right eye to have some interference. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the system; however, the site did not have a back up camera cable. The patient had been under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the technical service engineer determined that the customer reported issue was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the camera cable. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of june 16, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160498 | 2133041 | 2030017 | 2955842201100189 | 1 | NAY | Manufacturer | ForeignDistributor | Physician | Initial | 2011-05-20 | 2011-06-17 | No | Yes | Health Professional | IS1200 43P9 | No | Yes | 2011-05-20 | No | No | Unkown | 2011
40018010 | 2011-05-17 | Malfunction |  | It was reported that after a da vinci s hysterectomy procedure, two holes were observed to have been burned in the monopolar curved scissors (mcs) instrument tip cover accessory installed on the monopolar curved scissors instrument. No arcing was observed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed two holes burned though the silicone. The silicone exhibits localized melting around the hole which is indicative of an arcing event. The hole sizes are. 055 and. 045 x. 075 with varying shapes (round and oblong), and are located. 075 and. 320 above the silicone to sleeve interface. An additional investigation is being conducted to determine root cause of this event. If additional information becomes available, a follow-up mdr will be provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2159038 | 2131904 | 2028880 | 2955842201100177 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-17 | 2011-06-16 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-17 | Device Not Retu | Yes | Initial | 2011
420184 | 2011-05-17 | Malfunction |  | It was reported that after performing bivalve of the uterus during a hysterectomy procedure, the surgical staff attempted to remove the permanent cautery spatula instrument; however, the instrument's wrist was pitched to one side, preventing the instrument from being removed from the cannula. The surgeon then used a fenestrated bipolar instrument to straighten out the wrist of the permanent cautery spatula causing the spatula instrument to break. It was reported that three pieces of the instrument broke off and two were retrieved. Isi (b)(4) discussed the risks of instrument materials being left behind and as of the date of this mdr, the site has not reported the patient as experiencing any post operative issues. (b)(4). | Due to the instrument not returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings to handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2159051 | 2131918 | 2007640 | 2955842201100178 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-05-17 | 2011-06-16 | No | Yes | Health Professional | 420184 | No | No | 2011-05-17 | No | No | Unkown | 2011
400077 | 2011-06-01 | Malfunction |  | Patient with leiomyomatous uterus underwent robotic assisted hysterectomy (myomectomy only) and during surgery, the port seals leaked. Seals were switched/replaced and surgery continued with no apparent delays; no issue with patient. |  | 8MM CANNULA SEAL 1012MM CANNULA SEAL | ENDOWRIST INSTRUMENT 8MM ACCES | 2165261 | 2137606 | 2034582 | 2137606 | 1 | NAY | User Facility |  |  | Initial | 2011-06-16 | 2011-06-16 | No | Yes | Physician | 400077 | Yes | No Answer Provided |  |  |  |  | 2011
400169 | 2011-05-26 | Malfunction |  | Patient undergoing prostatectomy via robotic assistance. Surgeon was using the harmonic curved shears insert for the davinci when, during the procedure, he noticed the tip of the harmonic had come off. Harmonic tip noted to be off at approximately 3 hours into the surgery. It was all retrieved and surgery was complete. Patient sent to recovery. |  | HARMONIC CURVED SHEARS INSERT | ENDOWRIST INSTRUMENT 8MM ACCES | 2165264 | 2137607 | 2034583 | 2137607 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-06-16 | No | Yes | Physician | 400169 | No | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-05-16 | Other | Other | It was reported that during a da vinci s low anterior resection procedure, the surgeon side console (ssc) experienced a power loss. With the assistance of an isi technical support engineer, the site restarted the system, however, after homing the ssc continued to run off of the battery back-up system. The site tried multiple outlets in the operating room, however, the ssc continued to run from the battery back-up system. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found the power issue experienced by the customer was associated with the surgeon side console primary power supply (pps2). The pps2 is a 230v power supply with battery backup. The system was repaired by replacing the pps2 and the power supply fan and the battery pack module as a precaution. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2151351 | 2123830 | 2020806 | 2955842201100176 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-05-16 | 2011-06-14 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-05-16 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-13 | Malfunction |  | It was reported that while dissecting the pelvic lymph nodes during a da vinci si hysterectomy, a tear was observed on the monopolar curved scissors (mcs) tip cover accessory (installed on the mcs instrument) and arcing occurred. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and one hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is approximately. 020 in diameter and located. 090 above the silicone to sleeve interface. The silicone also exhibits scratches on the outer surface and a tear adjacent to the hole. Other various mechanical tears in the silicone were also observed in the areas not near the hole. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2147491 | 2119809 | 2016785 | 2955842201100174 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-13 | 2011-06-10 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-13 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P6 | 2011-05-12 | Other | Required Intervention | It was reported that while performing a da vinci si hysterectomy, the left master controller did not have full control of the maryland bipolar forceps instrument during preparation of the uterine tube, resulting in non-intuitive motion and causing a small bleed on the patient's uterine tube. Coagulation was used to stop the bleed and the planned surgical procedure was completed without further incident. No post surgical complications were experienced by the patient. | Investigation was conducted by a field service engineer (fse) who went onsite and evaluated the system. System verification testing was performed and passed. The fse was unable to replicate the customer reported failure mode and the root cause of the non-intuitive motion experienced by the customer cannot be determined. As a precaution, the fse recalibrated the left master tool manipulator (mtm). The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand and one to his right. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2147713 | 2120175 | 2017151 | 2955842201100175 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-05-12 | 2011-06-10 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-05-12 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Other | Other | It was reported that after a da vinci si surgical procedure, the patient was found to have what appeared to be a skin burn at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. Also see mdr 2955842-2011-00172. | An isi representative has attempted to gather additional information from the initial reporter related to this event without success. The sites robotics coordinator advised that often time, a bovie generator is used to create port incisions, which would give the appearance of a port burn. A follow-up mdr will be submitted if additional information is received. As of june 9, 2011, the site has continued to use the system and has not reported a recurrence of the issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2146622 | 2119068 | 2016044 | 2955842201100173 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-10 | 2011-06-09 | Yes | Yes | Health Professional | IS3000 A60P5 | No | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Other | Other | It was reported that after a da vinci si surgical procedure, the patient was found to have what appeared to be a skin burn at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. Also see mdr 2955842-2011-00173. | An isi representative has attempted to gather additional information from the initial reporter related to this event without success. The sites robotics coordinator advised that often time, a bovie generator is used to create port incisions, which would give the appearance of a port burn. A follow-up mdr will be submitted if additional information is received. As of june 9, 2011, the site has continued to use the system and has not reported a recurrence of the issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2146593 | 2119139 | 2016112 | 2955842201100172 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-10 | 2011-06-09 | Yes | Yes | Health Professional | IS3000 A60P5 | No | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-05-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the shaft of the monopolar curved scissors instrument bent and arching occurred. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned, however, a photographic image of the instrument was received and evaluated. Per the customer reported complaint, engineering observed a broken tube extension at the main tube interface with no evidence of arching seen. Since the instrument was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. --/-- Additional Manufacturer Narrative: --/-- (b)(4) - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension dislodged from the main tube and the entire tube extension wall circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. An arc track was found at the breakage location, confirming that an arcing event occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2145340 | 2117324 | 2014302 | 2955842201100171 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-05-09 | 2011-06-08 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-06-20 | Device Not Retu | No |  | 2011
| 2011-05-31 | Malfunction |  | Surgeon was using the device to dissect and it quit working so he removed it from the field and retrieved another one. After surgery, it was noted that a piece of the plastic tip was missing. This device has been reprocessed 4 times. Device was flushed in ultrasonic washer, sent through automated washer, inspected and packaged, and steam sterilized. |  | DA VINCI S ENDOWRIST | PERMANENT CAUTERY HOOK | 2159626 | 2132509 | 2029485 | 2132509 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-06-07 | 2011-06-07 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-05-05 | Other | Other | It was reported that approximately 3 hours into a da vinci si hysterectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, the site observed arcing to the patient's pelvic sidewall. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. No repair was required and the planned surgical procedure was completed. No patient harm or adverse outcome was reported. The following medwatch report was submitted to the fda in relation to this event: mfr report 2955842-2011-00166. | The tip cover accessory was returned and evaluated. Engineering observed various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole sizes are approximately. 010 inches to. 035 inches with varying shapes (round and oblong) and are located from. 070 inches to. 340 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2141327 | 2114010 | 2010934 | 2955842201100168 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-06 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | Yes | Initial | 2011
400169 | 2011-04-26 | Malfunction |  | (b)(4). Event desc: during a robotic laparoscopic supracervical hysterectomy, the harmonic ace curved shears insert's insulation piece broke off while being used. The piece was recovered without harm to patient. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | Multiple attempts with the site have been made, however, at the time of this report, the accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned and/or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2142680 | 2114813 | 1976010 | 2955842201100169 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-09 | 2011-06-06 | No | Yes | Health Professional | 400169 | No | Yes | 2011-05-09 | No | Yes | Unkown | 2011
420184 | 2011-02-22 | Malfunction |  | On may 12, 2011, maude event report (b)(4) was received: patient was undergoing robotic lh bilateral salpingo-oophorectomy and mccall's monarch anterior repair. The plastic part of the spatula broke off during the procedure. This has occurred previously with the same operator, according to operating room nurse. Operator has been re-trained by the company rep on previous occasions. Diagnosis or reason for use: surgery for hysterectomy and bladder suspension. | On may 17, 2011, (b)(6) risk manager, stated that the hospital no longer has the permanent cautery spatula instrument. She also stated that the surgical staff believes that a piece of the instrument fell into the patient, however, no pieces were retrieved during the procedure and the patient has not returned to have any fragments removed since. The operating room director stated that this surgeon has had recurring instrument breakage. Multiple clinical sales representatives had re-trained the surgeon in order to avoid future breakage events. Due to the instrument not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2142658 | 2114815 | 2011791 | 2955842201100170 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-05-12 | 2011-06-06 | No | Yes | Health Professional | 420184 | No | No | 2011-05-12 | No | No | Unkown | 2011
IS1200 A43P9 | 2011-05-06 | Other | Other | It was reported during a da vinci prostatectomy procedure, the bifurcated light guide cable melted. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found three bifurcated light guide cables to be burnt at the distal ends. The lumen light from the illuminator was tested and confirmed within specification. The bifurcated light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The system was repaired by replacing the illuminator as a precaution. The three bifurcated light guide cables were returned to isi for additional evaluation. Engineering observed two of the cables as having evidence of thermal damage at the glass tip on the proximal end connectors which is exhibiting melting. These cables also show signs of unauthorized re-work. The third cable had a red substance at the glass tip and metal connector which likely led to reduced performance. The illuminator was also returned to isi for evaluation and found to be within specification. As of june 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2141262 | 2113940 | 2010916 | 2955842201100165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-05 | Other | Other | It was reported that approximately 90 minutes into a da vinci si hysterectomy procedure, while dissecting the patient's lymph nodes and using the monopolar curved scissors instrument with the tip cover accessory installed, the site observed arcing to the patient's bowel. No repair was required and the planned surgical procedure continued. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. No patient harm or adverse outcome was reported. The following medwatch report was submitted to the fda in relation to this event: mfr report 2955842-2011-00167. | The tip cover accessory was returned and evaluated. Engineering observed various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole sizes are approximately. 010 inches in diameter and are located from. 060 inches to. 210 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes and a long tear at the distal tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2141249 | 2113958 | 2010934 | 2955842201100166 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-03 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | Yes | Initial | 2011
42018305 | 2011-04-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si sigmoid colectomy the tip of the permanent cautery hook instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the hook broken off of the instrument. The hook had fractured within the yaw pulley, at the cross section of the hole feature. The detached hook was not returned. Engineering also observed that one side of the ceramic sleeve is broken off and missing. The size of the missing piece is approximately. 155 x. 075. It was determined that the damage to the hook and sleeve may be due to overloading at the tip or other mishandling/misuse. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 2139543 | 2111686 | 2008662 | 2955842201100164 | 1 | NAY | Manufacturer | ForeignDistributor | Other | FollowupInitial | 2011-05-03 | 2011-06-02 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2011-06-08 | Device Not Retu | No |  | 2011
IS3000 A60P6 | 2011-05-02 | Other | Other | It was reported that when the surgeon released the energy control foot pedal and removed their head out of the stereo viewer during a da vinci si hysterectomy, the surgeon heard the electrosurgical unit (esu) still releasing energy. The surgeon put their head back into the stereo viewer and noticed that the monopolar curved scissors instrument, which was connected to the esu, was releasing energy. The surgeon took control of the instrument and inadvertently touched fatty tissue. The surgeon quickly pressed the foot pedal to stop the instrument from releasing energy. It was reported that the patient suffered a small superficial burn. | Investigation was conducted by a field service engineer (fse) who went onsite and evaluated the system. Visual and mechanical inspections were performed with no issues found. The system logs were reviewed and no errors were observed. System verification testing was performed and passed. The fse demonstrated to the site that the instrument would cease fire the instant the head is removed from the stereo viewer. The root cause of the issue experienced by the customer cannot be determined. As a precaution, the fse swapped the food pedal assembly. Additionally, the site's biomedical department will replace the electrosurgical unit and energy activation cable. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2138371 | 2110566 | 2007542 | 2955842201100163 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-02 | 2011-06-01 | Yes | No | Health Professional | IS3000 A60P6 | Yes | Yes | 2011-05-02 | Device Not Retu | No | Reuse | 2011
420179 | 2011-05-27 | Malfunction |  | While separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. |  | HOTSHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2142904 | 2115178 | 2012154 | 2115178 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-05-31 | No | Yes | Physician | 420179 | No | No Answer Provided |  |  |  |  | 2011
420179 | 2011-05-27 | Malfunction |  | While separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. |  | HOTSHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2142904 | 2115178 | 2012154 | 2115178 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-05-31 | No | Yes | Physician | 420179 | No | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 |  | Death | Death | It was reported that on (b)(6), 2011 a patient passed away 18 months after a da vinci s procedure to remove the lower lobe of their left lung due to bronchioloalveolar carcinoma. According to the patient's husband, during the procedure, the patient's stomach and diaphragm were punctured resulting in sepsis. The patient underwent a 5 month hospital stay and eighteen months as a near invalid. To date, (b)(6) has not provided the patient's husband with information regarding this event. Intuitive has contacted the sites risk management department, who advised that they would investigate this event and respond back to intuitive within the upcoming weeks. | Based on the information provided, it was undetermined as to how this event occured. The site has been contacted for additional information and once this information has been received, a follow up report will be sent to the fda. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2134370 | 2105804 | 2002780 | 2955842201100160 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2011-05-06 | 2011-05-27 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2011-05-06 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-05-03 | Other | Other | It was reported that after a da vinci surgical procedure, the patient was found to have charring at the 8mm port incision site. The patient did not require any additional medical treatment. | During our investigation with the site's rn, (b)(6), she advised that the port incisions reported by our isi field service engineer may have been intentionally caused during creation of the port incisions using a bovi unit. She advised that she would follow up with the surgical staff to obtain additional information. Despite following up with the site on several occasions, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2134351 | 2105813 | 2002789 | 2955842201100161 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-03 | 2011-05-27 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2011-05-03 | Device Not Retu | No | Reuse | 2011
400183 | 2011-03-14 | Malfunction |  | On may 3, 2011, medwatch (b)(4) was received with the following information: event desc: patient undergoing a robotic sparing whipple and cholecystectomy for a pancreatic mass. During the procedure a piece of the hook cautery was noted to be broken off and in the patient's abdominal cavity. Copious irrigation and suction were performed. It is unknown whether the piece was retained or removed. No additional information has been received despite several attempts to contact the site. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2134352 | 2105814 | 1968844 | 2955842201100162 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-05-03 | 2011-05-27 | No | Yes | Health Professional | 400183 | No | No | 2011-05-03 | No | No | Unkown | 2011
42018405 | 2011-04-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the permanent cautery spatula instrument broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PERMANENT CAUTERY SPATUAL INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2131566 | 2103542 | 2000518 | 2955842201100157 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-27 | 2011-05-26 | No | Yes | Health Professional | 42018405 | No | Yes | 2011-04-27 | No | No | Unkown | 2011
40018010 | 2011-05-24 | Other | Other | (b)(4): while the surgeon was using the da vinci robot, the monopolar scissors split a hole in three protective tip covers. When the third tip cover tore, a small burn was noted in the patient's bowel. What was the original intended procedure? total abdominal hysterectomy with bilateral salpingo-oophorectomy. Device usage problem: device malfunction - that is, the device did not do what it was supposed to do. The user facility had already reported this event to intuitive surgical (isi) on (b)(6) 2011. It is the hospital's belief that the patient injury's was caused by the monopolar curved scissors (mcs) instrument. Repair to the patient's damaged bowel was not required and no postoperative complications have been experienced by the patient. | The instrument was returned and evaluated. Engineering found no signs of arcing on the instrument tube extension, clevis, scissor tips, nor the tube extension to main tube connection. The instrument was driven on a system and functioned properly with no trouble found. The hospital informed isi that the three tip cover accessories would not be returned for evaluation and isi's request for photographic images was denied. The hospital did allow an isi representative to evaluate the tip covers at the site. Although tears were noticed on all of the tip covers, no evidence of charring/arcing was observed. As no evidence of arcing was observed on the instrument and tip covers, the root cause of the patient bowel burn cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2132127 | 2104010 | 2000986 | 2955842201100158 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-26 | 2011-05-26 | Yes | No | Health Professional | 40018010 | Yes | No | 2011-04-26 | Device Not Retu | Yes | Unkown | 2011
42009308 | 2011-04-26 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the instruments installed on both instrument arms did not have sufficient closing force on the instrument jaws to grasp tissue or throw a stitch. The surgical staff reseated the instruments, with no change. The instruments were replaced, with no change. The sterile adapters were then reseated, which resolved the issue. Later in the procedure, the wrist cables of the prograsp forceps instrument were stuck on the cannula tip, thus preventing removal of the instrument. The instrument arm was undocked from the patient, which allowed the surgical staff to withdraw the instrument from the cannula. After the instrument was removed, there was some difficulty redocking the instrument arm to the patient. Due to the instrument issues experienced by the surgeon, the decision was made to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing and the clevis hub exhibits a wear mark adjacent to the cable groove. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2132478 | 2104261 | 2001237 | 2955842201100159 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-26 | 2011-05-26 | Yes | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2011-04-26 | Device Not Retu | No | Reuse | 2011
420184 | 2011-04-27 | Malfunction |  | Patient underwent gynecological surgery utilizing da vinci robotics. An intuitive cautery spatula was utilized during the surgery. While moving the camera and the cautery spatula within the abdominal cavity, the two objects came in contact and a small plastic part was dislodged from the tip of the cautery spatula. It was not possible to identify the piece and the surgeon elected to leave it in the abdominal cavity. This was the instrument's final use. ======================manufacturer response for endowrist monopolar cautery instrument, permanent cautery spatula (per site reporter)======================the representative who works in the hospital will let his supervisor know about the issue. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST MONOPOLAR CAUTERY IN | 2138433 | 2110633 | 2007609 | 2110633 | 1 | NAY | User Facility |  |  | Initial | 2011-05-26 | 2011-05-26 | No | Yes | Physician | 420184 |  | No Answer Provided |  |  |  |  | 2011
| 2011-05-24 | Injury | Disability | During robotic prostatectomy, it was noted that the large needle driver wasn't holding and the suture was slipping as the physician was tying. The needle holder was removed and a new on placed. The physician then identified a small dark piece of material in the lower abdomen. It was removed and noted to be a small piece from the tip of the needle driver. No injury to the pt. |  | LARGE NEEDLE DRIVER 8MM IS2000 | LARGE NEEDLE DRIVER | 2144394 | 2113223 | 2010199 | MW5020815 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-05-26 | 2011-05-26 | No | Yes | Health Professional |  | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
| 2011-05-11 | Malfunction |  | No energy delivered to robotic monopolar device via the energy cable. This is the second cable from the manufactuer that has failed to function properly, in the last several weeks. The staff firmly believe the failure is due to a design flaw. ======================manufacturer response for energy source cable, energy activation cable, valleylab (per site reporter)======================the informal response from the engineer and the sales rep is that it is "user error". |  | ENERGY ACTIVATION CABLE VALLEYLAB | ENERGY SOURCE CABLE | 2132912 | 2104769 | 2001745 | 2104769 | 1 | GEI | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2011
420093 | 2011-02-24 | Malfunction |  | Patient underwent a da vinci-assisted laparoscopic bilateral pelvic lymph node dissection, extraction of bladder stones times four and radical prostatectomy, nonnerve-sparing. During the procedure, a prograsp forcep broke within a trocar. It was not possible to remove the device until the instrument within the trocar had been straightened with one of the other robotic arms. It was noted that there was a small foreign body loose within the abdomen after the forcep broke. Before the end of the procedure, the foreign body was located and removed. This was the 9th use with expected use of 10 times. ======================manufacturer response for endowrist grasper instrument, prograsp forceps 8 mm (per site reporter)======================none. |  | PROGRASP FORCEPS | ENDOWRIST GRASPERS INSTRUMENT | 2137915 | 2110066 | 2007042 | 2110066 | 1 | NAY | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Physician | 420093 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420183 | 2011-05-09 | Malfunction |  | Patient undergoing robotic myomectomy. The da vinci permanent cautery hook broke into three pieces in the intra-abdominal cavity. A scope was utilized and the cautery hook pieces were recovered. This was the 7th time this instrument had been used - expected life is ten uses. ======================manufacturer response for endowrist monopolar cautery instrument, permanent cautery hook (per site reporter)======================manufacturer has not responded as of the date of this notice. |  | PERMANENT CAUTERY HOOK | ENDOWRIST MONOPOLAR CAUTERY I | 2137917 | 2110067 | 2007043 | 2110067 | 1 | NAY | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Physician | 420183 | Yes | No Answer Provided |  |  |  |  | 2011
420184 | 2011-05-17 | Malfunction |  | When removing endowrist instrument from cannula, spatula was not coming out. Tip of instrument missing a piece on visualization. Md notified and inspected op site. Company representative was in room and is aware of issue. Called company. Device not radiolucent. Risk md from company said tip will not harm patient. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST MONOPOLAR CAUTERY IN | 2138482 | 2110664 | 2007640 | 2110664 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-23 | 2011-05-23 | No | Yes | Physician | 420184 | No | No Answer Provided |  |  |  |  | 2011
42018304 | 2011-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the permanent cautery hook (pch) instrument broke and a piece of the plastic clevis fell into the patient. The fragment was retrieved and the pch instrument was replaced with another pch instrument. The replacement instrument also broke and fragments fell into the patient. The surgical staff believes that all fragments were found, but cannot be certain. Postoperatively, the patient is reportedly doing well and has not experienced any complications. This report is for the initial pch instrument used during the surgical procedure. Medwatch mfr report 2955842-2011-00151 was submitted for the second pch instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis has one ear broken off at its base. The ear that is located on the conductor wire side broke off. The broken piece was not returned. The size of the missing ear piece is approximately. 280 x. 220. As a result of the breakage, the yaw pulley is dislodged and no longer axially aligned with the shaft. Engineering also observed that the conductor cap is missing, exposing the conductor wire. The yaw pulley boss that mates with the hole in the cap is also broken. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2126128 | 2097326 | 1994302 | 2955842201100150 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the permanent cautery hook (pch) instrument broke and a piece of the plastic clevis fell into the patient. The fragment was retrieved and the pch instrument was replaced with another pch instrument. The replacement instrument also broke and fragments fell into the patient. The surgical staff believes that all fragments were found, but cannot be certain. Postoperatively, the patient is reportedly doing well and has not experienced any complications. This report is for the second pch instrument used during the surgical procedure. Medwatch mfr report 2955842-2011-00150 was submitted for the initial pch instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off on the conductor wire side. The broken ear piece and conductor cap were returned detached from the distal end. The size of the broken ear is approximately. 285 x. 240. As a result of the breakage, the yaw pulley is dislodged and no longer axially aligned with the shaft. The yaw pulley boss feature that holds the cap in place is also broken. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2126130 | 2097327 | 1994302 | 2955842201100151 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | No | Reuse | 2011
420142 | 2011-03-08 | Other | Other | It was reported that during a da vinci s radical tonsillectomy procedure, the surgeon noticed the 5mm cautery instrument was not energizing at the tips as intended. The site checked the electrical surgical unit (esu) setting and cables and decided to increase the level of cautery. The surgical staff then noticed that cautery was transferring to the tongue retractor and dissipating throughout the patient's mouth. Unintended tissue was 'charred'; however, no significant injury was reported and the patient did not require any additional treatment or procedure. The planned surgical procedure was completed using the same instrument, however, the site created an insulation barrier out of operating room materials to cover the instrument snake wrist and ensure cautery leakage did not recur. No additional patient injury was reported. | Multiple attempts have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | 5MM MONOPOLAR CAUTERY INSTRUME | ENDOSCOPIC ELECTROSURGICAL INS | 2126689 | 2097936 | 1994912 | 2955842201100149 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-05-20 | Yes | Yes | Health Professional | 420142 | No | No | 2011-03-08 | No | No | Unkown | 2011
IS120043P9 | 2011-04-26 | Other | Other | It was reported that prior to starting a da vinci pelvic floor repair procedure, the surgical staff had difficulty aligning the endoscope. With the assistance of an isi technical support engineer (tse), the site exchanged the camera head and was then successfully able to align the endoscope. Approximately 1 hour later, the site called back stating the surgeon was experiencing double vision in the left eye of the surgeon side console (ssc) viewer. The tse had the site exchange and recalibrate the endoscope and restart the vision side cart. The image appeared normal and the surgeon continued with the planned procedure. Approximately 30 minutes later, the site called back stating the surgeon was experiencing black bars in the left eye, causing double vision. The patient was under anesthesia for approximately 2 hours when the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was related to the synchronizer. The synchronizer coordinates the left eye and right eye video streams so that events are displayed simultaneously to each of the surgeon's eyes. The system was repaired by replacing the affected synchronizer. The synchronizer was returned to isi for failure analysis investigation and engineering was able to replicate the customer reported issue. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2126690 | 2097940 | 1994916 | 2955842201100152 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-26 | 2011-05-20 | Yes | Yes | Health Professional | IS120043P9 | Device Returned To Manufacturer | Yes | 2011-04-26 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-04-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda in relation to this case: 2955842-2011-00154. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover to have multiple holes burned through the silicone, located along one parting line. The silicone exhibits localized melting around the holes indicative of arcing. The holes are irregularly shaped, ranging from. 045 inches to. 065 inches and are located. 280 inches and. 365 inches above the silicone to sleeve interface. The monopolar curved scissors (mcs) instrument used with the tip cover accessory in this procedure was also returned. With the tip cover accessory installed on the mcs instrument, the hole positions line up with the edge of the proximal clevis ear and the distal idler pulley. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2126691 | 2097945 | 1994931 | 2955842201100153 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-04-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The mcs instrument and tip cover accessory were replaced, however, the site experienced arcing from the replacement mcs instrument with a tip cover accessory installed. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda in relation to this case: 2955842-2011-00153. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2126704 | 2097955 | 1994931 | 2955842201100154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 40018010 | No | Yes | 2011-04-21 | No | Yes | Unkown | 2011
42019408 | 2011-04-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy, the pad from the tip of the mega needle driver instrument, fell into the patient and was retrieved. The planned surgical procedure was completed with no further incidents. No injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering's inspection found the instrument to be missing one carbide insert from grip tip. The surface of the separation showed no residue of adhesive material. No damage was found on clevis or grip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2126787 | 2098018 | 1994994 | 2955842201100155 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-28 | 2011-05-20 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2011-04-28 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-04-21 | Other | Other | It was reported that during set up for a da vinci s nephrectomy procedure, the surgical staff experienced the left eye on the surgeon side console to be gray with a line down the middle of the screen. With the assistance of an isi technical support engineer, the site verified their settings and tested the camera cable which was determined to need replacement. The patient was under anesthesia when the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of may 26, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2126788 | 2098019 | 1994995 | 2955842201100156 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-21 | 2011-05-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-04-21 | Device Not Retu | No | Reuse | 2011
420227 | 2011-05-18 | Malfunction |  | The forceps worked during the surgery. The thought is that the defect occurred during the cleaning process. There is a small black wire that broke during what the staff believe when cleaned. The forceps will not open and have been taken out of service. |  | PK DISSECTING FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2135783 | 2107330 | 2004306 | 2107330 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-20 | 2011-05-20 | No | Yes | Invalid Data | 420227 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-04-19 | Other | Other | It was reported that approximately 90 minutes into a da vinci s diverticulum surgical procedure, the site experienced system error code 19 followed by system error code 45049. With the assistance of an isi technical support engineer (tse) the site restarted the system, however, upon power up system error code 281 appeared. At this time, the site stated that a small amount of water had been spilled on the base of the patient side cart. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes experienced by the customer were associated with two remote arm controller (rac) pca boards and the patient cart controller (pcc). The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The pcc is a board embedded in the patient cart that provides all user interface, communications control, and patient cart motor control for the patient side system. The system was repaired by replacing the affected rac pca boards and the pcc board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error codes reported occur when the da vinci s safety system determines a hardware fault reaction logic occurring on a local rac. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The rac3, rac4, and pcc were returned to isi for failure analysis investigation. Engineering evaluation confirmed fluid contamination as reported by the site. The da vinci s surgical system user manual specifically states: caution: equipment is not designated for exposure to liquids. Care should be taken to ensure liquids do not flow into electronic equipment. On (b)(4) 2011, field service engineering had extensive conversations with the site regarding proper cleaning of the da vinci s system to eliminate future occurrences of this event. As of may 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2122927 | 2091980 | 1988956 | 2955842201100148 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-19 | 2011-05-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-04-19 | Device Not Retu | No | Reuse | 2011
42020703 | 2011-03-02 | Malfunction |  | (b)(4). Using the da vinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. On (b)(4) 2011, the initial reporter informed intuitive surgical that all broken pieces were retrieved and there was no injury to the patient. | The instrument was returned with the proximal clevis detached from the distal end. The proximal clevis was also returned, in three broken pieces. Engineering evaluation observed that the main tube interface wall has collapsed. As a result of the breakage, the wrist is loose and motion is non-intuitive. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2121184 | 2089892 | 1986868 | 2955842201100146 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-18 | 2011-05-16 | No | Yes | Health Professional | 42020703 | Device Returned To Manufacturer | No | 2011-04-18 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-04-19 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced system error code 21008. The system was restarted, however, the error code recurred. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 21008 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controller which provides the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 21008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. The mtmr was returned for failure analysis investigation. Engineering was able to replicate the customer reported failure mode and determined that an axis potentiometer and motor/encoder had malfunctioned. As of may 12, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2121185 | 2089893 | 1986869 | 2955842201100147 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-04-19 | 2011-05-16 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-04-19 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-04-11 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the site observed arcing from the shaft of the monopolar curved scissors (mcs) instrument. The patient experienced a superficial burn to their bowel which did not require any repair or treatment. The planned surgical procedure was successfully completed without incident. The patient was reported as doing well with no post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed traces of arcing at the joint between the tube extension and main tube. The material was slightly melted and the tube extension is dislodged from the main tube. No trace of arcing was found at clevis or scissor. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2120513 | 2089035 | 1986011 | 2955842201100144 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-13 | 2011-05-13 | No | Yes | Health Professional | 42017910 | No | No | 2011-04-13 | No | Yes | Initial | 2011
42018304 | 2011-04-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, a piece of the permanent cautery hook instrument fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2120523 | 2089044 | 1986020 | 2955842201100145 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-15 | 2011-05-13 | No | Yes | Health Professional | 42018304 | No | No | 2011-04-15 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-01-20 | Other | Other | It was reported that on (b)(6) 2011, a da vinci s pancreatoduodenectomy was successfully performed with no system malfunctions. On (b)(6) 2011, the patient's drain amylase level rose to 10,480. Additionally, pancreatic fistula from the anastomosis site between the pancreas and the jejunum was found. As the drainage during the procedure was proper, additional treatment was not performed. The patient was discharged on (b)(6) 2011. | The console surgeon who performed the procedure indicated that the postsurgical symptoms experienced by the patient can develop regardless of the modality used to perform the surgery and believes that the patient's symptoms are unrelated to the use of the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunctions occurred, and the system did not malfunction in a way that may have caused or contributed to the patient's postsurgical symptoms. Intuitive surgical concluded that the patient's symptoms are unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2115107 | 2083893 | 1980869 | 2955842201100141 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-04-11 | 2011-05-11 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-04-11 | No | No | Reuse | 2011
400180 | 2011-04-12 | Malfunction | Other | It was reported that after the right side lymphectomy dissection was completed during a da vinci s hysterectomy procedure, arcing was observed on the tip cover accessory that was installed on the monopolar curved scissors (mcs) instrument. No patient harm occurred during the arcing event and the planned surgical procedure was completed. | The tip cover accessory was discarded by the site, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2115109 | 2083909 | 1980885 | 2955842201100142 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-12 | 2011-05-11 | No | Yes | Health Professional | 400180 | No | No | 2011-04-12 | No | Yes | Unkown | 2011
42017207 | 2011-04-12 | Injury | Other | It was reported that while dissecting the right side "lymphectomy" during a da vinci s hysterectomy procedure a tear was noticed halfway through the patient's right iliac vein. The vein was repaired laparoscopically using 12 clips and floseal. The planned surgical procedure was completed and it was reported that no postoperative complications have been experienced by the patient. | A video recording of the procedure was reviewed by an intuitive surgical clinical sales representative and the console surgeon. The video revealed that the surgeon inadvertently grasped the patient's right iliac vein with a maryland bipolar forceps instrument. It is the surgeon's belief that this action caused the damage to the patient's vein. No issues were observed with the maryland bipolar forceps instrument. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2115110 | 2083910 | 1980886 | 2955842201100143 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-12 | 2011-05-11 | Yes | No | Health Professional | 42017207 | No | No | 2011-04-12 | No | No | Reuse | 2011
420290 | 2011-05-05 | Malfunction |  | Patient with ureteral obstruction on left side has ureteral stent in place and unexplained hydroureter w/confirmed obstruction. Proceed with ureteral implantation. Surgeon had robotic left ureteral reimplantation/robotic nephroureterectomy and during case the green cannula seals were leaking and needed to be changed out. They came from davinci kit accessory 3-arm which was opened for the procedure. Two individual green cannula seals were open to replace the ones leaking. |  | DISPOSABLE ACCESSORY KIT | DISPOSABLE ACCESSORY KIT 3ARM | 2124457 | 2094959 | 1991935 | 2094959 | 1 | NAY | User Facility |  |  | Initial | 2011-05-10 | 2011-05-10 | No | Yes | Physician | 420290 | No | No Answer Provided |  |  |  |  | 2011
40016002 | 2011-04-12 | Malfunction |  | Event Description: --/-- It was reported that while resecting tissue during a da vinci si tongue base resection procedure, the surgeon observed a flame on the 5mm cautery spatula tip accessory that was installed on the 5mm monopolar cautery instrument. The instrument and tip accessory were immediately removed and replaced and the planned surgical procedure was completed with a replacement instrument and tip accessory. No patient harm was reported. | Manufacturer Narrative: --/-- A video of the surgical procedure was provided to intuitive surgical by the hospital. Engineering reviewed the video and observed that on multiple occasions the surgeon activated the energy for the 5mm monopolar cautery instrument with a bipolar instrument in close proximity to its spatula tip. As a result, the spatula tip arcs to the bipolar instrument at various times, consequently degrading the insulation barrier on the spatula and leading to the malfunction experienced by the customer. The instruments and accessories user manual specifically states in the 5mm monopolar cautery instructions for use: warning: as with any cautery device, it is possible for energy to discharge in an area other than the instrument tip. It is important to exercise caution when using an energized endowrist 5 mm monopolar cautery instrument to help avoid unintended contact with tissue adjacent to the area to be cauterized. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip accessory were returned and evaluated. Per the customer reported complaint, engineering observed localized melting with material removed on one side of the plastic sleeve of the spatula tip. The damage indicates that an arcing event occurred. The instrument was also inspected and no physical damage was observed. Engineering performed functional tests with the instrument using a new spatula tip and the cautery performed as intended with no occurrence of arcing. Additional investigation was performed by a field service engineer, who went onsite and evaluated the system. No errors were found when reviewing the system error logs and no issues were observed during inspection of the system arm that the instrument was installed on during the surgical procedure. The electrocautery cables were also inspected and no damage was found. --/-- Additional Manufacturer Narrative: --/-- Correction: a non-energy instrument was used with the 5mm monopolar cautery spatula instrument when the incident occurred. Medwatch mfr report 2955842-2011-00134, follow-up 01, indicated that an energy (bipolar) instrument was used. Additional information: further evaluation of the surgical procedure video was performed by engineering. While reviewing the video, engineering observed that the non-energy instrument was used repeatedly to scrape off/clean the spatula tip, which could have contributed to the tip damage. Prior to the incident occurring, multiple arcing events were observed and chunks of silicone from the tip had fallen off of the spatula and were visible in the surgeon's field of view. The instrument damage can clearly be seen, however, the surgeon continues to operate. The instruments and accessories user manual specifically states under general precautions and warnings: - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM CAUTERY SPATULA TIP ACCESSORY | Initial | 2113145 | 2082089 | 1979065 | 2955842201100134 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-05-16 | 2011-05-09 | No | Yes | Health Professional | 40016002 | Device Returned To Manufacturer | No | 2011-06-17 | Device Not Retu | Yes |  | 2011
IS1200 A43P9 | 2011-04-07 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure, while the surgeon was starting to dissect the patient's bladder, the site experienced a system error code 212. With the assistance of an isi technical support engineer, the restarted the system and then performed an emergency power off (epo) however, the system error code 212 reoccurred. The patient was under anesthesia and port incisions had been placed when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator (mtml) arm. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering was able to confirm the occurrence of system error code 212, thus confirming the reported failure mode experienced by the customer. The mtml main harness and axis 2 motor were replaced. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111887 | 2080463 | 1977440 | 2955842201100135 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-07 | 2011-05-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-04-07 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-04-13 | Other | Other | It was reported that during the set up for a da vinci si hysterectomy procedure the site experienced system error 23027. The surgical staff pressed fault over ride and continued with the system set up. The system was docked to the patient and the surgeon began to work from the surgeon side console when the system error 23027 reoccurred. With the assistance of an isi technical support engineer the site back drove the master tool manipulators (mtms) with the system on and then off, however the system error 23027 continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23027 experienced by the customer was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23027 system error code appears when the da vinci si safety system determines that one of the health checks on a secondary joint sensor on the arm did not pass. Upon determining these conditions, the safety systems put da vinci si in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering determined that the flex circuit going from the axis 2 potentiometer to the embedded serializer (esmb) was loose, thus generating the system error code experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111889 | 2080464 | 1977441 | 2955842201100136 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2011-04-13 | 2011-05-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-04-13 | Device Not Retu | No | Reuse | 2011
40027202 | 2011-04-05 | Malfunction |  | It was reported that during a da vinci s colorectal procedure, while the surgeon was using the harmonic curved shears instrument with the harmonic ace insert, the tip broke and fell into the patient. The fragment was retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found the curved blade to be broken at the distal end with a 0. 415 inch by. 085 inch piece detached. The blade fractured slightly above the straight to curved transition point. The broken piece exhibits various scratch and gouge marks on the side that make contact with the clamp arm. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: "handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. " do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2111890 | 2080467 | 1977444 | 2955842201100137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-11 | 2011-05-06 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-04-11 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-04-08 | Other | Other | It was reported that during a da vinci s low anterior resection procedure, the site experienced a red image at the surgeon side console. The surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted found that the vision system was out of tolerance due to the camera controller unit (ccu) not calibrating. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111867 | 2080475 | 1977451 | 2955842201100138 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-04-08 | 2011-05-06 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-04-08 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-04-11 | Other | Other | It was reported that while starting a da vinci prostatectomy procedure, the site found a broken cable on a patient side manipulator (psm) arm. The patient was under anesthesia when the surgical staff decided to abort the planned procedure. No patient harm, injury or adverse outcome was reported. | The investigation conducted by field service engineering (fse) confirmed a broken cable on the patient side manipulator (psm) 2 arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the axis 6 cable to be broken. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111894 | 2080493 | 1977470 | 2955842201100139 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-04-11 | 2011-05-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-04-11 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-04-07 | Malfunction |  | It was reported that during a da vinci s sigmoid colectomy, an unidentified item from the harmonic curved shears insert instrument fell into the patient and was immediately retrieved. There has been no additional information provided to date. No patient harm, adverse outcome, or injury was reported. | Only the insert was returned and evaluated, but not the clamp arm. Per the customer reported complaint, engineering found the instrument clamp arm missing from the distal end. Clamp arm pins were bent backwards. The cause of the user damage was likely due to the user mishandling and/or over extension of the arm. There was no other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL | 2112073 | 2080766 | 1977743 | 2955842201100140 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-07 | 2011-05-06 | No | Yes | Other | 40016907 | Device Returned To Manufacturer | No | 2011-04-13 | Device Not Retu | No | Reuse | 2011
| 2011-04-25 | Malfunction |  | While the surgeon was using the davinci robot, the monopolar scissors split a hole in three protective tip covers. When the third tip cover tore, a small burn was noted in the patient's bowel. |  | HOT SHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2117890 | 2086793 | 1983769 | 2086793 | 1 | NAY | User Facility |  |  | Initial | 2011-05-04 | 2011-05-04 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420206 | 2011-05-04 | Malfunction |  | Event Description: --/-- Patient underwent robotic assisted hysterectomy. After placing trocars in patient, the robotic instruments were introduced. One of the green robotic seals began to leak as soon as the instrument was put in. Instrument was removed, seal reseated, instrument put back in, continued to have a significant leak of c02, which was audible. Seal was removed and replaced with a new one. Leak was resolved with new seal. --/-- Additional Event Description: --/-- Patient underwent robotic assisted hysterectomy. After placing trocars in patient, the robotic instruments were introduced. One of the green robotic seals began to leak as soon as the instrument was put in. Instrument was removed, seal reseated, instrument put back in, continued to have a significant leak of co2, which was audible. Seal was removed and replaced with a new one. Leak was resolved with new seal. | Manufacturer Narrative: --/-- | CANNULA SEAL 12MM |  | 2120834 | 2089511 | 1986487 | 2089511 | 1 | NAY | User Facility |  |  | Initial | 2011-05-04 | 2011-05-04 | No | Yes | Physician | 420206 | No | No Answer Provided |  |  |  |  | 2011
420205 | 2011-04-06 | Malfunction |  | Device never worked. Or staff will contact rep to come and get device. |  | FENESTRATED BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2117898 | 2086784 | 1983760 | 2086784 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-02 | 2011-05-02 | No | Yes | Physician | 420205 | Yes | No Answer Provided |  |  |  |  | 2011
420110 | 2011-04-11 | Malfunction |  | Could not get device to work. Same type of instrument (new) was used to complete procedure. No harm to patient. Or staff will call rep to come get device. |  | PRECISE BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2117901 | 2086787 | 1983763 | 2086787 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-02 | 2011-05-02 | No | Yes | Physician | 420110 | Yes | No Answer Provided |  |  |  |  | 2011
42020706 | 2011-03-23 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a piece of 'wire' from the tenaculum forceps instrument broke and fell into the patient. The fragment was retrieved, and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the pitch down cable to be broken at the distal clevis hub. The cable segment that contains the crimp was returned detached from the instrument. Other cables at the wrist are not damaged and the clevis does not exhibit any damage or wear. No other damage found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2103931 | 2072701 | 1969677 | 2955842201100127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-01 | 2011-04-29 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2011-04-01 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-03-30 | Other | Other | It was reported that while starting a da vinci si prostatectomy procedure, the site experienced a system error code 25580. The patient was under anesthesia, and port incisions had been placed when the surgical staff made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the customer was associated with a remote arm controller (rac) pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25580 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of april 28, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103915 | 2072702 | 1969678 | 2955842201100129 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-30 | 2011-04-29 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-03-30 | Device Not Retu | No | Reuse | 2011
42019407 | 2011-04-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was suturing and using the mega needle driver instrument, a break occurred in the instrument shaft and 'plastic' fragments fell into the patient's abdomen. One piece was retrieved, however, tiny fragments were unable to be removed and remain in the patient. The instrument was replaced with another instrument of the same type and the planned surgical procedure was completed. No injury was reported. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and/or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2103908 | 2072729 | 1969705 | 2955842201100126 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-05 | 2011-04-29 | No | Yes | Health Professional | 42019407 | No | Yes | 2011-04-05 | No | No | Unkown | 2011
42017909 | 2011-03-31 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the 'jaw' of the monopolar curved scissors instrument partially broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a segment of the right blade to be broken off from the instrument. The broken piece was not returned. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2103910 | 2072730 | 1969706 | 2955842201100128 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2011-03-31 | 2011-04-29 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2011-03-31 | Device Not Retu | No | Reuse | 2011
IS2000 51P7 | 2011-03-31 | Other | Other | It was reported that during a da vinci s hysterectomy procedure whenever the surgeon moved a specific patient side manipulator (psm) arm, system error 23008 would appear. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the right master tool manipulator (mtmr) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a nonrecoverable safe state. The mtmr gimbal was returned to isi for evaluation. Engineering confirmed that the customer reported failure mode was associated with a defective sensor on axis 5 and the sensor was replaced. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103916 | 2072734 | 1969710 | 2955842201100130 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-31 | 2011-04-29 | Yes | Yes | Health Professional | IS2000 51P7 | Device Returned To Manufacturer | No | 2011-03-31 | Device Not Retu | No | Reuse | 2011
42020505 | 2011-03-30 | Malfunction |  | It was reported that while resecting during a da vinci s hysterectomy procedure, the fenestrated bipolar forceps instrument tip broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one instrument grip broken off near the top of the yaw pulley. The broken grip appears to be bent slightly backwards when the fracture planes are matched up, indicating that the damage may have been caused by mishandling. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2103970 | 2072779 | 1969755 | 2955842201100131 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-03-30 | 2011-04-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2011-03-30 | Device Not Retu | No | Reuse | 2011
42017207 | 2011-03-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure a piece of the maryland bipolar forceps instrument fell into the patient. The fragmented piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a piece of the yaw pulley cover detached from the instrument. The broken piece is approximately. 240 x. 200 in size. The broken piece had fractured around the distal pin and the yaw pulley cover has two broken sections still within the clevis. The belleville washer on the broken side has no load on one side and as a result of the breakage, is tilted within the clevis. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Initial | 2103972 | 2072780 | 1969756 | 2955842201100132 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2011-03-30 | 2011-04-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2011-04-29 | Device Not Retu | No |  | 2011
IS2000 A51P8 | 2011-03-30 | Injury | Required Intervention | It was reported that during a da vinci s prostatectomy procedure a very small amount of blood was observed on the patient's descending colon, around the sigmoid. Another surgeon was brought in to assess the issue and confirmed that the patient's bowel had been perforated. Since the damage was located on the underside of the patient's bowel, visibility was limited and the surgeon made the decision to convert to traditional open surgical techniques to repair the damaged bowel and complete the planned procedure. Postoperatively, 4 units of blood were given to the patient to help with their bowel recovery and bleeding from the deep dorsal vein complex. It was reported that the patient is in stable condition. | Additional information was provided by the console surgeon, who indicated that an instrument may have touched the patient's bowel during an instrument exchange, however, no issues were observed during the exchange. Based on the information provided, it was determined that the da vinci s surgical system worked as intended. No information has been received which suggests that the da vinci s surgical system malfunctioned in a way that may have caused or contributed to the patient's injury. Intuitive surgical concluded that the patient injury was unrelated to the performance of the da vinci s surgical system. The da vinci s surgical system user's manual specifically states: caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the patient. Warning: the instrument may not be immediately visible when being moved from the cannula into the patient. Use appropriate caution when manually inserting instruments into the patient. Caution: removal of instruments during a procedure should be done very carefully and only with the knowledge of and full view of the surgeon console operator. Prior to instrument removal, the surgeon console operator should do the following: ensure the instrument is free and clear of any patient anatomy. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103976 | 2072781 | 1969757 | 2955842201100133 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-03-30 | 2011-04-29 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-30 | No | No | Reuse | 2011
| 2011-04-14 | Malfunction |  | Robotic large hemolock clip applier wire broke during surgical procedure. |  | INTUITIVE | LARGE HEMOLOCK CLIP APPLIER | 2109427 | 2075961 | 1972938 | MW5020447 | 1 | NAY | Voluntary |  | Nurse | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
400169 | 2011-04-26 | Malfunction |  | During a robotic laparoscopic supercervical hysterectomy, the harmonic ace curved shears insert's insulation piece broke off while being used. The piece was recovered without harm to patient. |  | HARMONIC CURVED SHEARS INSERT | HARMONIC CURVED SHEARS INSERT | 2110232 | 2079033 | 1976010 | 2079033 | 1 | NAY | User Facility |  |  | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2011
COR31 12X150MM | 2011-04-25 | Malfunction |  | During a robotic surgery, the physician moved the camera around inside the abdomen and noted tissue up against it, obstructing the view. So, the scrub tech went to move the port and advance it when they heard a noise and noted the port had cracked and a piece or two flew away from the patient. They had been losing gas; having problems maintaining pneumoperitoneum during the whole case and believe that there was a crack in the port the entire time, creating the loss of gas. |  | COR31 12X150MM KII OPTICAL ACCESS SYSTEM | TROCAR PORT | 2113698 | 2082589 | 1979565 | 2082589 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Invalid Data | COR31 12X150MM |  | No Answer Provided |  |  |  |  | 2011
42017207 | 2011-03-29 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the maryland bipolar forceps instrument, the tip broke off and fell into the patient. The fragment was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. (b)(4). | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available or the instrument is returned for evaluation, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2098284 | 2067674 | 1964650 | 2955842201100118 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-29 | 2011-04-25 | No | Yes | Health Professional | 42017207 | No | Yes | 2011-03-29 | No | No | Unkown | 2011
42000606 | 2011-03-14 | Malfunction |  | It was reported that during a da vinci si left renal diverticulectomy, while the surgeon was transferring a needle and using the large needle driver instrument, the needle became dislodged from the instrument and fell into the patient. Approximately one week later, the patient returned to the hospital for an x-ray to locate the needle and operation to have it removed. No adverse event was reported. | The instrument was returned and evaluated. Engineer was unable to confirm the customer reported complaint as the instrument was found to be fully functional. The instrument was placed on a test system and moved intuitively with full range of motion in all directions. The grips opened and closed properly and sample sutures were successfully made. No damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2098285 | 2067675 | 1964651 | 2955842201100120 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-29 | 2011-04-25 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2011-03-29 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-03-28 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the inner aspects of the mega needle driver instrument came off and fell into the patient. The broken pieces were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the carbide inserts on both grip tips detached from the metal surface. The metal surface is smooth and no residue of adhesive material was found. No damage was observed on the clevis and cable. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2099404 | 2069126 | 1966102 | 2955842201100125 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-28 | 2011-04-25 | No | Yes | Other | 42019408 | Device Returned To Manufacturer | No | 2011-03-28 | Device Not Retu | No | Reuse | 2011
| 2011-03-22 | Injury | HospitalizationOtherLife ThreateningRequired Intervention | A (b)(6) obese patient with high grade endometrial cancer underwent a robotic-assisted total laparoscopic hysterectomy, bso, bilateral pelvic lymphadenectomy using the davinci si surgical system. The robotic hot shears (monopolar scissors) were used in the right instrument port and bipolar forceps in the left instrument port. During mobilization and removal of the right obturator lymph nodes, arcing of energy through the protective sheath of the shaft of the hot shears was noted, causing direct thermal damage to the distal right external iliac vein. A laparotomy was performed and a vascular surgeon ligated the right external iliac vein secondary to the extent of damage caused, limited exposure and the location of injury. Postoperatively, the patient developed a right lower extremity deep venous thrombosis and was placed on anticoagulation. Dates of use: |  | HOT SHEARS | MONOPOLAR ROBOTIC SCISSORS | 2105292 | 2071831 | 1968807 | MW5020376 | 1 | NAY | Voluntary |  | Physician | Initial | 2011-04-23 | 2011-04-23 | Yes | Yes | Health Professional |  | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
400169 | 2011-03-03 | Malfunction |  | On (b)(4) 2011, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from california pacific medical center. The following problem description was provided on the report: during a total abdominal hysterectomy, the jaw of robotic harmonic scalpel instrument broke at joint. Loose piece was retrieved without incident. Complete match to non-damaged instrument when visually compared. Fracture of instrument monitored via robot camera. No foreign material remains in the patient per surgeon and assistant. Fracture occurred while surgeon using instrument to dissect tissue. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2096202 | 2065915 | 1962891 | 2955842201100112 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-03-25 | 2011-04-21 | No | Yes | Health Professional | 400169 | No | No | 2011-03-25 | No | Yes | Unkown | 2011
40018010 | 2011-03-22 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, while performing the right side pelvic lymph node dissection, and while using the monopolar curved scissors instrument with a tip cover accessory installed, the patient's external iliac vein was punctured. A vascular surgeon was called to assist and elected to convert to traditional open surgical techniques to repair the patient's iliac injury. The patient required a blood transfusion but was reported as stable and recovering well. | The tip cover accessory has not been returned for evaluation, however the monopolar curved scissors (mcs) instrument said to be used with the tip cover accessory during this procedure was returned. Per the customer reported complaint engineering observed the mcs instrument proximal clevis exhibits a char mark adjacent to the rectangular window that exposes the seal, approximately. 130 inches above the interface with the tube extension. The seal has a section of material removed at one corner, adjacent to the char mark. The char mark is located near the silicone to sleeve interface on a tip cover when installed on the mcs instrument. No other damage was found. The isi representative onsite during the procedure did not witness any arcing from the mcs tip cover. Additionally, the tip cover accessory was inspected after the procedure and noted to not contain any holes or tears. Tip cover arcing recommendations for prevention and precautions were provided to the account on (b)(6) 2011. As of the date of this report, no similar events have been reported by this site. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2096203 | 2065916 | 1958137 | 2955842201100113 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-22 | 2011-04-21 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-03-22 | No | Yes | Unkown | 2011
IS3000 A60P6 | 2011-03-24 | Other | Other | It was reported that while beginning a da vinci si thoracic sympathectomy procedure, and using an ethicon disposable 12mm trocar in the endoscopic camera manipulator (ecm) port, the trocar melted at the tip. The site replaced the trocar, however the replacement trocar melted at the tip. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The isi representative onsite for this procedure stated that the surgical staff was holding the camera backed up into the trocar for an extended period of time. The da vinci si surgical system user manual, setting up the vision system specifically states: caution: to avoid possible thermal damage to the cannula, do not leave the endoscope tip inside a plastic endoscope cannula (trocar) for a prolonged period of time while the lamp is on. The patient was reported to have recovered well post operatively with no complications. The site has continued to perform procedures and as (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096204 | 2065917 | 1962893 | 2955842201100114 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-03-24 | No | No | Reuse | 2011
42024902 | 2011-02-17 | Malfunction |  | It was reported that near the end of a da vinci si hysterectomy procedure shards of the 8mm dual blade retractor instrument fell into the patient. The patient was irrigated and the fragments were retrieved. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed the distal end of the instrument main tube exhibits scratches and abrasions with material removed and a rough surface finish. The damaged area is located approximately 2 inches above the proximal clevis. The abrasions are axially aligned with the tube and approximately 1. 3 inches long and. 085 inches wide. No other damage was found. In (b)(6) 2011, an isi representative performed additional training for prevention and detection of damaged cannula accessories at this hospital. A dropped or mishandled cannula can result in a dent of the cylindrical portion of the cannula and can become out-of-round which may result in scraping of the instrument shaft when inserted through the cannula. Drops may occur during cleaning or rough handling. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | 8MM DUAL BLADE RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2096205 | 2065918 | 1962894 | 2955842201100115 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-23 | 2011-04-21 | No | Yes | Health Professional | 42024902 | Device Returned To Manufacturer | Yes | 2011-03-23 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-03-18 | Death | Death | It was reported that post a successful da vinci si hysterectomy procedure the patient suffered a perioperative left hemisphere cerebrovascular accident (cva) secondary to complete occlusion of left carotid artery (embolic). Postoperatively noted to be completely aphagic and right hemiplegic. A clot was identified via mra imaging from bifurcation of aorta. No h/o hypercoagulability. Fibromuscular hyperplasia (fmh) significant for mother suffering fatal cerebrovascular accident (cva). Reportedly, the patient's demise was unrelated to the da vinci s surgical system. | No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. As of (b)(4), 2011, no similar instances have been reported to (b)(4). | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096208 | 2065919 | 1962895 | 2955842201100116 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-03-25 | 2011-04-21 | Yes | No | Health Professional | IS3000 A60P5 | No | Yes | 2011-03-25 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-24 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the site did not have 3d vision from the surgeon side console (ssc). With the assistance of an isi technical support engineer the site re-performed cameral calibrations and 3d image was restored. The site later reported that the vision issue recurred and the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the reported vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096612 | 2066364 | 1963340 | 2955842201100122 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-03-24 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2008-09-09 | Other | Other | On (b)(6) 2011 intuitive surgical received a legal summons stating that during a da vinci s myomectomy procedure the patient slid cephalad on the operating table and multiple repositions of the robot, repositioning of the da vinci port and re-docking were necessary to complete the procedure. The summons states that the patient injuries are permanent but no further details or information was provided. | Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunction was reported to have occurred and the system did not cause or contribute to the patient's injury. Multiple attempts for additional information have been made, however as of the date of this report no response from the patient or her representative has been received. If additional information is received, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096615 | 2066367 | 1963343 | 2955842201100121 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-24 | No | No | Reuse | 2011
42018405 | 2011-01-10 | Malfunction |  | It was reported that during the beginning of a da vinci s mitral valve procedure, the permanent cautery spatula broke while the surgeon was removing an excessive amount of paracardial fat in order to access the pericardium. Pieces of the broken instrument fell into the patient and were removed. The instrument was replaced and the procedure was completed using replacement instrumentation. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2096671 | 2066419 | 1963393 | 2955842201100124 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2011-01-26 | 2011-04-21 | No | Yes | Health Professional | 42018405 | No | No | 2011-03-25 | Device Not Retu |  | Reuse | 2011
IS2000 A51P8 | 2011-03-24 | Death | Death | It was reported that after a da vinci s mitral valve repair procedure, the patients blood pressure had dropped, the surgeon performed a mini thoracotomy to identify the sourse of the blood pressure loss. A leak in the mitral valve was found and a sternotomy was performed to repair of the valve; however, the patient expired before the valve could be repaired. The site has been contacted for additional information related to this event; however, no new information has been received as of the date of this report. | The investigation conducted by the field service engineer found the system to be made to specification with no issues found and the hospital has continued to use the system to perform surgery on patients. The site has been contacted for additional information related to this event. If additional information becomes available, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096683 | 2066439 | 1963415 | 2955842201100123 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-25 | 2011-04-21 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-03-25 | No | No | Reuse | 2011
400183 | 2011-03-14 | Malfunction |  | Patient undergoing a robotic sparing whipple and cholecystectomy for a pancreatic mass. During the procedure a piece of the hook cautery was noted to be broken off and in the patient's abdominal cavity. Copious irrigation and suction were performed. It is unknown whether the piece was retained or removed. |  | PERMANENT CAUTERY HOOK | ENDOWRIST MONOPOLAR CAUTERY IN | 2102648 | 2071868 | 1968844 | 2071868 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-19 | 2011-04-19 | No | Yes | Physician | 400183 | No | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-03-22 | Malfunction |  | It was reported that during a da vinci si procedure, sparking was observed at the tip cover accessory that was installed on the monopolar curved scissors (mcs) instrument. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicating arcing events occurred. The holes are approximately. 030 to. 020 in size and located. 425 above the silicone to sleeve interface. One hole exhibits a miniscule tear on one edge. Various small mechanical tears were noticed in the general vicinity of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2091963 | 2061161 | 1958137 | 2955842201100111 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-22 | 2011-04-18 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-03-22 | Device Not Retu | Yes | Initial | 2011
| 2011-04-12 | Malfunction | Other | A pt with a history of oropharyngeal squamous cell carcinoma was undergoing transoral resection of a base of tongue mass with davinci si robotic assistance. Several hours into the surgery, a small flame was visible at the spatula tip of the monopolar cautery of the robotic device. The instrument was withdrawn and the flame extinguished simultaneously by flushing the field with sterile saline. The insulation on the spatula cautery tip hub appeared damaged. The robotic arm and tip were replaced and surgery completed. There was no injury or respiratory compromise to the pt or harm to employees. |  | DAVINCI VERSION 2 5MM CAUTERY SPATULA TIP | MONOPOLAR TIP | 2096441 | 2065924 | 1962900 | MW5020319 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-04-18 | 2011-04-18 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
40017908 | 2011-03-18 | Malfunction |  | It was reported that after a successful da vinci hysterectomy procedure, the surgical staff found the monopolar curved scissors instrument tip was broken and missing a piece. The instrument fragment was found via x-ray and removed using traditional laparoscopic instrumentation. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the right scissor blade detached from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the crimp is exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2090732 | 2059822 | 1956798 | 2955842201100110 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-03-18 | 2011-04-15 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2011-03-18 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-15 | Injury | Other | It was reported that on (b)(6) 2011, the patient underwent a successful da vinci s fundectomy procedure, however, 3 hours post procedure, a pancreatic fistula was found. Drainage was performed and amylase levels were tested. On (b)(6) 2011, the doctor determined that the inserts used for drainage was not allowing sufficient drainage and as a result the surgeon cleaned the patient endoscopically and a drainage tube was placed. On (b)(6) 2011, the patient's respiratory status became worse and a ct scan showed an abscess inside the mediastinum and the chest cavity. It was found that the patient also experienced mediastinitis and empyema. Drainage was performed and the patient was put on a respirator. As of (b)(6) 2011, the site reported that the patient is recovering with no fever. The patient is said to be recovering well with all symptoms improving. | The surgeon commented that the system functioned properly and no anomalies were noted during the da vinci s surgical procedure performed on (b)(6) 2011. The surgeon advised that it could not determine as to how the fistula developed and suggested that the pancreas could have been unintentionally damaged during the procedure as the patient was reported as having visceral fat. He also suggested that damage to the pancrease could have been caused by a third party endoscopic forceps or ligasure instruments inserted through the assistant port. The surgeon advised that this type of surgical complication is a known risk of fundectomy procedures regardless of or surgical technique or tools used. Per the information provided, it was determined that the da vinci s surgical system, instrument and accessories worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2089300 | 2058205 | 1955181 | 2955842201100109 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-03-15 | 2011-04-14 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-15 | No | No | Reuse | 2011
420184 | 2011-03-29 | Malfunction |  | Davinci robotic instrument "permanent cautery spatula" used in thoracic surgery. Side piece broke off while spatula inside the chest cavity. Procedure went open for better exposure of surgical site, multiple irrigations, piece never found. Item on use # 3 out of manufacturer recommended 10 uses before disposal. Reason for use: cautery for nsclc. |  | DAVINCI | PERMANENT CAUTERY SPATULA | 2098576 | 2065679 | 1962655 | MW5020303 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-04-14 | 2011-04-14 | No | Yes | Health Professional | 420184 | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
420207 | 2011-03-02 | Malfunction |  | Using the davinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. |  | TENACULUM FORCEPS | ENDOWRIST GRASPERS FOR DAVINCI | 2096618 | 2066370 | 1963346 | 2066370 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-09 | 2011-04-09 | No | Yes | Physician | 420207 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
40018010 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The mcs instrument was removed and the tip cover was replaced. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is approximately. 032 inches by. 018 inches in size and is located. 35 inches above the silicone to sleeve interface. The hole exhibits a tear on one edge. On (b)(4) 2011, the robotics coordinator stated that she was unaware of the electrical surgical unit (esu) setting used during the case as well as the pre-procedure tip cover and cannula inspection process. On (b)(4) 2011, an (b)(4) representative provided the surgeon and (b)(4) coordinator with tip cover damage or arcing recommendations for prevention. Since the recommendations for prevention were provided, as of (b)(4) 2011, no further arcing events have been reported by (b)(4) hospital. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076649 | 2045758 | 1942735 | 2955842201100103 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-03-11 | Device Not Retu | Yes | Initial | 2011
42018404 | 2011-03-11 | Malfunction |  | It was reported that during a da vinci s lobectomy procedure, while using the permanent cautery spatula instrument, the surgeon witnessed brown fragments in the patient. The fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient injury, harm, or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed, the instrument distal clevis ear was completely broken at the base and missing. The broken piece was not returned with the instrument. Engineering concluded the instrument breakage was likely caused by overloading at the tip. Per the customer, all of the fragments were retrieved from the patient. The broken piece removed was compared to the instrument by the hospital and it was determined no other fragments were missing. | PERMANENT CAUTERY SPAULTA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2076650 | 2045759 | 1942736 | 2955842201100104 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | Yes | 2011-03-11 | Device Not Retu | No | Reuse | 2011
400180 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. On march 18, 2011 an isi representative provided the surgeon and the site's robotics coordinator with tip cover damage or arcing recommendations for prevention. Multiple attempts have been made to obtain additional information from the site regarding this event, however, as of the date of this report, no additional information has been provided. Since the recommendations for prevention were provided, as of april 7, 2011, no further arcing events have been reported by (b)(6) hospital. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076651 | 2045760 | 1942737 | 2955842201100105 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 400180 | No | No | 2011-03-11 | No | Yes | Unkown | 2011
400180 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. On march 18, 2011, an isi representative provided the surgeon and the site's robotics coordinator with tip cover damage or arcing recommendations for prevention. Multiple attempts have been made to obtain additional information from the site regarding this event, however, as of the date of this report, no additional information has been provided. Since the recommendations for prevention were provided, as of april 7, 2011, no further arcing events have been reported by (b)(6). | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076652 | 2045761 | 1942738 | 2955842201100106 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 400180 | No | No | 2011-03-11 | No | Yes | Unkown | 2011
IS3000 A60P5 | 2011-03-15 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced system error 297. With the assistance of an isi technical support engineer, the site attempted to troubleshoot by checking cable connections and recycling power; however, the system error code persisted. The patient was under anesthesia and the port incisions had been made when the surgeon decided to convert the planned procedure to traditional open techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) could not reproduce the system error; however, the camera cable and camera head were found to be loose. Based on the system error logs, it was determined that the system error experienced by the site may be associated with a defective doco unit. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco and tightening the loose camera cable. The doco was returned to isi for failure analysis and the customer reported complaint was not confirmed as the doco functioned properly with no problems found. As a result, it was determined that the system error code was associated with the loose camera cable found on site during the fse's visit. As of april 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2077285 | 2046159 | 1962895 | 2955842201100107 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-15 | 2011-04-08 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-03-15 | Device Not Retu | No | Reuse | 2011
IS2000 A51P7 | 2011-03-11 | Other |  | It was reported that during a da vinci s prostatectomy procedure, while performing urethrovesical anastamosis, the site experienced system error code 25004. With the assistance of an isi technical support engineer, the system was rebooted. Upon restart, system error code 25588 was experienced. An isi field service engineer (fse) arrived on site and attempted to troubleshoot the problem; however, the system issue persisted. At this time, the surgeon decided to convert to a laproscopic approach; however, converting to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering confirmed the occurrence of system error code 25004 as well as system error code 25588. Engineering concluded that system error codes 25004 and 25588 were associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the ecm to have a defective motor/encoder assembly which was replaced to repair the ecm. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2077542 | 2046429 | 1943407 | 2955842201100108 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-11 | 2011-04-08 | Yes | Yes | Health Professional | IS2000 A51P7 | Device Returned To Manufacturer | Yes | 2011-03-11 | Device Not Retu | No | Reuse | 2011
42019408 | 2001-03-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure and after insertion of the mega needle driver instrument for use, a piece of metal wire was found in the patient's abdomen. The surgeon removed the piece and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering's inspection found no damage on clevis or cable. Performance testing was conducted and the instrument was found to be within specification with no issues found. The recognition and engagement testing passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Based on these results, several attempts have been made to obtain additional information from the initial reporter without success. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2075466 | 2044570 | 1941546 | 2955842201100102 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-08 | 2011-04-07 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2010-10-19 | Device Not Retu | No | Reuse | 2011
420207 | 2011-03-02 | Malfunction |  | Using the da vinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. |  | TENACULUM FORCEPS | DA VINCI S SURGICAL SYSTEM END | 2096617 | 2066369 | 1963345 | 2066369 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-05 | 2011-04-07 | No | Yes | Physician | 420207 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
| 2011-04-05 | Malfunction |  | Cannula seals were leaking - we switched them both out and then they functioned fine. The seals came in the 3 arm drape pack (2nd time in a week similar occurrence, similar devices, similar leaking). Female with nonfunctioning kidney with chronic ureteropelvic junction obstruction and severe hydronephrosis. The nonfunctioning aspect of her kidney was confirmed with nephrostomy tube drainiage, which showed a creatinine clearance of 0 and very little urine output in spite of being very patent. Patient sent to or, was prepped and draped. Surgeon unable to obtain insuflation with a veress needle and never felt confident to actually be in the peritoneal cavity. Patient had a very extremely thick abdominal wall, secondary to obesity. Surgeon ended up doing a direct vision placement of a trocar with the longest - a 12 mm direct visiport - that we have. He was barely able to get this to go into the peritoneum. After insufflation, able to re-adjust it and pull it back enough that we would be able to be docked. Two 8mm robotic ports were placed in the usual fashion. We had some trouble with the superior port because of it pulling back and difficulty with visulazation until we were finally able to get this to be replaced. A 12 mm inferior trocar near the midline was inserted for assistant port. Surgery continued and finalized. Trocars removed sequentially under direct vision, closed and patient recovered from anesthetic and returned to recovery room in good condition. ======================manufacturer response for da vinci surgical system single use seal, intuitive surgical======================field representative is aware, has been in contact with or, however, any reports not yet available at time of this report. |  | 8MM CANNULA SEAL | DA VINCI SURGICAL SYSTEM 8MM A | 2092727 | 2061812 | 1958788 | 2061812 | 1 | NAY | User Facility |  |  | Initial | 2011-04-06 | 2011-04-06 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420225 | 2011-03-07 | Other | Other | It was reported that during a da vinci s oropharyngectomy procedure, fibers from the 5fr instrument introducer fell into the patient's mouth cavity. The surgeon was able to remove some of the fibers from the patient, however, decided to convert the procedure to traditional open surgical techniques to ensure all of the fragments were retrieved. The procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Multiple attempts to gather additional information have been made; however, as of the date of this report, the site has not provided any additional details. If additional information is received, a follow-up mdr will be submitted. | INSTRUMENT INTRODUCER 5FR | ENDOSCOPIC INSTRUMENT | 2069432 | 2037983 | 1934959 | 2955842201100097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-07 | 2011-04-04 | Yes | Yes | Health Professional | 420225 | No | No | 2011-03-07 | No | No | Unkown | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have epidermis burns at the cannula port location. The patient did not required any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded; however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069436 | 2037985 | 1934961 | 2955842201100098 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have epidermis burns at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded, however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069457 | 2037986 | 1934962 | 2955842201100100 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have burn marks at the cannula port location. Due to scarring, the patient required a second unspecified procedure. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded; however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069441 | 2038009 | 1934985 | 2955842201100099 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
40016906 | 2011-03-11 | Malfunction |  | It was reported that during a da vinci s sigmoid colectomy procedure, the surgical staff observed that a piece of the harmonic curved shears insert accessory was missing when it was inserted into another port. The missing piece was located in the patient and successfully retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the insert is missing the clamp arm. One side of the hinge feature that mates with the clamp arm pins is severely bent backwards. The damage appears to be consistent with hyperextension of the clamp arm, leading engineering to conclude that the breakage was most likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2067807 | 2036545 | 1933521 | 2955842201100089 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-11 | 2011-04-01 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | No | 2011-03-11 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-03-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, there was no video displayed on the left side monitor of the surgeon console. The surgical staff power cycled the system, reseated the vision cable, cycled the breaker and emergency powered off the surgeon console, however, these actions were unable to resolve the issue. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. The hrsv was returned to intuitive surgical for evaluation. Engineering discovered that the left side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. As of march 31, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2067816 | 2036547 | 1933523 | 2955842201100091 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-07 | 2011-04-01 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-03-07 | Device Not Retu | No | Reuse | 2011
42018304 | 2011-02-18 | Malfunction |  | It was reported that during a da vinci s thoracic sympathectomy procedure, a piece of the permanent cautery hook instrument fell into the patient's chest. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor cap was missing from the distal end and the yaw pulley hub had a broken off boss feature (the boss feature mates with the hole in the conductor cap and keeps the cap from slipping). Engineering concluded that the instrument hook may have been overloaded, causing the boss feature to break and allowing the conductor cap to fall out. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2067792 | 2036558 | 1933534 | 2955842201100090 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-17 | 2011-04-01 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2011-03-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-03-08 | Other | Other | It was reported that during the beginning of a da vinci si low anterior resection procedure, the surgical staff experienced double vision. The system was restarted twice and multiple endoscopes were tried, however, the issue could not be resolved. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | As part of intuitive surgical's failure analysis investigation, the system was evaluated at the site by a field service engineer (fse). System testing was performed using the customer's 30 degree endoscope and 3d calibration tool. The fse was unable to replicate the customer reported malfunction and the root cause of the customer reported failure mode cannot be determined. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2067821 | 2036560 | 1933536 | 2955842201100092 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-03 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure, the surgeon gripped tissue and cauterization automatically occurred, causing a superficial burn to the patient's sigmoid colon. A general surgeon was called in to review the injury and placed a couple stitches as a precaution. The electrical surgical unit (esu) and cable were disconnected and replaced as well as the instrument. The system was restarted and the planned surgical procedure was completed. No further patient harm or adverse events were reported. | The investigation conducted by the field service engineer (fse) discovered that the electrical surgical unit (esu) used during the procedure was set on the automatic setting, therefore, as soon as the surgeon gripped tissue with the instrument, the esu cauterized without the foot pedal being pressed. The fse notified the surgical staff of the incorrect esu setting used. The staff reported that the surgeon had a procedure directly following this event and the case was completed with the same equipment and there were no issues. The patient was reported as recovering fine with no post operative complications. As of april 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2068013 | 2036620 | 1933596 | 2955842201100093 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-03 | 2011-04-01 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-03 | Device Not Retu | No | Reuse | 2011
400169 | 2011-02-04 | Malfunction |  | On (b)(6), 2011, user facility medwatch (b)(4) was received from the fda reporting the following event. Event desc: during the procedure, the distal portion of a robotic surgical instrument - harmonic curved shears - used for davinci surgery detached from the instrument for unknown reason. A thorough search of the patient's uterus was performed. One piece of the shears were found but the other piece of the shears was not found - even on x-ray. | Multiple attempts with the site have been made; however, at the time of this report, the instrument has not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2068523 | 2037070 | 1895144 | 2955842201100095 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-03-04 | 2011-04-01 | No | Yes | Health Professional | 400169 | Yes | No | 2011-03-04 | No | Yes | Initial | 2011
40018010 | 2011-02-25 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The monopolar curved scissors instrument accessory has as not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2068527 | 2037074 | 1934050 | 2955842201100094 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-02 | 2011-04-01 | No | Yes | Health Professional | 40018010 | No | Yes | 2011-03-02 | Device Not Retu | Yes | Unkown | 2011
IS2000 | 2011-01-27 | Other | Hospitalization | It was reported that the da vinci s surgical system was used on (b)(6) 2011, to perform a prostatectomy procedure. The surgery was successful and there were no system malfunctions during the case. The patient was discharged on (b)(6) 2011, however, returned to the hospital on the same day due to hematuria and difficulty in urinating. A clot was detected and removed and electrical coagulation was performed transurethrally (not using the da vinci s surgical system). The patient recovered and was discharged from the hospital on (b)(6) 2011. | Information was provided by the surgeon, who suspects that the crust of the anastomosis site between the urethra and vesica came off, resulting in the blood in the patient's urine. The surgeon indicated that the patient's post surgical symptoms were unrelated to the da vinci s surgical system and are part of the complications related to prostatectomy procedures, regardless of the modality used to perform the surgery. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunctions occurred, and the system did not cause or contribute to the patient's post surgical complications. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2068515 | 2037093 | 1934069 | 2955842201100096 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2011-03-03 | 2011-04-01 | Yes | No | Health Professional | IS2000 | No | No | 2011-03-03 | No | No | Reuse | 2011
400180 | 2011-02-24 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The monopolar curved instrument (mcs) was returned with the ultem sleeve from the mcs tip cover accessory installed; however, but the silicone part had been removed. The distal end components of the instrument were visually inspected and no char marks or other signs of arcing were observed. Electrical continuity passed and the instrument had fully been used with no lives remaining. No physical damage found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2066674 | 2035530 | 1932505 | 2955842201100088 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-01 | 2011-03-31 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | No | 2011-03-01 | Device Not Retu | Yes | Initial | 2011
| 2011-03-24 | Malfunction |  | During robotic prostatectomy, the team had problems with laproscopic c02 pressure being low. Attempted a number of things to fix the problem; started by changing out the disposable trocars, but c02 still leaked. It appeared that the leak was coming from the seals on the divinci trocars. All seals were switched out and it continued to leak. Then, vaseline gauze was wrapped around the instruments to block the leak a little bit, it helped for the time being. ======================manufacturer response for trocar seals for davinci, da vinci surgical system======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for ethicon trocar bladeless endopath xcel, ethicon trocar bladeless endopath xcel w/stability sleeve======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for ethicon trocar bladeless endopath xcel, ethicon trocar bladeless endopath xcel w/stability sleeve======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. |  | DA VINCI SURGICAL SYSTEM | TROCAR SEALS FOR DAVINCI | 2074107 | 2043364 | 1940340 | 2043364 | 1 | NAY | User Facility |  |  | Initial | 2011-03-31 | 2011-03-31 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2011
420172 | 2011-03-29 | Malfunction |  | During surgery, a piece of the maryland bipolar endowrist broke off inside patient. Broken piece of device was retrieved by doctor (md). Entire device was inspected and then taken out of operating room (or). |  | MARYLAND BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2092720 | 2061778 | 1958754 | 2061778 | 1 | NAY | User Facility |  |  | Initial | 2011-03-31 | 2011-03-31 | No | Yes | Invalid Data | 420172 | Yes | No Answer Provided |  |  |  |  | 2011
420227 | 2011-02-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the clamps were not closing. At this time, the surgeon observed a plastic part on the pk dissecting instrument was broken. When the instrument was removed and inspected, a piece of plastic was found to be missing. The surgeon reported that there have been no reports of patient harm, adverse outcome or injury as a result of the piece falling into the patient. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2065574 | 2034496 | 1931471 | 2955842201100086 | 1 | LLZ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-02-28 | 2011-03-30 | No | Yes | Health Professional | 420227 | No | Yes | 2011-02-28 | No | No | Unkown | 2011
42000306 | 2011-02-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the small clip applier instrument's tip became detached and fell into the patient. The piece was retrieved and accounted for from the patient. The procedure was successfully completed with a replacement instrument of the same type. The site reported that there was no patient injury or adverse event experienced. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to be returned with a section of one grip detached from the distal end. The size of piece is roughly. 150 x. 050. The grip fractured near the base of the clip groove and the grip has a slight bend below the fracture surface. Additional observation not reported by site is a bent grip. The intact grip is bent at roughly a 30 degree angle. Location of bending is near the base of the clip groove, similar to where the other grip fractured. The location of fracture/bending of the grips matches closely with the depth the grips get installed into the clip holding fixture. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage found. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2065619 | 2034510 | 1931486 | 2955842201100087 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-28 | 2011-03-30 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2011-02-28 | Device Not Retu | No | Initial | 2011
| 2011-03-18 | Malfunction |  | Jaws of scissors weren't working; the jaws did not open or close and they never worked. |  | ENDOWRIST | MONOPOLAR CAUTERY INSTRUMENTS | 2068407 | 2036945 | 1933921 | 2036945 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2011-03-25 | 2011-03-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-02-22 | Other | Other | It was reported that during a da vinci s nephrectomy procedure, the site experienced difficulty with the system recognizing the instrument. The site attempted to reseat the patient side manipulator (psm) arm; however, the instrument still was not recognized by the system. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the clinical sales representative (csr) found that the instrument being used was out of lives. The site was made aware of this. As of march 22, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2059715 | 2028765 | 1925743 | 2955842201100085 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-02-22 | 2011-03-24 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-02-22 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-02-21 | Other | Other | It was reported that during a da vinci s cystectomy procedure, the surgical staff experienced multiple system error code 212. The site was able to override the error code on several occasions; however, the error returned. Prior to calling into isi for technical assistance, the site undocked and rebooted the system. The system powered up and homed with no issues. The site then called to advise an isi technical support engineer (tse) and to gain additional information regarding the system error. At this time, the system error code 212 returned. The site was able to override the fault; however, it returned shortly after clearing. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator and multiple servo driver (msd). The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system contains two multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected mtmr and the msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The msd and mtmr were returned to isi for failure analysis investigation. Engineering evaluation found that the msd was within manufacturing specifications with no issues noted. Engineering also evaluated the returned mtmr and found that the main harness and an axis motor had malfunctioned thus generating the system error code experienced by the customer. As of march 22, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2058602 | 2028112 | 1925088 | 2955842201100084 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-02-21 | 2011-03-23 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-02-21 | Device Not Retu | No | Reuse | 2011
420174 | 2011-03-17 | Malfunction |  | Using robotic harmonic scalpel during robotically assisted thymectomy. Robotic harmonic scalpel includes reusable hand piece, reusable robotic sleeve, and disposable harmonic blade insert. After the surgeon began to use the insert, one of the jaws broke off. |  | HARMONIC CURVED SHEARS | ENDOWRIST ULTRASONIC ENERGY IN | 2068263 | 2036803 | 1933779 | 2036803 | 1 | NAY | User Facility |  |  | Initial | 2011-03-21 | 2011-03-23 | No | Yes | Physician | 420174 |  | No Answer Provided |  |  |  |  | 2011
420174 | 2011-03-17 | Malfunction |  | Using robotic harmonic scalpel during robotically assisted thymectomy. Robotic harmonic scalpel includes reusable hand piece, reusable robotic sleeve, and disposable harmonic blade insert. After the surgeon began to use the insert, one of the jaws broke off. |  | HARMONIC CURVED SHEARS | ENDOWRIST ULTRASONIC ENERGY IN | 2068263 | 2036803 | 1933779 | 2036803 | 1 | NAY | User Facility |  |  | Initial | 2011-03-21 | 2011-03-23 | No | Yes | Physician | 420174 |  | No Answer Provided |  |  |  |  | 2011
400193 | 2011-03-03 | Malfunction |  | At the end of the davinci robotic procedure, the balloon port for the procedure was removed and after inspection by the scrub rn to confirm integrity of the balloon port, the sheath was noted to have a circumferential crack. ======================manufacturer response for disposable, balloon port======================no comments. |  | BALLOON PORT | ENDOWRIST ACCESSORIES FOR DAVI | 2068238 | 2036726 | 1933702 | 2036726 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-03-21 | 2011-03-21 | No | Yes | Invalid Data | 400193 | Yes | No Answer Provided |  |  |  |  | 2011
IS3000 A60P5 |  | Death | Death | It was reported that eight days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient expired due to uncontrollable post operative bleeding. | On (b)(6) 2011, additional information regarding the event was provided by dr. (b)(6), console surgeon who performed the tors procedure. Dr. (b)(6) indicated that the patient underwent surgery for a benign mass at the base of the tongue, which had been causing obstructive symptoms for a year and there was no other way to approach the lesion. Due to the location of the mass, dr. (b)(6) was close to a branch of the lingual artery. Dr. (b)(6) cauterized the artery with bipolar energy and put tisseel over it. The case went well, and dr. (b)(6) didn't have to go too deep and had good visualization. Eight days postop, the patient began to bleed. Dr. (b)(6) mentioned that typically a bleed within 24-48 hours postop is considered a technical problem and after 48 hours, it is due to a wound healing issue which expose the vessels and make them more likely to bleed. Dr. (b)(6) indicated that this complication is a known risk with a transoral approach and was not at all related to any limitations of or problems with the da vinci si surgical system. Based on the information provided by dr. (b)(6), it was determined that the patient's demise was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's death. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2053945 | 2023904 | 1920894 | 2955842201100082 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-17 | 2011-03-18 | Yes | No | Health Professional | IS3000 A60P5 | No | No | 2011-02-17 | No | No | Reuse | 2011
IS3000 | 2011-02-11 | Other | Other | It was reported that after a da vinci si hysterectomy procedure was completed, the patient developed a 103 degree fever and their heart rate jumped to 130. The surgeon performed a ct scan and a blood clot was observed. No other information was received. | Multiple attempts have been made to obtain additional information from the surgeon, however, the surgeon has been unavailable. To date, no information has been received which suggests that the da vinci si surgical system may have caused or contributed to the patient's post surgical complications. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2054373 | 2024247 | 1921257 | 2955842201100083 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-16 | 2011-03-18 | Yes | No | Health Professional | IS3000 | No | No | 2011-02-16 | No | No | Reuse | 2011
400184 | 2011-02-15 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed that the amber plastic part of the permanent cautery spatula instrument was broken at the tip after it was inserted into the patient. The break was not observed and no pieces from the instrument were found in the patient. It is unclear if the peace broke inside of the patient; however, prior to use, the instrument was inspected when anti-char solution was applied. No issues with the instrument were noted at this time. As of the date of this report, no patient harm, adverse outcome or injury has been reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2052948 | 2022563 | 1919537 | 2955842201100081 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2011-02-15 | 2011-03-17 | No | Yes | Health Professional | 400184 | No | No | 2011-02-15 | No | No | Unkown | 2011
IS3000 A60P3 | 2010-02-15 | Other | Other | It was reported that during set up for a da vinci si hysterectomy procedure, the site experienced a non recoverable system error fault 22003. With the assistance of an isi technical support engineer, the site was able to apply various troubleshooting techniques to override the system error to continue with the procedure. The site called back reporting a re-occurrence of system error 22003 and elected to go forward using the system while disabling one of the three patient side manipulators (psm) arm. When the site docked to the patient, the site could not follow on matching grip with the psm 3. At this point, the surgical staff elected to convert the planned procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes 322003 found in the error logs were associated with a setup joint (suj) arm. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from its own joint positions to the rac. Errors in this communication at system start-up would induce the system error codes in this case. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The system was repaired by replacing the affected suj. The suj was returned to isi for failure analysis investigation. Engineering confirmed the customer reported complaint as they were able to duplicate the system error codes. As a result, the suj was repaired by replacing several harness cables. As of march 15, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2052978 | 2022592 | 1919582 | 2955842201100079 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2010-02-15 | 2011-03-17 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | Yes | 2009-02-15 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-01-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) with a mcs tip cover accessory installed. A small thermal injury was observed on the patient's bowel and was repaired via stitch. The planned surgical procedure was completed with a replacement mcs instrument tip cover accessory. No patient recovered well with no post operative complications or prolonged hospitalization. The delay of this report is due to the site's initial reporter going out on medical leave prior to submitting the report as initially intended. Upon her return, several weeks later, the report was filed. An isi representative was onsite during the procedure; however, believed the initial reporter would initiate the complaint record with isi customer service. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2052979 | 2022593 | 1919583 | 2955842201100080 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-18 | 2011-03-17 | No | Yes | Health Professional | 40018010 | No | No | 2011-01-18 | No | No | Unkown | 2011
42017909 | 2011-02-03 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure a piece of the monopolar curved scissors instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. The delay in reporting is the result of an administrative error by a complaint handling team member. This error has been addressed with the applicable team member. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the right blade was broken off and missing from the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2053232 | 2022859 | 1919849 | 2955842201100071 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2011-02-03 | 2011-03-17 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2011-03-21 | Device Not Retu | No |  | 2011
40027202 | 2011-02-14 | Malfunction |  | It was reported that during a da vinci s abdominoperineal resection procedure the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that a section of the curved blade had broken off. A piece of the curved blade, measuring approximately. 385 x. 080, was missing and not returned with the insert. The fracture surface of the blade is not level and outer surface of blade has a couple of scratch marks below fracture. It was determined that this damage may have contributed to the blade breakage since the broken piece has scratch marks in the axial direction, consistent with this motion. Although intuitive surgical uses the term blade above, there are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. The da vinci harmonic ace curved shears instructions for use specifically states: inspection prior to use examine the device prior to use. In particular, examine the following components for cracks or flaws: the clamp arm and blade on the insert, the shaft of the insert, the reusable instrument housing caution: if cracks or flaws are observed, do not use the device. Note: scratches on the blade may lead to premature blade failure. Avoid accidental contact with other instruments during use. Do not use any means other than the blade wrench to attach or detach the instrument from the hand piece. General precautions and warnings: do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2051671 | 2021232 | 1918228 | 2955842201100078 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-14 | 2011-03-16 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-02-14 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P6 | 2011-02-18 | Other | Other | It was reported that when the surgical staff was getting ready to dock the da vinci si surgical system to the patient for a myomectomy procedure, system error code 25521 occurred. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 25521 was associated with an endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart and provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25521 appears when the system detects that it has lost communication of sensor information between two printed circuit assemblies in the arm control system. In the event of these communication issues, the system records the fault and transitions to a safe state. The ecm was returned for failure analysis investigation. Engineering was able to replicate the reported failure and observed that both flat flex cables in and out of a small board embedded in the insertion axis of the ecm were not well connected, thus generating the system error code experienced by the customer. As of march 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2048128 | 2018315 | 1915311 | 2955842201100076 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-18 | 2011-03-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-02-18 | Device Not Retu | No | Reuse | 2011
42017207 | 2011-02-18 | Malfunction |  | Event Description: --/-- It was reported that while cauterizing in the chest wall area during a da vinci si lobectomy procedure, the smell of burning plastic was detected by the surgical staff and arcing was observed from the jaws of the maryland bipolar forceps instrument to the endowrist. An intuitive surgical representative in attendance during the surgical procedure noticed that a loose gia staple was caught in the jaws of the instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- (b)(4) - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on both yaw pulleys with material removed at the grip base. No char marks were observed on the sides of the yaw pulleys or the metal wrist components. It was determined that the damage may have been caused by a breach in the insulation or carbonized tissue which created a conductive path. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 2048155 | 2018338 | 1915333 | 2955842201100072 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-02-18 | 2011-03-14 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2011-04-11 | Device Not Retu | No |  | 2011
40027202 | 2011-02-14 | Malfunction |  | It was reported that during a da vinci s abdominoperineal resection procedure, the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that a section of the curved blade had broken off. A piece of the curved blade, measuring approximately. 645 x. 085, was returned with the insert. The fracture surface of the blade is not leveled and one of the mating features that slides in and out of the shaft when the clamp arm opens/closes is bent slightly inwards and rubs against one side of the blade. It was determined that this damage may have contributed to the blade breakage since the broken piece has wearing/scratching in the axial direction, consistent with this motion. Although intuitive surgical uses the term blade above, there are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. The da vinci harmonic ace curved shears instructions for use specifically states: inspection prior to use, examine the device prior to use. In particular, examine the following components for cracks or flaws: the clamp arm and blade on the insert - the shaft of the insert - the reusable instrument housing. Caution: if cracks or flaws are observed, do not use the device. Note: scratches on the blade may lead to premature blade failure. Avoid accidental contact with other instruments during use. Do not use any means other than the blade wrench to attach or detach the instrument from the hand piece. General precautions and warnings - do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2048173 | 2018343 | 1915338 | 2955842201100073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-14 | 2011-03-14 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-02-14 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P5 | 2011-02-18 | Other | Other | It was reported that during a da vinci si vesicolithotomy procedure, the surgical staff experienced non-recoverable system error code 31030. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The system logs were reviewed during the investigation conducted by the field service engineer (fse) and multiple occurrences of system error code 23 were observed along with system error code 31030. The fse concluded that system error codes 23 & 31030 were associated with a patient side manipulator (psm) setup-joint axis. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The setup-joint (suj) is a non-motorized support arm which holds a patient side manipulator (psm) during surgery. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the psm. System error code 31030 appears when the da vinci si safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining this condition, the system records the fault and transitions to a safe state. The suj was returned for failure analysis evaluation. Engineering was able to replicate the reported failure and observed damages on the cable harness, thus generating the system error codes experienced by the customer. As of march 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2048182 | 2018348 | 1915343 | 2955842201100077 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-18 | 2011-03-14 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-02-18 | Device Not Retu | No | Reuse | 2011
400077 | 2011-03-09 | Malfunction |  | Patient undergoing lysis of adhesions and total laparoscopic hysterectomy with robotic davinci assist. The davinci has 4 arms; 3 of which use 8mm cannula seals on them (the 4th is for a camera and the seal for this arm is built into the introducer). At beginning of procedure it was noted 3 of the 8mm cannula seals utilized on the arms of the davinci machine would not seal properly, staff could hear air coming through the seals. All 3 seals were from the same lot number and were removed from the machine when complete seal could not be obtained. Staff obtained new seals from a completely different lot number and encountered no further issues. No injury to patient, just time delay in troubleshooting issues. |  | 8 MM CANNULA SEAL | DAVINCI ENDOWRIST 8MM ACCESSOR | 2064380 | 2033149 | 1930125 | 2033149 | 1 | NAY | User Facility |  |  | Initial | 2011-03-14 | 2011-03-14 | No | Yes | Physician | 400077 | Yes | No Answer Provided |  |  |  |  | 2011
42017908 | 2011-02-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found a piece of the left scissor blade to be broken off from the instrument, the broken piece was returned. The blade fractured at the cross section of the grip cable crimp resulting in half of the cable crimp to be exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2046357 | 2016298 | 1913294 | 2955842201100075 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-05 | 2011-03-11 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2011-02-10 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-02-08 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgical staff experienced the left eye from the surgeon side console to be pink. With the assistance of an isi technical support engineer, the site verified their settings and it was then determined that the camera cable needed to be replaced. The patient was under anesthesia when the surgeon staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of march 8, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2043254 | 2013313 | 1910309 | 2955842201100066 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-08 | 2011-03-10 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-02-08 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-02-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The tip cover accessory has not been returned for evaluation, however, the monopolar curved scissors instrument used with the tip cover during the procedure was returned and evaluated. Per the customer reported complaint engineering found no char marks or signs of arcing on the instrument. Electrical continuity passed and no damage was found. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2043255 | 2013314 | 1910310 | 2955842201100067 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-09 | 2011-03-10 | No | Yes | Health Professional | 40018010 | No | No | 2011-02-09 | No | Yes | Unkown | 2011
400180 | 2011-02-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. It was noted that the electrical surgical unit (esu) setting was above the isi recommended range. | The tip cover accessory has not been returned for evaluation. Multiple attempts for additional information have been made, however as of the time of this report, no additional information has been provided by the site. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2043256 | 2013315 | 1910311 | 2955842201100069 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-17 | 2011-03-10 | No | Yes | Health Professional | 400180 | No | No | 2011-02-17 | No | Yes | Unkown | 2011
IS3000 A60P5 | 2011-02-11 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure the site was unable to get the system to power on. The site reconnected all and changed all power outlet connections, however the system continued to not gain power. With the assistance of an isi technical support engineer the site reset the main system breaker, however, the system continued to not power on. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm or injury was reported. | The investigation conducted by the field service engineer concluded that system power issue was associated the core power distribution (ipd) board. The ipd is a pca (printed circuit assembly) board that converts the incoming electrical voltage to the appropriate levels to power the other circuit boards and regulates the current supplied to them. The system was repaired by replacing the affected ipd pca board. The ipd pca board was returned to isi for failure analysis investigation. Engineering was able to replicate the customer reported complaint and found a defective module on the ipd pca board. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2043368 | 2013395 | 1910391 | 2955842201100070 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-11 | 2011-03-10 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2011-02-11 | Device Not Retu | No | Reuse | 2011
| 2011-02-22 | Malfunction |  | Patient was undergoing robotic lh bilateral salpingo-oophorectomy and mccall's monarch anterior repair. The plastic part of the spatula broke off during the procedure. This has occurred previously with the same operator, according to operating room nurse operator has been re-trained by the company rep on previous occasions. Diagnosis or reason for use: surgery for hysterectomy and bladder suspension. |  | INTUITIVE SURGICAL | DAVINCI CAUTERY SPATULA | 2051848 | 2018713 | 1915708 | MW5019768 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-03-09 | 2011-03-09 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
IS3000 60P6 | 2011-02-07 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci si prostatectomy procedure, the surgical staff was unable to calibrate the camera and the system diagnostics generated a message indicating that a possible cause was a dirty scope. The site power cycled the system, however, this did not resolve the issue. The patient had been under anesthesia for 20 minutes when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the camera cable instead of the endoscope. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The camera cable was returned and evaluated. Engineering was able to replicate the reported failure mode and concluded that the issue experienced by the customer was a result of conductor damage. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of march 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2041079 | 2011052 | 1908048 | 2955842201100068 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-07 | 2011-03-08 | Yes | Yes | Health Professional | IS3000 60P6 | Device Returned To Manufacturer | Yes | 2011-02-07 | Device Not Retu | No | Reuse | 2011
| 2011-01-31 | Malfunction |  | Md was suturing mesh to the vaginal wall using the endowrist (a mega suturecut needle driver used with the davinci robot). The circulator noticed that a wire was hanging from the endowrist. Surgery department opened up another endowrist instrument and within 5 minutes the same thing happened to this instrument. The circulator describes the wire as a braided wire cable and that the cable was dangling. Another instrument was opened which functioned appropriately and the md was able to finish the procedure. However, there was the potential for a fragment to fall into the patient. ======================manufacturer response for suturing device for davinci robot, endowrist mega suture cut needle driver======================unknown. |  | MEGA SUTURECUT NEEDLE DRIVER | ENDOWRIST INSTRUMENT FOR DAVIN | 2054420 | 2024310 | 1921300 | 2024310 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-03-07 | 2011-03-07 | No | Yes | Invalid Data |  | No | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2011-02-03 | Other |  | It was reported that during a da vinci prostatectomy procedure, the surgical staff was unable to focus the camera. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue by changing the camera cable; however, the vision issue remained. The site could hear the motor running and attempted to work with the focus assembly without success. The patient was under anesthesia approximately 1 hour and the port incisions were made when the surgeon made the decision to abort the planned surgical procedure. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera head was returned to the original equipment manufacturer (oem) for evaluation. The oem discovered that a mechanical shock to the camera head damaged the stage assembly, thus resulting in the vision issue experienced by the customer. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2038249 | 2008222 | 1905218 | 2955842201100064 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-03 | 2011-03-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-02-03 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-02-03 | Other | Other | It was reported that the beginning of a da vinci prostatectomy procedure, the surgical staff experienced system error 153. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue by checking cables and restarting the system after undocking from the patient. Once powered back on, a system error message indicated to check the green cable. The site powered down the system, checked the green cable cable and restarted. The system successfully started up and the procedure continued. Two hours later, an isi field service engineer called in stating that system error code 153 reoccurred and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that system error code 153 was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. The ria pca performs numerous tasks related to the function, control and hardware fault monitoring for a specific arm. There is a ria pca assigned to each system arm. The system was repaired by replacing the affected ria pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 153 is reported by hardware to denote a high side switch fault. This indicates that there was a short on the board that triggered this. In the event of these power supply issues, the system records the fault and transitions to a safe state. As of march 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2038250 | 2008223 | 1905219 | 2955842201100065 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-03 | 2011-03-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-02-03 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-03 | Injury | Life Threatening | It was reported that during a da vinci s rectopexy procedure the surgeon experienced visualization issues due the patient's extremely large uterus. Due to the poor vision, the surgeon inadvertently hit a vessel over the sacral promontory and uncontrollable bleeding occurred. As a result, the surgeon converted to traditional open surgical techniques to maintain the bleeding and complete the planned procedure. The delay in reporting is a result of the complaint handling team mistakenly believing that the complaint information had already been recorded in the complaint database. | (b)(4). The initial reporter confirmed that the system and associated instruments/accessories did not malfunction during the surgical procedure and did not cause or contribute to the patient injury. The initial reporter also indicated that the patient is recovering well and no post surgical complications were reported. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no device failure occurred, and the patient injury was unrelated to the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2033191 | 2003418 | 1900414 | 2955842201100061 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-06 | 2011-03-01 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-12-06 | No | No | Reuse | 2011
IS3000 A60P5 | 2011-02-01 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced a loss of video in the right eye of the surgeon side console. The site reseated and then replaced the camera cable, however, the vision loss continued to occur. With the assistance of a technical support engineer, the site restarted the system, however, the vision loss was still present. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) and review of the system error logs determined that the vision issue experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The doco was returned to isi for failure analysis and the customer reported complaint was confirmed as the doco did not function properly. As of march 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2033522 | 2003912 | 1900908 | 2955842201100062 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-01 | 2011-03-01 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-02-01 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-02-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the end piece of the permanent cautery spatula (pcs) instrument broke off and fell into the patient. The broken piece was retrieved and the instrument was replaced with another pcs instrument. The end piece of the second pcs instrument also broke off during use and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a third pcs instrument. No patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00059 was submitted for the initial pcs instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base. The ear had broken off on the opposite side of the conductor wire and cap. The broken piece was not returned. The conductor cap is also missing. Engineering also observed a broken proximal clevis at the main tube interface. As a result, the clevis is dislodged from the main tube. The proximal clevis and main tube are missing pieces. A. 185 x. 175 piece of the proximal clevis was returned with the instrument, however, the clevis has a remaining broken section that was not returned. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2032792 | 2002988 | 1900008 | 2955842201100060 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-02 | 2011-02-28 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-02-02 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-02-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the end piece of the permanent cautery spatula (pcs) instrument broke off and fell into the patient. The broken piece was retrieved and the instrument was replaced with another pcs instrument. The end piece of the second pcs instrument also broke off during use and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a third pcs instrument. No patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00060 was submitted for the second pcs instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base. The size of the broken piece is approximately. 285 x. 235. The ear had broken off on the opposite side of the conductor wire and cap. The broken ear piece was not returned, however, the conductor cap was returned taped to the housing. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2032789 | 2003012 | 1900008 | 2955842201100059 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-02 | 2011-02-28 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-02-02 | Device Not Retu | No | Reuse | 2011
42018405 | 2011-01-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a brown piece from the distal end of the permanent cautery spatula had fallen into the patient. The piece was successfully retrieved and the procedure completed with no patient injury. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument returned with a piece of the distal clevis broken. One distal clevis ear is broken off at its base. The broken piece was not returned; however, the missing piece was roughly. 300 x. 235 and was located on the opposite side of the conductor wire and cap. Engineering has concluded that the instrument breakage may be due to overloading the tip. No other damage found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2031948 | 2001957 | 1898953 | 2955842201100057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-26 | 2011-02-27 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2010-11-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P4 | 2011-01-26 | Other | Other | It was reported that during a da vinci si mitral valve replacement procedure, the site experienced a black image in the left eye of the surgeon side console. With the assistance of an isi technical support engineer, the surgical staff reseated the camera cable; however, the vision issue was not resolved. At this time, the surgical staff made the decision to convert the planned procedure to traditional open surgical techniques to successfully complete the procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of february 24, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2031964 | 2001958 | 1898954 | 2955842201100058 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-26 | 2011-02-27 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2011-01-26 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-21 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure, the surgical staff noticed flickering and wavy lines in the image viewed through the surgeon side console (ssc). The site attempted to power down and restart the system, however, they were unable to get the system to power back on. With the assistance of an isi technical support engineer, the site reseated the power cord, reset the ssc power breaker and moved the power cord to an isolated circuit, however, the ssc would not regain power. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering confirmed that the issue experienced by the customer was related to the surgeon side console (ssc) primary power supply (pps). The pps is a +115v power supply with battery backup and contains logic and protection circuits to provide monitoring of under voltage, over voltage, over current, ac line status, over temperature, power supply fan faults, and detection of battery operation. The system was repaired by replacing the pps. As of february 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2026492 | 1996704 | 1893700 | 2955842201100053 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-21 | 2011-02-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-21 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-27 | Other | Other | It was reported that during a da vinci s surgical procedure, the site was unable to focus the vision settings from the surgeon side console (ssc) and the vision cart. With the assistance of an isi technical support engineer, the site replaced the camera cable and restarted the system. The surgeon was then able to focus the image as expected and continued with the case. The site later reported that the vision issues recurred and the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the reported vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of february 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2026494 | 1996706 | 1893702 | 2955842201100054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-27 | 2011-02-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-27 | Device Not Retu | No | Reuse | 2011
400006 | 2011-02-16 | Malfunction |  | Davinci needle driver would not close properly. The instrument had 8 out of 10 uses left. The instrument was replaced in the field and sent down to decontam decontamination to be sent to risk management and then sent back for credit from intuitive surgical. |  | LARGE NEEDLE DRIVER FOR DAVINCI | ENDOWRIST NEEDLE DRIVER | 2033833 | 2004223 | 1901219 | 2004223 | 1 | NAY | User Facility |  |  | Initial | 2011-02-18 | 2011-02-18 | No | Yes | Physician | 400006 |  | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-01-17 | Other | Required Intervention | It was reported that during a da vinci s lobectomy procedure, an injury to a branch of the patient's main pulmonary artery occurred during its dissection, requiring that the surgeon convert to traditional open surgical techniques to complete the planned procedure. | On (b)(6) 2011, the surgeon stated via email, the injury was not caused by a robotic instrument or arm malfunction. The patient recovered completely from the operation and has been home for several weeks. Per the information provided by the surgeon, intuitive surgical determined that the patient injury was unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2023628 | 1993942 | 1890938 | 2955842201100050 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-01-17 | 2011-02-16 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-01-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-01-18 | Other | Required Intervention | It was reported that during a da vinci si surgical procedure, the instrument became stuck on the system arm and the surgical staff was unable to get the instrument to release the tissue that it was grasping using the emergency release wrench. The surgeon cut out the portion of the tissue that was being grasped in order to remove the instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported issue; however, a link port switch and the cannula mount were replaced as a precaution. As of february 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2023632 | 1993945 | 1890941 | 2955842201100051 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-18 | 2011-02-16 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2011-01-18 | Device Not Retu | No | Unkown | 2011
42017910 | 2010-12-21 | Malfunction |  | It was reported that during a da vinci si ovarian cystectomy procedure, sparking was observed from a source other than the tips of the monopolar curved scissors instrument(mcs) with a tip cover accessory installed. The mcs instrument and tip cover accessory were removed and replaced with an instrument of the same type. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of sparking as the instrument passed electrical continuity testing and did not exhibit any signs of charring or arcing. The instrument was placed on an in house test system and passed recognition and engagement tests as well as a cut test. The instrument moved intuitively with full range of motion in all directions. No damage was found. Multiple attempts have been made to retrieve additional information from the hospital regarding this event, however, as of the date of this report, no response has been received. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2014825 | 1985162 | 1882158 | 2955842201100045 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-13 | 2011-02-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2011-01-13 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-01-11 | Injury | Other | It was reported that approximately 35 to 50 minutes into a da vinci s sacrocolpopexy procedure, while the surgeon was exploring the surgical field, the permanent cautery hook instrument broke and bumped the patient's bifurcation of the inferior vena cava. The procedure was converted to traditional open surgical techniques, sutures were placed to repair the iliac perforation, and the patient was provided units of blood as a result of the bleeding. On (b)(6) 2011, medwatch uf. Importer report (b)(4) was received. | On (b)(6) 2011, dr. (b)(6) stated that the procedure would have likely converted to open due to the patient's size and extent of adhesions. He stated that the patient was recovering fine and there were no additional complications. On (b)(6) 2011, dr. (b)(6) contacted isi regarding the possibility that instrument fragments may have fallen into the patient during the procedure, and requested biocompatibility information for the permanent cautery hook instrument. An x-ray and ct scan were performed and were found to be inconclusive in confirming if any fragments were retained in the patient. On (b)(6) 2011, (b)(6) hospital risk management representative, (b)(6), confirmed the patient has recovered well and has not had any post surgical complications. The instrument will not be returned for evaluation, however, the hospital did provide photographic images of the instrument for isi evaluation. Per the images provided, the instrument has a broken distal clevis ear with the ear having broken off at the base. The conductor cap is missing due to the ear breakage. A follow-up mdr will be submitted if the instrument is returned and / or if more information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2014844 | 1985183 | 1882179 | 2955842201100044 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-01-12 | 2011-02-11 | Yes | Yes | Health Professional | 42018305 | No | Yes | 2011-01-12 | No | No | Unkown | 2011
IS2000 A51P8 |  | No Answer Provided | Other | The patient posted a comment on an isi website, called patient stories, where she reported the following story. After a da vinci hysterectomy procedure, the patient was reported having a long recovery and post operative uncomfort. Over a week post op, the patient underwent a pelvic exam which resulted in the discover of a tear and an infection. A second surgery was required to repair the tear; however, one month after the second surgery, the patient was experiencing pain during intercourse which resulted in an emergency room visit. | The site's risk management department has been contacted to obtain additional details related to this event. As of the date of this report, additional information has not been provided to support the reported event. Isi will continue to make attempts to gain additional information and will provide the fda with a follow up report once the additional information becomes available. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015557 | 1985870 | 1882866 | 2955842201100046 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-13 | 2011-02-11 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-01-13 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-01-12 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the surgeon experienced double vision with the image displayed on the surgeon console. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the vision issue experienced by the customer was due to incorrectly connected cables at the rear of the surgeon console. The issue was resolved by swapping the cables to their correct location. As of february 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015939 | 1986227 | 1883223 | 2955842201100047 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2011-01-12 | 2011-02-11 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2011-01-12 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-12-07 | Malfunction |  | It was reported that during a da vinci si nephroureterectomy procedure, arcing was observed at the side of the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The electrosurgical unit was not activated when the arcing was observed. The tip cover was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and five holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicating multiple arcing events occurred. The hole sizes range from under. 010 to. 025 with varying shapes (round and oblong) and are located. 290 to. 415 above the silicone to sleeve interface. The various mechanical tears are in the general vicinity of the holes and there is an axially oriented tear at the distal tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2015955 | 1986243 | 1883239 | 2955842201100048 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-12 | 2011-02-11 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-01-12 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-01-17 | Other | Other | It was reported that during a da vinci s procedure, the surgical staff experienced issues with the patient side cart intermittently switching to battery power. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the power issue experienced by the customer was due a patient side cart power cord malfunction. The system was repaired by replacing the affected power cord. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015956 | 1986244 | 1883240 | 2955842201100049 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-17 | 2011-02-11 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-01-17 | Device Not Retu | No | Reuse | 2011
400169 | 2011-02-04 | Malfunction |  | During the procedure, the distal portion of a robotic surgical instrument - harmonic curved shears - used for davinci surgery detached from the instrument for unknown reason. A thorough search of the patient's uterus was performed. One piece of the shears were found, but the other piece of the shears was not found - even on x-ray. |  | HARMONIC CURVED SHEARS INSERT | ULTRASONIC ENERGY INSTRUMENT | 2027847 | 1998148 | 1895144 | 1998148 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-02-10 | 2011-02-10 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-01-10 | Other | Other | It was reported that prior to a da vinci s prostatectomy procedure, the site experienced system error code 23013 and 23008 when attempting to attach the sterile adapter to the patient side manipulator (psm). An isi technical support engineer (tse) attempted to assist the site to override the system error; however, the site made the decision to abort the planned procedure. The patient was under anesthesia at the time the decision was made not to continue with the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer (fse) concluded that system error code 23013 and 23008 were associated patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23013 and 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The psm was returned to isi for evaluation. Engineering determined that the potentiometer had malfunctioned on the psm, thus generating the system error codes experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2012932 | 1982927 | 1879922 | 2955842201100043 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2011-01-10 | 2011-02-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-01-10 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-10 | Other | Required InterventionHospitalization | On (b)(6) 2011, intuitive surgical received a legal summons and complaint filed by a pt alleging that they suffered significant injury and damages due to the surgeon's inability to provide reasonably prudent medical care after a da vinci s excision of endometriosis including a partial thickness sigmoid colon resection procedure was performed. The surgery took approx 10 hours to complete and post operatively, the pt had ct findings consistent with a perforated viscus and developing abscess. The complaint further alleges that the pt has undergone numerous medical procedures (including add'l surgeries) due to the surgeon's failure to properly, adequately or timely monitor, manage, diagnose, refer, consult, inform and treat the pt's medical condition. | No allegations of deficiency with the da vinci s surgical system were reported. Based on the info received, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2039096 | 2001786 | 1898782 | 2955842201100042 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-01-10 | 2011-02-09 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-01-10 | Device Not Retu | No | Reuse | 2011
42011704 | 2010-12-15 | Malfunction |  | It was reported that while suturing during a da vinci si single port hysterectomy procedure the jaw of the 5 mm needle driver instrument broke and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. The rptr indicated that there was likely extra stress on the instrument due to the close proximity of the instrument arms while performing the single port procedure. No pt harm, adverse outcome or injury was reported. The delay in reporting is due to an isi rep who was in attendance during the surgical procedure when the malfunction occurred but did not report the event within the time period required. The isi rep has been reminded of isi's reporting timeframe requirements. | The broken instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A f/u medwatch report will be submitted if the instrument is returned or if add'l info is rec'd. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 2043733 | 2003155 | 1900151 | 2955842201100038 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-15 | 2011-02-09 | No | Yes | Health Professional | 42011704 | No | No | 2010-12-15 | Device Not Retu | No | Unkown | 2011
42017909 | 2010-10-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if more information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2011486 | 1981964 | 1878960 | 2955842201100041 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-11 | 2011-02-08 | No | Yes | Health Professional | 42017909 | No | Yes | 2011-01-11 | No | No | Unkown | 2011
42018304 | 2010-12-10 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff observed small fragments in the patient's pelvis area. The end of the permanent cautery hook instrument was noted to be misaligned with pieces of the sheath end missing. The fragments were removed and the procedure was completed with a replacement instrument of the same type. No patient harm was reported. The delay in reporting is due to an administrative oversight where the complaint handling representative misfiled the report for reporting the week after it was due. | The instrument was returned and evaluated. Confirming the customer's reported complaint, engineering found the instrument's wrist cables and back end cables to be tight and intact. Additional observations found the instrument's main tube to be broken at the junction of the proximal clevis and main tube. The main tube feature which mates with the proximal clevis is broken off and missing. Based on these findings, engineering has concluded that the damage is likely due to misuse and/or mishandling. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2009414 | 1980338 | 1877334 | 2955842201100039 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-30 | 2011-02-07 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2010-12-30 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Death | Death | It was reported that 10 days after a da vinci si transoral resection procedure was successfully performed, the patient experienced bleeding in the surgical area and expired. The date of the surgical procedure and date of death were not received. Pre surgery, the patient's cancer was thought to be at stage t2, however, during the procedure, it was discovered that the cancer was at stage t3. There were no system malfunctions during the procedure and the surgeon indicated that the patient's demise was unrelated to the da vinci si surgical system. The surgeon believes that the post surgical bleeding may have been related to a delay in healing and that a scab may have come off. The delay in reporting is due to a misunderstanding about when intuitive surgical first became aware of the event. | Per the information provided by the surgeon, it was determined that the patient's demise was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and no system malfunctions occurred. On (b)(6) 2011, attempts were made to obtain additional information from the surgeon, however, no response has been received to date. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2010063 | 1980770 | 1877766 | 2955842201100040 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-12-13 | 2011-02-07 | Yes | No | Health Professional | IS3000 A60P5 | No | Yes | 2010-12-13 | No | No | Reuse | 2011
40018010 | 2011-01-05 | Malfunction |  | Event Description: --/-- It was reported that while cauterizing the surgical stump during a da vinci si hysterectomy procedure, the surgeon noticed that the fully mobilized/detached uterus was being blanched by the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed one hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating that an arcing event occurred. The hole size is approximately. 020 in diameter and the hole center is located. 215 above the silicone to sleeve interface. --/-- Additional Manufacturer Narrative: --/-- During evaluation of the device, it was also observed that a section of the silicone was not properly adhering at the shoulder of the ultem. This additional information is being provided per a retrospective review that was performed. | 8MM MCS TIP COVER ACCESSORY | Initial | 2008014 | 1979008 | 1876004 | 2955842201100037 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-01-05 | 2011-02-04 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-01-27 | Device Not Retu | Yes |  | 2011
IS3000 A60P4 | 2011-01-06 | Other | Other | It was reported that approximately 45 minutes during a da vinci si pulmonary lobectomy procedure, an instrument could not be recognized on a patient side manipulator (psm) arm. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer (fse) found the issue experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported issue, however, found the psm wrist input assembly to be below specification. All four wrist assemblies and the cannula mount were replaced. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2008105 | 1979115 | 1876111 | 2955842201100036 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-06 | 2011-02-04 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2011
| 2011-01-27 | Malfunction |  | Patient underwent robotic internal mammary takedown. Da vinci balloon port was placed in the patient's right anterior chest wall to facilitate port to introduce robotic instrumentation for internal mammary artery (ima) takedown. After successful completion of harvest, port was removed. At this point, request was made to the anesthesiologist to inflate the lung. The bedside md assistant placed his finger in the incision site where the port was to feel the lung inflation and he noted a hard object with the chest wall. He retreived the object and the tip of the da vinci balloon port (approx. 1/2 inch) was retreived. When the balloon port was inspected by the surgical team, it was then noted that a circumfrential 1/2 inch length of the port broke off. Both edges were examined and a "clean break" was noted with no rough edges. Procedure completed in normal fashion. |  | 12 MM BALLOON PORT | DAVINCI BALLOON CAMERA PORT | 2022076 | 1992036 | 1889032 | 1992036 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-02-04 | 2011-02-04 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2011-01-03 | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure, the surgeon experienced difficulty with the grips opening and closing for the instruments controlled by the left master tool manipulator (mtm). The patient had been under anesthesia for approximately one hour and 45 minutes when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtml gimbal. The mtml gimbal was returned to isi for failure analysis investigation. Engineering evaluation of the mtml gimbal found the gimbal grip levers had excessive play, thus confirming the reported complaint experienced by the customer. The gimbal grip levers were replaced. As of (b)(4), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGCAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005852 | 1977151 | 1874147 | 2955842201100033 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-03 | 2011-02-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-01-03 | Device Not Retu | No | Reuse | 2011
42017909 | 2011-12-03 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure and while performing a dissection, the tip of the monopolar curved scissors instrument broke and fell into the patient. The piece was retrieved and the procedure was successfully completed using a replacement instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a piece of the left blade had broken off and is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2005876 | 1977175 | 1874171 | 2955842201100034 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-01-03 | 2011-02-02 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-03 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-04 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a non recoverable system error code 23. With the assistance of an isi technical support engineer (tse), the site powered down the system; however, the system error persisted. The site then powered down the system again, cycled the breakers, and applied emergency power off (epo) before powering on the system. The system was working for approximately 10 minutes when the system error re-occurred. The patient had been under anesthesia for approximately 1. 5 hours and docked to the system for 45 minutes when the surgeon made the decision to abort the planned procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes 23 experienced by the customer was associated with an endoscopic camera manipulator (ecm). Endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The ecm was returned to isi for failure analysis investigation. Engineering could not duplicate the reported system error code; however, as a precaution the ecm harness and embedded serializer patient manipulator (espm) board were replaced. As of february 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005877 | 1977176 | 1874172 | 2955842201100035 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-04 | 2011-02-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-04 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-01-07 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the v-care uterine manipulator had been placed incorrectly and perforated the patient's fibroid. While resecting the broad ligament fibroid, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument was unintentionally closed over the v-care uterine manipulator. Energy was then applied and the top jaw of the hcs insert broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the insert is missing the clamp arm. One side of the hinge feature that mates with the clamp arm pins is severely bent backwards. The damage appears to be consistent with hyperextension of the clamp arm and engineering concluded that the breakage was most likely caused by mishandling/misuse. The broken tip was returned with the device. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2005913 | 1977212 | 1874208 | 2955842201100031 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-07 | 2011-02-02 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2011-01-07 | Device Not Retu | Yes | Initial | 2011
IS1200 A43P9 | 2011-01-06 | Other | Other | It was reported that prior to using the system to perform any tasks for a da vinci surgical procedure, the surgical staff experienced text/words on the left side image displayed in the surgeon console. The surgical staff power cycled the system and checked the settings on the synchronizer, however, they were unable to resolve the issue. The patient had been under anesthesia for 45 minutes when the surgeon made the decision to abort the planned procedure and reschedule it for a later date. No patient harm was reported. | Investigation was conducted by a field service engineer (fse). Prior to evaluating the system, the fse was informed that something had spilt on the bottom area of the vision cart (vc), and the vc had been cleaned before the scheduled case. During evaluation of the system, the fse discovered that the left camera control unit (ccu) options menu had been opened and the camera id had been activated and set to a random series of letters and numbers. The fse concluded that these buttons were most likely pressed during cleaning, thus causing the vision issues experienced by the customer. The fse repaired the system by resetting the ccu settings. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005915 | 1977214 | 1874210 | 2955842201100032 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-06 | 2011-02-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2011-01-06 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-03 | Malfunction |  | On (b)(4) 2011, medwatch uf / importer report (b)(4) was received. It was reported by the hospital that during a da vinci s radical vaginectomy procedure, when the surgeon activated the monopolar cautery foot pedal, the bipolar instrument became activated as well as the monopolar instrument. It was noted by the surgical staff that the monopolar and bipolar cautery cables from the electrical surgical unit (esu) to the da vinci instruments were intertwined. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | On site investigation conducted by the field service engineer (fse) was unable to replicate the customer reported monopolar and bipolar cautery activation issues. Inspection found the bracket on the back of the monopolar connection to be broken, resulting in a loose connection to the electrical surgical unit (esu) and that the cautery cables from the esu to the monopolar and bipolar instruments were intertwined during the procedure. Engineering review indicates that these conditions are unlikely to have led to the cautery activation issue reported by the surgeon. The monopolar instrument drives current into the patient body which returns to the electrical ground through the path of least resistance, usually the grounding pad connected to the patient. It is possible for the bi-polar instrument to be the path of least resistance, due to close tip proximity or insufficient contact with the grounding pad, causing current to flow through the bipolar tips and make them appear to be activated. It is the view of isi engineering that this is the most likely cause of the reported issue. This is an issue inherent in the use of energy instruments and not a flaw in the da vinci surgical system. The surgeon referred to the wires between the esu and the instruments as having been bundled. Since these are non-shielded cables carrying high current, it is conceivable that current in one cable could have induced current in the other. Since this is part of the esu setup, it is outside of intuitive's expertise and would be best addressed by the esu manufacturer. For precautionary purposes, the fse replaced the ssc footrest cautery cable. The footrest cautery cable connects the ssc footswitch panel to the esu unit, causing cauterization to occur when the surgeon activates the applicable mono or bipolar foot pedal. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1999713 | 1971475 | 1868471 | 2955842201100030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-01-06 | 2011-01-27 | No | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2011-01-06 | Device Not Retu | No | Reuse | 2011
42017909 | 2010-11-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed that the monopolar curved scissors (mcs) instrument was cracked just above the endowrist. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. The customer reported malfunction does not in itself constitute a fda reportable event, however, engineering discovered evidence of arcing during internal evaluation of the instrument. | The instrument was returned and evaluated on december 29, 2010. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube and the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. It was determined that the damage was most likely due to excessive side loading or other type of mishandling. Engineering also observed charring on the tube extension. The tube extension has a section that exhibits charring/thermal damage directly below the breakage. The instrument was disassembled to find a light char mark present on the inner surface of the tube reinforcement ring. The tip cover accessory returned with the instrument also exhibits a char mark on the proximal end of the plastic sleeve, likely caused by energy escaping from the tube extension breakage. On january 11, 2011, the initial reporter provided additional information and indicated that arcing from the mcs instrument was observed during the surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1994426 | 1966554 | 1863550 | 2955842201100028 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-29 | 2011-01-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-12-29 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-11-17 | Other | Required Intervention | It was reported that on (b)(6) 2010, the patient underwent a successful da vinci s prostatectomy procedure, however, on (b)(6) 2010, the patient became febrile and began experiencing low abdominal tenderness. On (b)(6) 2010, the patient had a fever and blood in his stool. On (b)(6) 2010, a ct scan revealed that the patient had possibly developed generalized peritonitis. A re-operation was performed and a sigmoid colon perforation was discovered. Open drainage was performed and an artificial anus was created. Based on the laparotomy findings, the symptom seemed to be delayed fistulation. The patient recovered and was discharged on (b)(6) 2010. | The surgeon commented that the system functioned properly and no anomalies were noted during the da vinci s surgical procedure performed on (b)(6) 2010. Although the system was used in the area where the perforation was discovered, the surgeon indicated that no perforations were observed during the surgical procedure. Per the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1994490 | 1966750 | 1863746 | 2955842201100029 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-11-24 | 2011-01-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-12-24 | No | No | Reuse | 2011
42000606 | 2010-12-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the large needle driver instrument broke while the surgeon was suturing and an instrument piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned with one carbide insert detached from its grip. Per the customer reported complaint, engineering evaluation observed that the grip brazing surface has various spots where the filler material is removed and the insert brazing surface has matching spots with the filler material added. A couple of scratch marks were found on the outer surface of the grip. No other damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1993556 | 1965939 | 1862935 | 2955842201100027 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-21 | 2011-01-20 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2010-12-21 | Device Not Retu | No | Reuse | 2011
400309 | 2011-01-11 | Malfunction |  | Mega suturecut needle driver inserted into trocar, appeared ok. When surgeon went to use it, one of the threads broke making the instrument unusable. |  | MEGA SUTURECUT NEEDLE DRIVER | DAVINCI ENDOWRIST NEEDLE DRIVE | 2003856 | 1975612 | 1872608 | 1975612 | 1 | NAY | User Facility |  |  | Initial | 2011-01-20 | 2011-01-20 | No | Yes | Invalid Data | 400309 |  | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2010-12-17 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced a non-recoverable system error. The site power cycled the system, however, the error continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found system error codes 20003 and 20013 in the system error log. These errors are associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error code 20013 is generated when an error is detected as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The system error code 20003 indicated that another fault occurred during the system start up. As of january 18, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1990416 | 1963102 | 1860098 | 2955842201100026 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-12-19 | 2011-01-18 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-12-19 | Device Not Retu | No | Reuse | 2011
40017208 | 2010-12-08 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure a fragment broke off of the maryland bipolar forceps instrument and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the main tube and distal tip detached from the housing. Per the customer reported complaint, engineering evaluation observed that the distal end drive cables and conductor wires have been cut at the interface between the metal proximal clevis and the main tube. The distal end of the main tube is broken on one side. The evidence is not conclusive, however, a likely failure mode is a dislodged proximal clevis due to excessive overloading at the tip, which resulted in the inability to remove the instrument from the cannula and required the cables/wires to be cut. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1988928 | 1961770 | 1858766 | 2955842201100018 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-15 | 2011-01-14 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2010-12-15 | Device Not Retu | No | Reuse | 2011
42018405 | 2010-12-22 | Malfunction |  | It was reported that during a da vinci si right lobectomy procedure, the permanent cautery spatula (pcs) instrument made contact with another endowrist instrument, causing a yellow piece on the pcs instrument to break off and fall into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that yellow colored components (proximal clevis and distal clevis) at the distal end are intact and undamaged, however, the brown ceramic sleeve on the instrument has a piece broken off at the distal end. The size of the missing piece is approximately. 090 x. 075. The ceramic sleeve also exhibits a couple of scratch marks below the breakage. While the evidence is inconclusive, engineering has concluded that damage to the instrument is most likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1988963 | 1961785 | 1858781 | 2955842201100020 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-22 | 2011-01-14 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2010-12-22 | Device Not Retu | No | Initial | 2011
42011704 | 2010-12-06 | Malfunction |  | It was reported that during a da vinci si ventral hernia repair procedure and while suturing mesh to repair a ventral hernia, the tip of the 5mm needle driver instrument broke and fell into the patient. The surgeon proceeded with the procedure using a replacement 5mm needle driver instrument. The piece was not retrieved from the patient; however, no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to have one of the grip tips broken at the grip base. The broken surface was flat and no sign of damage was found on grip cables. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1988988 | 1961795 | 1858791 | 2955842201100023 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-17 | 2011-01-14 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-12-17 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-15 | Death | Death | It was reported that while removing tissue during a da vinci s transoral resection procedure performed on (b)(6) 2010, the surgeon encountered an unusual group of veins and patient bleeding occurred. The surgeon was able to stop the bleeding and the planned surgical procedure was completed. There were no system malfunctions during the procedure. Five days post-op, the patient experienced bleeding and expired on (b)(6) 2010. | Per the surgeon, the patient's demise was unrelated to the da vinci s surgical system and the bleeding experienced by the patient would have occurred regardless of the modality used to perform the procedure. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1988993 | 1961834 | 1858830 | 2955842201100022 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-12-16 | 2011-01-14 | Yes | No | Service Personnel | IS2000 A51P8 | No | Yes | 2010-12-16 | No | No | Reuse | 2011
IS3000 A60P5 | 2010-12-17 | Other | Other | It was reported that approximately 2 hours during a da vinci si rectopexy procedure, while an instrument change was in progress, the surgeon removed his hands from the master tool manipulators (mtm); however; left his head in the surgeon side stereo viewer. This caused the instrument being exchanged in the guided tool change mode, to move slightly forward and bump the patient's colon. Slight bleeding of the patient's colon occurred; however, no repair, blood transfusion or intervention was necessary. The planned procedure was completed and no adverse outcome was reported. | The investigation conducted by field service engineering concluded that mtm movement experienced by the customer was associated with use error. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The guided tool change is an efficient and safe method for instrument insertion, providing assistance to the patient side cart operator by guiding the instrument into the patient. The system was found to be within specification, however, gravity calibration was performed on the mtms as a precaution. On (b)(6) 2010, additional training was conducted with the surgeon to ensure that in future cases, he will remove his head from the surgeon side viewer when his hands are not engaged on the mtms. The da vinci si user manual specifically states: note: for guided tool change, the surgeon should remove their head from the stereo viewer to take all instruments out of following before removing and installing an instrument. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1989021 | 1961862 | 1858858 | 2955842201100019 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2010-12-17 | 2011-01-14 | Yes | No | Health Professional | IS3000 A60P5 | Yes | No | 2010-12-17 | Device Not Retu | No | Reuse | 2011
42017206 | 2010-12-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the maryland bipolar instrument tip became detached from the shaft and could not be removed through the cannula accessory. With the assistance of a technical support engineer (tse), the site disengaged the sterile adapter from the patient side manipulator (psm) arm and removed the instrument with the cannula attached. A larger port incision was made to the patient to allow the cannula and instrument to be removed. After the maryland bipolar instrument was removed, the surgical staff found that wires had fallen into the patient. The affected area was irrigated and suctioned to remove the fragments. The site replaced the psm drape, sterile adapter and instrument and completed the planned procedure. No adverse outcome was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On (b)(4) 2010, a field service engineer inspected the da vinci s surgical system and found the system to have no issue. On (b)(4) 2010, (b)(4) stated that the patient underwent an x-ray and no foreign material was detected. If the instrument is returned or if further information is received, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1989042 | 1961869 | 1858865 | 2955842201100024 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-15 | 2011-01-14 | Yes | Yes | Health Professional | 42017206 | No | No | 2010-12-15 | No | No | Unkown | 2011
40018010 | 2010-12-20 | Other | Other | Event Description: --/-- It was reported that approximately 30 minutes into a da vinci s prostatectomy procedure, the surgeon noticed a hole had developed in the tip cover accessory installed on the monopolar curved scissors instrument which was believed to have caused a burn to the patient's right external iliac vein. The surgeon repaired the iliac vein burn with 5-0 prolene. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. The planned procedure was complete and no adverse outcome was reported. | Manufacturer Narrative: --/-- The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found various mechanical tears and a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicating of arcing. The hole is approximately. 030 in by. 020 and is located. 425 inches above the silicone to sleeve interface. The tip cover also has various small mechanical tears in the general vicinity of the holes. A large tear, approximately. 195 inches long, is located 90 degrees from the hole, on one edge. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip cover accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument and/or tip cover accessory are returned (post-evaluation) and/or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | Initial | 1989068 | 1961907 | 1858903 | 2955842201100025 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2011-03-28 | 2011-01-14 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-03-28 | Device Not Retu | Yes |  | 2011
8MM CAUTERY HOOK | 2011-01-11 | Injury | Required Intervention | The pt is a (b)(6) female admitted to the medical center for surgery to repair a vaginal prolapse. The surgeon stated that he was performing a robotic sacrocolpopexy. During the procedure the j-hook da vinci instrument malfunctioned. The j-hook is a surgical instrument that is attached to the da vinci robot used to cauterize tissue (permanent cautery hook). One side of the yoke that secures the j-hook fractured; this caused the j-hook to move involuntarily perforating the bifurcation of the inferior vena cava. The surgery at that point had to be converted to an open procedure. Sutures were used to repair the two small perforations in the vein. This complication occurred at a point when the surgeon had paused during the surgery. His hands were not on the pedal controls for the j-hook. The pt did require units of blood as a result of the bleeding. |  | PERMANENT CAUTERY HOOK | PERMANENT CAUTERY HOOK | 1997605 | 1966432 | 1863428 | 1966432 | 1 | NAY | User Facility | Unknown | Other | Initial | 2011-01-13 | 2011-01-14 | Yes | Yes | Health Professional | 8MM CAUTERY HOOK | Yes | No |  | No Answer Provi |  | Invalid Data | 2011
| 2010-11-30 | No Answer Provided | Other | It was reported that 10 hours post successful da vinci s hysterectomy procedure, the patient experienced a fever of 103. A pan culture was taken and was found to be negative with no fever indicated. The patient underwent an unknown treatment and discharged on schedule and without prolonged hospitalization. The patient was reported as doing well with no additional post operative issues. | During our investigation of this event at the site, it was determined that retraining was needed to advise the site on proper cleaning and sterilization processes for da vinci instrument and accessories. During the retraining the cleaning and sterilization staff reported that there may have been one or more steps in the recommended process that they were not performing. Since the retraining and change in the sites cleaning and sterilization process, there have not been any further reports of post operative temperatures. | DA VINCI S INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1987793 | 1960556 | 1857601 | 2955842201100017 | 1 | NAY | Manufacturer | User facilityCompany Representative | RISK MANAGER | Initial | 2010-12-14 | 2011-01-13 | Yes | No | Service Personnel |  | No | No | 2010-12-14 | No |  | Unkown | 2011
400180 | 2010-12-17 | Other | Other | It was reported that approximately one hour during a da vinci s hysterectomy procedure, the surgical staff witnessed an arc of energy from the monopolar curved scissors instrument with the tip cover accessory installed, to the patient's uterus. No intervention or repair was required as the uterus was being removed. Approximately 2 hours later in the procedure, the surgeon decided to convert the procedure to traditional open surgical techniques due to the patient's anatomy. The procedure was completed and no patient harm or adverse outcome was reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however, the isi representative present during the event stated that throughout the procedure, the monopolar curved scissors instrument with the tip cover accessory installed were being compromised due to the movement of the patient's large organ. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987769 | 1960583 | 1857579 | 2955842201100011 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-17 | 2011-01-13 | Yes | Yes | Health Professional | 400180 | No | No | 2010-12-17 | No | Yes | Unkown | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage and therefore only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda in regarding these events: 2955842-2011-00013 and 2955842-2011-00014. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987770 | 1960584 | 1857580 | 2955842201100012 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
40018402 |  | Malfunction |  | Event Description: --/-- On (b)(6) 2010, medwatch uf/importer report (b)(4) was received: during the surgical procedure, the operating physician noticed a small foreign body in the operating field. Upon further inspection, it was noted that a piece of plastic had broken off the intuitive endowrist monopolar spatula. The instrument was removed and replaced by a new instrument. | Manufacturer Narrative: --/-- Multiple attempts with the site have been made; however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering observed the ceramic sleeve to be broken at the distal end and missing a piece approximately. 060 inches by. 090 inches. The ceramic sleeve exhibits scratches below the broken section. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 1987773 | 1960587 | 1857583 | 2955842201100010 | 1 | NAY | Manufacturer | Other | RISK MANAGER | InitialFollowup | 2011-04-01 | 2011-01-13 | No | Yes | Health Professional | 40018402 | Device Returned To Manufacturer | No | 2011-04-01 | Device Not Retu | No |  | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage and therefore only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda regarding these events: 2955842-2011-00012 and 2955842-2011-00014. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987774 | 1960588 | 1857580 | 2955842201100013 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage, and therefore, only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification, and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda in regarding these events: 2955842-2011-00012, 2955842-2011-00013. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987738 | 1960589 | 1857580 | 2955842201100014 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
400180 | 2010-12-08 | Other | Other | It was reported that during a da vinci s surgical procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument, to the patient's vessel. The surgeon repaired the vessel with a few stitches and the planned procedure was completed. No adverse outcome was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however, it was noted that the esu setting used during the procedure was above the isi recommended range. On (b)(6) 2011, the isi representative was told by the surgical technician that reported the event, that the patient recovered well and there was no adverse outcome. Additional information has been requested multiple times, however, as of the date of this report, the site has not provided any further information. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987775 | 1960590 | 1857586 | 2955842201100015 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-17 | 2011-01-13 | Yes | Yes | Health Professional | 400180 | No | No | 2010-12-17 | No | Yes | Unkown | 2011
| 2010-12-01 | No Answer Provided | Other | It was reported that 10 hours post successful da vinci s hysterectomy procedure, the patient experienced a fever of 103. A pan culture was taken and was found to be negative with no fever indicated. The patient underwent an unknown treatment and discharged on schedule, without prolonged hospitalization. The patient was reported as doing well with no additional post operative issues. | During our investigation of this event at the site, it was determined that an retraining was needed to advise the site on proper cleaning and sterilization processes for da vinci instrument and accessories. During the retraining, the cleaning and sterilization staff reported that there may have been one or more steps in the recommended process that they were not performing. Since the retraining and change in the site's cleaning and sterilization process, there have not been any further reports of post operative temperatures. | DA VINCI S INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1987794 | 1960605 | 1857601 | 2955842201100016 | 1 | NAY | Manufacturer | User facilityCompany Representative | RISK MANAGER | Initial | 2010-12-14 | 2011-01-13 | Yes | No | Service Personnel |  | No | No | 2010-12-14 | No |  | Unkown | 2011
| 2010-11-12 | Malfunction |  | The davinci tip cover accessory developed holes in it intra-operatively. |  | DAVINCI TIP COVER ACCESSORY | DAVINCI TIP COVER ACCESSORY | 1999514 | 1971338 | 1868334 | 1971338 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-01-10 | 2011-01-10 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
42017910 | 2010-11-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure sparking was seen coming from the tips of the monopolar curved scissors instrument with the tip cover accessory installed. The surgeon decided to continue with the case, however shortly after a second sparking event was noticed. At this time the surgical staff removed the monopolar curved scissors instrument and tip cover accessory and replaced them with another instrument and accessory of the same type. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The mcs instrument and tip cover accessory were returned and evaluated. Per the customer reported complaint engineering found no visual evidence of arcing or char marks on the instrument. The instrument was inspected and found to pass electrical continuity testing and did not have any burrs or abnormally sharp edges that could contribute to arcing. Visual inspection of the tip cover accessory showed no thermal or mechanical damage to the silicone. On (b)(6) 2010 (b)(6) reported that the instrument and accessory were inspected prior to use and after removed from this procedure, however, no evidence of malfunction could be found. The esu setting used was noted to be within isi's recommended range. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1979926 | 1952946 | 1849941 | 2955842201100005 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-13 | 2011-01-07 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes |  | Device Not Retu | No | Reuse | 2011
42011704 | 2010-12-08 | Malfunction |  | It was reported that during a da vinci s tubal reanastomosis procedure, while the surgeon was suturing, the tip of the 5mm needle driver instrument broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse events, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one grip to be detached from the distal end of the 5mm needle driver instrument. The carbide insert was found to still be attached to the grip. The grip fractured at the junction with the hub feature, where there is a 90 degree angle, and hairline cracking is exhibited. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1979906 | 1952970 | 1849965 | 2955842201100006 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-12-09 | 2011-01-07 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | It was reported that during a da vinci s mitral valve replacement procedure, an instrument became stuck on a patient side manipulator (psm) arm. With the assistance of an isi technical support engineer, the psm was undocked from the patient and the cannula and instrument were removed. The psm was re-draped and the site attempted to re-engage a different instrument on the arm, however, the engagement problem continued to occur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found the instrument engagement issue reported by the customer was associated with the right spring plunger on the affected patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The spring plungers are mechanical elements used to help secure and position the sterile adapter interface between the psm and the sterile instrument. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found one of the psm's spring plungers, between axis 4-7 on top of link 6, to be bent and difficult to depress. The 4 affected spring plungers were replaced. As of january 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1979990 | 1953007 | 1850002 | 2955842201100007 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | It was reported that during a da vinci s myomectomy procedure, system error code 212 occurred and could not be overridden. The surgical staff power cycled the system, however, system error code 212 recurred immediately after power up. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci s in a recoverable safe state. The mtml was returned to isi for failure analysis investigation. Engineering evaluation found that the harness conductor had malfunctioned, thus generating the system error code experienced by the customer. As of january 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1980062 | 1953031 | 1850027 | 2955842201100008 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | Event Description: --/-- It was reported that during a da vinci s surgical procedure, the surgical staff experienced a green tint on the left side image displayed in the surgeon console. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the vision cable. The vision cable connects the vision cart (vc) to the surgeon console (sc) and passes the video signals between the sc and vc. The system was repaired by replacing the affected vision cable. The vision cable was returned to isi for failure analysis investigation. Engineering observed contamination on the cable, which likely caused the image issue experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the vision cable. The vision cable connects the vision cart (vc) to the surgeon console (sc) and passes the video signals between the sc and vc. The system was repaired by replacing the affected vision cable. The vision cable was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure during evaluation and concluded that the image issue experienced by the customer may have been caused by the vision cable not being properly connected. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | Reuse | 1980305 | 1953271 | 1850266 | 2955842201100009 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No |  | 2011
IS1200 A43P9 | 2010-12-06 | Other | Other | It was reported that at system start up for a da vinci prostatectomy procedure, the site experienced multiple occurrences of system error code 20105. With the assistance of an isi technical support engineer, the cable connections were checked; however, the system error persisted. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system error code 20105 was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards, endoscopic camera manipulator (ecm), and multiple servo driver (msd). An ria pca is assigned to each system arm and performs numerous functions related to the function, control and hardware fault monitoring for its assigned arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system contains two multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected ria pca, ecm and msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20105 appears when an auxiliary voltage error is detected during self-test. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of january 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1976403 | 1949858 | 1846854 | 2955842201100002 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-06 | 2011-01-05 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2010-12-06 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgical staff heard a pop sound and the system would not start up. The patient was under anesthesia and port incisions had been made before the surgical staff made the decision to abort the procedure and reschedule for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the power supply. The system was repaired by replacing the affected power supply. As of january 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1976405 | 1949860 | 1846856 | 2955842201100003 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-06 | 2011-01-05 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
42017206 | 2010-09-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument stopped working. The site also reported that fragment(s) were removed from the patient and there was no harm to the patient. Despite attempts to gather additional information, the site has not provided any additional event details. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument grip to be misaligned. The grip tip was found with 0. 035 offset at the tips. The bent grip has separation of the yaw pulley and pulley cover at the glue joint. Engineering has concluded that this failure may be the result of overloading at the tip. Electrical continuity passed. Engineering also observed pitch down cable to be broken at the distal clevis hub. The cable segment that contains the crimp is missing from the clevis. The clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1976370 | 1949881 | 1846877 | 2955842201100004 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-06 | 2011-01-05 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-04 | Other |  | It was reported that during a da vinci s ureteroplasty procedure, the customer experienced system error code 23017. With the assistance of a technical support engineer (tse), the site performed an emergency power off (epo) to restart the system multiples times, however, system error recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering confirmed the occurrence of system error code 23017 as well as system error code 25211. Engineering concluded that system error codes 23017 and 25211 were associated with the endoscopic camera manipulator (ecm) arm and remote arm controller (rac). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The remote arm controller (rac) consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected ecm and rac. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. System error code 25511 appears when the da vinci s safety system determines that a calibration error has occurred. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1972539 | 1945108 | 1842104 | 2955842201100001 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-04 | 2011-01-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-04 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-05-30 | Other | Other | It was reported that while starting a da vinci si hysterectomy procedure the site experienced a system error code 25326. With the assistance of an isi technical support engineer the site restarted the system. The system restarted without error and the site decided to continue with the planned surgical procedure. During anastomosis the surgeon had difficulty engaging an instrument installed on a patient side manipulator (psm) arm. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the system error code 25326 experienced by the customer was associated with a remote arm controller (rac)pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The fse was unable to replicate the instrument engagement issue reported by the customer. Multiple insertions of the instrument used in the procedure were made by the fe without any issue. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25326 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put davinci in a recoverable safe state. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2169222 | 2141421 | 2038397 | 2955842201100203 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-30 | 2011-06-27 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-30 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-25 | Malfunction |  | On (b)(6) 2011, a copy of medwatch uf/importer report (b)(4) was received from (b)(6) hospital with the following information: during robotic hysterectomy, guard on tip of davinci mega needle driver fell off inside patient. Guard was able to be retrieved. | The hospital indicated that the instrument would not be returned for evaluation. An intuitive surgical representative was able to obtain a photograph of the affected instrument, which was then evaluated by engineering. Per the photograph, engineering observed that the carbide insert had detached from one of the instrument grips. Since the instrument was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2168329 | 2140068 | 2037043 | 2955842201100197 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-02 | 2011-06-25 | No | Yes | Health Professional | 42019408 | No | Yes | 2011-06-02 | No | No | Unkown | 2011
42022703 | 2011-06-01 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while using the pk dissecting forceps instrument, a plastic piece fell into the patient. The site was unable to locate the fragment or confirm if the fragment was still retained in the patient. The planned surgical procedure was completed and no patient injury, harm or adverse outcome was reported. | Multiple attempts have been made, however as of the date of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2167913 | 2139660 | 2036636 | 2955842201100192 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-06-01 | 2011-06-24 | No | Yes | Health Professional | 42022703 | No | No | 2011-06-01 | No | No | Unkown | 2011
40027202 | 2011-05-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory was returned and evaluated. Per the customer reported complaint engineering observed the teflon pad is detached from the clamp arm. The middle section of the teflon pad is still attached to the clamp arm, but a u shaped piece of the clamp arm is detached. An isi performed a site visit on june 14, 2011 to observed cases, surgeon and surgical staff technique, and provided feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2167914 | 2139661 | 2036637 | 2955842201100193 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2011-05-26 | Device Not Retu | Yes | Initial | 2011
40027202 | 2011-05-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed the insert to be returned with a broken piece approximately. 670 inches by. 085 inches in size. The blade fractured below the clamp arm, within the shaft, and the clamp arm is intact. An isi product manager performed a site visit on (b)(6) 2011 to observe cases, surgeon and surgical staff technique, and provide feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2167915 | 2139662 | 2036638 | 2955842201100194 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2011-05-26 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-06-08 | Other | Other | It was reported that during the beginning of a da vinci s prostatectomy procedure, system error code 23008 occurred. In an attempt to clear the error, the surgical staff power cycled the system and back drove the left master tool manipulator. The system was then restarted, however, the same error code recurred after restart. The surgeon made the decision to abort the planned procedure and reschedule it for a later date. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. The mtml was returned for evaluation. Engineering was able to replicate the reported failure mode and determined that the main cable harness had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2167946 | 2139693 | 2036669 | 2955842201100195 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-06-08 | 2011-06-24 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-06-08 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-06-06 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the permanent cautery spatula instrument broke and the surgical staff had difficulty removing the instrument from the cannula due to the angle of the instrument wrist. As a result, the system arm was undocked from the patient with the instrument still installed and inserted in the cannula. While undocking, a piece of the instrument fell into the patient and was retrieved. The site was able to remove the broken instrument, the system arm was redocked, and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the reported complaint, engineering observed the yaw pulley and spatula assembly detached from the distal end of the instrument. The conductor wire has been cut at the distal end. One of the distal clevis ears is broken off at its base. The broken ear piece was not returned, with the missing piece approximately. 280 x. 235 in size. The conductor wire cap is also missing. The yaw pulley has the extruded boss feature at the center of the hub and one side of the cable crimp pocket broken off. It was determined that the initial failure was likely a broke clevis ear and the spatula/yaw pulley assembly was likely detached to remove the instrument from the cannula. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2167947 | 2139709 | 2036664 | 2955842201100196 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-06-06 | 2011-06-24 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-06-06 | Device Not Retu | No | Reuse | 2011
42011704 | 2011-05-24 | Malfunction |  | It was reported that during a da vinci s nissen paraesophageal repair procedure the tip of the 5mm needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one needle driver grip broken off where the grip arm meets the grip hub. The grip has a 90 degree angle at the fracture location. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2167951 | 2139712 | 2036687 | 2955842201100198 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-06-02 | 2011-06-24 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2011-06-02 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-05-25 | Malfunction |  | It was reported that while coagulating some tissue during a da vinci si hysterectomy procedure, the monopolar curved scissors instrument sparked. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a slight char mark on the tube extension connection point with the main tube. The connection is broken throughout the circumference. Some material is melted at the char mark, indicating an arcing event occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2167952 | 2139713 | 2036688 | 2955842201100199 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-25 | 2011-06-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-05-25 | Device Not Retu | No | Reuse | 2011
IS2000 51P8 | 2011-05-27 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff was unable to insert an instrument on the patient side manipulator (psm). The site attempted to troubleshoot by rebooting the system; however, the issue persisted. The patient had been under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the psm. The psm was returned and evaluated. Engineering was not able to replicate the reported failure mode; however, suspect components were replaced to repair the psm. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2168049 | 2139779 | 2036755 | 2955842201100201 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-27 | 2011-06-24 | Yes | Yes | Health Professional | IS2000 51P8 | Device Returned To Manufacturer | Yes | 2011-05-27 | Device Not Retu | No | Reuse | 2011
42000606 | 2011-05-24 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a fragment from the large needle driver instrument fell into the patient and was retrieved. The planned procedure was successfully completed with no harm to the patient. | The instrument was returned and evaluated. Engineering observed the instrument to be returned with the carbide insert detached from one grip. The brazing surface on the grip exhibits various voids in the filler material. The outer surface of grips do not exhibit damage. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2168035 | 2139794 | 2036769 | 2955842201100200 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-25 | 2011-06-24 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2011-05-25 | Device Not Retu | No | Reuse | 2011
40017908 | 2011-05-20 | Malfunction |  | It was reported that during a da vinci nephrectomy procedure, the surgical staff observed the tip of the monopolar curved scissors instrument break resulting in a piece falling into the patient. The piece was removed and the procedure was successfully completed with no harm to the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the left blade to be broken off of the instrument. The broken segment was not returned. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2168060 | 2139807 | 2036782 | 2955842201100202 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-26 | 2011-06-24 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2011-05-26 | Device Not Retu | No | Reuse | 2011
42017207 | 405928 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument broke and a fragment fell into the patient. While trying to locate the broken piece, there was a delay in the procedure. The broken piece was successfully retrieved and the planned surgical procedure was completed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip broken off at the base of the clevis. As a result, there is separation of the yaw pulley and pulley cover at the glue joint. The broken grip was returned with the instrument. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2164546 | 2136853 | 2033829 | 2955842201100191 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-05-23 | 2011-06-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 405928 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-12 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was suturing and using the mega needle driver instrument, the serrated teeth plate fell into the patient. The fragment was retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | Multiple attempts with the site have been made, however, as the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2161916 | 2134479 | 2031455 | 2955842201100190 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-12 | 2011-06-20 | No | Yes | Health Professional | 42019408 | No | No | 2011-05-12 | No | No | Unkown | 2011
P5300 | 2011-05-24 | Injury | Required Intervention | While docking robot, a non-recoverable fault alarm sounded. Troubleshooting measures did not work. Manufacturer contacted. This problem resulted in converting from laparoscopic to open surgery. Another case scheduled for that day had to be cancelled. Eval by manufacturer found 16 faults and no power to arm 2. The set up joint was replaced and equipment was tested. |  | DAVINA ROBOT | ROBOT | 2168624 | 2138245 | 2035221 | MW5021068 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-06-14 | 2011-06-20 | No | Yes | Health Professional | P5300 | Yes | Yes |  |  |  |  | 2011
400272 | 2011-04-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, when the surgeon began to amputate the uterus from the cervix, the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic ace insert accessory has not been returned for evaluation, however, an isi product manager performed a site visit on (b)(6) 2011 to observe cases, surgeon and surgical staff technique, and provide feedback to the site regarding frequent harmonic ace insert breakage. It was observed that the breakage the site is experiencing is occurring when the active blade comes in contact with metal or the blade is damaged due to intra operative collisions or other physical contact with other instruments. The da vinci harmonic ace curved shears instructions for use specifically states: general precautions and warnings: avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. The root cause of the customer reported failure mode cannot be determined due to the harmonic ace insert accessory not being returned for evaluation. If additional information is received, a follow-up mdr will be submitted. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2159836 | 2132658 | 2029632 | 2955842201100179 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-18 | 2011-06-17 | No | Yes | Health Professional | 400272 | No | Yes | 2011-05-18 | No | Yes | Unkown | 2011
40018302 | 2011-05-13 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, arcing was observed coming from the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a char mark on the yaw pulley and yaw pulley exit. The conductor wire has surface damage but is not broken. The ceramic sleeve also exhibits damage. No other damage found. The charring indicates an arcing event occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2159842 | 2132662 | 2029636 | 2955842201100180 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-05-19 | 2011-06-17 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | No | 2011-05-19 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-05-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy, the plate inside tip of the mega needle driver instrument came off, fell into the patient and was retrieved. The planned surgical procedure was completed with no patient harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering noticed the carbide insert detached from one of the instrument grips and no adhesive residue was observed on the grip tip surface. The detached insert was returned with the instrument. No damage was found on the insert on the other grip tip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2160398 | 2132954 | 2029928 | 2955842201100181 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-23 | 2011-06-17 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2011-05-23 | Device Not Retu | No | Initial | 2011
IS2000 | 2011-05-20 | Other | Life ThreateningOther | It was reported that while isolating the pulmonary vessel during a da vinci s thoracic sympathectomy procedure, a tear on the underside of the patient's pulmonary vessel occurred when the surgeon pulled on a calcified lymph node. Due to the location of the tear, the surgeon was unable to angle the system camera to view the affected area, therefore, unable to repair the artery using the system. The surgeon made the decision to convert to traditional open surgical techniques to repair the tear and complete the planned procedure. The da vinci procedure had been in progress for approx 2 hours when the event occurred. No patient is reportedly doing well and no post surgical complications have been reported. | An intuitive surgical field service engineer (fse) went onsite and evaluated the system. Visual and mechanical inspections were performed with no issues found. System verification testing was performed and passed. Based on the information provided, it was determined that the da vinci s surgical system worked as intended. No information has been received which suggests that the da vinci s surgical system malfunctioned in a way that may have caused or contributed to the patient's injury. Intuitive surgical concluded that the patient injury was unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160402 | 2132958 | 2029932 | 2955842201100183 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-20 | 2011-06-17 | Yes | No | Health Professional | IS2000 | Yes | No | 2011-05-20 | Device Not Retu | Yes | Reuse | 2011
IS3000 A60P6 | 2011-05-20 | Other | Other | It was reported that during a da vinci si nephrectomy procedure, the site experienced a loss of vision in the right eye image of the high stereo resolution viewer (hrsv). With the assistance of an isi technical support engineer the surgical staff reseated the vision cable on both ends and restarted the system, however, the vision issue was not resolved. The surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160410 | 2132966 | 2029940 | 2955842201100182 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-20 | 2011-06-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-20 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-05-24 | Other | Other | It was reported that during a da vinci si surgical procedure the site experienced system error code 16 and lost power on a patient side manipulator (psm) arm. With the assistance of an isi technical support engineer the site restarted the system and tried to manually reposition the psm, however the psm did not regain power. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system error code 16 was associated with a patient side manipulator (psm) setup-joint axis. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The setup-joint (suj) is a non-motorized support arm which holds a patient side manipulator (psm) during surgery. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 16 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the psm. Upon determining this condition, the system records the fault and transitions to a safe state. The suj was returned for failure analysis evaluation. Engineering was able to replicate the reported failure and replaced all internal electrical cables to address the system error codes experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160412 | 2132968 | 2029942 | 2955842201100184 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-24 | 2011-06-17 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-05-24 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-05-24 | Other | Other | It was reported that during a da vinci s hysterectomy the surgeon experienced a depth perception issue between the left and right eye of the surgeon console. The system was restarted, however, this action was unable to resolve the vision issue. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The camera head was returned and evaluated by the original equipment manufacturer (oem). The oem observed that the camera stage was damaged, thus causing the vision issue experienced by the customer. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160415 | 2132971 | 2029945 | 2955842201100185 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-05-24 | 2011-06-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-05-24 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-05-20 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s nissen fundoplication procedure, the plastic sheeting of the permanent cautery hook instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | Unkown | 2160493 | 2133036 | 2030012 | 2955842201100186 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-05-20 | 2011-06-17 | No | Yes | Health Professional | 42018305 | Yes | Yes | 2011-05-20 | Device Not Retu | No |  | 2011
42017910 | 2011-05-19 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the site observed the tip of the monopolar curved scissors (mcs) instrument to have broken. The piece was immediately retrieved and the instrument was replaced before the planned surgery was successfully completed without incident. The patient was reported as doing well with no post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension to be broken and missing a. 280 x. 210 piece at the prox clevis interface. The scissor blades are intact and the clevis is dislodged from tube extension. Evidence not conclusive, but tube extension likely broke due to excessive side loading. Electrical continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2160495 | 2133038 | 2030014 | 2955842201100187 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-05-19 | 2011-06-17 | No | Yes | Health Professional | 42017910 | No | No | 2011-04-13 | No | Yes | Initial | 2011
IS2000 51P8 | 2011-05-18 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff was unable to focus the camera. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the system; however they did not have a back up camera cable or camera head. The patient had been under anesthesia for 2 hours when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the camera cable, camera head and video digitizer. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images (such as the user interface icons and text or an auxilliary image from an ultrasound or other source). The system was repaired by replacing the camera cable, camera head and video digitizer. The components were returned and evaluated. Engineering not able to replicate the reported failure mode with the camera cable and digitizer; however, did reproduce the reported failure mode on the returned camera head. The camera head showed evidence of mechanical shock resulting in damaged stage assembly thus causing the reported complaint. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160496 | 2133039 | 2030015 | 2955842201100188 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-18 | 2011-06-17 | Yes | Yes | Health Professional | IS2000 51P8 | Device Returned To Manufacturer | Yes | 2011-05-18 | Device Not Retu | No | Reuse | 2011
IS1200 43P9 | 2011-05-20 | Malfunction |  | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, the surgical staff experienced the left eye of the surgeon side consoles high resolution stereo viewer (hrsv) to be black and the right eye to have some interference. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the system; however, the site did not have a back up camera cable. The patient had been under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the technical service engineer determined that the customer reported issue was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the camera cable. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of june 16, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2160498 | 2133041 | 2030017 | 2955842201100189 | 1 | NAY | Manufacturer | ForeignDistributor | Physician | Initial | 2011-05-20 | 2011-06-17 | No | Yes | Health Professional | IS1200 43P9 | No | Yes | 2011-05-20 | No | No | Unkown | 2011
40018010 | 2011-05-17 | Malfunction |  | It was reported that after a da vinci s hysterectomy procedure, two holes were observed to have been burned in the monopolar curved scissors (mcs) instrument tip cover accessory installed on the monopolar curved scissors instrument. No arcing was observed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed two holes burned though the silicone. The silicone exhibits localized melting around the hole which is indicative of an arcing event. The hole sizes are. 055 and. 045 x. 075 with varying shapes (round and oblong), and are located. 075 and. 320 above the silicone to sleeve interface. An additional investigation is being conducted to determine root cause of this event. If additional information becomes available, a follow-up mdr will be provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2159038 | 2131904 | 2028880 | 2955842201100177 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-17 | 2011-06-16 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-17 | Device Not Retu | Yes | Initial | 2011
420184 | 2011-05-17 | Malfunction |  | It was reported that after performing bivalve of the uterus during a hysterectomy procedure, the surgical staff attempted to remove the permanent cautery spatula instrument; however, the instrument's wrist was pitched to one side, preventing the instrument from being removed from the cannula. The surgeon then used a fenestrated bipolar instrument to straighten out the wrist of the permanent cautery spatula causing the spatula instrument to break. It was reported that three pieces of the instrument broke off and two were retrieved. Isi (b)(4) discussed the risks of instrument materials being left behind and as of the date of this mdr, the site has not reported the patient as experiencing any post operative issues. (b)(4). | Due to the instrument not returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings to handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2159051 | 2131918 | 2007640 | 2955842201100178 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-05-17 | 2011-06-16 | No | Yes | Health Professional | 420184 | No | No | 2011-05-17 | No | No | Unkown | 2011
400077 | 2011-06-01 | Malfunction |  | Patient with leiomyomatous uterus underwent robotic assisted hysterectomy (myomectomy only) and during surgery, the port seals leaked. Seals were switched/replaced and surgery continued with no apparent delays; no issue with patient. |  | 8MM CANNULA SEAL 1012MM CANNULA SEAL | ENDOWRIST INSTRUMENT 8MM ACCES | 2165261 | 2137606 | 2034582 | 2137606 | 1 | NAY | User Facility |  |  | Initial | 2011-06-16 | 2011-06-16 | No | Yes | Physician | 400077 | Yes | No Answer Provided |  |  |  |  | 2011
400169 | 2011-05-26 | Malfunction |  | Patient undergoing prostatectomy via robotic assistance. Surgeon was using the harmonic curved shears insert for the davinci when, during the procedure, he noticed the tip of the harmonic had come off. Harmonic tip noted to be off at approximately 3 hours into the surgery. It was all retrieved and surgery was complete. Patient sent to recovery. |  | HARMONIC CURVED SHEARS INSERT | ENDOWRIST INSTRUMENT 8MM ACCES | 2165264 | 2137607 | 2034583 | 2137607 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-06-16 | No | Yes | Physician | 400169 | No | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-05-16 | Other | Other | It was reported that during a da vinci s low anterior resection procedure, the surgeon side console (ssc) experienced a power loss. With the assistance of an isi technical support engineer, the site restarted the system, however, after homing the ssc continued to run off of the battery back-up system. The site tried multiple outlets in the operating room, however, the ssc continued to run from the battery back-up system. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found the power issue experienced by the customer was associated with the surgeon side console primary power supply (pps2). The pps2 is a 230v power supply with battery backup. The system was repaired by replacing the pps2 and the power supply fan and the battery pack module as a precaution. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2151351 | 2123830 | 2020806 | 2955842201100176 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-05-16 | 2011-06-14 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-05-16 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-13 | Malfunction |  | It was reported that while dissecting the pelvic lymph nodes during a da vinci si hysterectomy, a tear was observed on the monopolar curved scissors (mcs) tip cover accessory (installed on the mcs instrument) and arcing occurred. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and one hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is approximately. 020 in diameter and located. 090 above the silicone to sleeve interface. The silicone also exhibits scratches on the outer surface and a tear adjacent to the hole. Other various mechanical tears in the silicone were also observed in the areas not near the hole. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2147491 | 2119809 | 2016785 | 2955842201100174 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-13 | 2011-06-10 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-13 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P6 | 2011-05-12 | Other | Required Intervention | It was reported that while performing a da vinci si hysterectomy, the left master controller did not have full control of the maryland bipolar forceps instrument during preparation of the uterine tube, resulting in non-intuitive motion and causing a small bleed on the patient's uterine tube. Coagulation was used to stop the bleed and the planned surgical procedure was completed without further incident. No post surgical complications were experienced by the patient. | Investigation was conducted by a field service engineer (fse) who went onsite and evaluated the system. System verification testing was performed and passed. The fse was unable to replicate the customer reported failure mode and the root cause of the non-intuitive motion experienced by the customer cannot be determined. As a precaution, the fse recalibrated the left master tool manipulator (mtm). The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand and one to his right. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2147713 | 2120175 | 2017151 | 2955842201100175 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-05-12 | 2011-06-10 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-05-12 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Other | Other | It was reported that after a da vinci si surgical procedure, the patient was found to have what appeared to be a skin burn at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. Also see mdr 2955842-2011-00172. | An isi representative has attempted to gather additional information from the initial reporter related to this event without success. The sites robotics coordinator advised that often time, a bovie generator is used to create port incisions, which would give the appearance of a port burn. A follow-up mdr will be submitted if additional information is received. As of june 9, 2011, the site has continued to use the system and has not reported a recurrence of the issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2146622 | 2119068 | 2016044 | 2955842201100173 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-10 | 2011-06-09 | Yes | Yes | Health Professional | IS3000 A60P5 | No | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Other | Other | It was reported that after a da vinci si surgical procedure, the patient was found to have what appeared to be a skin burn at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. Also see mdr 2955842-2011-00173. | An isi representative has attempted to gather additional information from the initial reporter related to this event without success. The sites robotics coordinator advised that often time, a bovie generator is used to create port incisions, which would give the appearance of a port burn. A follow-up mdr will be submitted if additional information is received. As of june 9, 2011, the site has continued to use the system and has not reported a recurrence of the issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2146593 | 2119139 | 2016112 | 2955842201100172 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-05-10 | 2011-06-09 | Yes | Yes | Health Professional | IS3000 A60P5 | No | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-05-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the shaft of the monopolar curved scissors instrument bent and arching occurred. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned, however, a photographic image of the instrument was received and evaluated. Per the customer reported complaint, engineering observed a broken tube extension at the main tube interface with no evidence of arching seen. Since the instrument was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. --/-- Additional Manufacturer Narrative: --/-- (b)(4) - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension dislodged from the main tube and the entire tube extension wall circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. An arc track was found at the breakage location, confirming that an arcing event occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2145340 | 2117324 | 2014302 | 2955842201100171 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-05-09 | 2011-06-08 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2011-06-20 | Device Not Retu | No |  | 2011
| 2011-05-31 | Malfunction |  | Surgeon was using the device to dissect and it quit working so he removed it from the field and retrieved another one. After surgery, it was noted that a piece of the plastic tip was missing. This device has been reprocessed 4 times. Device was flushed in ultrasonic washer, sent through automated washer, inspected and packaged, and steam sterilized. |  | DA VINCI S ENDOWRIST | PERMANENT CAUTERY HOOK | 2159626 | 2132509 | 2029485 | 2132509 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-06-07 | 2011-06-07 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-05-05 | Other | Other | It was reported that approximately 3 hours into a da vinci si hysterectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, the site observed arcing to the patient's pelvic sidewall. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. No repair was required and the planned surgical procedure was completed. No patient harm or adverse outcome was reported. The following medwatch report was submitted to the fda in relation to this event: mfr report 2955842-2011-00166. | The tip cover accessory was returned and evaluated. Engineering observed various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole sizes are approximately. 010 inches to. 035 inches with varying shapes (round and oblong) and are located from. 070 inches to. 340 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2141327 | 2114010 | 2010934 | 2955842201100168 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-06 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | Yes | Initial | 2011
400169 | 2011-04-26 | Malfunction |  | (b)(4). Event desc: during a robotic laparoscopic supracervical hysterectomy, the harmonic ace curved shears insert's insulation piece broke off while being used. The piece was recovered without harm to patient. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | Multiple attempts with the site have been made, however, at the time of this report, the accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned and/or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2142680 | 2114813 | 1976010 | 2955842201100169 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-09 | 2011-06-06 | No | Yes | Health Professional | 400169 | No | Yes | 2011-05-09 | No | Yes | Unkown | 2011
420184 | 2011-02-22 | Malfunction |  | On may 12, 2011, maude event report (b)(4) was received: patient was undergoing robotic lh bilateral salpingo-oophorectomy and mccall's monarch anterior repair. The plastic part of the spatula broke off during the procedure. This has occurred previously with the same operator, according to operating room nurse. Operator has been re-trained by the company rep on previous occasions. Diagnosis or reason for use: surgery for hysterectomy and bladder suspension. | On may 17, 2011, (b)(6) risk manager, stated that the hospital no longer has the permanent cautery spatula instrument. She also stated that the surgical staff believes that a piece of the instrument fell into the patient, however, no pieces were retrieved during the procedure and the patient has not returned to have any fragments removed since. The operating room director stated that this surgeon has had recurring instrument breakage. Multiple clinical sales representatives had re-trained the surgeon in order to avoid future breakage events. Due to the instrument not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2142658 | 2114815 | 2011791 | 2955842201100170 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-05-12 | 2011-06-06 | No | Yes | Health Professional | 420184 | No | No | 2011-05-12 | No | No | Unkown | 2011
IS1200 A43P9 | 2011-05-06 | Other | Other | It was reported during a da vinci prostatectomy procedure, the bifurcated light guide cable melted. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found three bifurcated light guide cables to be burnt at the distal ends. The lumen light from the illuminator was tested and confirmed within specification. The bifurcated light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The system was repaired by replacing the illuminator as a precaution. The three bifurcated light guide cables were returned to isi for additional evaluation. Engineering observed two of the cables as having evidence of thermal damage at the glass tip on the proximal end connectors which is exhibiting melting. These cables also show signs of unauthorized re-work. The third cable had a red substance at the glass tip and metal connector which likely led to reduced performance. The illuminator was also returned to isi for evaluation and found to be within specification. As of june 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2141262 | 2113940 | 2010916 | 2955842201100165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-05-05 | Other | Other | It was reported that approximately 90 minutes into a da vinci si hysterectomy procedure, while dissecting the patient's lymph nodes and using the monopolar curved scissors instrument with the tip cover accessory installed, the site observed arcing to the patient's bowel. No repair was required and the planned surgical procedure continued. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. No patient harm or adverse outcome was reported. The following medwatch report was submitted to the fda in relation to this event: mfr report 2955842-2011-00167. | The tip cover accessory was returned and evaluated. Engineering observed various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole sizes are approximately. 010 inches in diameter and are located from. 060 inches to. 210 inches above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes and a long tear at the distal tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2141249 | 2113958 | 2010934 | 2955842201100166 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-06 | 2011-06-03 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-05-06 | Device Not Retu | Yes | Initial | 2011
42018305 | 2011-04-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si sigmoid colectomy the tip of the permanent cautery hook instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the hook broken off of the instrument. The hook had fractured within the yaw pulley, at the cross section of the hole feature. The detached hook was not returned. Engineering also observed that one side of the ceramic sleeve is broken off and missing. The size of the missing piece is approximately. 155 x. 075. It was determined that the damage to the hook and sleeve may be due to overloading at the tip or other mishandling/misuse. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 2139543 | 2111686 | 2008662 | 2955842201100164 | 1 | NAY | Manufacturer | ForeignDistributor | Other | FollowupInitial | 2011-05-03 | 2011-06-02 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2011-06-08 | Device Not Retu | No |  | 2011
IS3000 A60P6 | 2011-05-02 | Other | Other | It was reported that when the surgeon released the energy control foot pedal and removed their head out of the stereo viewer during a da vinci si hysterectomy, the surgeon heard the electrosurgical unit (esu) still releasing energy. The surgeon put their head back into the stereo viewer and noticed that the monopolar curved scissors instrument, which was connected to the esu, was releasing energy. The surgeon took control of the instrument and inadvertently touched fatty tissue. The surgeon quickly pressed the foot pedal to stop the instrument from releasing energy. It was reported that the patient suffered a small superficial burn. | Investigation was conducted by a field service engineer (fse) who went onsite and evaluated the system. Visual and mechanical inspections were performed with no issues found. The system logs were reviewed and no errors were observed. System verification testing was performed and passed. The fse demonstrated to the site that the instrument would cease fire the instant the head is removed from the stereo viewer. The root cause of the issue experienced by the customer cannot be determined. As a precaution, the fse swapped the food pedal assembly. Additionally, the site's biomedical department will replace the electrosurgical unit and energy activation cable. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2138371 | 2110566 | 2007542 | 2955842201100163 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-05-02 | 2011-06-01 | Yes | No | Health Professional | IS3000 A60P6 | Yes | Yes | 2011-05-02 | Device Not Retu | No | Reuse | 2011
420179 | 2011-05-27 | Malfunction |  | While separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. |  | HOTSHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2142904 | 2115178 | 2012154 | 2115178 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-05-31 | No | Yes | Physician | 420179 | No | No Answer Provided |  |  |  |  | 2011
420179 | 2011-05-27 | Malfunction |  | While separating the plane of the vagina from the rectum, the surgeon noticed a small microburn on the superficial layer of the bowel. To check device, the surgeon used the instrument again on a non-critical area of the bladder that was being removed and an arc/spark was noted. The monopolar scissors were taken out, the tip was inspected, and 2 small holes were noted in the tip cover accessory. Another tip was opened, inspected, and placed on the same scissors; another arc was noted. |  | HOTSHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2142904 | 2115178 | 2012154 | 2115178 | 1 | NAY | User Facility |  |  | Initial | 2011-05-31 | 2011-05-31 | No | Yes | Physician | 420179 | No | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 |  | Death | Death | It was reported that on (b)(6), 2011 a patient passed away 18 months after a da vinci s procedure to remove the lower lobe of their left lung due to bronchioloalveolar carcinoma. According to the patient's husband, during the procedure, the patient's stomach and diaphragm were punctured resulting in sepsis. The patient underwent a 5 month hospital stay and eighteen months as a near invalid. To date, (b)(6) has not provided the patient's husband with information regarding this event. Intuitive has contacted the sites risk management department, who advised that they would investigate this event and respond back to intuitive within the upcoming weeks. | Based on the information provided, it was undetermined as to how this event occured. The site has been contacted for additional information and once this information has been received, a follow up report will be sent to the fda. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2134370 | 2105804 | 2002780 | 2955842201100160 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2011-05-06 | 2011-05-27 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2011-05-06 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-05-03 | Other | Other | It was reported that after a da vinci surgical procedure, the patient was found to have charring at the 8mm port incision site. The patient did not require any additional medical treatment. | During our investigation with the site's rn, (b)(6), she advised that the port incisions reported by our isi field service engineer may have been intentionally caused during creation of the port incisions using a bovi unit. She advised that she would follow up with the surgical staff to obtain additional information. Despite following up with the site on several occasions, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2134351 | 2105813 | 2002789 | 2955842201100161 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-05-03 | 2011-05-27 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2011-05-03 | Device Not Retu | No | Reuse | 2011
400183 | 2011-03-14 | Malfunction |  | On may 3, 2011, medwatch (b)(4) was received with the following information: event desc: patient undergoing a robotic sparing whipple and cholecystectomy for a pancreatic mass. During the procedure a piece of the hook cautery was noted to be broken off and in the patient's abdominal cavity. Copious irrigation and suction were performed. It is unknown whether the piece was retained or removed. No additional information has been received despite several attempts to contact the site. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2134352 | 2105814 | 1968844 | 2955842201100162 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-05-03 | 2011-05-27 | No | Yes | Health Professional | 400183 | No | No | 2011-05-03 | No | No | Unkown | 2011
42018405 | 2011-04-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the permanent cautery spatula instrument broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | PERMANENT CAUTERY SPATUAL INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2131566 | 2103542 | 2000518 | 2955842201100157 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-27 | 2011-05-26 | No | Yes | Health Professional | 42018405 | No | Yes | 2011-04-27 | No | No | Unkown | 2011
40018010 | 2011-05-24 | Other | Other | (b)(4): while the surgeon was using the da vinci robot, the monopolar scissors split a hole in three protective tip covers. When the third tip cover tore, a small burn was noted in the patient's bowel. What was the original intended procedure? total abdominal hysterectomy with bilateral salpingo-oophorectomy. Device usage problem: device malfunction - that is, the device did not do what it was supposed to do. The user facility had already reported this event to intuitive surgical (isi) on (b)(6) 2011. It is the hospital's belief that the patient injury's was caused by the monopolar curved scissors (mcs) instrument. Repair to the patient's damaged bowel was not required and no postoperative complications have been experienced by the patient. | The instrument was returned and evaluated. Engineering found no signs of arcing on the instrument tube extension, clevis, scissor tips, nor the tube extension to main tube connection. The instrument was driven on a system and functioned properly with no trouble found. The hospital informed isi that the three tip cover accessories would not be returned for evaluation and isi's request for photographic images was denied. The hospital did allow an isi representative to evaluate the tip covers at the site. Although tears were noticed on all of the tip covers, no evidence of charring/arcing was observed. As no evidence of arcing was observed on the instrument and tip covers, the root cause of the patient bowel burn cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2132127 | 2104010 | 2000986 | 2955842201100158 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-26 | 2011-05-26 | Yes | No | Health Professional | 40018010 | Yes | No | 2011-04-26 | Device Not Retu | Yes | Unkown | 2011
42009308 | 2011-04-26 | Other | Other | It was reported that during a da vinci si myomectomy procedure, the instruments installed on both instrument arms did not have sufficient closing force on the instrument jaws to grasp tissue or throw a stitch. The surgical staff reseated the instruments, with no change. The instruments were replaced, with no change. The sterile adapters were then reseated, which resolved the issue. Later in the procedure, the wrist cables of the prograsp forceps instrument were stuck on the cannula tip, thus preventing removal of the instrument. The instrument arm was undocked from the patient, which allowed the surgical staff to withdraw the instrument from the cannula. After the instrument was removed, there was some difficulty redocking the instrument arm to the patient. Due to the instrument issues experienced by the surgeon, the decision was made to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing and the clevis hub exhibits a wear mark adjacent to the cable groove. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2132478 | 2104261 | 2001237 | 2955842201100159 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-26 | 2011-05-26 | Yes | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2011-04-26 | Device Not Retu | No | Reuse | 2011
420184 | 2011-04-27 | Malfunction |  | Patient underwent gynecological surgery utilizing da vinci robotics. An intuitive cautery spatula was utilized during the surgery. While moving the camera and the cautery spatula within the abdominal cavity, the two objects came in contact and a small plastic part was dislodged from the tip of the cautery spatula. It was not possible to identify the piece and the surgeon elected to leave it in the abdominal cavity. This was the instrument's final use. ======================manufacturer response for endowrist monopolar cautery instrument, permanent cautery spatula (per site reporter)======================the representative who works in the hospital will let his supervisor know about the issue. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST MONOPOLAR CAUTERY IN | 2138433 | 2110633 | 2007609 | 2110633 | 1 | NAY | User Facility |  |  | Initial | 2011-05-26 | 2011-05-26 | No | Yes | Physician | 420184 |  | No Answer Provided |  |  |  |  | 2011
| 2011-05-24 | Injury | Disability | During robotic prostatectomy, it was noted that the large needle driver wasn't holding and the suture was slipping as the physician was tying. The needle holder was removed and a new on placed. The physician then identified a small dark piece of material in the lower abdomen. It was removed and noted to be a small piece from the tip of the needle driver. No injury to the pt. |  | LARGE NEEDLE DRIVER 8MM IS2000 | LARGE NEEDLE DRIVER | 2144394 | 2113223 | 2010199 | MW5020815 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-05-26 | 2011-05-26 | No | Yes | Health Professional |  | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
| 2011-05-11 | Malfunction |  | No energy delivered to robotic monopolar device via the energy cable. This is the second cable from the manufactuer that has failed to function properly, in the last several weeks. The staff firmly believe the failure is due to a design flaw. ======================manufacturer response for energy source cable, energy activation cable, valleylab (per site reporter)======================the informal response from the engineer and the sales rep is that it is "user error". |  | ENERGY ACTIVATION CABLE VALLEYLAB | ENERGY SOURCE CABLE | 2132912 | 2104769 | 2001745 | 2104769 | 1 | GEI | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2011
420093 | 2011-02-24 | Malfunction |  | Patient underwent a da vinci-assisted laparoscopic bilateral pelvic lymph node dissection, extraction of bladder stones times four and radical prostatectomy, nonnerve-sparing. During the procedure, a prograsp forcep broke within a trocar. It was not possible to remove the device until the instrument within the trocar had been straightened with one of the other robotic arms. It was noted that there was a small foreign body loose within the abdomen after the forcep broke. Before the end of the procedure, the foreign body was located and removed. This was the 9th use with expected use of 10 times. ======================manufacturer response for endowrist grasper instrument, prograsp forceps 8 mm (per site reporter)======================none. |  | PROGRASP FORCEPS | ENDOWRIST GRASPERS INSTRUMENT | 2137915 | 2110066 | 2007042 | 2110066 | 1 | NAY | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Physician | 420093 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420183 | 2011-05-09 | Malfunction |  | Patient undergoing robotic myomectomy. The da vinci permanent cautery hook broke into three pieces in the intra-abdominal cavity. A scope was utilized and the cautery hook pieces were recovered. This was the 7th time this instrument had been used - expected life is ten uses. ======================manufacturer response for endowrist monopolar cautery instrument, permanent cautery hook (per site reporter)======================manufacturer has not responded as of the date of this notice. |  | PERMANENT CAUTERY HOOK | ENDOWRIST MONOPOLAR CAUTERY I | 2137917 | 2110067 | 2007043 | 2110067 | 1 | NAY | User Facility |  |  | Initial | 2011-05-25 | 2011-05-25 | No | Yes | Physician | 420183 | Yes | No Answer Provided |  |  |  |  | 2011
420184 | 2011-05-17 | Malfunction |  | When removing endowrist instrument from cannula, spatula was not coming out. Tip of instrument missing a piece on visualization. Md notified and inspected op site. Company representative was in room and is aware of issue. Called company. Device not radiolucent. Risk md from company said tip will not harm patient. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST MONOPOLAR CAUTERY IN | 2138482 | 2110664 | 2007640 | 2110664 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-23 | 2011-05-23 | No | Yes | Physician | 420184 | No | No Answer Provided |  |  |  |  | 2011
42018304 | 2011-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the permanent cautery hook (pch) instrument broke and a piece of the plastic clevis fell into the patient. The fragment was retrieved and the pch instrument was replaced with another pch instrument. The replacement instrument also broke and fragments fell into the patient. The surgical staff believes that all fragments were found, but cannot be certain. Postoperatively, the patient is reportedly doing well and has not experienced any complications. This report is for the initial pch instrument used during the surgical procedure. Medwatch mfr report 2955842-2011-00151 was submitted for the second pch instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis has one ear broken off at its base. The ear that is located on the conductor wire side broke off. The broken piece was not returned. The size of the missing ear piece is approximately. 280 x. 220. As a result of the breakage, the yaw pulley is dislodged and no longer axially aligned with the shaft. Engineering also observed that the conductor cap is missing, exposing the conductor wire. The yaw pulley boss that mates with the hole in the cap is also broken. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2126128 | 2097326 | 1994302 | 2955842201100150 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the permanent cautery hook (pch) instrument broke and a piece of the plastic clevis fell into the patient. The fragment was retrieved and the pch instrument was replaced with another pch instrument. The replacement instrument also broke and fragments fell into the patient. The surgical staff believes that all fragments were found, but cannot be certain. Postoperatively, the patient is reportedly doing well and has not experienced any complications. This report is for the second pch instrument used during the surgical procedure. Medwatch mfr report 2955842-2011-00150 was submitted for the initial pch instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off on the conductor wire side. The broken ear piece and conductor cap were returned detached from the distal end. The size of the broken ear is approximately. 285 x. 240. As a result of the breakage, the yaw pulley is dislodged and no longer axially aligned with the shaft. The yaw pulley boss feature that holds the cap in place is also broken. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2126130 | 2097327 | 1994302 | 2955842201100151 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | No | Reuse | 2011
420142 | 2011-03-08 | Other | Other | It was reported that during a da vinci s radical tonsillectomy procedure, the surgeon noticed the 5mm cautery instrument was not energizing at the tips as intended. The site checked the electrical surgical unit (esu) setting and cables and decided to increase the level of cautery. The surgical staff then noticed that cautery was transferring to the tongue retractor and dissipating throughout the patient's mouth. Unintended tissue was 'charred'; however, no significant injury was reported and the patient did not require any additional treatment or procedure. The planned surgical procedure was completed using the same instrument, however, the site created an insulation barrier out of operating room materials to cover the instrument snake wrist and ensure cautery leakage did not recur. No additional patient injury was reported. | Multiple attempts have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | 5MM MONOPOLAR CAUTERY INSTRUME | ENDOSCOPIC ELECTROSURGICAL INS | 2126689 | 2097936 | 1994912 | 2955842201100149 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-05-20 | Yes | Yes | Health Professional | 420142 | No | No | 2011-03-08 | No | No | Unkown | 2011
IS120043P9 | 2011-04-26 | Other | Other | It was reported that prior to starting a da vinci pelvic floor repair procedure, the surgical staff had difficulty aligning the endoscope. With the assistance of an isi technical support engineer (tse), the site exchanged the camera head and was then successfully able to align the endoscope. Approximately 1 hour later, the site called back stating the surgeon was experiencing double vision in the left eye of the surgeon side console (ssc) viewer. The tse had the site exchange and recalibrate the endoscope and restart the vision side cart. The image appeared normal and the surgeon continued with the planned procedure. Approximately 30 minutes later, the site called back stating the surgeon was experiencing black bars in the left eye, causing double vision. The patient was under anesthesia for approximately 2 hours when the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was related to the synchronizer. The synchronizer coordinates the left eye and right eye video streams so that events are displayed simultaneously to each of the surgeon's eyes. The system was repaired by replacing the affected synchronizer. The synchronizer was returned to isi for failure analysis investigation and engineering was able to replicate the customer reported issue. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2126690 | 2097940 | 1994916 | 2955842201100152 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-26 | 2011-05-20 | Yes | Yes | Health Professional | IS120043P9 | Device Returned To Manufacturer | Yes | 2011-04-26 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-04-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda in relation to this case: 2955842-2011-00154. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover to have multiple holes burned through the silicone, located along one parting line. The silicone exhibits localized melting around the holes indicative of arcing. The holes are irregularly shaped, ranging from. 045 inches to. 065 inches and are located. 280 inches and. 365 inches above the silicone to sleeve interface. The monopolar curved scissors (mcs) instrument used with the tip cover accessory in this procedure was also returned. With the tip cover accessory installed on the mcs instrument, the hole positions line up with the edge of the proximal clevis ear and the distal idler pulley. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2126691 | 2097945 | 1994931 | 2955842201100153 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-04-21 | Device Not Retu | Yes | Initial | 2011
40018010 | 2011-04-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The mcs instrument and tip cover accessory were replaced, however, the site experienced arcing from the replacement mcs instrument with a tip cover accessory installed. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda in relation to this case: 2955842-2011-00153. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2126704 | 2097955 | 1994931 | 2955842201100154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-21 | 2011-05-20 | No | Yes | Health Professional | 40018010 | No | Yes | 2011-04-21 | No | Yes | Unkown | 2011
42019408 | 2011-04-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy, the pad from the tip of the mega needle driver instrument, fell into the patient and was retrieved. The planned surgical procedure was completed with no further incidents. No injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering's inspection found the instrument to be missing one carbide insert from grip tip. The surface of the separation showed no residue of adhesive material. No damage was found on clevis or grip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2126787 | 2098018 | 1994994 | 2955842201100155 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-28 | 2011-05-20 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2011-04-28 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-04-21 | Other | Other | It was reported that during set up for a da vinci s nephrectomy procedure, the surgical staff experienced the left eye on the surgeon side console to be gray with a line down the middle of the screen. With the assistance of an isi technical support engineer, the site verified their settings and tested the camera cable which was determined to need replacement. The patient was under anesthesia when the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of may 26, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2126788 | 2098019 | 1994995 | 2955842201100156 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-21 | 2011-05-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-04-21 | Device Not Retu | No | Reuse | 2011
420227 | 2011-05-18 | Malfunction |  | The forceps worked during the surgery. The thought is that the defect occurred during the cleaning process. There is a small black wire that broke during what the staff believe when cleaned. The forceps will not open and have been taken out of service. |  | PK DISSECTING FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2135783 | 2107330 | 2004306 | 2107330 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-20 | 2011-05-20 | No | Yes | Invalid Data | 420227 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-04-19 | Other | Other | It was reported that approximately 90 minutes into a da vinci s diverticulum surgical procedure, the site experienced system error code 19 followed by system error code 45049. With the assistance of an isi technical support engineer (tse) the site restarted the system, however, upon power up system error code 281 appeared. At this time, the site stated that a small amount of water had been spilled on the base of the patient side cart. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes experienced by the customer were associated with two remote arm controller (rac) pca boards and the patient cart controller (pcc). The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The pcc is a board embedded in the patient cart that provides all user interface, communications control, and patient cart motor control for the patient side system. The system was repaired by replacing the affected rac pca boards and the pcc board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error codes reported occur when the da vinci s safety system determines a hardware fault reaction logic occurring on a local rac. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The rac3, rac4, and pcc were returned to isi for failure analysis investigation. Engineering evaluation confirmed fluid contamination as reported by the site. The da vinci s surgical system user manual specifically states: caution: equipment is not designated for exposure to liquids. Care should be taken to ensure liquids do not flow into electronic equipment. On (b)(4) 2011, field service engineering had extensive conversations with the site regarding proper cleaning of the da vinci s system to eliminate future occurrences of this event. As of may 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2122927 | 2091980 | 1988956 | 2955842201100148 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-19 | 2011-05-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-04-19 | Device Not Retu | No | Reuse | 2011
42020703 | 2011-03-02 | Malfunction |  | (b)(4). Using the da vinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. On (b)(4) 2011, the initial reporter informed intuitive surgical that all broken pieces were retrieved and there was no injury to the patient. | The instrument was returned with the proximal clevis detached from the distal end. The proximal clevis was also returned, in three broken pieces. Engineering evaluation observed that the main tube interface wall has collapsed. As a result of the breakage, the wrist is loose and motion is non-intuitive. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2121184 | 2089892 | 1986868 | 2955842201100146 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-18 | 2011-05-16 | No | Yes | Health Professional | 42020703 | Device Returned To Manufacturer | No | 2011-04-18 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-04-19 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced system error code 21008. The system was restarted, however, the error code recurred. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 21008 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controller which provides the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 21008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the system records the fault and transitions to a safe state. The mtmr was returned for failure analysis investigation. Engineering was able to replicate the customer reported failure mode and determined that an axis potentiometer and motor/encoder had malfunctioned. As of may 12, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2121185 | 2089893 | 1986869 | 2955842201100147 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-04-19 | 2011-05-16 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-04-19 | Device Not Retu | No | Reuse | 2011
42017910 | 2011-04-11 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the site observed arcing from the shaft of the monopolar curved scissors (mcs) instrument. The patient experienced a superficial burn to their bowel which did not require any repair or treatment. The planned surgical procedure was successfully completed without incident. The patient was reported as doing well with no post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed traces of arcing at the joint between the tube extension and main tube. The material was slightly melted and the tube extension is dislodged from the main tube. No trace of arcing was found at clevis or scissor. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2120513 | 2089035 | 1986011 | 2955842201100144 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-13 | 2011-05-13 | No | Yes | Health Professional | 42017910 | No | No | 2011-04-13 | No | Yes | Initial | 2011
42018304 | 2011-04-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, a piece of the permanent cautery hook instrument fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2120523 | 2089044 | 1986020 | 2955842201100145 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-15 | 2011-05-13 | No | Yes | Health Professional | 42018304 | No | No | 2011-04-15 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-01-20 | Other | Other | It was reported that on (b)(6) 2011, a da vinci s pancreatoduodenectomy was successfully performed with no system malfunctions. On (b)(6) 2011, the patient's drain amylase level rose to 10,480. Additionally, pancreatic fistula from the anastomosis site between the pancreas and the jejunum was found. As the drainage during the procedure was proper, additional treatment was not performed. The patient was discharged on (b)(6) 2011. | The console surgeon who performed the procedure indicated that the postsurgical symptoms experienced by the patient can develop regardless of the modality used to perform the surgery and believes that the patient's symptoms are unrelated to the use of the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunctions occurred, and the system did not malfunction in a way that may have caused or contributed to the patient's postsurgical symptoms. Intuitive surgical concluded that the patient's symptoms are unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2115107 | 2083893 | 1980869 | 2955842201100141 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-04-11 | 2011-05-11 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-04-11 | No | No | Reuse | 2011
400180 | 2011-04-12 | Malfunction | Other | It was reported that after the right side lymphectomy dissection was completed during a da vinci s hysterectomy procedure, arcing was observed on the tip cover accessory that was installed on the monopolar curved scissors (mcs) instrument. No patient harm occurred during the arcing event and the planned surgical procedure was completed. | The tip cover accessory was discarded by the site, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2115109 | 2083909 | 1980885 | 2955842201100142 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-12 | 2011-05-11 | No | Yes | Health Professional | 400180 | No | No | 2011-04-12 | No | Yes | Unkown | 2011
42017207 | 2011-04-12 | Injury | Other | It was reported that while dissecting the right side "lymphectomy" during a da vinci s hysterectomy procedure a tear was noticed halfway through the patient's right iliac vein. The vein was repaired laparoscopically using 12 clips and floseal. The planned surgical procedure was completed and it was reported that no postoperative complications have been experienced by the patient. | A video recording of the procedure was reviewed by an intuitive surgical clinical sales representative and the console surgeon. The video revealed that the surgeon inadvertently grasped the patient's right iliac vein with a maryland bipolar forceps instrument. It is the surgeon's belief that this action caused the damage to the patient's vein. No issues were observed with the maryland bipolar forceps instrument. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2115110 | 2083910 | 1980886 | 2955842201100143 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-04-12 | 2011-05-11 | Yes | No | Health Professional | 42017207 | No | No | 2011-04-12 | No | No | Reuse | 2011
420290 | 2011-05-05 | Malfunction |  | Patient with ureteral obstruction on left side has ureteral stent in place and unexplained hydroureter w/confirmed obstruction. Proceed with ureteral implantation. Surgeon had robotic left ureteral reimplantation/robotic nephroureterectomy and during case the green cannula seals were leaking and needed to be changed out. They came from davinci kit accessory 3-arm which was opened for the procedure. Two individual green cannula seals were open to replace the ones leaking. |  | DISPOSABLE ACCESSORY KIT | DISPOSABLE ACCESSORY KIT 3ARM | 2124457 | 2094959 | 1991935 | 2094959 | 1 | NAY | User Facility |  |  | Initial | 2011-05-10 | 2011-05-10 | No | Yes | Physician | 420290 | No | No Answer Provided |  |  |  |  | 2011
40016002 | 2011-04-12 | Malfunction |  | Event Description: --/-- It was reported that while resecting tissue during a da vinci si tongue base resection procedure, the surgeon observed a flame on the 5mm cautery spatula tip accessory that was installed on the 5mm monopolar cautery instrument. The instrument and tip accessory were immediately removed and replaced and the planned surgical procedure was completed with a replacement instrument and tip accessory. No patient harm was reported. | Manufacturer Narrative: --/-- A video of the surgical procedure was provided to intuitive surgical by the hospital. Engineering reviewed the video and observed that on multiple occasions the surgeon activated the energy for the 5mm monopolar cautery instrument with a bipolar instrument in close proximity to its spatula tip. As a result, the spatula tip arcs to the bipolar instrument at various times, consequently degrading the insulation barrier on the spatula and leading to the malfunction experienced by the customer. The instruments and accessories user manual specifically states in the 5mm monopolar cautery instructions for use: warning: as with any cautery device, it is possible for energy to discharge in an area other than the instrument tip. It is important to exercise caution when using an energized endowrist 5 mm monopolar cautery instrument to help avoid unintended contact with tissue adjacent to the area to be cauterized. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip accessory were returned and evaluated. Per the customer reported complaint, engineering observed localized melting with material removed on one side of the plastic sleeve of the spatula tip. The damage indicates that an arcing event occurred. The instrument was also inspected and no physical damage was observed. Engineering performed functional tests with the instrument using a new spatula tip and the cautery performed as intended with no occurrence of arcing. Additional investigation was performed by a field service engineer, who went onsite and evaluated the system. No errors were found when reviewing the system error logs and no issues were observed during inspection of the system arm that the instrument was installed on during the surgical procedure. The electrocautery cables were also inspected and no damage was found. --/-- Additional Manufacturer Narrative: --/-- Correction: a non-energy instrument was used with the 5mm monopolar cautery spatula instrument when the incident occurred. Medwatch mfr report 2955842-2011-00134, follow-up 01, indicated that an energy (bipolar) instrument was used. Additional information: further evaluation of the surgical procedure video was performed by engineering. While reviewing the video, engineering observed that the non-energy instrument was used repeatedly to scrape off/clean the spatula tip, which could have contributed to the tip damage. Prior to the incident occurring, multiple arcing events were observed and chunks of silicone from the tip had fallen off of the spatula and were visible in the surgeon's field of view. The instrument damage can clearly be seen, however, the surgeon continues to operate. The instruments and accessories user manual specifically states under general precautions and warnings: - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM CAUTERY SPATULA TIP ACCESSORY | Initial | 2113145 | 2082089 | 1979065 | 2955842201100134 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-05-16 | 2011-05-09 | No | Yes | Health Professional | 40016002 | Device Returned To Manufacturer | No | 2011-06-17 | Device Not Retu | Yes |  | 2011
IS1200 A43P9 | 2011-04-07 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure, while the surgeon was starting to dissect the patient's bladder, the site experienced a system error code 212. With the assistance of an isi technical support engineer, the restarted the system and then performed an emergency power off (epo) however, the system error code 212 reoccurred. The patient was under anesthesia and port incisions had been placed when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator (mtml) arm. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering was able to confirm the occurrence of system error code 212, thus confirming the reported failure mode experienced by the customer. The mtml main harness and axis 2 motor were replaced. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111887 | 2080463 | 1977440 | 2955842201100135 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-04-07 | 2011-05-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-04-07 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-04-13 | Other | Other | It was reported that during the set up for a da vinci si hysterectomy procedure the site experienced system error 23027. The surgical staff pressed fault over ride and continued with the system set up. The system was docked to the patient and the surgeon began to work from the surgeon side console when the system error 23027 reoccurred. With the assistance of an isi technical support engineer the site back drove the master tool manipulators (mtms) with the system on and then off, however the system error 23027 continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23027 experienced by the customer was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 23027 system error code appears when the da vinci si safety system determines that one of the health checks on a secondary joint sensor on the arm did not pass. Upon determining these conditions, the safety systems put da vinci si in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering determined that the flex circuit going from the axis 2 potentiometer to the embedded serializer (esmb) was loose, thus generating the system error code experienced by the customer. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111889 | 2080464 | 1977441 | 2955842201100136 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2011-04-13 | 2011-05-06 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2011-04-13 | Device Not Retu | No | Reuse | 2011
40027202 | 2011-04-05 | Malfunction |  | It was reported that during a da vinci s colorectal procedure, while the surgeon was using the harmonic curved shears instrument with the harmonic ace insert, the tip broke and fell into the patient. The fragment was retrieved and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, injury or adverse outcome was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found the curved blade to be broken at the distal end with a 0. 415 inch by. 085 inch piece detached. The blade fractured slightly above the straight to curved transition point. The broken piece exhibits various scratch and gouge marks on the side that make contact with the clamp arm. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: "handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. " do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2111890 | 2080467 | 1977444 | 2955842201100137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-11 | 2011-05-06 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-04-11 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-04-08 | Other | Other | It was reported that during a da vinci s low anterior resection procedure, the site experienced a red image at the surgeon side console. The surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted found that the vision system was out of tolerance due to the camera controller unit (ccu) not calibrating. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111867 | 2080475 | 1977451 | 2955842201100138 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-04-08 | 2011-05-06 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-04-08 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-04-11 | Other | Other | It was reported that while starting a da vinci prostatectomy procedure, the site found a broken cable on a patient side manipulator (psm) arm. The patient was under anesthesia when the surgical staff decided to abort the planned procedure. No patient harm, injury or adverse outcome was reported. | The investigation conducted by field service engineering (fse) confirmed a broken cable on the patient side manipulator (psm) 2 arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the axis 6 cable to be broken. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2111894 | 2080493 | 1977470 | 2955842201100139 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-04-11 | 2011-05-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-04-11 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-04-07 | Malfunction |  | It was reported that during a da vinci s sigmoid colectomy, an unidentified item from the harmonic curved shears insert instrument fell into the patient and was immediately retrieved. There has been no additional information provided to date. No patient harm, adverse outcome, or injury was reported. | Only the insert was returned and evaluated, but not the clamp arm. Per the customer reported complaint, engineering found the instrument clamp arm missing from the distal end. Clamp arm pins were bent backwards. The cause of the user damage was likely due to the user mishandling and/or over extension of the arm. There was no other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL | 2112073 | 2080766 | 1977743 | 2955842201100140 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-04-07 | 2011-05-06 | No | Yes | Other | 40016907 | Device Returned To Manufacturer | No | 2011-04-13 | Device Not Retu | No | Reuse | 2011
| 2011-04-25 | Malfunction |  | While the surgeon was using the davinci robot, the monopolar scissors split a hole in three protective tip covers. When the third tip cover tore, a small burn was noted in the patient's bowel. |  | HOT SHEARS MONOPOLAR CURVED SCISSORS | ENDOWRIST MONOPOLAR CAUTERY IN | 2117890 | 2086793 | 1983769 | 2086793 | 1 | NAY | User Facility |  |  | Initial | 2011-05-04 | 2011-05-04 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420206 | 2011-05-04 | Malfunction |  | Event Description: --/-- Patient underwent robotic assisted hysterectomy. After placing trocars in patient, the robotic instruments were introduced. One of the green robotic seals began to leak as soon as the instrument was put in. Instrument was removed, seal reseated, instrument put back in, continued to have a significant leak of c02, which was audible. Seal was removed and replaced with a new one. Leak was resolved with new seal. --/-- Additional Event Description: --/-- Patient underwent robotic assisted hysterectomy. After placing trocars in patient, the robotic instruments were introduced. One of the green robotic seals began to leak as soon as the instrument was put in. Instrument was removed, seal reseated, instrument put back in, continued to have a significant leak of co2, which was audible. Seal was removed and replaced with a new one. Leak was resolved with new seal. | Manufacturer Narrative: --/-- | CANNULA SEAL 12MM |  | 2120834 | 2089511 | 1986487 | 2089511 | 1 | NAY | User Facility |  |  | Initial | 2011-05-04 | 2011-05-04 | No | Yes | Physician | 420206 | No | No Answer Provided |  |  |  |  | 2011
420205 | 2011-04-06 | Malfunction |  | Device never worked. Or staff will contact rep to come and get device. |  | FENESTRATED BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2117898 | 2086784 | 1983760 | 2086784 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-02 | 2011-05-02 | No | Yes | Physician | 420205 | Yes | No Answer Provided |  |  |  |  | 2011
420110 | 2011-04-11 | Malfunction |  | Could not get device to work. Same type of instrument (new) was used to complete procedure. No harm to patient. Or staff will call rep to come get device. |  | PRECISE BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2117901 | 2086787 | 1983763 | 2086787 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-05-02 | 2011-05-02 | No | Yes | Physician | 420110 | Yes | No Answer Provided |  |  |  |  | 2011
42020706 | 2011-03-23 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a piece of 'wire' from the tenaculum forceps instrument broke and fell into the patient. The fragment was retrieved, and the planned surgical procedure was completed. No patient harm, injury or adverse event was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the pitch down cable to be broken at the distal clevis hub. The cable segment that contains the crimp was returned detached from the instrument. Other cables at the wrist are not damaged and the clevis does not exhibit any damage or wear. No other damage found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2103931 | 2072701 | 1969677 | 2955842201100127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-04-01 | 2011-04-29 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2011-04-01 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-03-30 | Other | Other | It was reported that while starting a da vinci si prostatectomy procedure, the site experienced a system error code 25580. The patient was under anesthesia, and port incisions had been placed when the surgical staff made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system error code experienced by the customer was associated with a remote arm controller (rac) pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25580 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of april 28, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103915 | 2072702 | 1969678 | 2955842201100129 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-30 | 2011-04-29 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-03-30 | Device Not Retu | No | Reuse | 2011
42019407 | 2011-04-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was suturing and using the mega needle driver instrument, a break occurred in the instrument shaft and 'plastic' fragments fell into the patient's abdomen. One piece was retrieved, however, tiny fragments were unable to be removed and remain in the patient. The instrument was replaced with another instrument of the same type and the planned surgical procedure was completed. No injury was reported. | Multiple attempts with the site have been made, however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and/or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2103908 | 2072729 | 1969705 | 2955842201100126 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-04-05 | 2011-04-29 | No | Yes | Health Professional | 42019407 | No | Yes | 2011-04-05 | No | No | Unkown | 2011
42017909 | 2011-03-31 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the 'jaw' of the monopolar curved scissors instrument partially broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a segment of the right blade to be broken off from the instrument. The broken piece was not returned. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2103910 | 2072730 | 1969706 | 2955842201100128 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2011-03-31 | 2011-04-29 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2011-03-31 | Device Not Retu | No | Reuse | 2011
IS2000 51P7 | 2011-03-31 | Other | Other | It was reported that during a da vinci s hysterectomy procedure whenever the surgeon moved a specific patient side manipulator (psm) arm, system error 23008 would appear. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the right master tool manipulator (mtmr) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a nonrecoverable safe state. The mtmr gimbal was returned to isi for evaluation. Engineering confirmed that the customer reported failure mode was associated with a defective sensor on axis 5 and the sensor was replaced. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103916 | 2072734 | 1969710 | 2955842201100130 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-31 | 2011-04-29 | Yes | Yes | Health Professional | IS2000 51P7 | Device Returned To Manufacturer | No | 2011-03-31 | Device Not Retu | No | Reuse | 2011
42020505 | 2011-03-30 | Malfunction |  | It was reported that while resecting during a da vinci s hysterectomy procedure, the fenestrated bipolar forceps instrument tip broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one instrument grip broken off near the top of the yaw pulley. The broken grip appears to be bent slightly backwards when the fracture planes are matched up, indicating that the damage may have been caused by mishandling. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2103970 | 2072779 | 1969755 | 2955842201100131 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-03-30 | 2011-04-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2011-03-30 | Device Not Retu | No | Reuse | 2011
42017207 | 2011-03-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure a piece of the maryland bipolar forceps instrument fell into the patient. The fragmented piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a piece of the yaw pulley cover detached from the instrument. The broken piece is approximately. 240 x. 200 in size. The broken piece had fractured around the distal pin and the yaw pulley cover has two broken sections still within the clevis. The belleville washer on the broken side has no load on one side and as a result of the breakage, is tilted within the clevis. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Initial | 2103972 | 2072780 | 1969756 | 2955842201100132 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2011-03-30 | 2011-04-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2011-04-29 | Device Not Retu | No |  | 2011
IS2000 A51P8 | 2011-03-30 | Injury | Required Intervention | It was reported that during a da vinci s prostatectomy procedure a very small amount of blood was observed on the patient's descending colon, around the sigmoid. Another surgeon was brought in to assess the issue and confirmed that the patient's bowel had been perforated. Since the damage was located on the underside of the patient's bowel, visibility was limited and the surgeon made the decision to convert to traditional open surgical techniques to repair the damaged bowel and complete the planned procedure. Postoperatively, 4 units of blood were given to the patient to help with their bowel recovery and bleeding from the deep dorsal vein complex. It was reported that the patient is in stable condition. | Additional information was provided by the console surgeon, who indicated that an instrument may have touched the patient's bowel during an instrument exchange, however, no issues were observed during the exchange. Based on the information provided, it was determined that the da vinci s surgical system worked as intended. No information has been received which suggests that the da vinci s surgical system malfunctioned in a way that may have caused or contributed to the patient's injury. Intuitive surgical concluded that the patient injury was unrelated to the performance of the da vinci s surgical system. The da vinci s surgical system user's manual specifically states: caution: ensure all installed instruments are visible in the surgeon console view before proceeding to prevent inadvertent harm to the patient. Warning: the instrument may not be immediately visible when being moved from the cannula into the patient. Use appropriate caution when manually inserting instruments into the patient. Caution: removal of instruments during a procedure should be done very carefully and only with the knowledge of and full view of the surgeon console operator. Prior to instrument removal, the surgeon console operator should do the following: ensure the instrument is free and clear of any patient anatomy. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2103976 | 2072781 | 1969757 | 2955842201100133 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-03-30 | 2011-04-29 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-30 | No | No | Reuse | 2011
| 2011-04-14 | Malfunction |  | Robotic large hemolock clip applier wire broke during surgical procedure. |  | INTUITIVE | LARGE HEMOLOCK CLIP APPLIER | 2109427 | 2075961 | 1972938 | MW5020447 | 1 | NAY | Voluntary |  | Nurse | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
400169 | 2011-04-26 | Malfunction |  | During a robotic laparoscopic supercervical hysterectomy, the harmonic ace curved shears insert's insulation piece broke off while being used. The piece was recovered without harm to patient. |  | HARMONIC CURVED SHEARS INSERT | HARMONIC CURVED SHEARS INSERT | 2110232 | 2079033 | 1976010 | 2079033 | 1 | NAY | User Facility |  |  | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2011
COR31 12X150MM | 2011-04-25 | Malfunction |  | During a robotic surgery, the physician moved the camera around inside the abdomen and noted tissue up against it, obstructing the view. So, the scrub tech went to move the port and advance it when they heard a noise and noted the port had cracked and a piece or two flew away from the patient. They had been losing gas; having problems maintaining pneumoperitoneum during the whole case and believe that there was a crack in the port the entire time, creating the loss of gas. |  | COR31 12X150MM KII OPTICAL ACCESS SYSTEM | TROCAR PORT | 2113698 | 2082589 | 1979565 | 2082589 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2011-04-28 | 2011-04-28 | No | Yes | Invalid Data | COR31 12X150MM |  | No Answer Provided |  |  |  |  | 2011
42017207 | 2011-03-29 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while using the maryland bipolar forceps instrument, the tip broke off and fell into the patient. The fragment was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. (b)(4). | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available or the instrument is returned for evaluation, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2098284 | 2067674 | 1964650 | 2955842201100118 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-29 | 2011-04-25 | No | Yes | Health Professional | 42017207 | No | Yes | 2011-03-29 | No | No | Unkown | 2011
42000606 | 2011-03-14 | Malfunction |  | It was reported that during a da vinci si left renal diverticulectomy, while the surgeon was transferring a needle and using the large needle driver instrument, the needle became dislodged from the instrument and fell into the patient. Approximately one week later, the patient returned to the hospital for an x-ray to locate the needle and operation to have it removed. No adverse event was reported. | The instrument was returned and evaluated. Engineer was unable to confirm the customer reported complaint as the instrument was found to be fully functional. The instrument was placed on a test system and moved intuitively with full range of motion in all directions. The grips opened and closed properly and sample sutures were successfully made. No damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2098285 | 2067675 | 1964651 | 2955842201100120 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-29 | 2011-04-25 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2011-03-29 | Device Not Retu | No | Reuse | 2011
42019408 | 2011-03-28 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the inner aspects of the mega needle driver instrument came off and fell into the patient. The broken pieces were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the carbide inserts on both grip tips detached from the metal surface. The metal surface is smooth and no residue of adhesive material was found. No damage was observed on the clevis and cable. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2099404 | 2069126 | 1966102 | 2955842201100125 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-28 | 2011-04-25 | No | Yes | Other | 42019408 | Device Returned To Manufacturer | No | 2011-03-28 | Device Not Retu | No | Reuse | 2011
| 2011-03-22 | Injury | HospitalizationOtherLife ThreateningRequired Intervention | A (b)(6) obese patient with high grade endometrial cancer underwent a robotic-assisted total laparoscopic hysterectomy, bso, bilateral pelvic lymphadenectomy using the davinci si surgical system. The robotic hot shears (monopolar scissors) were used in the right instrument port and bipolar forceps in the left instrument port. During mobilization and removal of the right obturator lymph nodes, arcing of energy through the protective sheath of the shaft of the hot shears was noted, causing direct thermal damage to the distal right external iliac vein. A laparotomy was performed and a vascular surgeon ligated the right external iliac vein secondary to the extent of damage caused, limited exposure and the location of injury. Postoperatively, the patient developed a right lower extremity deep venous thrombosis and was placed on anticoagulation. Dates of use: |  | HOT SHEARS | MONOPOLAR ROBOTIC SCISSORS | 2105292 | 2071831 | 1968807 | MW5020376 | 1 | NAY | Voluntary |  | Physician | Initial | 2011-04-23 | 2011-04-23 | Yes | Yes | Health Professional |  | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
400169 | 2011-03-03 | Malfunction |  | On (b)(4) 2011, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from california pacific medical center. The following problem description was provided on the report: during a total abdominal hysterectomy, the jaw of robotic harmonic scalpel instrument broke at joint. Loose piece was retrieved without incident. Complete match to non-damaged instrument when visually compared. Fracture of instrument monitored via robot camera. No foreign material remains in the patient per surgeon and assistant. Fracture occurred while surgeon using instrument to dissect tissue. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2096202 | 2065915 | 1962891 | 2955842201100112 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-03-25 | 2011-04-21 | No | Yes | Health Professional | 400169 | No | No | 2011-03-25 | No | Yes | Unkown | 2011
40018010 | 2011-03-22 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, while performing the right side pelvic lymph node dissection, and while using the monopolar curved scissors instrument with a tip cover accessory installed, the patient's external iliac vein was punctured. A vascular surgeon was called to assist and elected to convert to traditional open surgical techniques to repair the patient's iliac injury. The patient required a blood transfusion but was reported as stable and recovering well. | The tip cover accessory has not been returned for evaluation, however the monopolar curved scissors (mcs) instrument said to be used with the tip cover accessory during this procedure was returned. Per the customer reported complaint engineering observed the mcs instrument proximal clevis exhibits a char mark adjacent to the rectangular window that exposes the seal, approximately. 130 inches above the interface with the tube extension. The seal has a section of material removed at one corner, adjacent to the char mark. The char mark is located near the silicone to sleeve interface on a tip cover when installed on the mcs instrument. No other damage was found. The isi representative onsite during the procedure did not witness any arcing from the mcs tip cover. Additionally, the tip cover accessory was inspected after the procedure and noted to not contain any holes or tears. Tip cover arcing recommendations for prevention and precautions were provided to the account on (b)(6) 2011. As of the date of this report, no similar events have been reported by this site. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2096203 | 2065916 | 1958137 | 2955842201100113 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-22 | 2011-04-21 | Yes | Yes | Health Professional | 40018010 | No | No | 2011-03-22 | No | Yes | Unkown | 2011
IS3000 A60P6 | 2011-03-24 | Other | Other | It was reported that while beginning a da vinci si thoracic sympathectomy procedure, and using an ethicon disposable 12mm trocar in the endoscopic camera manipulator (ecm) port, the trocar melted at the tip. The site replaced the trocar, however the replacement trocar melted at the tip. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The isi representative onsite for this procedure stated that the surgical staff was holding the camera backed up into the trocar for an extended period of time. The da vinci si surgical system user manual, setting up the vision system specifically states: caution: to avoid possible thermal damage to the cannula, do not leave the endoscope tip inside a plastic endoscope cannula (trocar) for a prolonged period of time while the lamp is on. The patient was reported to have recovered well post operatively with no complications. The site has continued to perform procedures and as (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096204 | 2065917 | 1962893 | 2955842201100114 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2011-03-24 | No | No | Reuse | 2011
42024902 | 2011-02-17 | Malfunction |  | It was reported that near the end of a da vinci si hysterectomy procedure shards of the 8mm dual blade retractor instrument fell into the patient. The patient was irrigated and the fragments were retrieved. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed the distal end of the instrument main tube exhibits scratches and abrasions with material removed and a rough surface finish. The damaged area is located approximately 2 inches above the proximal clevis. The abrasions are axially aligned with the tube and approximately 1. 3 inches long and. 085 inches wide. No other damage was found. In (b)(6) 2011, an isi representative performed additional training for prevention and detection of damaged cannula accessories at this hospital. A dropped or mishandled cannula can result in a dent of the cylindrical portion of the cannula and can become out-of-round which may result in scraping of the instrument shaft when inserted through the cannula. Drops may occur during cleaning or rough handling. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | 8MM DUAL BLADE RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2096205 | 2065918 | 1962894 | 2955842201100115 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-23 | 2011-04-21 | No | Yes | Health Professional | 42024902 | Device Returned To Manufacturer | Yes | 2011-03-23 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-03-18 | Death | Death | It was reported that post a successful da vinci si hysterectomy procedure the patient suffered a perioperative left hemisphere cerebrovascular accident (cva) secondary to complete occlusion of left carotid artery (embolic). Postoperatively noted to be completely aphagic and right hemiplegic. A clot was identified via mra imaging from bifurcation of aorta. No h/o hypercoagulability. Fibromuscular hyperplasia (fmh) significant for mother suffering fatal cerebrovascular accident (cva). Reportedly, the patient's demise was unrelated to the da vinci s surgical system. | No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. As of (b)(4), 2011, no similar instances have been reported to (b)(4). | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096208 | 2065919 | 1962895 | 2955842201100116 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-03-25 | 2011-04-21 | Yes | No | Health Professional | IS3000 A60P5 | No | Yes | 2011-03-25 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-24 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the site did not have 3d vision from the surgeon side console (ssc). With the assistance of an isi technical support engineer the site re-performed cameral calibrations and 3d image was restored. The site later reported that the vision issue recurred and the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the reported vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096612 | 2066364 | 1963340 | 2955842201100122 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-03-24 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2008-09-09 | Other | Other | On (b)(6) 2011 intuitive surgical received a legal summons stating that during a da vinci s myomectomy procedure the patient slid cephalad on the operating table and multiple repositions of the robot, repositioning of the da vinci port and re-docking were necessary to complete the procedure. The summons states that the patient injuries are permanent but no further details or information was provided. | Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunction was reported to have occurred and the system did not cause or contribute to the patient's injury. Multiple attempts for additional information have been made, however as of the date of this report no response from the patient or her representative has been received. If additional information is received, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096615 | 2066367 | 1963343 | 2955842201100121 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-03-24 | 2011-04-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-24 | No | No | Reuse | 2011
42018405 | 2011-01-10 | Malfunction |  | It was reported that during the beginning of a da vinci s mitral valve procedure, the permanent cautery spatula broke while the surgeon was removing an excessive amount of paracardial fat in order to access the pericardium. Pieces of the broken instrument fell into the patient and were removed. The instrument was replaced and the procedure was completed using replacement instrumentation. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2096671 | 2066419 | 1963393 | 2955842201100124 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2011-01-26 | 2011-04-21 | No | Yes | Health Professional | 42018405 | No | No | 2011-03-25 | Device Not Retu |  | Reuse | 2011
IS2000 A51P8 | 2011-03-24 | Death | Death | It was reported that after a da vinci s mitral valve repair procedure, the patients blood pressure had dropped, the surgeon performed a mini thoracotomy to identify the sourse of the blood pressure loss. A leak in the mitral valve was found and a sternotomy was performed to repair of the valve; however, the patient expired before the valve could be repaired. The site has been contacted for additional information related to this event; however, no new information has been received as of the date of this report. | The investigation conducted by the field service engineer found the system to be made to specification with no issues found and the hospital has continued to use the system to perform surgery on patients. The site has been contacted for additional information related to this event. If additional information becomes available, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2096683 | 2066439 | 1963415 | 2955842201100123 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-25 | 2011-04-21 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-03-25 | No | No | Reuse | 2011
400183 | 2011-03-14 | Malfunction |  | Patient undergoing a robotic sparing whipple and cholecystectomy for a pancreatic mass. During the procedure a piece of the hook cautery was noted to be broken off and in the patient's abdominal cavity. Copious irrigation and suction were performed. It is unknown whether the piece was retained or removed. |  | PERMANENT CAUTERY HOOK | ENDOWRIST MONOPOLAR CAUTERY IN | 2102648 | 2071868 | 1968844 | 2071868 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-19 | 2011-04-19 | No | Yes | Physician | 400183 | No | No Answer Provided |  |  |  |  | 2011
40018010 | 2011-03-22 | Malfunction |  | It was reported that during a da vinci si procedure, sparking was observed at the tip cover accessory that was installed on the monopolar curved scissors (mcs) instrument. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicating arcing events occurred. The holes are approximately. 030 to. 020 in size and located. 425 above the silicone to sleeve interface. One hole exhibits a miniscule tear on one edge. Various small mechanical tears were noticed in the general vicinity of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2091963 | 2061161 | 1958137 | 2955842201100111 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-22 | 2011-04-18 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-03-22 | Device Not Retu | Yes | Initial | 2011
| 2011-04-12 | Malfunction | Other | A pt with a history of oropharyngeal squamous cell carcinoma was undergoing transoral resection of a base of tongue mass with davinci si robotic assistance. Several hours into the surgery, a small flame was visible at the spatula tip of the monopolar cautery of the robotic device. The instrument was withdrawn and the flame extinguished simultaneously by flushing the field with sterile saline. The insulation on the spatula cautery tip hub appeared damaged. The robotic arm and tip were replaced and surgery completed. There was no injury or respiratory compromise to the pt or harm to employees. |  | DAVINCI VERSION 2 5MM CAUTERY SPATULA TIP | MONOPOLAR TIP | 2096441 | 2065924 | 1962900 | MW5020319 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-04-18 | 2011-04-18 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
40017908 | 2011-03-18 | Malfunction |  | It was reported that after a successful da vinci hysterectomy procedure, the surgical staff found the monopolar curved scissors instrument tip was broken and missing a piece. The instrument fragment was found via x-ray and removed using traditional laparoscopic instrumentation. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the right scissor blade detached from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the crimp is exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2090732 | 2059822 | 1956798 | 2955842201100110 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2011-03-18 | 2011-04-15 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2011-03-18 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-15 | Injury | Other | It was reported that on (b)(6) 2011, the patient underwent a successful da vinci s fundectomy procedure, however, 3 hours post procedure, a pancreatic fistula was found. Drainage was performed and amylase levels were tested. On (b)(6) 2011, the doctor determined that the inserts used for drainage was not allowing sufficient drainage and as a result the surgeon cleaned the patient endoscopically and a drainage tube was placed. On (b)(6) 2011, the patient's respiratory status became worse and a ct scan showed an abscess inside the mediastinum and the chest cavity. It was found that the patient also experienced mediastinitis and empyema. Drainage was performed and the patient was put on a respirator. As of (b)(6) 2011, the site reported that the patient is recovering with no fever. The patient is said to be recovering well with all symptoms improving. | The surgeon commented that the system functioned properly and no anomalies were noted during the da vinci s surgical procedure performed on (b)(6) 2011. The surgeon advised that it could not determine as to how the fistula developed and suggested that the pancreas could have been unintentionally damaged during the procedure as the patient was reported as having visceral fat. He also suggested that damage to the pancrease could have been caused by a third party endoscopic forceps or ligasure instruments inserted through the assistant port. The surgeon advised that this type of surgical complication is a known risk of fundectomy procedures regardless of or surgical technique or tools used. Per the information provided, it was determined that the da vinci s surgical system, instrument and accessories worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2089300 | 2058205 | 1955181 | 2955842201100109 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-03-15 | 2011-04-14 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-15 | No | No | Reuse | 2011
420184 | 2011-03-29 | Malfunction |  | Davinci robotic instrument "permanent cautery spatula" used in thoracic surgery. Side piece broke off while spatula inside the chest cavity. Procedure went open for better exposure of surgical site, multiple irrigations, piece never found. Item on use # 3 out of manufacturer recommended 10 uses before disposal. Reason for use: cautery for nsclc. |  | DAVINCI | PERMANENT CAUTERY SPATULA | 2098576 | 2065679 | 1962655 | MW5020303 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-04-14 | 2011-04-14 | No | Yes | Health Professional | 420184 | Device Returned To Manufacturer | Yes |  |  |  |  | 2011
420207 | 2011-03-02 | Malfunction |  | Using the davinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. |  | TENACULUM FORCEPS | ENDOWRIST GRASPERS FOR DAVINCI | 2096618 | 2066370 | 1963346 | 2066370 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-09 | 2011-04-09 | No | Yes | Physician | 420207 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
40018010 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The mcs instrument was removed and the tip cover was replaced. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is approximately. 032 inches by. 018 inches in size and is located. 35 inches above the silicone to sleeve interface. The hole exhibits a tear on one edge. On (b)(4) 2011, the robotics coordinator stated that she was unaware of the electrical surgical unit (esu) setting used during the case as well as the pre-procedure tip cover and cannula inspection process. On (b)(4) 2011, an (b)(4) representative provided the surgeon and (b)(4) coordinator with tip cover damage or arcing recommendations for prevention. Since the recommendations for prevention were provided, as of (b)(4) 2011, no further arcing events have been reported by (b)(4) hospital. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076649 | 2045758 | 1942735 | 2955842201100103 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-03-11 | Device Not Retu | Yes | Initial | 2011
42018404 | 2011-03-11 | Malfunction |  | It was reported that during a da vinci s lobectomy procedure, while using the permanent cautery spatula instrument, the surgeon witnessed brown fragments in the patient. The fragments were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient injury, harm, or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed, the instrument distal clevis ear was completely broken at the base and missing. The broken piece was not returned with the instrument. Engineering concluded the instrument breakage was likely caused by overloading at the tip. Per the customer, all of the fragments were retrieved from the patient. The broken piece removed was compared to the instrument by the hospital and it was determined no other fragments were missing. | PERMANENT CAUTERY SPAULTA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2076650 | 2045759 | 1942736 | 2955842201100104 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | Yes | 2011-03-11 | Device Not Retu | No | Reuse | 2011
400180 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. On march 18, 2011 an isi representative provided the surgeon and the site's robotics coordinator with tip cover damage or arcing recommendations for prevention. Multiple attempts have been made to obtain additional information from the site regarding this event, however, as of the date of this report, no additional information has been provided. Since the recommendations for prevention were provided, as of april 7, 2011, no further arcing events have been reported by (b)(6) hospital. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076651 | 2045760 | 1942737 | 2955842201100105 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 400180 | No | No | 2011-03-11 | No | Yes | Unkown | 2011
400180 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed arcing coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. On march 18, 2011, an isi representative provided the surgeon and the site's robotics coordinator with tip cover damage or arcing recommendations for prevention. Multiple attempts have been made to obtain additional information from the site regarding this event, however, as of the date of this report, no additional information has been provided. Since the recommendations for prevention were provided, as of april 7, 2011, no further arcing events have been reported by (b)(6). | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2076652 | 2045761 | 1942738 | 2955842201100106 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-11 | 2011-04-08 | No | Yes | Health Professional | 400180 | No | No | 2011-03-11 | No | Yes | Unkown | 2011
IS3000 A60P5 | 2011-03-15 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced system error 297. With the assistance of an isi technical support engineer, the site attempted to troubleshoot by checking cable connections and recycling power; however, the system error code persisted. The patient was under anesthesia and the port incisions had been made when the surgeon decided to convert the planned procedure to traditional open techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) could not reproduce the system error; however, the camera cable and camera head were found to be loose. Based on the system error logs, it was determined that the system error experienced by the site may be associated with a defective doco unit. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco and tightening the loose camera cable. The doco was returned to isi for failure analysis and the customer reported complaint was not confirmed as the doco functioned properly with no problems found. As a result, it was determined that the system error code was associated with the loose camera cable found on site during the fse's visit. As of april 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2077285 | 2046159 | 1962895 | 2955842201100107 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-15 | 2011-04-08 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-03-15 | Device Not Retu | No | Reuse | 2011
IS2000 A51P7 | 2011-03-11 | Other |  | It was reported that during a da vinci s prostatectomy procedure, while performing urethrovesical anastamosis, the site experienced system error code 25004. With the assistance of an isi technical support engineer, the system was rebooted. Upon restart, system error code 25588 was experienced. An isi field service engineer (fse) arrived on site and attempted to troubleshoot the problem; however, the system issue persisted. At this time, the surgeon decided to convert to a laproscopic approach; however, converting to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering confirmed the occurrence of system error code 25004 as well as system error code 25588. Engineering concluded that system error codes 25004 and 25588 were associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The ecm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the ecm to have a defective motor/encoder assembly which was replaced to repair the ecm. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2077542 | 2046429 | 1943407 | 2955842201100108 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-11 | 2011-04-08 | Yes | Yes | Health Professional | IS2000 A51P7 | Device Returned To Manufacturer | Yes | 2011-03-11 | Device Not Retu | No | Reuse | 2011
42019408 | 2001-03-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure and after insertion of the mega needle driver instrument for use, a piece of metal wire was found in the patient's abdomen. The surgeon removed the piece and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering's inspection found no damage on clevis or cable. Performance testing was conducted and the instrument was found to be within specification with no issues found. The recognition and engagement testing passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Based on these results, several attempts have been made to obtain additional information from the initial reporter without success. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2075466 | 2044570 | 1941546 | 2955842201100102 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-08 | 2011-04-07 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2010-10-19 | Device Not Retu | No | Reuse | 2011
420207 | 2011-03-02 | Malfunction |  | Using the da vinci tenaculum to hold the fibroid/uterus, it was noticed that a piece of metal was coming loose and falling off the instrument. The instrument was carefully removed and any visible pieces of metal were removed from the surgical field. |  | TENACULUM FORCEPS | DA VINCI S SURGICAL SYSTEM END | 2096617 | 2066369 | 1963345 | 2066369 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-04-05 | 2011-04-07 | No | Yes | Physician | 420207 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
| 2011-04-05 | Malfunction |  | Cannula seals were leaking - we switched them both out and then they functioned fine. The seals came in the 3 arm drape pack (2nd time in a week similar occurrence, similar devices, similar leaking). Female with nonfunctioning kidney with chronic ureteropelvic junction obstruction and severe hydronephrosis. The nonfunctioning aspect of her kidney was confirmed with nephrostomy tube drainiage, which showed a creatinine clearance of 0 and very little urine output in spite of being very patent. Patient sent to or, was prepped and draped. Surgeon unable to obtain insuflation with a veress needle and never felt confident to actually be in the peritoneal cavity. Patient had a very extremely thick abdominal wall, secondary to obesity. Surgeon ended up doing a direct vision placement of a trocar with the longest - a 12 mm direct visiport - that we have. He was barely able to get this to go into the peritoneum. After insufflation, able to re-adjust it and pull it back enough that we would be able to be docked. Two 8mm robotic ports were placed in the usual fashion. We had some trouble with the superior port because of it pulling back and difficulty with visulazation until we were finally able to get this to be replaced. A 12 mm inferior trocar near the midline was inserted for assistant port. Surgery continued and finalized. Trocars removed sequentially under direct vision, closed and patient recovered from anesthetic and returned to recovery room in good condition. ======================manufacturer response for da vinci surgical system single use seal, intuitive surgical======================field representative is aware, has been in contact with or, however, any reports not yet available at time of this report. |  | 8MM CANNULA SEAL | DA VINCI SURGICAL SYSTEM 8MM A | 2092727 | 2061812 | 1958788 | 2061812 | 1 | NAY | User Facility |  |  | Initial | 2011-04-06 | 2011-04-06 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
420225 | 2011-03-07 | Other | Other | It was reported that during a da vinci s oropharyngectomy procedure, fibers from the 5fr instrument introducer fell into the patient's mouth cavity. The surgeon was able to remove some of the fibers from the patient, however, decided to convert the procedure to traditional open surgical techniques to ensure all of the fragments were retrieved. The procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Multiple attempts to gather additional information have been made; however, as of the date of this report, the site has not provided any additional details. If additional information is received, a follow-up mdr will be submitted. | INSTRUMENT INTRODUCER 5FR | ENDOSCOPIC INSTRUMENT | 2069432 | 2037983 | 1934959 | 2955842201100097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-03-07 | 2011-04-04 | Yes | Yes | Health Professional | 420225 | No | No | 2011-03-07 | No | No | Unkown | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have epidermis burns at the cannula port location. The patient did not required any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded; however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069436 | 2037985 | 1934961 | 2955842201100098 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have epidermis burns at the cannula port location. The patient did not require any additional medical treatment. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded, however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069457 | 2037986 | 1934962 | 2955842201100100 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 |  | Other | Other | It was reported that after a da vinci s hysterectomy procedure, during a follow up visit, the patient was found to have burn marks at the cannula port location. Due to scarring, the patient required a second unspecified procedure. No additional patient harm, adverse outcome or injury was reported. | An isi representative onsite discovered that the hospital was using an electrical surgical unit (esu) that is not a validated model for use with the da vinci s surgical system. The da vinci s user manual specifically states: general precautions and warnings: warning: do not use any esu other than those validated for use with the da vinci s system. Previous investigations of similar events have shown that it is possible for the electrosurgical unit currents to pass through the patient side manipulator arms to ground through the cannula accessory and to the patient if the patient is not properly grounded; however, the root cause of the customer reported event cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of april 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2069441 | 2038009 | 1934985 | 2955842201100099 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
40016906 | 2011-03-11 | Malfunction |  | It was reported that during a da vinci s sigmoid colectomy procedure, the surgical staff observed that a piece of the harmonic curved shears insert accessory was missing when it was inserted into another port. The missing piece was located in the patient and successfully retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the insert is missing the clamp arm. One side of the hinge feature that mates with the clamp arm pins is severely bent backwards. The damage appears to be consistent with hyperextension of the clamp arm, leading engineering to conclude that the breakage was most likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2067807 | 2036545 | 1933521 | 2955842201100089 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-11 | 2011-04-01 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | No | 2011-03-11 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-03-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, there was no video displayed on the left side monitor of the surgeon console. The surgical staff power cycled the system, reseated the vision cable, cycled the breaker and emergency powered off the surgeon console, however, these actions were unable to resolve the issue. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. The hrsv was returned to intuitive surgical for evaluation. Engineering discovered that the left side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. As of march 31, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2067816 | 2036547 | 1933523 | 2955842201100091 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-07 | 2011-04-01 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-03-07 | Device Not Retu | No | Reuse | 2011
42018304 | 2011-02-18 | Malfunction |  | It was reported that during a da vinci s thoracic sympathectomy procedure, a piece of the permanent cautery hook instrument fell into the patient's chest. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor cap was missing from the distal end and the yaw pulley hub had a broken off boss feature (the boss feature mates with the hole in the conductor cap and keeps the cap from slipping). Engineering concluded that the instrument hook may have been overloaded, causing the boss feature to break and allowing the conductor cap to fall out. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2067792 | 2036558 | 1933534 | 2955842201100090 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-03-17 | 2011-04-01 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2011-03-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P6 | 2011-03-08 | Other | Other | It was reported that during the beginning of a da vinci si low anterior resection procedure, the surgical staff experienced double vision. The system was restarted twice and multiple endoscopes were tried, however, the issue could not be resolved. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | As part of intuitive surgical's failure analysis investigation, the system was evaluated at the site by a field service engineer (fse). System testing was performed using the customer's 30 degree endoscope and 3d calibration tool. The fse was unable to replicate the customer reported malfunction and the root cause of the customer reported failure mode cannot be determined. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2067821 | 2036560 | 1933536 | 2955842201100092 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-08 | 2011-04-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2011-03-08 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-03-03 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure, the surgeon gripped tissue and cauterization automatically occurred, causing a superficial burn to the patient's sigmoid colon. A general surgeon was called in to review the injury and placed a couple stitches as a precaution. The electrical surgical unit (esu) and cable were disconnected and replaced as well as the instrument. The system was restarted and the planned surgical procedure was completed. No further patient harm or adverse events were reported. | The investigation conducted by the field service engineer (fse) discovered that the electrical surgical unit (esu) used during the procedure was set on the automatic setting, therefore, as soon as the surgeon gripped tissue with the instrument, the esu cauterized without the foot pedal being pressed. The fse notified the surgical staff of the incorrect esu setting used. The staff reported that the surgeon had a procedure directly following this event and the case was completed with the same equipment and there were no issues. The patient was reported as recovering fine with no post operative complications. As of april 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2068013 | 2036620 | 1933596 | 2955842201100093 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-03 | 2011-04-01 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-03-03 | Device Not Retu | No | Reuse | 2011
400169 | 2011-02-04 | Malfunction |  | On (b)(6), 2011, user facility medwatch (b)(4) was received from the fda reporting the following event. Event desc: during the procedure, the distal portion of a robotic surgical instrument - harmonic curved shears - used for davinci surgery detached from the instrument for unknown reason. A thorough search of the patient's uterus was performed. One piece of the shears were found but the other piece of the shears was not found - even on x-ray. | Multiple attempts with the site have been made; however, at the time of this report, the instrument has not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2068523 | 2037070 | 1895144 | 2955842201100095 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2011-03-04 | 2011-04-01 | No | Yes | Health Professional | 400169 | Yes | No | 2011-03-04 | No | Yes | Initial | 2011
40018010 | 2011-02-25 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The monopolar curved scissors instrument accessory has as not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2068527 | 2037074 | 1934050 | 2955842201100094 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-03-02 | 2011-04-01 | No | Yes | Health Professional | 40018010 | No | Yes | 2011-03-02 | Device Not Retu | Yes | Unkown | 2011
IS2000 | 2011-01-27 | Other | Hospitalization | It was reported that the da vinci s surgical system was used on (b)(6) 2011, to perform a prostatectomy procedure. The surgery was successful and there were no system malfunctions during the case. The patient was discharged on (b)(6) 2011, however, returned to the hospital on the same day due to hematuria and difficulty in urinating. A clot was detected and removed and electrical coagulation was performed transurethrally (not using the da vinci s surgical system). The patient recovered and was discharged from the hospital on (b)(6) 2011. | Information was provided by the surgeon, who suspects that the crust of the anastomosis site between the urethra and vesica came off, resulting in the blood in the patient's urine. The surgeon indicated that the patient's post surgical symptoms were unrelated to the da vinci s surgical system and are part of the complications related to prostatectomy procedures, regardless of the modality used to perform the surgery. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no system malfunctions occurred, and the system did not cause or contribute to the patient's post surgical complications. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2068515 | 2037093 | 1934069 | 2955842201100096 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2011-03-03 | 2011-04-01 | Yes | No | Health Professional | IS2000 | No | No | 2011-03-03 | No | No | Reuse | 2011
400180 | 2011-02-24 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The monopolar curved instrument (mcs) was returned with the ultem sleeve from the mcs tip cover accessory installed; however, but the silicone part had been removed. The distal end components of the instrument were visually inspected and no char marks or other signs of arcing were observed. Electrical continuity passed and the instrument had fully been used with no lives remaining. No physical damage found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2066674 | 2035530 | 1932505 | 2955842201100088 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-03-01 | 2011-03-31 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | No | 2011-03-01 | Device Not Retu | Yes | Initial | 2011
| 2011-03-24 | Malfunction |  | During robotic prostatectomy, the team had problems with laproscopic c02 pressure being low. Attempted a number of things to fix the problem; started by changing out the disposable trocars, but c02 still leaked. It appeared that the leak was coming from the seals on the divinci trocars. All seals were switched out and it continued to leak. Then, vaseline gauze was wrapped around the instruments to block the leak a little bit, it helped for the time being. ======================manufacturer response for trocar seals for davinci, da vinci surgical system======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for ethicon trocar bladeless endopath xcel, ethicon trocar bladeless endopath xcel w/stability sleeve======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for 12mm bladeless endopath xcel trocar w/stability sleeve, ethicon endo-surgery endopath excel bladeless trocar======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. ======================manufacturer response for ethicon trocar bladeless endopath xcel, ethicon trocar bladeless endopath xcel w/stability sleeve======================the representative is waiting for risk mgmt to complete the report - the device(s) will be released to representative. |  | DA VINCI SURGICAL SYSTEM | TROCAR SEALS FOR DAVINCI | 2074107 | 2043364 | 1940340 | 2043364 | 1 | NAY | User Facility |  |  | Initial | 2011-03-31 | 2011-03-31 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2011
420172 | 2011-03-29 | Malfunction |  | During surgery, a piece of the maryland bipolar endowrist broke off inside patient. Broken piece of device was retrieved by doctor (md). Entire device was inspected and then taken out of operating room (or). |  | MARYLAND BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 2092720 | 2061778 | 1958754 | 2061778 | 1 | NAY | User Facility |  |  | Initial | 2011-03-31 | 2011-03-31 | No | Yes | Invalid Data | 420172 | Yes | No Answer Provided |  |  |  |  | 2011
420227 | 2011-02-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the clamps were not closing. At this time, the surgeon observed a plastic part on the pk dissecting instrument was broken. When the instrument was removed and inspected, a piece of plastic was found to be missing. The surgeon reported that there have been no reports of patient harm, adverse outcome or injury as a result of the piece falling into the patient. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2065574 | 2034496 | 1931471 | 2955842201100086 | 1 | LLZ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-02-28 | 2011-03-30 | No | Yes | Health Professional | 420227 | No | Yes | 2011-02-28 | No | No | Unkown | 2011
42000306 | 2011-02-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the small clip applier instrument's tip became detached and fell into the patient. The piece was retrieved and accounted for from the patient. The procedure was successfully completed with a replacement instrument of the same type. The site reported that there was no patient injury or adverse event experienced. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to be returned with a section of one grip detached from the distal end. The size of piece is roughly. 150 x. 050. The grip fractured near the base of the clip groove and the grip has a slight bend below the fracture surface. Additional observation not reported by site is a bent grip. The intact grip is bent at roughly a 30 degree angle. Location of bending is near the base of the clip groove, similar to where the other grip fractured. The location of fracture/bending of the grips matches closely with the depth the grips get installed into the clip holding fixture. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage found. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2065619 | 2034510 | 1931486 | 2955842201100087 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-28 | 2011-03-30 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2011-02-28 | Device Not Retu | No | Initial | 2011
| 2011-03-18 | Malfunction |  | Jaws of scissors weren't working; the jaws did not open or close and they never worked. |  | ENDOWRIST | MONOPOLAR CAUTERY INSTRUMENTS | 2068407 | 2036945 | 1933921 | 2036945 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2011-03-25 | 2011-03-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-02-22 | Other | Other | It was reported that during a da vinci s nephrectomy procedure, the site experienced difficulty with the system recognizing the instrument. The site attempted to reseat the patient side manipulator (psm) arm; however, the instrument still was not recognized by the system. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the clinical sales representative (csr) found that the instrument being used was out of lives. The site was made aware of this. As of march 22, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2059715 | 2028765 | 1925743 | 2955842201100085 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2011-02-22 | 2011-03-24 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2011-02-22 | No | No | Reuse | 2011
IS2000 A51P8 | 2011-02-21 | Other | Other | It was reported that during a da vinci s cystectomy procedure, the surgical staff experienced multiple system error code 212. The site was able to override the error code on several occasions; however, the error returned. Prior to calling into isi for technical assistance, the site undocked and rebooted the system. The system powered up and homed with no issues. The site then called to advise an isi technical support engineer (tse) and to gain additional information regarding the system error. At this time, the system error code 212 returned. The site was able to override the fault; however, it returned shortly after clearing. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator and multiple servo driver (msd). The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system contains two multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected mtmr and the msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The msd and mtmr were returned to isi for failure analysis investigation. Engineering evaluation found that the msd was within manufacturing specifications with no issues noted. Engineering also evaluated the returned mtmr and found that the main harness and an axis motor had malfunctioned thus generating the system error code experienced by the customer. As of march 22, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2058602 | 2028112 | 1925088 | 2955842201100084 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-02-21 | 2011-03-23 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-02-21 | Device Not Retu | No | Reuse | 2011
420174 | 2011-03-17 | Malfunction |  | Using robotic harmonic scalpel during robotically assisted thymectomy. Robotic harmonic scalpel includes reusable hand piece, reusable robotic sleeve, and disposable harmonic blade insert. After the surgeon began to use the insert, one of the jaws broke off. |  | HARMONIC CURVED SHEARS | ENDOWRIST ULTRASONIC ENERGY IN | 2068263 | 2036803 | 1933779 | 2036803 | 1 | NAY | User Facility |  |  | Initial | 2011-03-21 | 2011-03-23 | No | Yes | Physician | 420174 |  | No Answer Provided |  |  |  |  | 2011
420174 | 2011-03-17 | Malfunction |  | Using robotic harmonic scalpel during robotically assisted thymectomy. Robotic harmonic scalpel includes reusable hand piece, reusable robotic sleeve, and disposable harmonic blade insert. After the surgeon began to use the insert, one of the jaws broke off. |  | HARMONIC CURVED SHEARS | ENDOWRIST ULTRASONIC ENERGY IN | 2068263 | 2036803 | 1933779 | 2036803 | 1 | NAY | User Facility |  |  | Initial | 2011-03-21 | 2011-03-23 | No | Yes | Physician | 420174 |  | No Answer Provided |  |  |  |  | 2011
400193 | 2011-03-03 | Malfunction |  | At the end of the davinci robotic procedure, the balloon port for the procedure was removed and after inspection by the scrub rn to confirm integrity of the balloon port, the sheath was noted to have a circumferential crack. ======================manufacturer response for disposable, balloon port======================no comments. |  | BALLOON PORT | ENDOWRIST ACCESSORIES FOR DAVI | 2068238 | 2036726 | 1933702 | 2036726 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-03-21 | 2011-03-21 | No | Yes | Invalid Data | 400193 | Yes | No Answer Provided |  |  |  |  | 2011
IS3000 A60P5 |  | Death | Death | It was reported that eight days after a da vinci si tongue base resection (tors) procedure was successfully performed, the patient expired due to uncontrollable post operative bleeding. | On (b)(6) 2011, additional information regarding the event was provided by dr. (b)(6), console surgeon who performed the tors procedure. Dr. (b)(6) indicated that the patient underwent surgery for a benign mass at the base of the tongue, which had been causing obstructive symptoms for a year and there was no other way to approach the lesion. Due to the location of the mass, dr. (b)(6) was close to a branch of the lingual artery. Dr. (b)(6) cauterized the artery with bipolar energy and put tisseel over it. The case went well, and dr. (b)(6) didn't have to go too deep and had good visualization. Eight days postop, the patient began to bleed. Dr. (b)(6) mentioned that typically a bleed within 24-48 hours postop is considered a technical problem and after 48 hours, it is due to a wound healing issue which expose the vessels and make them more likely to bleed. Dr. (b)(6) indicated that this complication is a known risk with a transoral approach and was not at all related to any limitations of or problems with the da vinci si surgical system. Based on the information provided by dr. (b)(6), it was determined that the patient's demise was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and did not malfunction in a way that would have led or contributed to the patient's death. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2053945 | 2023904 | 1920894 | 2955842201100082 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-17 | 2011-03-18 | Yes | No | Health Professional | IS3000 A60P5 | No | No | 2011-02-17 | No | No | Reuse | 2011
IS3000 | 2011-02-11 | Other | Other | It was reported that after a da vinci si hysterectomy procedure was completed, the patient developed a 103 degree fever and their heart rate jumped to 130. The surgeon performed a ct scan and a blood clot was observed. No other information was received. | Multiple attempts have been made to obtain additional information from the surgeon, however, the surgeon has been unavailable. To date, no information has been received which suggests that the da vinci si surgical system may have caused or contributed to the patient's post surgical complications. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2054373 | 2024247 | 1921257 | 2955842201100083 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-16 | 2011-03-18 | Yes | No | Health Professional | IS3000 | No | No | 2011-02-16 | No | No | Reuse | 2011
400184 | 2011-02-15 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed that the amber plastic part of the permanent cautery spatula instrument was broken at the tip after it was inserted into the patient. The break was not observed and no pieces from the instrument were found in the patient. It is unclear if the peace broke inside of the patient; however, prior to use, the instrument was inspected when anti-char solution was applied. No issues with the instrument were noted at this time. As of the date of this report, no patient harm, adverse outcome or injury has been reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2052948 | 2022563 | 1919537 | 2955842201100081 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2011-02-15 | 2011-03-17 | No | Yes | Health Professional | 400184 | No | No | 2011-02-15 | No | No | Unkown | 2011
IS3000 A60P3 | 2010-02-15 | Other | Other | It was reported that during set up for a da vinci si hysterectomy procedure, the site experienced a non recoverable system error fault 22003. With the assistance of an isi technical support engineer, the site was able to apply various troubleshooting techniques to override the system error to continue with the procedure. The site called back reporting a re-occurrence of system error 22003 and elected to go forward using the system while disabling one of the three patient side manipulators (psm) arm. When the site docked to the patient, the site could not follow on matching grip with the psm 3. At this point, the surgical staff elected to convert the planned procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes 322003 found in the error logs were associated with a setup joint (suj) arm. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from its own joint positions to the rac. Errors in this communication at system start-up would induce the system error codes in this case. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The system was repaired by replacing the affected suj. The suj was returned to isi for failure analysis investigation. Engineering confirmed the customer reported complaint as they were able to duplicate the system error codes. As a result, the suj was repaired by replacing several harness cables. As of march 15, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2052978 | 2022592 | 1919582 | 2955842201100079 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2010-02-15 | 2011-03-17 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | Yes | 2009-02-15 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-01-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) with a mcs tip cover accessory installed. A small thermal injury was observed on the patient's bowel and was repaired via stitch. The planned surgical procedure was completed with a replacement mcs instrument tip cover accessory. No patient recovered well with no post operative complications or prolonged hospitalization. The delay of this report is due to the site's initial reporter going out on medical leave prior to submitting the report as initially intended. Upon her return, several weeks later, the report was filed. An isi representative was onsite during the procedure; however, believed the initial reporter would initiate the complaint record with isi customer service. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2052979 | 2022593 | 1919583 | 2955842201100080 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-18 | 2011-03-17 | No | Yes | Health Professional | 40018010 | No | No | 2011-01-18 | No | No | Unkown | 2011
42017909 | 2011-02-03 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure a piece of the monopolar curved scissors instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. The delay in reporting is the result of an administrative error by a complaint handling team member. This error has been addressed with the applicable team member. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the right blade was broken off and missing from the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2053232 | 2022859 | 1919849 | 2955842201100071 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2011-02-03 | 2011-03-17 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2011-03-21 | Device Not Retu | No |  | 2011
40027202 | 2011-02-14 | Malfunction |  | It was reported that during a da vinci s abdominoperineal resection procedure the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that a section of the curved blade had broken off. A piece of the curved blade, measuring approximately. 385 x. 080, was missing and not returned with the insert. The fracture surface of the blade is not level and outer surface of blade has a couple of scratch marks below fracture. It was determined that this damage may have contributed to the blade breakage since the broken piece has scratch marks in the axial direction, consistent with this motion. Although intuitive surgical uses the term blade above, there are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. The da vinci harmonic ace curved shears instructions for use specifically states: inspection prior to use examine the device prior to use. In particular, examine the following components for cracks or flaws: the clamp arm and blade on the insert, the shaft of the insert, the reusable instrument housing caution: if cracks or flaws are observed, do not use the device. Note: scratches on the blade may lead to premature blade failure. Avoid accidental contact with other instruments during use. Do not use any means other than the blade wrench to attach or detach the instrument from the hand piece. General precautions and warnings: do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2051671 | 2021232 | 1918228 | 2955842201100078 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-14 | 2011-03-16 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-02-14 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P6 | 2011-02-18 | Other | Other | It was reported that when the surgical staff was getting ready to dock the da vinci si surgical system to the patient for a myomectomy procedure, system error code 25521 occurred. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 25521 was associated with an endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart and provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25521 appears when the system detects that it has lost communication of sensor information between two printed circuit assemblies in the arm control system. In the event of these communication issues, the system records the fault and transitions to a safe state. The ecm was returned for failure analysis investigation. Engineering was able to replicate the reported failure and observed that both flat flex cables in and out of a small board embedded in the insertion axis of the ecm were not well connected, thus generating the system error code experienced by the customer. As of march 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2048128 | 2018315 | 1915311 | 2955842201100076 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-18 | 2011-03-14 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2011-02-18 | Device Not Retu | No | Reuse | 2011
42017207 | 2011-02-18 | Malfunction |  | Event Description: --/-- It was reported that while cauterizing in the chest wall area during a da vinci si lobectomy procedure, the smell of burning plastic was detected by the surgical staff and arcing was observed from the jaws of the maryland bipolar forceps instrument to the endowrist. An intuitive surgical representative in attendance during the surgical procedure noticed that a loose gia staple was caught in the jaws of the instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- (b)(4) - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on both yaw pulleys with material removed at the grip base. No char marks were observed on the sides of the yaw pulleys or the metal wrist components. It was determined that the damage may have been caused by a breach in the insulation or carbonized tissue which created a conductive path. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 2048155 | 2018338 | 1915333 | 2955842201100072 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2011-02-18 | 2011-03-14 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2011-04-11 | Device Not Retu | No |  | 2011
40027202 | 2011-02-14 | Malfunction |  | It was reported that during a da vinci s abdominoperineal resection procedure, the tip of the harmonic ace curved shears insert broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that a section of the curved blade had broken off. A piece of the curved blade, measuring approximately. 645 x. 085, was returned with the insert. The fracture surface of the blade is not leveled and one of the mating features that slides in and out of the shaft when the clamp arm opens/closes is bent slightly inwards and rubs against one side of the blade. It was determined that this damage may have contributed to the blade breakage since the broken piece has wearing/scratching in the axial direction, consistent with this motion. Although intuitive surgical uses the term blade above, there are no components in the harmonic ace curved shears instruments or inserts that meet the definition of medical sharps. The da vinci harmonic ace curved shears instructions for use specifically states: inspection prior to use, examine the device prior to use. In particular, examine the following components for cracks or flaws: the clamp arm and blade on the insert - the shaft of the insert - the reusable instrument housing. Caution: if cracks or flaws are observed, do not use the device. Note: scratches on the blade may lead to premature blade failure. Avoid accidental contact with other instruments during use. Do not use any means other than the blade wrench to attach or detach the instrument from the hand piece. General precautions and warnings - do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2048173 | 2018343 | 1915338 | 2955842201100073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-14 | 2011-03-14 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2011-02-14 | Device Not Retu | Yes | Unkown | 2011
IS3000 A60P5 | 2011-02-18 | Other | Other | It was reported that during a da vinci si vesicolithotomy procedure, the surgical staff experienced non-recoverable system error code 31030. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The system logs were reviewed during the investigation conducted by the field service engineer (fse) and multiple occurrences of system error code 23 were observed along with system error code 31030. The fse concluded that system error codes 23 & 31030 were associated with a patient side manipulator (psm) setup-joint axis. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The setup-joint (suj) is a non-motorized support arm which holds a patient side manipulator (psm) during surgery. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the psm. System error code 31030 appears when the da vinci si safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining this condition, the system records the fault and transitions to a safe state. The suj was returned for failure analysis evaluation. Engineering was able to replicate the reported failure and observed damages on the cable harness, thus generating the system error codes experienced by the customer. As of march 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2048182 | 2018348 | 1915343 | 2955842201100077 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-18 | 2011-03-14 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-02-18 | Device Not Retu | No | Reuse | 2011
400077 | 2011-03-09 | Malfunction |  | Patient undergoing lysis of adhesions and total laparoscopic hysterectomy with robotic davinci assist. The davinci has 4 arms; 3 of which use 8mm cannula seals on them (the 4th is for a camera and the seal for this arm is built into the introducer). At beginning of procedure it was noted 3 of the 8mm cannula seals utilized on the arms of the davinci machine would not seal properly, staff could hear air coming through the seals. All 3 seals were from the same lot number and were removed from the machine when complete seal could not be obtained. Staff obtained new seals from a completely different lot number and encountered no further issues. No injury to patient, just time delay in troubleshooting issues. |  | 8 MM CANNULA SEAL | DAVINCI ENDOWRIST 8MM ACCESSOR | 2064380 | 2033149 | 1930125 | 2033149 | 1 | NAY | User Facility |  |  | Initial | 2011-03-14 | 2011-03-14 | No | Yes | Physician | 400077 | Yes | No Answer Provided |  |  |  |  | 2011
42017908 | 2011-02-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found a piece of the left scissor blade to be broken off from the instrument, the broken piece was returned. The blade fractured at the cross section of the grip cable crimp resulting in half of the cable crimp to be exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2046357 | 2016298 | 1913294 | 2955842201100075 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-05 | 2011-03-11 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2011-02-10 | No | No | Unkown | 2011
IS2000 A51P8 | 2011-02-08 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgical staff experienced the left eye from the surgeon side console to be pink. With the assistance of an isi technical support engineer, the site verified their settings and it was then determined that the camera cable needed to be replaced. The patient was under anesthesia when the surgeon staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of march 8, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2043254 | 2013313 | 1910309 | 2955842201100066 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-08 | 2011-03-10 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-02-08 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-02-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while using the monopolar curved scissors instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. | The tip cover accessory has not been returned for evaluation, however, the monopolar curved scissors instrument used with the tip cover during the procedure was returned and evaluated. Per the customer reported complaint engineering found no char marks or signs of arcing on the instrument. Electrical continuity passed and no damage was found. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2043255 | 2013314 | 1910310 | 2955842201100067 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-09 | 2011-03-10 | No | Yes | Health Professional | 40018010 | No | No | 2011-02-09 | No | Yes | Unkown | 2011
400180 | 2011-02-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. No patient injury, harm or adverse outcome was reported. It was noted that the electrical surgical unit (esu) setting was above the isi recommended range. | The tip cover accessory has not been returned for evaluation. Multiple attempts for additional information have been made, however as of the time of this report, no additional information has been provided by the site. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2043256 | 2013315 | 1910311 | 2955842201100069 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2011-02-17 | 2011-03-10 | No | Yes | Health Professional | 400180 | No | No | 2011-02-17 | No | Yes | Unkown | 2011
IS3000 A60P5 | 2011-02-11 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure the site was unable to get the system to power on. The site reconnected all and changed all power outlet connections, however the system continued to not gain power. With the assistance of an isi technical support engineer the site reset the main system breaker, however, the system continued to not power on. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm or injury was reported. | The investigation conducted by the field service engineer concluded that system power issue was associated the core power distribution (ipd) board. The ipd is a pca (printed circuit assembly) board that converts the incoming electrical voltage to the appropriate levels to power the other circuit boards and regulates the current supplied to them. The system was repaired by replacing the affected ipd pca board. The ipd pca board was returned to isi for failure analysis investigation. Engineering was able to replicate the customer reported complaint and found a defective module on the ipd pca board. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2043368 | 2013395 | 1910391 | 2955842201100070 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-11 | 2011-03-10 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2011-02-11 | Device Not Retu | No | Reuse | 2011
| 2011-02-22 | Malfunction |  | Patient was undergoing robotic lh bilateral salpingo-oophorectomy and mccall's monarch anterior repair. The plastic part of the spatula broke off during the procedure. This has occurred previously with the same operator, according to operating room nurse operator has been re-trained by the company rep on previous occasions. Diagnosis or reason for use: surgery for hysterectomy and bladder suspension. |  | INTUITIVE SURGICAL | DAVINCI CAUTERY SPATULA | 2051848 | 2018713 | 1915708 | MW5019768 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2011-03-09 | 2011-03-09 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2011
IS3000 60P6 | 2011-02-07 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci si prostatectomy procedure, the surgical staff was unable to calibrate the camera and the system diagnostics generated a message indicating that a possible cause was a dirty scope. The site power cycled the system, however, this did not resolve the issue. The patient had been under anesthesia for 20 minutes when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the customer reported issue was associated with the camera cable instead of the endoscope. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The camera cable was returned and evaluated. Engineering was able to replicate the reported failure mode and concluded that the issue experienced by the customer was a result of conductor damage. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of march 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2041079 | 2011052 | 1908048 | 2955842201100068 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-07 | 2011-03-08 | Yes | Yes | Health Professional | IS3000 60P6 | Device Returned To Manufacturer | Yes | 2011-02-07 | Device Not Retu | No | Reuse | 2011
| 2011-01-31 | Malfunction |  | Md was suturing mesh to the vaginal wall using the endowrist (a mega suturecut needle driver used with the davinci robot). The circulator noticed that a wire was hanging from the endowrist. Surgery department opened up another endowrist instrument and within 5 minutes the same thing happened to this instrument. The circulator describes the wire as a braided wire cable and that the cable was dangling. Another instrument was opened which functioned appropriately and the md was able to finish the procedure. However, there was the potential for a fragment to fall into the patient. ======================manufacturer response for suturing device for davinci robot, endowrist mega suture cut needle driver======================unknown. |  | MEGA SUTURECUT NEEDLE DRIVER | ENDOWRIST INSTRUMENT FOR DAVIN | 2054420 | 2024310 | 1921300 | 2024310 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-03-07 | 2011-03-07 | No | Yes | Invalid Data |  | No | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2011-02-03 | Other |  | It was reported that during a da vinci prostatectomy procedure, the surgical staff was unable to focus the camera. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue by changing the camera cable; however, the vision issue remained. The site could hear the motor running and attempted to work with the focus assembly without success. The patient was under anesthesia approximately 1 hour and the port incisions were made when the surgeon made the decision to abort the planned surgical procedure. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera head was returned to the original equipment manufacturer (oem) for evaluation. The oem discovered that a mechanical shock to the camera head damaged the stage assembly, thus resulting in the vision issue experienced by the customer. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2038249 | 2008222 | 1905218 | 2955842201100064 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-03 | 2011-03-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-02-03 | Device Not Retu | No | Reuse | 2011
IS1200 A43P9 | 2011-02-03 | Other | Other | It was reported that the beginning of a da vinci prostatectomy procedure, the surgical staff experienced system error 153. With the assistance of an isi technical support engineer, the site attempted to troubleshoot the issue by checking cables and restarting the system after undocking from the patient. Once powered back on, a system error message indicated to check the green cable. The site powered down the system, checked the green cable cable and restarted. The system successfully started up and the procedure continued. Two hours later, an isi field service engineer called in stating that system error code 153 reoccurred and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that system error code 153 was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. The ria pca performs numerous tasks related to the function, control and hardware fault monitoring for a specific arm. There is a ria pca assigned to each system arm. The system was repaired by replacing the affected ria pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 153 is reported by hardware to denote a high side switch fault. This indicates that there was a short on the board that triggered this. In the event of these power supply issues, the system records the fault and transitions to a safe state. As of march 3, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2038250 | 2008223 | 1905219 | 2955842201100065 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-02-03 | 2011-03-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2011-02-03 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-03 | Injury | Life Threatening | It was reported that during a da vinci s rectopexy procedure the surgeon experienced visualization issues due the patient's extremely large uterus. Due to the poor vision, the surgeon inadvertently hit a vessel over the sacral promontory and uncontrollable bleeding occurred. As a result, the surgeon converted to traditional open surgical techniques to maintain the bleeding and complete the planned procedure. The delay in reporting is a result of the complaint handling team mistakenly believing that the complaint information had already been recorded in the complaint database. | (b)(4). The initial reporter confirmed that the system and associated instruments/accessories did not malfunction during the surgical procedure and did not cause or contribute to the patient injury. The initial reporter also indicated that the patient is recovering well and no post surgical complications were reported. Based on the information provided, it was determined that the da vinci s surgical system worked as intended, no device failure occurred, and the patient injury was unrelated to the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2033191 | 2003418 | 1900414 | 2955842201100061 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-06 | 2011-03-01 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-12-06 | No | No | Reuse | 2011
IS3000 A60P5 | 2011-02-01 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced a loss of video in the right eye of the surgeon side console. The site reseated and then replaced the camera cable, however, the vision loss continued to occur. With the assistance of a technical support engineer, the site restarted the system, however, the vision loss was still present. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) and review of the system error logs determined that the vision issue experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The doco was returned to isi for failure analysis and the customer reported complaint was confirmed as the doco did not function properly. As of march 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2033522 | 2003912 | 1900908 | 2955842201100062 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-02-01 | 2011-03-01 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | Yes | 2011-02-01 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-02-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the end piece of the permanent cautery spatula (pcs) instrument broke off and fell into the patient. The broken piece was retrieved and the instrument was replaced with another pcs instrument. The end piece of the second pcs instrument also broke off during use and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a third pcs instrument. No patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00059 was submitted for the initial pcs instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base. The ear had broken off on the opposite side of the conductor wire and cap. The broken piece was not returned. The conductor cap is also missing. Engineering also observed a broken proximal clevis at the main tube interface. As a result, the clevis is dislodged from the main tube. The proximal clevis and main tube are missing pieces. A. 185 x. 175 piece of the proximal clevis was returned with the instrument, however, the clevis has a remaining broken section that was not returned. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2032792 | 2002988 | 1900008 | 2955842201100060 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-02 | 2011-02-28 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-02-02 | Device Not Retu | No | Reuse | 2011
42018404 | 2011-02-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the end piece of the permanent cautery spatula (pcs) instrument broke off and fell into the patient. The broken piece was retrieved and the instrument was replaced with another pcs instrument. The end piece of the second pcs instrument also broke off during use and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a third pcs instrument. No patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2011-00060 was submitted for the second pcs instrument used during the procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base. The size of the broken piece is approximately. 285 x. 235. The ear had broken off on the opposite side of the conductor wire and cap. The broken ear piece was not returned, however, the conductor cap was returned taped to the housing. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2032789 | 2003012 | 1900008 | 2955842201100059 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-02-02 | 2011-02-28 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-02-02 | Device Not Retu | No | Reuse | 2011
42018405 | 2011-01-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a brown piece from the distal end of the permanent cautery spatula had fallen into the patient. The piece was successfully retrieved and the procedure completed with no patient injury. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument returned with a piece of the distal clevis broken. One distal clevis ear is broken off at its base. The broken piece was not returned; however, the missing piece was roughly. 300 x. 235 and was located on the opposite side of the conductor wire and cap. Engineering has concluded that the instrument breakage may be due to overloading the tip. No other damage found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2031948 | 2001957 | 1898953 | 2955842201100057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-26 | 2011-02-27 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2010-11-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P4 | 2011-01-26 | Other | Other | It was reported that during a da vinci si mitral valve replacement procedure, the site experienced a black image in the left eye of the surgeon side console. With the assistance of an isi technical support engineer, the surgical staff reseated the camera cable; however, the vision issue was not resolved. At this time, the surgical staff made the decision to convert the planned procedure to traditional open surgical techniques to successfully complete the procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of february 24, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2031964 | 2001958 | 1898954 | 2955842201100058 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-26 | 2011-02-27 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2011-01-26 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-21 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure, the surgical staff noticed flickering and wavy lines in the image viewed through the surgeon side console (ssc). The site attempted to power down and restart the system, however, they were unable to get the system to power back on. With the assistance of an isi technical support engineer, the site reseated the power cord, reset the ssc power breaker and moved the power cord to an isolated circuit, however, the ssc would not regain power. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering confirmed that the issue experienced by the customer was related to the surgeon side console (ssc) primary power supply (pps). The pps is a +115v power supply with battery backup and contains logic and protection circuits to provide monitoring of under voltage, over voltage, over current, ac line status, over temperature, power supply fan faults, and detection of battery operation. The system was repaired by replacing the pps. As of february 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2026492 | 1996704 | 1893700 | 2955842201100053 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-21 | 2011-02-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-21 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-27 | Other | Other | It was reported that during a da vinci s surgical procedure, the site was unable to focus the vision settings from the surgeon side console (ssc) and the vision cart. With the assistance of an isi technical support engineer, the site replaced the camera cable and restarted the system. The surgeon was then able to focus the image as expected and continued with the case. The site later reported that the vision issues recurred and the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the reported vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of february 17, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2026494 | 1996706 | 1893702 | 2955842201100054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-27 | 2011-02-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-27 | Device Not Retu | No | Reuse | 2011
400006 | 2011-02-16 | Malfunction |  | Davinci needle driver would not close properly. The instrument had 8 out of 10 uses left. The instrument was replaced in the field and sent down to decontam decontamination to be sent to risk management and then sent back for credit from intuitive surgical. |  | LARGE NEEDLE DRIVER FOR DAVINCI | ENDOWRIST NEEDLE DRIVER | 2033833 | 2004223 | 1901219 | 2004223 | 1 | NAY | User Facility |  |  | Initial | 2011-02-18 | 2011-02-18 | No | Yes | Physician | 400006 |  | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-01-17 | Other | Required Intervention | It was reported that during a da vinci s lobectomy procedure, an injury to a branch of the patient's main pulmonary artery occurred during its dissection, requiring that the surgeon convert to traditional open surgical techniques to complete the planned procedure. | On (b)(6) 2011, the surgeon stated via email, the injury was not caused by a robotic instrument or arm malfunction. The patient recovered completely from the operation and has been home for several weeks. Per the information provided by the surgeon, intuitive surgical determined that the patient injury was unrelated to the performance of the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2023628 | 1993942 | 1890938 | 2955842201100050 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2011-01-17 | 2011-02-16 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2011-01-17 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2011-01-18 | Other | Required Intervention | It was reported that during a da vinci si surgical procedure, the instrument became stuck on the system arm and the surgical staff was unable to get the instrument to release the tissue that it was grasping using the emergency release wrench. The surgeon cut out the portion of the tissue that was being grasped in order to remove the instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported issue; however, a link port switch and the cannula mount were replaced as a precaution. As of february 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2023632 | 1993945 | 1890941 | 2955842201100051 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-18 | 2011-02-16 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2011-01-18 | Device Not Retu | No | Unkown | 2011
42017910 | 2010-12-21 | Malfunction |  | It was reported that during a da vinci si ovarian cystectomy procedure, sparking was observed from a source other than the tips of the monopolar curved scissors instrument(mcs) with a tip cover accessory installed. The mcs instrument and tip cover accessory were removed and replaced with an instrument of the same type. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of sparking as the instrument passed electrical continuity testing and did not exhibit any signs of charring or arcing. The instrument was placed on an in house test system and passed recognition and engagement tests as well as a cut test. The instrument moved intuitively with full range of motion in all directions. No damage was found. Multiple attempts have been made to retrieve additional information from the hospital regarding this event, however, as of the date of this report, no response has been received. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2014825 | 1985162 | 1882158 | 2955842201100045 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-13 | 2011-02-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2011-01-13 | Device Not Retu | No | Reuse | 2011
42018305 | 2011-01-11 | Injury | Other | It was reported that approximately 35 to 50 minutes into a da vinci s sacrocolpopexy procedure, while the surgeon was exploring the surgical field, the permanent cautery hook instrument broke and bumped the patient's bifurcation of the inferior vena cava. The procedure was converted to traditional open surgical techniques, sutures were placed to repair the iliac perforation, and the patient was provided units of blood as a result of the bleeding. On (b)(6) 2011, medwatch uf. Importer report (b)(4) was received. | On (b)(6) 2011, dr. (b)(6) stated that the procedure would have likely converted to open due to the patient's size and extent of adhesions. He stated that the patient was recovering fine and there were no additional complications. On (b)(6) 2011, dr. (b)(6) contacted isi regarding the possibility that instrument fragments may have fallen into the patient during the procedure, and requested biocompatibility information for the permanent cautery hook instrument. An x-ray and ct scan were performed and were found to be inconclusive in confirming if any fragments were retained in the patient. On (b)(6) 2011, (b)(6) hospital risk management representative, (b)(6), confirmed the patient has recovered well and has not had any post surgical complications. The instrument will not be returned for evaluation, however, the hospital did provide photographic images of the instrument for isi evaluation. Per the images provided, the instrument has a broken distal clevis ear with the ear having broken off at the base. The conductor cap is missing due to the ear breakage. A follow-up mdr will be submitted if the instrument is returned and / or if more information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2014844 | 1985183 | 1882179 | 2955842201100044 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-01-12 | 2011-02-11 | Yes | Yes | Health Professional | 42018305 | No | Yes | 2011-01-12 | No | No | Unkown | 2011
IS2000 A51P8 |  | No Answer Provided | Other | The patient posted a comment on an isi website, called patient stories, where she reported the following story. After a da vinci hysterectomy procedure, the patient was reported having a long recovery and post operative uncomfort. Over a week post op, the patient underwent a pelvic exam which resulted in the discover of a tear and an infection. A second surgery was required to repair the tear; however, one month after the second surgery, the patient was experiencing pain during intercourse which resulted in an emergency room visit. | The site's risk management department has been contacted to obtain additional details related to this event. As of the date of this report, additional information has not been provided to support the reported event. Isi will continue to make attempts to gain additional information and will provide the fda with a follow up report once the additional information becomes available. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015557 | 1985870 | 1882866 | 2955842201100046 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-13 | 2011-02-11 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-01-13 | No | No | Reuse | 2011
IS1200 A43P9 | 2011-01-12 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the surgeon experienced double vision with the image displayed on the surgeon console. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the vision issue experienced by the customer was due to incorrectly connected cables at the rear of the surgeon console. The issue was resolved by swapping the cables to their correct location. As of february 10, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015939 | 1986227 | 1883223 | 2955842201100047 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2011-01-12 | 2011-02-11 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2011-01-12 | Device Not Retu | No | Reuse | 2011
40018010 | 2011-12-07 | Malfunction |  | It was reported that during a da vinci si nephroureterectomy procedure, arcing was observed at the side of the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The electrosurgical unit was not activated when the arcing was observed. The tip cover was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and five holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicating multiple arcing events occurred. The hole sizes range from under. 010 to. 025 with varying shapes (round and oblong) and are located. 290 to. 415 above the silicone to sleeve interface. The various mechanical tears are in the general vicinity of the holes and there is an axially oriented tear at the distal tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2015955 | 1986243 | 1883239 | 2955842201100048 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-12 | 2011-02-11 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-01-12 | Device Not Retu | Yes | Initial | 2011
IS2000 A51P8 | 2011-01-17 | Other | Other | It was reported that during a da vinci s procedure, the surgical staff experienced issues with the patient side cart intermittently switching to battery power. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that the power issue experienced by the customer was due a patient side cart power cord malfunction. The system was repaired by replacing the affected power cord. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2015956 | 1986244 | 1883240 | 2955842201100049 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-17 | 2011-02-11 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2011-01-17 | Device Not Retu | No | Reuse | 2011
400169 | 2011-02-04 | Malfunction |  | During the procedure, the distal portion of a robotic surgical instrument - harmonic curved shears - used for davinci surgery detached from the instrument for unknown reason. A thorough search of the patient's uterus was performed. One piece of the shears were found, but the other piece of the shears was not found - even on x-ray. |  | HARMONIC CURVED SHEARS INSERT | ULTRASONIC ENERGY INSTRUMENT | 2027847 | 1998148 | 1895144 | 1998148 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-02-10 | 2011-02-10 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2011
IS2000 A51P8 | 2011-01-10 | Other | Other | It was reported that prior to a da vinci s prostatectomy procedure, the site experienced system error code 23013 and 23008 when attempting to attach the sterile adapter to the patient side manipulator (psm). An isi technical support engineer (tse) attempted to assist the site to override the system error; however, the site made the decision to abort the planned procedure. The patient was under anesthesia at the time the decision was made not to continue with the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer (fse) concluded that system error code 23013 and 23008 were associated patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23013 and 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The psm was returned to isi for evaluation. Engineering determined that the potentiometer had malfunctioned on the psm, thus generating the system error codes experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2012932 | 1982927 | 1879922 | 2955842201100043 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2011-01-10 | 2011-02-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2011-01-10 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-10 | Other | Required InterventionHospitalization | On (b)(6) 2011, intuitive surgical received a legal summons and complaint filed by a pt alleging that they suffered significant injury and damages due to the surgeon's inability to provide reasonably prudent medical care after a da vinci s excision of endometriosis including a partial thickness sigmoid colon resection procedure was performed. The surgery took approx 10 hours to complete and post operatively, the pt had ct findings consistent with a perforated viscus and developing abscess. The complaint further alleges that the pt has undergone numerous medical procedures (including add'l surgeries) due to the surgeon's failure to properly, adequately or timely monitor, manage, diagnose, refer, consult, inform and treat the pt's medical condition. | No allegations of deficiency with the da vinci s surgical system were reported. Based on the info received, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2039096 | 2001786 | 1898782 | 2955842201100042 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-01-10 | 2011-02-09 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2011-01-10 | Device Not Retu | No | Reuse | 2011
42011704 | 2010-12-15 | Malfunction |  | It was reported that while suturing during a da vinci si single port hysterectomy procedure the jaw of the 5 mm needle driver instrument broke and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. The rptr indicated that there was likely extra stress on the instrument due to the close proximity of the instrument arms while performing the single port procedure. No pt harm, adverse outcome or injury was reported. The delay in reporting is due to an isi rep who was in attendance during the surgical procedure when the malfunction occurred but did not report the event within the time period required. The isi rep has been reminded of isi's reporting timeframe requirements. | The broken instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A f/u medwatch report will be submitted if the instrument is returned or if add'l info is rec'd. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 2043733 | 2003155 | 1900151 | 2955842201100038 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-15 | 2011-02-09 | No | Yes | Health Professional | 42011704 | No | No | 2010-12-15 | Device Not Retu | No | Unkown | 2011
42017909 | 2010-10-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned and / or if more information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2011486 | 1981964 | 1878960 | 2955842201100041 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-11 | 2011-02-08 | No | Yes | Health Professional | 42017909 | No | Yes | 2011-01-11 | No | No | Unkown | 2011
42018304 | 2010-12-10 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff observed small fragments in the patient's pelvis area. The end of the permanent cautery hook instrument was noted to be misaligned with pieces of the sheath end missing. The fragments were removed and the procedure was completed with a replacement instrument of the same type. No patient harm was reported. The delay in reporting is due to an administrative oversight where the complaint handling representative misfiled the report for reporting the week after it was due. | The instrument was returned and evaluated. Confirming the customer's reported complaint, engineering found the instrument's wrist cables and back end cables to be tight and intact. Additional observations found the instrument's main tube to be broken at the junction of the proximal clevis and main tube. The main tube feature which mates with the proximal clevis is broken off and missing. Based on these findings, engineering has concluded that the damage is likely due to misuse and/or mishandling. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2009414 | 1980338 | 1877334 | 2955842201100039 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-30 | 2011-02-07 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2010-12-30 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 |  | Death | Death | It was reported that 10 days after a da vinci si transoral resection procedure was successfully performed, the patient experienced bleeding in the surgical area and expired. The date of the surgical procedure and date of death were not received. Pre surgery, the patient's cancer was thought to be at stage t2, however, during the procedure, it was discovered that the cancer was at stage t3. There were no system malfunctions during the procedure and the surgeon indicated that the patient's demise was unrelated to the da vinci si surgical system. The surgeon believes that the post surgical bleeding may have been related to a delay in healing and that a scab may have come off. The delay in reporting is due to a misunderstanding about when intuitive surgical first became aware of the event. | Per the information provided by the surgeon, it was determined that the patient's demise was unrelated to the da vinci si surgical system. It was concluded that the da vinci si surgical system worked as intended and no system malfunctions occurred. On (b)(6) 2011, attempts were made to obtain additional information from the surgeon, however, no response has been received to date. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2010063 | 1980770 | 1877766 | 2955842201100040 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-12-13 | 2011-02-07 | Yes | No | Health Professional | IS3000 A60P5 | No | Yes | 2010-12-13 | No | No | Reuse | 2011
40018010 | 2011-01-05 | Malfunction |  | Event Description: --/-- It was reported that while cauterizing the surgical stump during a da vinci si hysterectomy procedure, the surgeon noticed that the fully mobilized/detached uterus was being blanched by the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed one hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating that an arcing event occurred. The hole size is approximately. 020 in diameter and the hole center is located. 215 above the silicone to sleeve interface. --/-- Additional Manufacturer Narrative: --/-- During evaluation of the device, it was also observed that a section of the silicone was not properly adhering at the shoulder of the ultem. This additional information is being provided per a retrospective review that was performed. | 8MM MCS TIP COVER ACCESSORY | Initial | 2008014 | 1979008 | 1876004 | 2955842201100037 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2011-01-05 | 2011-02-04 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-01-27 | Device Not Retu | Yes |  | 2011
IS3000 A60P4 | 2011-01-06 | Other | Other | It was reported that approximately 45 minutes during a da vinci si pulmonary lobectomy procedure, an instrument could not be recognized on a patient side manipulator (psm) arm. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer (fse) found the issue experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported issue, however, found the psm wrist input assembly to be below specification. All four wrist assemblies and the cannula mount were replaced. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2008105 | 1979115 | 1876111 | 2955842201100036 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2011-01-06 | 2011-02-04 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2011
| 2011-01-27 | Malfunction |  | Patient underwent robotic internal mammary takedown. Da vinci balloon port was placed in the patient's right anterior chest wall to facilitate port to introduce robotic instrumentation for internal mammary artery (ima) takedown. After successful completion of harvest, port was removed. At this point, request was made to the anesthesiologist to inflate the lung. The bedside md assistant placed his finger in the incision site where the port was to feel the lung inflation and he noted a hard object with the chest wall. He retreived the object and the tip of the da vinci balloon port (approx. 1/2 inch) was retreived. When the balloon port was inspected by the surgical team, it was then noted that a circumfrential 1/2 inch length of the port broke off. Both edges were examined and a "clean break" was noted with no rough edges. Procedure completed in normal fashion. |  | 12 MM BALLOON PORT | DAVINCI BALLOON CAMERA PORT | 2022076 | 1992036 | 1889032 | 1992036 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-02-04 | 2011-02-04 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2011-01-03 | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure, the surgeon experienced difficulty with the grips opening and closing for the instruments controlled by the left master tool manipulator (mtm). The patient had been under anesthesia for approximately one hour and 45 minutes when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtml gimbal. The mtml gimbal was returned to isi for failure analysis investigation. Engineering evaluation of the mtml gimbal found the gimbal grip levers had excessive play, thus confirming the reported complaint experienced by the customer. The gimbal grip levers were replaced. As of (b)(4), 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGCAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005852 | 1977151 | 1874147 | 2955842201100033 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-03 | 2011-02-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2011-01-03 | Device Not Retu | No | Reuse | 2011
42017909 | 2011-12-03 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure and while performing a dissection, the tip of the monopolar curved scissors instrument broke and fell into the patient. The piece was retrieved and the procedure was successfully completed using a replacement instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a piece of the left blade had broken off and is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2005876 | 1977175 | 1874171 | 2955842201100034 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2011-01-03 | 2011-02-02 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-03 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2011-01-04 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a non recoverable system error code 23. With the assistance of an isi technical support engineer (tse), the site powered down the system; however, the system error persisted. The site then powered down the system again, cycled the breakers, and applied emergency power off (epo) before powering on the system. The system was working for approximately 10 minutes when the system error re-occurred. The patient had been under anesthesia for approximately 1. 5 hours and docked to the system for 45 minutes when the surgeon made the decision to abort the planned procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes 23 experienced by the customer was associated with an endoscopic camera manipulator (ecm). Endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The ecm was returned to isi for failure analysis investigation. Engineering could not duplicate the reported system error code; however, as a precaution the ecm harness and embedded serializer patient manipulator (espm) board were replaced. As of february 1, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005877 | 1977176 | 1874172 | 2955842201100035 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2011-01-04 | 2011-02-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2011-01-04 | Device Not Retu | No | Reuse | 2011
40016907 | 2011-01-07 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the v-care uterine manipulator had been placed incorrectly and perforated the patient's fibroid. While resecting the broad ligament fibroid, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument was unintentionally closed over the v-care uterine manipulator. Energy was then applied and the top jaw of the hcs insert broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the insert is missing the clamp arm. One side of the hinge feature that mates with the clamp arm pins is severely bent backwards. The damage appears to be consistent with hyperextension of the clamp arm and engineering concluded that the breakage was most likely caused by mishandling/misuse. The broken tip was returned with the device. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2005913 | 1977212 | 1874208 | 2955842201100031 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-07 | 2011-02-02 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2011-01-07 | Device Not Retu | Yes | Initial | 2011
IS1200 A43P9 | 2011-01-06 | Other | Other | It was reported that prior to using the system to perform any tasks for a da vinci surgical procedure, the surgical staff experienced text/words on the left side image displayed in the surgeon console. The surgical staff power cycled the system and checked the settings on the synchronizer, however, they were unable to resolve the issue. The patient had been under anesthesia for 45 minutes when the surgeon made the decision to abort the planned procedure and reschedule it for a later date. No patient harm was reported. | Investigation was conducted by a field service engineer (fse). Prior to evaluating the system, the fse was informed that something had spilt on the bottom area of the vision cart (vc), and the vc had been cleaned before the scheduled case. During evaluation of the system, the fse discovered that the left camera control unit (ccu) options menu had been opened and the camera id had been activated and set to a random series of letters and numbers. The fse concluded that these buttons were most likely pressed during cleaning, thus causing the vision issues experienced by the customer. The fse repaired the system by resetting the ccu settings. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2005915 | 1977214 | 1874210 | 2955842201100032 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2011-01-06 | 2011-02-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2011-01-06 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-03 | Malfunction |  | On (b)(4) 2011, medwatch uf / importer report (b)(4) was received. It was reported by the hospital that during a da vinci s radical vaginectomy procedure, when the surgeon activated the monopolar cautery foot pedal, the bipolar instrument became activated as well as the monopolar instrument. It was noted by the surgical staff that the monopolar and bipolar cautery cables from the electrical surgical unit (esu) to the da vinci instruments were intertwined. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | On site investigation conducted by the field service engineer (fse) was unable to replicate the customer reported monopolar and bipolar cautery activation issues. Inspection found the bracket on the back of the monopolar connection to be broken, resulting in a loose connection to the electrical surgical unit (esu) and that the cautery cables from the esu to the monopolar and bipolar instruments were intertwined during the procedure. Engineering review indicates that these conditions are unlikely to have led to the cautery activation issue reported by the surgeon. The monopolar instrument drives current into the patient body which returns to the electrical ground through the path of least resistance, usually the grounding pad connected to the patient. It is possible for the bi-polar instrument to be the path of least resistance, due to close tip proximity or insufficient contact with the grounding pad, causing current to flow through the bipolar tips and make them appear to be activated. It is the view of isi engineering that this is the most likely cause of the reported issue. This is an issue inherent in the use of energy instruments and not a flaw in the da vinci surgical system. The surgeon referred to the wires between the esu and the instruments as having been bundled. Since these are non-shielded cables carrying high current, it is conceivable that current in one cable could have induced current in the other. Since this is part of the esu setup, it is outside of intuitive's expertise and would be best addressed by the esu manufacturer. For precautionary purposes, the fse replaced the ssc footrest cautery cable. The footrest cautery cable connects the ssc footswitch panel to the esu unit, causing cauterization to occur when the surgeon activates the applicable mono or bipolar foot pedal. As of (b)(6) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1999713 | 1971475 | 1868471 | 2955842201100030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2011-01-06 | 2011-01-27 | No | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2011-01-06 | Device Not Retu | No | Reuse | 2011
42017909 | 2010-11-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed that the monopolar curved scissors (mcs) instrument was cracked just above the endowrist. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. The customer reported malfunction does not in itself constitute a fda reportable event, however, engineering discovered evidence of arcing during internal evaluation of the instrument. | The instrument was returned and evaluated on december 29, 2010. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube and the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. It was determined that the damage was most likely due to excessive side loading or other type of mishandling. Engineering also observed charring on the tube extension. The tube extension has a section that exhibits charring/thermal damage directly below the breakage. The instrument was disassembled to find a light char mark present on the inner surface of the tube reinforcement ring. The tip cover accessory returned with the instrument also exhibits a char mark on the proximal end of the plastic sleeve, likely caused by energy escaping from the tube extension breakage. On january 11, 2011, the initial reporter provided additional information and indicated that arcing from the mcs instrument was observed during the surgical procedure. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1994426 | 1966554 | 1863550 | 2955842201100028 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-29 | 2011-01-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-12-29 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-11-17 | Other | Required Intervention | It was reported that on (b)(6) 2010, the patient underwent a successful da vinci s prostatectomy procedure, however, on (b)(6) 2010, the patient became febrile and began experiencing low abdominal tenderness. On (b)(6) 2010, the patient had a fever and blood in his stool. On (b)(6) 2010, a ct scan revealed that the patient had possibly developed generalized peritonitis. A re-operation was performed and a sigmoid colon perforation was discovered. Open drainage was performed and an artificial anus was created. Based on the laparotomy findings, the symptom seemed to be delayed fistulation. The patient recovered and was discharged on (b)(6) 2010. | The surgeon commented that the system functioned properly and no anomalies were noted during the da vinci s surgical procedure performed on (b)(6) 2010. Although the system was used in the area where the perforation was discovered, the surgeon indicated that no perforations were observed during the surgical procedure. Per the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1994490 | 1966750 | 1863746 | 2955842201100029 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-11-24 | 2011-01-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-12-24 | No | No | Reuse | 2011
42000606 | 2010-12-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the large needle driver instrument broke while the surgeon was suturing and an instrument piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned with one carbide insert detached from its grip. Per the customer reported complaint, engineering evaluation observed that the grip brazing surface has various spots where the filler material is removed and the insert brazing surface has matching spots with the filler material added. A couple of scratch marks were found on the outer surface of the grip. No other damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1993556 | 1965939 | 1862935 | 2955842201100027 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-21 | 2011-01-20 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2010-12-21 | Device Not Retu | No | Reuse | 2011
400309 | 2011-01-11 | Malfunction |  | Mega suturecut needle driver inserted into trocar, appeared ok. When surgeon went to use it, one of the threads broke making the instrument unusable. |  | MEGA SUTURECUT NEEDLE DRIVER | DAVINCI ENDOWRIST NEEDLE DRIVE | 2003856 | 1975612 | 1872608 | 1975612 | 1 | NAY | User Facility |  |  | Initial | 2011-01-20 | 2011-01-20 | No | Yes | Invalid Data | 400309 |  | No Answer Provided |  |  |  |  | 2011
IS1200 A43P9 | 2010-12-17 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced a non-recoverable system error. The site power cycled the system, however, the error continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found system error codes 20003 and 20013 in the system error log. These errors are associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error code 20013 is generated when an error is detected as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer). Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The system error code 20003 indicated that another fault occurred during the system start up. As of january 18, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1990416 | 1963102 | 1860098 | 2955842201100026 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-12-19 | 2011-01-18 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-12-19 | Device Not Retu | No | Reuse | 2011
40017208 | 2010-12-08 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure a fragment broke off of the maryland bipolar forceps instrument and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the main tube and distal tip detached from the housing. Per the customer reported complaint, engineering evaluation observed that the distal end drive cables and conductor wires have been cut at the interface between the metal proximal clevis and the main tube. The distal end of the main tube is broken on one side. The evidence is not conclusive, however, a likely failure mode is a dislodged proximal clevis due to excessive overloading at the tip, which resulted in the inability to remove the instrument from the cannula and required the cables/wires to be cut. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1988928 | 1961770 | 1858766 | 2955842201100018 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-15 | 2011-01-14 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2010-12-15 | Device Not Retu | No | Reuse | 2011
42018405 | 2010-12-22 | Malfunction |  | It was reported that during a da vinci si right lobectomy procedure, the permanent cautery spatula (pcs) instrument made contact with another endowrist instrument, causing a yellow piece on the pcs instrument to break off and fall into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that yellow colored components (proximal clevis and distal clevis) at the distal end are intact and undamaged, however, the brown ceramic sleeve on the instrument has a piece broken off at the distal end. The size of the missing piece is approximately. 090 x. 075. The ceramic sleeve also exhibits a couple of scratch marks below the breakage. While the evidence is inconclusive, engineering has concluded that damage to the instrument is most likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1988963 | 1961785 | 1858781 | 2955842201100020 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-22 | 2011-01-14 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2010-12-22 | Device Not Retu | No | Initial | 2011
42011704 | 2010-12-06 | Malfunction |  | It was reported that during a da vinci si ventral hernia repair procedure and while suturing mesh to repair a ventral hernia, the tip of the 5mm needle driver instrument broke and fell into the patient. The surgeon proceeded with the procedure using a replacement 5mm needle driver instrument. The piece was not retrieved from the patient; however, no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to have one of the grip tips broken at the grip base. The broken surface was flat and no sign of damage was found on grip cables. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1988988 | 1961795 | 1858791 | 2955842201100023 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-17 | 2011-01-14 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-12-17 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-15 | Death | Death | It was reported that while removing tissue during a da vinci s transoral resection procedure performed on (b)(6) 2010, the surgeon encountered an unusual group of veins and patient bleeding occurred. The surgeon was able to stop the bleeding and the planned surgical procedure was completed. There were no system malfunctions during the procedure. Five days post-op, the patient experienced bleeding and expired on (b)(6) 2010. | Per the surgeon, the patient's demise was unrelated to the da vinci s surgical system and the bleeding experienced by the patient would have occurred regardless of the modality used to perform the procedure. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1988993 | 1961834 | 1858830 | 2955842201100022 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-12-16 | 2011-01-14 | Yes | No | Service Personnel | IS2000 A51P8 | No | Yes | 2010-12-16 | No | No | Reuse | 2011
IS3000 A60P5 | 2010-12-17 | Other | Other | It was reported that approximately 2 hours during a da vinci si rectopexy procedure, while an instrument change was in progress, the surgeon removed his hands from the master tool manipulators (mtm); however; left his head in the surgeon side stereo viewer. This caused the instrument being exchanged in the guided tool change mode, to move slightly forward and bump the patient's colon. Slight bleeding of the patient's colon occurred; however, no repair, blood transfusion or intervention was necessary. The planned procedure was completed and no adverse outcome was reported. | The investigation conducted by field service engineering concluded that mtm movement experienced by the customer was associated with use error. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The guided tool change is an efficient and safe method for instrument insertion, providing assistance to the patient side cart operator by guiding the instrument into the patient. The system was found to be within specification, however, gravity calibration was performed on the mtms as a precaution. On (b)(6) 2010, additional training was conducted with the surgeon to ensure that in future cases, he will remove his head from the surgeon side viewer when his hands are not engaged on the mtms. The da vinci si user manual specifically states: note: for guided tool change, the surgeon should remove their head from the stereo viewer to take all instruments out of following before removing and installing an instrument. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1989021 | 1961862 | 1858858 | 2955842201100019 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2010-12-17 | 2011-01-14 | Yes | No | Health Professional | IS3000 A60P5 | Yes | No | 2010-12-17 | Device Not Retu | No | Reuse | 2011
42017206 | 2010-12-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the maryland bipolar instrument tip became detached from the shaft and could not be removed through the cannula accessory. With the assistance of a technical support engineer (tse), the site disengaged the sterile adapter from the patient side manipulator (psm) arm and removed the instrument with the cannula attached. A larger port incision was made to the patient to allow the cannula and instrument to be removed. After the maryland bipolar instrument was removed, the surgical staff found that wires had fallen into the patient. The affected area was irrigated and suctioned to remove the fragments. The site replaced the psm drape, sterile adapter and instrument and completed the planned procedure. No adverse outcome was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On (b)(4) 2010, a field service engineer inspected the da vinci s surgical system and found the system to have no issue. On (b)(4) 2010, (b)(4) stated that the patient underwent an x-ray and no foreign material was detected. If the instrument is returned or if further information is received, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1989042 | 1961869 | 1858865 | 2955842201100024 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-15 | 2011-01-14 | Yes | Yes | Health Professional | 42017206 | No | No | 2010-12-15 | No | No | Unkown | 2011
40018010 | 2010-12-20 | Other | Other | Event Description: --/-- It was reported that approximately 30 minutes into a da vinci s prostatectomy procedure, the surgeon noticed a hole had developed in the tip cover accessory installed on the monopolar curved scissors instrument which was believed to have caused a burn to the patient's right external iliac vein. The surgeon repaired the iliac vein burn with 5-0 prolene. The monopolar curved scissors instrument was removed and the tip cover accessory was replaced. The planned procedure was complete and no adverse outcome was reported. | Manufacturer Narrative: --/-- The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint engineering found various mechanical tears and a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicating of arcing. The hole is approximately. 030 in by. 020 and is located. 425 inches above the silicone to sleeve interface. The tip cover also has various small mechanical tears in the general vicinity of the holes. A large tear, approximately. 195 inches long, is located 90 degrees from the hole, on one edge. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip cover accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument and/or tip cover accessory are returned (post-evaluation) and/or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | Initial | 1989068 | 1961907 | 1858903 | 2955842201100025 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2011-03-28 | 2011-01-14 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2011-03-28 | Device Not Retu | Yes |  | 2011
8MM CAUTERY HOOK | 2011-01-11 | Injury | Required Intervention | The pt is a (b)(6) female admitted to the medical center for surgery to repair a vaginal prolapse. The surgeon stated that he was performing a robotic sacrocolpopexy. During the procedure the j-hook da vinci instrument malfunctioned. The j-hook is a surgical instrument that is attached to the da vinci robot used to cauterize tissue (permanent cautery hook). One side of the yoke that secures the j-hook fractured; this caused the j-hook to move involuntarily perforating the bifurcation of the inferior vena cava. The surgery at that point had to be converted to an open procedure. Sutures were used to repair the two small perforations in the vein. This complication occurred at a point when the surgeon had paused during the surgery. His hands were not on the pedal controls for the j-hook. The pt did require units of blood as a result of the bleeding. |  | PERMANENT CAUTERY HOOK | PERMANENT CAUTERY HOOK | 1997605 | 1966432 | 1863428 | 1966432 | 1 | NAY | User Facility | Unknown | Other | Initial | 2011-01-13 | 2011-01-14 | Yes | Yes | Health Professional | 8MM CAUTERY HOOK | Yes | No |  | No Answer Provi |  | Invalid Data | 2011
| 2010-11-30 | No Answer Provided | Other | It was reported that 10 hours post successful da vinci s hysterectomy procedure, the patient experienced a fever of 103. A pan culture was taken and was found to be negative with no fever indicated. The patient underwent an unknown treatment and discharged on schedule and without prolonged hospitalization. The patient was reported as doing well with no additional post operative issues. | During our investigation of this event at the site, it was determined that retraining was needed to advise the site on proper cleaning and sterilization processes for da vinci instrument and accessories. During the retraining the cleaning and sterilization staff reported that there may have been one or more steps in the recommended process that they were not performing. Since the retraining and change in the sites cleaning and sterilization process, there have not been any further reports of post operative temperatures. | DA VINCI S INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1987793 | 1960556 | 1857601 | 2955842201100017 | 1 | NAY | Manufacturer | User facilityCompany Representative | RISK MANAGER | Initial | 2010-12-14 | 2011-01-13 | Yes | No | Service Personnel |  | No | No | 2010-12-14 | No |  | Unkown | 2011
400180 | 2010-12-17 | Other | Other | It was reported that approximately one hour during a da vinci s hysterectomy procedure, the surgical staff witnessed an arc of energy from the monopolar curved scissors instrument with the tip cover accessory installed, to the patient's uterus. No intervention or repair was required as the uterus was being removed. Approximately 2 hours later in the procedure, the surgeon decided to convert the procedure to traditional open surgical techniques due to the patient's anatomy. The procedure was completed and no patient harm or adverse outcome was reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however, the isi representative present during the event stated that throughout the procedure, the monopolar curved scissors instrument with the tip cover accessory installed were being compromised due to the movement of the patient's large organ. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987769 | 1960583 | 1857579 | 2955842201100011 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-17 | 2011-01-13 | Yes | Yes | Health Professional | 400180 | No | No | 2010-12-17 | No | Yes | Unkown | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage and therefore only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda in regarding these events: 2955842-2011-00013 and 2955842-2011-00014. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987770 | 1960584 | 1857580 | 2955842201100012 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
40018402 |  | Malfunction |  | Event Description: --/-- On (b)(6) 2010, medwatch uf/importer report (b)(4) was received: during the surgical procedure, the operating physician noticed a small foreign body in the operating field. Upon further inspection, it was noted that a piece of plastic had broken off the intuitive endowrist monopolar spatula. The instrument was removed and replaced by a new instrument. | Manufacturer Narrative: --/-- Multiple attempts with the site have been made; however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering observed the ceramic sleeve to be broken at the distal end and missing a piece approximately. 060 inches by. 090 inches. The ceramic sleeve exhibits scratches below the broken section. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 1987773 | 1960587 | 1857583 | 2955842201100010 | 1 | NAY | Manufacturer | Other | RISK MANAGER | InitialFollowup | 2011-04-01 | 2011-01-13 | No | Yes | Health Professional | 40018402 | Device Returned To Manufacturer | No | 2011-04-01 | Device Not Retu | No |  | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage and therefore only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda regarding these events: 2955842-2011-00012 and 2955842-2011-00014. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987774 | 1960588 | 1857580 | 2955842201100013 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
400180 | 2010-12-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing was observed. The mcs instrument and tip cover were removed and replaced, and the procedure continued as planned. After an unspecified amount of time, the replacement mcs instrument with the tip cover accessory installed was observed to have arced. The mcs instrument and tip cover were removed, the mcs instrument was inspected and found to have no damage, and therefore, only the tip cover was replaced. The surgical procedure resumed as planned. A short time later, the surgical staff witnessed arcing from the mcs instrument with the tip cover accessory installed. The mcs instrument and tip cover were removed a third time and the tip cover was replaced. The procedure was completed with the fourth replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer (fse) was unable to replicate the customer reported arcing problem. The da vinci system was found to be within specification, and the esu settings used during the procedure were within the isi recommended range. Different mcs instruments and tip cover accessories were used during the procedure and these items were unavailable during fse's site visit. The tip cover accessories have been discarded by the site and monopolar curved scissors (mcs) instruments are not available for return, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of january 13, 2011, there have been no reported recurrences of the issue at this hospital. The following medwatch mfr reports were submitted to the fda in regarding these events: 2955842-2011-00012, 2955842-2011-00013. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987738 | 1960589 | 1857580 | 2955842201100014 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-16 | 2011-01-13 | No | Yes | Health Professional | 400180 | No | Yes | 2010-12-16 | No | Yes | Unkown | 2011
400180 | 2010-12-08 | Other | Other | It was reported that during a da vinci s surgical procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument, to the patient's vessel. The surgeon repaired the vessel with a few stitches and the planned procedure was completed. No adverse outcome was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however, it was noted that the esu setting used during the procedure was above the isi recommended range. On (b)(6) 2011, the isi representative was told by the surgical technician that reported the event, that the patient recovered well and there was no adverse outcome. Additional information has been requested multiple times, however, as of the date of this report, the site has not provided any further information. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1987775 | 1960590 | 1857586 | 2955842201100015 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-12-17 | 2011-01-13 | Yes | Yes | Health Professional | 400180 | No | No | 2010-12-17 | No | Yes | Unkown | 2011
| 2010-12-01 | No Answer Provided | Other | It was reported that 10 hours post successful da vinci s hysterectomy procedure, the patient experienced a fever of 103. A pan culture was taken and was found to be negative with no fever indicated. The patient underwent an unknown treatment and discharged on schedule, without prolonged hospitalization. The patient was reported as doing well with no additional post operative issues. | During our investigation of this event at the site, it was determined that an retraining was needed to advise the site on proper cleaning and sterilization processes for da vinci instrument and accessories. During the retraining, the cleaning and sterilization staff reported that there may have been one or more steps in the recommended process that they were not performing. Since the retraining and change in the site's cleaning and sterilization process, there have not been any further reports of post operative temperatures. | DA VINCI S INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1987794 | 1960605 | 1857601 | 2955842201100016 | 1 | NAY | Manufacturer | User facilityCompany Representative | RISK MANAGER | Initial | 2010-12-14 | 2011-01-13 | Yes | No | Service Personnel |  | No | No | 2010-12-14 | No |  | Unkown | 2011
| 2010-11-12 | Malfunction |  | The davinci tip cover accessory developed holes in it intra-operatively. |  | DAVINCI TIP COVER ACCESSORY | DAVINCI TIP COVER ACCESSORY | 1999514 | 1971338 | 1868334 | 1971338 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2011-01-10 | 2011-01-10 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2011
42017910 | 2010-11-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure sparking was seen coming from the tips of the monopolar curved scissors instrument with the tip cover accessory installed. The surgeon decided to continue with the case, however shortly after a second sparking event was noticed. At this time the surgical staff removed the monopolar curved scissors instrument and tip cover accessory and replaced them with another instrument and accessory of the same type. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The mcs instrument and tip cover accessory were returned and evaluated. Per the customer reported complaint engineering found no visual evidence of arcing or char marks on the instrument. The instrument was inspected and found to pass electrical continuity testing and did not have any burrs or abnormally sharp edges that could contribute to arcing. Visual inspection of the tip cover accessory showed no thermal or mechanical damage to the silicone. On (b)(6) 2010 (b)(6) reported that the instrument and accessory were inspected prior to use and after removed from this procedure, however, no evidence of malfunction could be found. The esu setting used was noted to be within isi's recommended range. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1979926 | 1952946 | 1849941 | 2955842201100005 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-12-13 | 2011-01-07 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes |  | Device Not Retu | No | Reuse | 2011
42011704 | 2010-12-08 | Malfunction |  | It was reported that during a da vinci s tubal reanastomosis procedure, while the surgeon was suturing, the tip of the 5mm needle driver instrument broke and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse events, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one grip to be detached from the distal end of the 5mm needle driver instrument. The carbide insert was found to still be attached to the grip. The grip fractured at the junction with the hub feature, where there is a 90 degree angle, and hairline cracking is exhibited. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1979906 | 1952970 | 1849965 | 2955842201100006 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-12-09 | 2011-01-07 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | It was reported that during a da vinci s mitral valve replacement procedure, an instrument became stuck on a patient side manipulator (psm) arm. With the assistance of an isi technical support engineer, the psm was undocked from the patient and the cannula and instrument were removed. The psm was re-draped and the site attempted to re-engage a different instrument on the arm, however, the engagement problem continued to occur. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found the instrument engagement issue reported by the customer was associated with the right spring plunger on the affected patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The spring plungers are mechanical elements used to help secure and position the sterile adapter interface between the psm and the sterile instrument. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found one of the psm's spring plungers, between axis 4-7 on top of link 6, to be bent and difficult to depress. The 4 affected spring plungers were replaced. As of january 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1979990 | 1953007 | 1850002 | 2955842201100007 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | It was reported that during a da vinci s myomectomy procedure, system error code 212 occurred and could not be overridden. The surgical staff power cycled the system, however, system error code 212 recurred immediately after power up. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci s in a recoverable safe state. The mtml was returned to isi for failure analysis investigation. Engineering evaluation found that the harness conductor had malfunctioned, thus generating the system error code experienced by the customer. As of january 6, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1980062 | 1953031 | 1850027 | 2955842201100008 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-09 | Other | Other | Event Description: --/-- It was reported that during a da vinci s surgical procedure, the surgical staff experienced a green tint on the left side image displayed in the surgeon console. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the vision cable. The vision cable connects the vision cart (vc) to the surgeon console (sc) and passes the video signals between the sc and vc. The system was repaired by replacing the affected vision cable. The vision cable was returned to isi for failure analysis investigation. Engineering observed contamination on the cable, which likely caused the image issue experienced by the customer. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the vision cable. The vision cable connects the vision cart (vc) to the surgeon console (sc) and passes the video signals between the sc and vc. The system was repaired by replacing the affected vision cable. The vision cable was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure during evaluation and concluded that the image issue experienced by the customer may have been caused by the vision cable not being properly connected. As of (b)(4) 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | Reuse | 1980305 | 1953271 | 1850266 | 2955842201100009 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-12-09 | 2011-01-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-09 | Device Not Retu | No |  | 2011
IS1200 A43P9 | 2010-12-06 | Other | Other | It was reported that at system start up for a da vinci prostatectomy procedure, the site experienced multiple occurrences of system error code 20105. With the assistance of an isi technical support engineer, the cable connections were checked; however, the system error persisted. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system error code 20105 was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards, endoscopic camera manipulator (ecm), and multiple servo driver (msd). An ria pca is assigned to each system arm and performs numerous functions related to the function, control and hardware fault monitoring for its assigned arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system contains two multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected ria pca, ecm and msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20105 appears when an auxiliary voltage error is detected during self-test. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of january 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1976403 | 1949858 | 1846854 | 2955842201100002 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-06 | 2011-01-05 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
IS3000 A60P5 | 2010-12-06 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgical staff heard a pop sound and the system would not start up. The patient was under anesthesia and port incisions had been made before the surgical staff made the decision to abort the procedure and reschedule for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the power supply. The system was repaired by replacing the affected power supply. As of january 4, 2011, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1976405 | 1949860 | 1846856 | 2955842201100003 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-06 | 2011-01-05 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
42017206 | 2010-09-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument stopped working. The site also reported that fragment(s) were removed from the patient and there was no harm to the patient. Despite attempts to gather additional information, the site has not provided any additional event details. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument grip to be misaligned. The grip tip was found with 0. 035 offset at the tips. The bent grip has separation of the yaw pulley and pulley cover at the glue joint. Engineering has concluded that this failure may be the result of overloading at the tip. Electrical continuity passed. Engineering also observed pitch down cable to be broken at the distal clevis hub. The cable segment that contains the crimp is missing from the clevis. The clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1976370 | 1949881 | 1846877 | 2955842201100004 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-06 | 2011-01-05 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-12-06 | Device Not Retu | No | Reuse | 2011
IS2000 A51P8 | 2010-12-04 | Other |  | It was reported that during a da vinci s ureteroplasty procedure, the customer experienced system error code 23017. With the assistance of a technical support engineer (tse), the site performed an emergency power off (epo) to restart the system multiples times, however, system error recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering confirmed the occurrence of system error code 23017 as well as system error code 25211. Engineering concluded that system error codes 23017 and 25211 were associated with the endoscopic camera manipulator (ecm) arm and remote arm controller (rac). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The remote arm controller (rac) consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected ecm and rac. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. System error code 25511 appears when the da vinci s safety system determines that a calibration error has occurred. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1972539 | 1945108 | 1842104 | 2955842201100001 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-12-04 | 2011-01-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-12-04 | Device Not Retu | No | Reuse | 2011
40018010 | 2010-11-19 | Injury | Other | Event Description: --/-- On (b)(6) 2010, medwatch (b)(6) was received: the bovie cautery unit cover had a defect in it. When the cautery was activated it arced through the defect causing a laceration of the patient's iliac vein. The vein was repaired. The cover device was inspected prior to use and no tears or punctures were noted, however, two additional events that did not result in patient harm occurred the day before this event with the same device. One looked like a small tear or puncture, and the other looked like a small slit in the material. The packaging was not retained, however, the lot from the storage area for the same product was c10222. Medwatch mfr reports 2955842-2010-00561 and 2955842-2010-00562 were also sent to the fda regarding these events. | Manufacturer Narrative: --/-- On (b)(6) 2011, (b)(6) stated that the tip cover accessories returned to isi were not used during the event reported on medwatch uf / importer report (b)(4). The site has decided to retain the tip cover accessory involved in this (b)(6) 2010 event. Due to this additional information, please disregard the failure analysis information as reported on isi mfr report 2955842-2010-00563. The tip cover accessory will not be returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The tip cover accessories returned by (b)(6) hospital will be reported via medwatch mfr reports: 2955842-2010-00561, 2955842-2010-00562. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was found to have a mechanical tear and various holes burned through the silicon. Two of the holes are adjacent to each other and the third hole is approximately 90 degrees away. The silicon exhibits localized melting around the holes, indicative of arcing. The hole sizes range from. 010 inches to. 030 inches with varying shapes (round and oblong), and are located. 225 inches to. 470 inches above the silicon to sleeve interface. The tip cover also has a mechanical tear adjacent to the edge of one of the holes. On (b)(6) 2010, (b)(6), stated that the patient was recovering well. | 8MM MCS TIP COVER ACCESSORY | Initial | 1970355 | 1942498 | 1839494 | 2955842201000563 | 1 | NAY | Manufacturer | OtherUser facility | RISK MANAGER | InitialFollowup | 2011-01-04 | 2010-12-30 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2011-01-04 | Device Not Retu | Yes |  | 2010
40016907 | 2010-11-10 | Malfunction |  | Event Description: --/-- On (b)(4), 2010, isi received uf/importer report (b)(4) with the following event description: end piece broke off and was retrieved. | Manufacturer Narrative: --/-- The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found that the clamp arm detached from the distal end. One side of the shaft feature that mates with clamp arm pins is bent backwards. This damage is indicative of clamp arm hyperextension. While the evidence is inconclusive, engineering concluded that damage to the insert accessory is most likely due to mishandling/misuse. On (b)(6), 2010 (b)(6), (b)(6) hospital risk manager stated that during a da vinci s colostomy closure procedure, a piece of the harmonic curved shear (hcs) insert broke off and fell into the patient. The broken piece was retrieved. She also stated that a piece from a replacement insert accessory also used during the surgical procedure broke off and fell into the patient. The broken piece was retrieved. The planned surgical procedure was completed and there was no patient harm, adverse outcome or injury. Med watch report 2955842-2010-00564 was also submitted regarding this event. --/-- Additional Manufacturer Narrative: --/-- | HARMONIC CURVED SHEARS INSERT | Unkown | 1970570 | 1942604 | 1839600 | 2955842201000565 | 1 | NAY | Manufacturer | OtherUser facility | RISK MANAGER | FollowupInitial | 2010-12-01 | 2010-12-30 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-12-01 | Device Not Retu | Yes |  | 2010
40016907 | 2010-11-10 | Malfunction |  | On (b)(4), 2010, isi received uf/importer report (b)(4) with the following event description: end piece of device broke off. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found that the clamp arm detached from the distal end. One side of the shaft feature that mates with clamp arm pins is slightly bent backwards. While the evidence inconclusive, engineering concluded that damage to the insert accessory is most likely due to mishandling/misuse. On (b)(6), 2010 (b)(6), (b)(6) hospital risk manager stated that during a da vinci s colostomy closure procedure, a piece of the harmonic curved shear (hcs) insert broke off and fell into the patient. The broken piece was retrieved. She also stated that a piece from a replacement insert accessory also used during the surgical procedure broke off and fell into the patient. The broken piece was retrieved. The planned surgical procedure was completed and there was no patient harm, adverse outcome or injury. Med watch report 2955842-2010-00564 was also submitted regarding this event. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1970572 | 1942605 | 1839600 | 2955842201000564 | 1 | NAY | Manufacturer | OtherUser facility | RISK MANAGER | Initial | 2010-12-01 | 2010-12-30 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-12-01 | Device Not Retu | Yes | Unkown | 2010
| 2010-04-13 | Malfunction |  | On (b)(6), 2010, isi received uf/importer report (b)(4) with the following event description: internal wire within da vinci s surgical robot system broke during the surgical procedure but was able to be retrieved easily from patient and was not a retained object. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2010, (b)(6), the director of materials management at (b)(6) medical center stated that during a da vinci s hysterectomy procedure, a wire from an endowrist instrument broke off and fell into the patient. The broken piece was retrieved and no patient harm, adverse outcome or injury was reported. No other information was provided. |  | ENDOSCOPIC INSTRUMENT | 1970581 | 1942606 | 1839602 | 2955842201000566 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-12-01 | 2010-12-30 | No | Yes | Health Professional |  | No | No | 2010-12-01 | No | No | Unkown | 2010
40018010 | 2010-11-18 | Malfunction |  | Event Description: --/-- On (b)(4), 2010, medwatch uf / importer report (b)(4) was received: the bovie cautery unit cover had a defect in it. When the cautery was activated it arced through the defect causing a laceration of the patient's ileac vein. The vein was repaired. The cover device was inspected prior to use and no tears or punctures were noted, however, two additional events that did not result in patient harm occurred the day before this event with the same device. One looked like a small tear or puncture, and the other looked like a small slit in the material. The packaging was not retained however, the lot from the storage area for the same product was c10222. Medwatch mfr reports 2955842-2010-00561 and 2955842-2010-00563 were also sent to the fda regarding these events. | Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was found to have a mechanical tear and various holes burned through the silicone. Two of the holes are adjacent to each other and the third hole is approximately 90 degrees away. The silicone exhibits localized melting around the holes, indicative of arcing. The hole sizes range from. 010 inches to. 030 inches with varying shapes (round and oblong), and are located. 225 inches to. 470 inches above the silicone to sleeve interface. The tip cover also has a mechanical tear adjacent to the edge of one of the holes. Failure analysis concerning the returned tip cover accessory for this report was initially provided in mfr report 2955842-2010-00563. On (b)(6) 2010, isi received 2 tip cover accessories, believed to have been used during events reported in mfr report s 2955842-2010-00563 and 2955842-2010-00561. On (b)(6) 2011, (b)(6), stated that the site has decided to retain the tip cover accessory used during the event reported in medwatch uf / importer report (b)(4) (mfr report 2955842-2010-00563). A follow-up mdr has been sent to the fda concerning mfr report 2955842-2010-00563. --/-- Additional Manufacturer Narrative: --/-- The tip cover accessory and mcs instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. On (b)(6), 2010 (b)(6), (b)(6) risk manager, stated that during a da vinci si surgical procedure, arcing from the monopolar curved scissors (mcs) instrument with the mcs tip cover installed was observed. There was no patient harm, adverse outcome or injury to the patient. No additional information was provided. | 8MM TIP COVER ACCESSORY | Unkown | 1970575 | 1942613 | 1839494 | 2955842201000562 | 1 | NAY | Manufacturer | OtherUser facility | RISK MANAGER | InitialFollowup | 2010-12-01 | 2010-12-30 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2010-01-03 | Device Not Retu | Yes |  | 2010
40018010 | 2010-11-18 | Malfunction |  | On (b)(4), 2010, medwatch uf / importer report (b)(4) was received: the bovie cautery unit cover had a defect in it. When the cautery was activated it arced through the defect causing a laceration of the patient's ileac vein. The vein was repaired. The cover device was inspected prior to use and no tears or punctures were noted, however, two additional events that did not result in patient harm occurred the day before this event with the same device. One looked like a small tear or puncture, and the other looked like a small slit in the material. The packaging was not retained however the lot from the storage area for the same product was c10222. Medwatch mfr reports 2955842-2010-00562 and 2955842-2010-00563 were also sent to the fda regarding these events. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was found to have a mechanical tear approximately. 080 long at the distal tip of the silicone. The tear is in the axial direction. The tip cover accessory does not exhibit any damage related to arcing. No other damage was found. On (b)(6), 2010 (b)(6), (b)(6) risk manager, stated that during a da vinci si surgical procedure, arcing from the monopolar curved scissors (mcs) instrument with the mcs tip cover installed was observed. There was no patient harm, adverse outcome or injury to the patient. No additional information was provided. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1970794 | 1943068 | 1839494 | 2955842201000561 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2010-12-01 | 2010-12-30 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2010-12-01 | No | Yes | Unkown | 2010
420172 | 2010-12-15 | Malfunction | Other | The robotic (b)(4) dissector did not have complete function during a robotic left nerve-sparing radical prostatectomy. This was the 10th and final usage of the dissector. The 8 am robotic sheath partially disengaged from the robotic arm. (b)(4). |  | BIPOLAR MARYLAND FORCEPS | ROBOTIC BIPOLAR DISSECTOR | 1981906 | 1951131 | 1837133 | MW5018860 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-12-22 | 2010-12-30 | No | Yes | Health Professional | 420172 | Yes | Yes |  |  |  |  | 2010
400180 | 2010-11-30 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument and a burn was then found on the patient's bladder. The procedure was completed and no additional patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On (b)(4) 2010, the isi clinical sales representative confirmed with the surgeon that the patient burn reported did not require intervention or repair. The surgeon stated that the burn occurred on the bladder reflection and did not result in any trauma to the patient. Multiple requests for additional information were made, however, as of the date of this report, no additional information was provided by the surgeon or the hospital. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1968355 | 1939631 | 1836627 | 2955842201000559 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-12-01 | 2010-12-28 | Yes | Yes | Health Professional | 400180 | No | Yes | 2010-12-01 | No | Yes | Unkown | 2010
IS3000 A60P5 | 2010-12-01 | Other |  | It was reported that while preparing for a da vinci si hysterectomy procedure, while the site was moving the patient side cart, the blue fiber cable was damaged and a system communication fault appeared. The patient had been under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering confirmed that the blue fiber cable was damaged as the connector was found to be separated from the cable. The patient side cart (psc) fiber cable connects the psc to the other da vinci system components, such as the vision side cart and the surgeon side cart. The system was repaired by replacing the affected psc fiber cable. As of dec 27, 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1968310 | 1939653 | 1836649 | 2955842201000560 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-12-01 | 2010-12-28 | Yes | Yes | Health Professional | IS3000 A60P5 | Yes | No | 2010-12-01 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-11-24 | Other | Other | It was reported that prior to using the da vinci surgical system to perform any tasks for a prostatectomy procedure, the surgical staff experienced no image on the right and left high resolution stereo viewer of the surgeon side console and on the assistance monitor. With the assistance of an isi representative, troubleshooting was performed; however, the site was unable to resolve the vision issue. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that vision issue experienced by the site was associated with a camera control unit (ccu). The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected ccu. The ccu was returned to isi for failure analysis investigation. Engineering discovered a blown fuse, thus generating the vision issue experienced by the customer. As of december 22, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1966240 | 1937212 | 1834208 | 2955842201000558 | 1 | NAY | Manufacturer | ForeignUser facility | BIOMEDICAL ENGINEER | Initial | 2010-11-24 | 2010-12-23 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2010-11-24 | Device Not Retu | No | Reuse | 2010
IS2000 | 2010-10-14 | Other | Required Intervention | It was reported that 17 hours after a da vinci s prostatectomy procedure was successfully completed, the patient's blood pressure decreased and bleeding was observed coming from the neurovascular bundle. A blood transfusion was performed and a hematoma evacuation was conducted by open surgery. The surgeon indicated that the system functioned normally during the surgical procedure and no anomalies with the system were noticed. | On (b)(6) 2010, the surgeon indicated that the post operative bleeding experienced by the patient was unrelated to the da vinci s surgical system and a post surgical complications related to prostatectomy procedures. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. Per the information provided above, the initial reporter informed intuitive surgical that this event does not require reporting to the (b)(4) medical device regulatory agency under their medical device surveillance system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1966239 | 1937246 | 1834242 | 2955842201000557 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-11-25 | 2010-12-23 | Yes | No | Health Professional | IS2000 | No | No | 2010-11-25 | No | No | Reuse | 2010
40018010 | 2010-11-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) with an mcs tip cover accessory installed. Inspection of the tip cover accessory by the surgical staff revealed that the mcs tip cover had a hole. The planned surgical procedure was completed with a replacement mcs instrument tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1964990 | 1936479 | 1833475 | 2955842201000556 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-22 | 2010-12-22 | No | Yes | Health Professional | 40018010 | No | Yes | 2010-11-22 | No | No | Unkown | 2010
420172 | 2010-12-15 | Malfunction | Other | The robotic maryland bipolar dissector did not have complete function during a robotic left nerve-sparing radical prostatectomy. This was the 10th and final usage of the dissection. The 8 mm robotic sheath partially disengaged from the robotic arm. |  | BIPOLAR MARYLAND FORCEPS | ROBOTIC BIPOLAR DISSECTOR 8MM | 1971737 | 1940137 | 1837133 | MW5018707 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-12-22 | 2010-12-22 | No | Yes | Health Professional | 420172 | Yes | Yes |  |  |  |  | 2010
420230 | 2010-12-15 | Malfunction |  | Device was working prior to failure. It failed to properly grip staple to allow for it to be advanced through trocar. |  | HEMOLOCK | CLIP APPLIER LARGE | 1969453 | 1941395 | 1838391 | 1941395 | 1 | FZP | User Facility |  | RISK MANAGER | Initial | 2010-12-22 | 2010-12-22 | No | Yes | Physician | 420230 | Yes | No Answer Provided |  |  |  |  | 2010
420179 | 2010-12-20 | Malfunction |  | Robotic total hysterectomy - monopolar curved scissors worked intermittently throughout case. Monopolar power cable was unplugged and replugged in to try and resolve issue. Power cord was very hot and later broke off. Monopolar scissor was also broken at the base. No known injury to the patient. This item is a reusable item. The staff is unsure if the cord was frayed before the case, or if it happened during the case. This event will be forwarded to risk management. |  | MONOPOLAR CURVED SCISSOR | DAVINCI ENDOWRIST INSTRUMENT | 1972797 | 1945461 | 1842457 | 1945461 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-12-21 | 2010-12-21 | No | Yes | Physician | 420179 |  | No Answer Provided |  |  |  |  | 2010
IS3000 A60P5 | 2010-11-22 | Other | Other | It was reported that during a da vinci si hysterectomy procedure the surgical staff experienced multiple occurrences of system error code 252. An intuitive surgical technical support engineer (tse) remotely reviewed the system error logs and found multiple entries of system error codes 307 and 252. The tse attempted to troubleshoot the issue via telephone, however, system error code 252 continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes 307 and 252 were associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 307 appears when one component of the system does complete its power up process at system startup. System error code 252 is a sympathetic error and appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the safety system puts da vinci si in a non-recoverable safe state. The doco was returned for evaluation, however, engineering was unable to replicate the customer reported failure mode during in-house testing. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1960637 | 1931637 | 1828630 | 2955842201000550 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-22 | 2010-12-17 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-11-22 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-11-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, and while cutting the bladder neck tissue, the scissor tip on the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1961003 | 1932264 | 1829257 | 2955842201000553 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-19 | 2010-12-17 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-11-19 | No | No | Unkown | 2010
42017909 | 2010-10-15 | Malfunction |  | It was reported that during a da vinci s salpingo-oophorectomy procedure, the scissor tip on the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade was returned detached from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1961011 | 1932272 | 1829265 | 2955842201000551 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-17 | 2010-12-17 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-11-17 | Device Not Retu | No | Reuse | 2010
400180 | 2010-11-23 | Malfunction |  | It was reported that during a da vinci s surgical procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with an mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover is returned for evaluation a follow-up mdr will be submitted. The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal clevis exhibits a series of char marks from approximately. 100 to. 140 above the interface with the tube extension. Engineering concluded that the char marks are most likely the areas where energy escaped through the tip cover. Although the mcs tip cover accessory used in conjunction with the monopolar curved scissors instrument was not returned for evaluation, engineering found that the char marks on the instrument are located where the tip cover accessory silicone and ultem sleeve meet. There are no burrs/damage on the instrument in the area where the char marks are located that would contribute to an arcing event. No other damage found. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1961013 | 1932274 | 1829267 | 2955842201000552 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-24 | 2010-12-17 | No | Yes | Health Professional | 400180 | No | No | 2010-11-23 | No | Yes | Unkown | 2010
400180 | 2010-11-17 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with an mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover is returned for evaluation, a follow-up mdr will be submitted. The instrument used in conjunction with the tip cover was returned and evaluated. Engineering evaluation found that the instrument does not exhibit evidence of arcing or damage. The tube extension and grip scissor tips do not have an arc track. The instrument also passed functional testing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1960973 | 1932278 | 1829271 | 2955842201000554 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2010-11-19 | 2010-12-17 | No | Yes | Health Professional | 400180 | No | Yes | 2010-11-19 | No | Yes | Unkown | 2010
42022703 | 2010-11-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, one of the tips on the pk dissecting forceps instrument broke off and fell into the patient. The site was unable to locate and remove the broken piece after 3 hours of searching. The planned surgical procedure was completed with an instrument of the same type. A post operative x-ray revealed that the broken piece remains in the patient. No patient harm, adverse outcome or injury was reported. On (b)(6), 2010, intuitive surgical, inc. , received a voluntary report for patient (b)(6) with the following event description: patient was in the process of having a da vinci laparoscopic hysterectomy when it was noted that the intuitive endowrist pk dissecting forcep was missing one of the forcep tips. A laparoscopic search of the abdomen and pelvis for the missing forcep tip was conducted but it could not be found. The tip was seen on the flat plate x-ray of the abdomen under the spleen. The forcep tip remains inside the patient and the patient is aware of this. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6), 2010, it was reported by the initial reporter that upon location of the broken instrument piece under x-ray, the surgeon made the decision not to remove the broken instrument piece. There was no harm, adverse outcome or injury experienced by the patient. The patient was discharged from the hospital on (b)(6) 2010 and to the best of his knowledge, the patient has not returned to the hospital due to any post operative complications. On (b)(4), 2010, the site's risk manager also confirmed that the hospital has not received any complaints from the patient regarding experiencing post operative complications. No additional information concerning the patient and the reported event were provided. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1960974 | 1932279 | 1829272 | 2955842201000555 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-18 | 2010-12-17 | No | Yes | Health Professional | 42022703 | No | No | 2010-11-18 | No | No | Unkown | 2010
400205 | 2010-12-06 | Malfunction |  | Surgeon noticed splintering on the shaft of the instrument during the robotic laparoscopy procedure. Instrument was replaced with a new one. Patient was not harmed. ====================== health professional's impression======================wear on the instrument, which is considered semi-disposable====================== manufacturer response for fenestrated bipolar forceps, intuitive======================we returned product per instructions for immediate replacement and reduced cost by company. Instrument is semi-disposable, to be used 10 times only. |  | FENESTRATED BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 1965050 | 1936510 | 1833506 | 1936510 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-12-14 | 2010-12-14 | No | Yes | Physician | 400205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
S2000 | 2010-12-03 | Malfunction |  | According to the md, "today during my robotic radical vaginectomy, i discovered a very dangerous problem with the davinci electrosurgical wiring. I noticed, quite by accident and not a little good fortune that when i activated the monopolar current to my right handed endoshears, my left handed bipolar device also fired. The tissue i was holding was just peritoneum, so no harm done. However, there could have been more serious consequences. What was realized was that the monopolar and bipolar cords at the patient end were intertwined. This created a coupling so that the monopolar current flowing through the old grey cord was creating a current in the poorly shielded blue bipolar cord. It's important to make sure these two cords are not close together or a thermal injury may occur. ". |  | DA VINCI SURGICAL SYSTEM | DA VINCI SURGICAL SYSTEM | 1972773 | 1945459 | 1842455 | 1945459 | 1 | NAY | User Facility |  | Nurse | Initial | 2010-12-14 | 2010-12-14 | No | Yes | Physician | S2000 | Yes | Yes |  |  |  |  | 2010
IS3000 A65P5 | 2010-11-16 | Other | Other | It was reported that two hours into a da vinci si hysterectomy procedure, the surgeon experienced image loss on the surgeon console. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The system was evaluated by a field service engineer (fse). The fse cycled the system multiple times and could not replicate the customer reported failure mode. The fse concluded that the image loss experienced by the customer may have been caused by a dirty vision cable or the vision cable not being properly connected. As of december 9, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1952472 | 1923273 | 1820269 | 2955842201000547 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-16 | 2010-12-13 | Yes | Yes | Health Professional | IS3000 A65P5 | Yes | No | 2010-11-16 | Device Not Retu | No | Reuse | 2010
IS3000 A60P5 | 2010-11-16 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff experienced system error code 802 when the setup joint clutch was pressed on one of the instrument arms. An intuitive surgical technical support engineer reviewed the system pop-up logs and observed multiple occurrences of system error code 23. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes 23 and 802 were associated with a setup joint (suj) arm. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from its own joint positions to the remote arm controller. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm, with system error code 802 being generated as a sympathetic fault. In the event of these communication issues, the system records the fault and transitions to a safe state. The suj was returned for evaluation. Engineering discovered a shortage in the suj's internal cable harness, thus generating the system error code experienced by the customer. As of december 9, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1952474 | 1923274 | 1820270 | 2955842201000549 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-11-16 | 2010-12-13 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-11-16 | Device Not Retu | No | Reuse | 2010
42018404 | 2010-08-27 | Malfunction |  | On (b)(6), 2010 , medwatch uf/importer report (b)(4) was received: event desc: surgeon noticed a piece of metal inside the patient's abdomen. Further investigation noted that part of the robotic permanent cautery spatula instrument was cracked and missing a piece. The piece was removed from the abdomen and taken odd the field. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the ceramic sleeve to be chipped ad the base of the spatula. No additional cracks were found on the sleeve and no arc trac was evident on the sleeve or clevis. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1952670 | 1923346 | 1820342 | 2955842201000548 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2010-11-17 | 2010-12-13 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2010-11-17 | Device Not Retu | No | Reuse | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1949354 | 1919778 | 1816774 | 2955842201000536 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949379 | 1919781 | 1816777 | 2955842201000537 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949380 | 1919782 | 1816778 | 2955842201000538 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949382 | 1919783 | 1816779 | 2955842201000539 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949389 | 1919785 | 1816781 | 2955842201000540 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949392 | 1919786 | 1816782 | 2955842201000541 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949395 | 1919798 | 1816794 | 2955842201000542 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949409 | 1919812 | 1816808 | 2955842201000544 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1949530 | 1919972 | 1816968 | 2955842201000543 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400169 | No | Yes | 2010-11-10 | No | Yes | Unkown | 2010
400180 | 2010-11-10 | Malfunction |  | It was reported that during a da vinci si surgical procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1949533 | 1919978 | 1816974 | 2955842201000545 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-11-10 | 2010-12-10 | No | Yes | Health Professional | 400180 | No | No | 2010-11-10 | No | Yes | Unkown | 2010
40016907 | 2010-10-22 | Malfunction |  | It was reported that during a da vinci s sigmoid resection procedure, a piece of the tip on the harmonic curved shears insert (hcs) accessory broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was detached from the distal end. The clamp arm was not returned. One side of the shaft features at the distal end that mates with the clamp arm pins is bent backwards. The damage appears to be consistent with hyperextension of the clamp arm and engineering concluded that the breakage was most likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARSINSERT ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1949877 | 1920322 | 1817318 | 2955842201000546 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-11 | 2010-12-10 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2010-11-11 | Device Not Retu | No | Reuse | 2010
42018305 | 2010-11-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the hook on the permanent cautery hook instrument broke off and fell into the patient. The broken piece was retrieved and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's hook is broken off within the yaw pulley, at the distal most hole of the shank portion. The broken hook piece was not returned. Engineering also observed that the ceramic sleeve is broken. The ceramic sleeve broke in half lengthwise and the broken piece, approximately,. 170 x. 125 in size was returned detached from the distal end. The intact piece of the sleeve exhibits scratch marks. While the evidence is inconclusive, engineering has concluded that damage to the hook and ceramic sleeve is most likely due to mishandling/misuse. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1946056 | 1916662 | 1813658 | 2955842201000533 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-08 | 2010-12-07 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2010-11-08 | Device Not Retu | No | Initial | 2010
40016907 | 2010-10-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the tip of the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was detached from the distal end. The distal end shaft features that mate with the clamp arm pins are bent backwards. The damage appears to be consistent with hyperextension of the clamp arm and engineering concluded that the breakage was most likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences. Do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1946112 | 1916718 | 1813714 | 2955842201000532 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-08 | 2010-12-07 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-11-08 | Device Not Retu | Yes | Unkown | 2010
IS1200 A43P9 | 2010-11-09 | Other | Other | It was reported that prior to using the da vinci surgical system to perform any tasks for a prostatectomy procedure, the surgical staff experienced homing issues with the system. When the surgical staff was able to home the system, the instrument sterile adapter did not engage. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the homing and engagement issues were associated with a patient side manipulator (psm), which had been physically damaged when the customer transported the system from one location to another. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned for failure analysis investigation. Engineering evaluation observed that the shroud was damaged, thus preventing normal movement for insertion and resulting in the system issues experienced by the customer. As of december 2, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1946113 | 1916719 | 1813715 | 2955842201000531 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-11-09 | 2010-12-07 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-11-09 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-11-08 | Other | Other | It was reported that during set up for a da vinci s prostatectomy procedure, the surgeon experienced a black image through both eyes of the high resolution stereo viewer which is located on the surgeon side console (ssc). Additionally, the touchscreen monitor located on the vision side cart (vsc) displayed color bars. With the assistance of an isi technical support engineer (tse), the site verified settings and pressed the bars button on the camera control unit (ccu) to repair the color bars on the touchscreen monitor. Additional troubleshooting to repair the image on the ssc consisted of reseating cables, verifying the green led on the rear of the digitizer and rebooting the system multiple times; however, the black image persisted. The patient was under anesthesia for 1. 5 hours and port incisions had been made before the surgical staff made the decision to abort the procedure and reschedule for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the video input output processor board (vio). The system was repaired by replacing the affected vio. As of december 06, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1946144 | 1916750 | 1813746 | 2955842201000535 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-08 | 2010-12-07 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-11-11 | Device Not Retu | No | Reuse | 2010
IS3000 A60P5 | 2010-11-08 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the site experienced system error 23008. With the assistance of an isi technical support engineer (tse), the site exercised the affected arm axis to recover from the system error and restarted the system; however, the fault immediately returned. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned surgical procedure. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtm gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a nonrecoverable safe state. The mtml gimbal was returned to isi for failure analysis investigation. Engineering confirmed the reported complaint and replaced the motor to repair the affected mtml. As of december 6, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1946146 | 1916752 | 1813748 | 2955842201000534 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-08 | 2010-12-07 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-11-08 | Device Not Retu | No | Reuse | 2010
| 2010-11-23 | Malfunction |  | Mega suturecut needle driver not gripping enough. Six lives left on item. No pt involvement. Uf# (b)(4). |  | MEGA SUTURECUT NEEDLE DRIVER | NONE | 1953366 | 1919705 | 1816701 | MW5018507 | 1 | NAY | Voluntary |  | Nurse | Initial | 2010-12-03 | 2010-12-07 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2010
IS3000 A60P5 | 2010-11-10 | Other | Other | It was reported that while preparing for a da vinci si procedure, the surgical staff detected a sulfur like smell coming from the patient side cart (psc) in the operating room. No system errors were reported and when moving the psc into another room, the odor seemed to follow the system. The patient was under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule for a later date. | Investigation conducted by the field service engineer (fe) was unable to reproduce the customer reported sulfur smell. Surgical staff then notified the fe that during a case the day prior to this event, a patient suffered severe diarrhea and that the smell was likely related to this event. The fe was unable to find any physical evidence or odor resulting from contamination or fluid intrusion to the patient side cart. Additionally, no system errors were found to have occurred. As of dec 2, 2010 the site has continued to perform procedures with the da vinci si surgical system and there have been no reported recurrences of this issue. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1944327 | 1915186 | 1812182 | 2955842201000524 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-10 | 2010-12-06 | Yes | No | Health Professional | IS3000 A60P5 | Yes | No | 2010-11-10 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-08-12 | Death | Death | Event Description: --/-- It was reported that post a successful da vinci hysterectomy procedure, the patient developed an infection and was not discharged as planned. Ten days post op, the patient was given a blood thinner and expired due to an arterial bleed. Reportedly, the patient's demise was unrelated to the da vinci surgical system. | Manufacturer Narrative: --/-- On november 30, 2010, an isi clinical sales representative spoke to the surgeon that performed the da vinci hysterectomy on the patient. The surgeon stated that the patient did have an autopsy performed which found that the patient had an injury to her left external iliac artery. The source of the injury was not confirmed. Based on the limited information provided, it is indeterminable if the da vinci system, instruments or accessories contributed to the patient's demise. As of december 3, 2010, no adverse events have been experienced with the site's system. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2012, isi received a legal summons indicating that the patient expired on (b)(6) 2010, post op a da vinci si hysterectomy with bilateral salpingoo-oophorectomy and pelvic node dissection procedure performed on (b)(6) 2010. Emergency surgical procedures performed on the patient on (b)(6) 2010, revealed that the patient had sustained a burn to the right external iliac artery, pumping blood in the body cavity, causing bowel ischemia incompatible with life. Following the surgical procedure, pain medication was continually administered to the patient without good affect due to significant pain experienced by the patient. | DA VINCI SURGICAL SYSTEM | Reuse | 1944331 | 1915208 | 1812204 | 2955842201000528 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2010-11-11 | 2010-12-06 | Yes | No | Health Professional | IS1200 A43P9 | No | Yes | 2012-04-16 | Device Not Retu | No |  | 2010
IS1200 A43P9 | 2010-11-08 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the site experienced multiple occurrences of system error code 21003. The patient was under anesthesia for approximately 2 hours when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | An isi representative arrived onsite shortly after the case was converted to open. With the assistance of an isi technical support engineer, the surgeon side console cables were inspected and reconnected. The system was rebooted and homed and no system errors occurred. Additional investigation was conducted by a field service engineer whom was unable to replicate the customer reported system error code 21003. It was concluded that system error code 21003 was associated with patient side manipulator (psm) arm collisions. On (b)(4), 2010, the isi representative confirmed with the surgeon that the case would have likely been converted due to the patient's anatomy. The patient had a large uterus, 16 to 18 weeks in size, and the surgeon was having difficulty completing the case robotically. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1944332 | 1915210 | 1812206 | 2955842201000530 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-08 | 2010-12-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-11-08 | Device Not Retu | No | Reuse | 2010
420183 | 2010-11-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, half of the yellow plastic piece on the permanent cautery hook instrument broke off and fell into the patient. The broken piece was not retrieved. The planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. On (b)(6) 2010, the site's risk manager reported that no additional procedure was required to locate and remove the broken instrument piece. The instrument fragment was reported to have fallen into fatty tissue and was unable to be located. The patient is aware that the broken instrument piece was not retrieved. There were no complications experienced by the patient and the patient is reported to be doing well. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1943582 | 1914593 | 1811589 | 2955842201000522 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2010-11-04 | 2010-12-03 | No | Yes | Health Professional | 420183 | No | No | 2010-11-04 | No | No | Unkown | 2010
420179 | 2010-11-05 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si ureteral reimplantation procedure, one of the blades on the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece of the left blade broken off. The broken piece was not returned. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1943605 | 1914596 | 1811592 | 2955842201000525 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2010-11-05 | 2010-12-03 | No | Yes | Health Professional | 420179 | Device Returned To Manufacturer | No | 2010-12-16 | Device Not Retu | No |  | 2010
42017910 | 2010-11-05 | Other | Other | It was reported that at the end of a da vinci s sacrocolpopexy procedure, while the surgical staff was removing the cannula from the patient, the surgeon observed charring like burns at one of the cannula port incision sites. The port incision was sutured close and the affected area did not require any additional procedure or treatment. No additional patient harm, adverse outcome or injury was reported. | The instrument used during the surgical procedure has not been returned for evaluation, however, the field service engineer performed an electrical safety check on the da vinci s surgical system used during the surgical procedure and found that the system is functioning within specifications. Investigation has shown that it is possible for electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of december 3, 2010, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1943607 | 1914598 | 1811594 | 2955842201000527 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-05 | 2010-12-03 | Yes | Yes | Health Professional | 42017910 | No | No | 2010-11-05 | No | No | Unkown | 2010
40016907 | 2010-09-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the blades on the harmonic shear insert separated from the instrument and fell into the patient. The broken pieces were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the insert was returned with the clamp arm detached and the curved blade broken off at distal end. One side of shaft feature at the distal end that mates with clamp arm pins is bent backwards. The curved blade fractured at a straight section, before the blade transitions to the curved profile. The fracture surface is not flat and the midpoint of the blade has a couple of gouges/indentations on the outer edge. While the evidence is inconclusive, engineering concluded that damage to the blade and clamp arm is most likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1943583 | 1914609 | 1811604 | 2955842201000523 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-05 | 2010-12-03 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-11-05 | Device Not Retu | Yes | Unkown | 2010
42018304 | 2010-11-04 | Other | Other | Event Description: --/-- It was reported that during a da vinci s myomectomy procedure, while attempting to remove the permanent cautery hook (pch) instrument from the patient side manipulator (psm), the instrument became stuck. With the assistance of an isi technical support engineer (tse), the site was instructed to reposition the psm to view the instrument tip. After repositioning the psm, the surgical staff observed that the tip of the pch instrument was positioned at an acute angle, which prevented removal of the instrument. After straightening the instrument tip, the site was able to remove the instrument. Upon inspection of the pch instrument by the surgical staff, it was observed that a yellow piece from the instrument was missing. The planned surgical procedure was converted to an open procedure to locate and remove the broken instrument piece. No patient complications were reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece of the distal clevis detached. The broken piece was also returned along with the instrument. One distal clevis ear is broken off at its base. The broken piece is approximately. 300 x. 235 and is on the opposite side as the conductor wire and cap. Engineering concluded that damage to the instrument was most likely due to tip overloading. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, however, upon follow up with the site by the field service engineer, the site reported that the tip of the instrument was in an open position while attempting to remove the instrument. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. It has been determined that damage to the instrument was caused by improper positioning of the instrument tip and excessive force while attempting to remove the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 1943586 | 1914612 | 1811607 | 2955842201000526 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2010-11-04 | 2010-12-03 | Yes | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-11-19 | Device Not Retu | No |  | 2010
42017909 |  | Malfunction |  | Event Description: --/-- It was reported that during a da vinci surgical procedure, one of the blades on the monopolar curved scissors (mcs) instrument broke off and fell into the patient. No additional information concerning the reported event was provided. | Manufacturer Narrative: --/-- (b)(4) - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece of the left blade broken. The broken piece was not returned. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of blade is nearly straight. No other damage found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1943612 | 1914617 | 1811612 | 2955842201000529 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2010-11-05 | 2010-12-03 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-12-20 | Device Not Retu | No |  | 2010
400184 |  | No Answer Provided |  | During the surgical procedure, the operating physician noticed a small foreign body in the operating field. Upon further inspection, it was noted that a piece of plastic had broken off the intuitive endowrist monopolar spatula. The instrument was removed and replaced by a new instrument. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST MONOPOLAR CAUTERY IN | 1956464 | 1927421 | 1824422 | 1927421 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-12-03 | 2010-12-03 | No |  | Invalid Data | 400184 |  | No Answer Provided |  |  |  |  | 2010
42017207 | 2010-09-23 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, smoke, and arcing was observed coming from the maryland bipolar forceps instrument. The instrument was removed and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both yaw pulleys exhibit charring and localized material removal at the grip base, indicating that an arcing event occurred. While the evidence is inconclusive, engineering concluded that charring on the instrument was most likely due to a breach in the insulation. The instrument passed electrical continuity testing. The grips have a slight offset approximately. 015 at the tips. No other damage was observed. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1942298 | 1913167 | 1810163 | 2955842201000520 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-02 | 2010-12-02 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2010-11-02 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-10-21 | Malfunction | HospitalizationLife ThreateningRequired Intervention | It was reported that during a da vinci s hysterectomy procedure, the surgeon had observed sparking coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. Two burn injuries were noticed to the patient's small bowel. The case was paused, the instrument was removed and the tip cover accessory was replaced. A colorectal surgeon was called to repair the patient's injury. No further information provided by site. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1942351 | 1913206 | 1810202 | 2955842201000521 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-11-02 | 2010-12-02 | Yes | Yes | Health Professional | 40018010 | No | No | 2010-11-02 | No | Yes | Unkown | 2010
42018304 | 2010-10-29 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the tip of the permanent cautery hook instrument fell into the patient. The planned surgical procedure was completed, however, the customer is unsure as to whether or not the broken tip was retrieved. No patient harm, adverse outcome or injury was reported. The site indicated that the instrument would not be returned for evaluation. | The instrument will not be returned for evaluation. As a result, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1940593 | 1911816 | 1803081 | 2955842201000519 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-11-01 | 2010-12-01 | No | Yes | Health Professional | 42018304 | No | Yes | 2010-11-01 | No | No | Unkown | 2010
| 2010-11-17 | Malfunction | Other | Pt was in the process of having a da vinci laparoscopic hysterectomy when it was noted that the intuitive endowrist pk dissecting forceps was missing one of the forceps tips. A laparoscopic search of the abdomen and pelvis for the missing forceps tip was conducted but it could not be found. The tip was seen on flat plate x-ray of the abdomen under the spleen. The forceps tip remains inside the pt and the pt is aware of this. |  | INTUITIVE | INTUITIVE ENDOWRIST PK DISSECT | 1948936 | 1916839 | 1813835 | MW5018401 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-11-19 | 2010-11-30 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2010
400172 | 2010-11-16 | Malfunction |  | Intuitive surgical maryland endowrist instrument ceased cauterizing during procedure. Eight out of ten uses left on this instrument. |  | DAVINCI FENESTRATED MARYLAND BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 1944095 | 1915078 | 1812074 | 1915078 | 1 | NAY | User Facility |  |  | Initial | 2010-11-26 | 2010-11-26 | No | Yes | Physician | 400172 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-11-05 | Other | Other | It was reported that prior to starting a da vinci si nephrectomy procedure, "the site a purple tint in the left eye of the surgeon side console. " with the assistance of an isi technical support engineer, the surgical staff reseated the camera cable, checked the camera controller settings and touch screen settings and reseated the fiber cable, however, the vision issue was not resolved. The patient was under anesthesia and the port incisions were made when the surgical staff made the decision to abort the procedure and reschedule. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci si surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci si system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1936476 | 1907110 | 1804106 | 2955842201000512 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-05 | 2010-11-24 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-11-05 | Device Not Retu | No | Reuse | 2010
IS2000 A43P9 | 2010-11-01 | Other | Other | It was reported that during a da vinci s pelvic lymphadenectomy procedure, the surgical staff experienced system error code 212 every time the right master tool manipulator (mtm) was raised towards the surgeon console monitors. The site restarted the system multiple times and exercised the right mtm, however, system error code 212 continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The mtmr was returned to isi for failure analysis investigation. Engineering evaluation found that the harness conductor had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1936986 | 1907860 | 1804856 | 2955842201000515 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2010-11-01 | 2010-11-24 | Yes | Yes | Health Professional | IS2000 A43P9 | Device Returned To Manufacturer | Yes | 2010-11-01 | Device Not Retu | No | Reuse | 2010
40016906 | 2010-10-28 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the tip of the harmonic curved shears insert which was installed on the harmonic curved shears instrument, fell into the pt's abdomen. The piece was retrieved from the pt and three additional harmonic curved shears inserts were used without success as they were said to be broken at the proximal clevis. At this time the surgeon decided to convert the planned procedure to traditional laparoscopic methods. The site reported that there was no pt injury or adverse event as a result from this event. | The harmonic curved shears insert accessory was not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available, a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1960501 | 1912281 | 1809277 | 2955842201000516 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-10-28 | 2010-11-24 | No | Yes | Health Professional | 40016906 | No | No | 2010-10-28 | Device Not Retu | Yes | Initial | 2010
400169 | 2010-10-29 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the tip of the harmonic curved shears insert which was installed on the harmonic curved shears instrument, fell into the pt. The procedure was successfully completed. The site did not report any pt injury as a result from this event. No additional information has been provided despite attempts to gather additional information. | The harmonic curved shears insert accessory was not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information becomes available, a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1960470 | 1912301 | 1809297 | 2955842201000518 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-10-29 | 2010-11-24 | No | Yes | Health Professional | 400169 | No | No | 2010-10-29 | Device Not Retu | Yes | Initial | 2010
IS3000 A60P5 | 2010-10-26 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the site experienced system error #16. The site restarted the system and encountered system error #31030 on start up. To attempt to troubleshoot the issue, the site powered down the system, emergency powered off (epo's) the pt side cart (psc) and reset the breaker on all the components. The ends of the system cables were swapped between the console and psc; however, system error #31030 reoccurred. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering (fse) discovered that the system issue experienced by the customer was associated with a faulty embedded serializer for setup-joint (essj), remote arm controller (rac), and the pt side manipulator (psm). The essj is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected essj, rac, and psm. The affected components were returned to isi for failure analysis investigation. Engineering found the essj to be within specification with no trouble found; however, engineering did confirm the customer reported complaint as they were able to duplicate the system error codes. As a result, the rac and psm were repaired by replacing affected sub-components. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1960864 | 1912358 | 1809354 | 2955842201000517 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-26 | 2010-11-24 | Yes | No | Health Professional | IS3000 A60P5 | Yes | No | 2010-10-26 | Yes | No | Reuse | 2010
| 2010-11-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the harmonic curved shears insert fell inside the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was detached from the distal end. One side of the distal end shaft feature that mates with the clamp arm pins is bent backwards. The evidence is not conclusive, however, it appears that the damage is likely due to mishandling/misuse. The detached clamp arm was returned with the insert. No other damage was found. The instruments and accessories user manual specifically states: avoid contact with any and all metal or plastic instruments or objects when the instrument is activated. Contact with staples, clips, or other instruments while the instrument is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1935273 | 1906075 | 1803071 | 2955842201000513 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-11-04 | 2010-11-23 | No | Yes | Health Professional |  | Device Returned To Manufacturer | No | 2010-11-04 | Device Not Retu | Yes | Unkown | 2010
42018305 | 2010-10-26 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy the plastic stabilizer that joins up the tip of the permanent cautery hook instrument snapped off and fell inside the patient. The surgeon was able to retrieve a small round plastic disc and brown plastic piece. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base on the conductor cap side. The conductor cap is cracked at its hole feature and the extruded boss feature on the yaw pulley that mates with the hole feature on the conductor cap is broken off, allowing the cap to dislodge. The broken distal clevis ear piece and detached conductor cap were returned with the instrument. No other damage was found. On (b)(6) 2010, the initial reporter indicated that although the extended boss feature was not retrieved, the surgeon felt that it was most likely suctioned along with the fluids due to its size. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1935278 | 1906085 | 1803081 | 2955842201000514 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-10-26 | 2010-11-23 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2010-10-26 | Device Not Retu | No | Reuse | 2010
42017908 | 2010-03-29 | Malfunction |  | It was reported that during a da vinci surgical procedure while using monopolar curved scissors (mcs) instrument, the tip broke and fell into the patient. The tip was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found a section of the left blade to be broken and missing. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1934508 | 1905065 | 1802061 | 2955842201000505 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | PHYSICIAN ASSISTANT | Initial | 2010-03-29 | 2010-11-22 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | Yes | 2010-03-29 | Device Not Retu | No | Reuse | 2010
40016907 | 2010-09-10 | Malfunction |  | It was reported that approximately 30 minutes during a da vinci si hysterectomy procedure, the harmonic curved shears insert accessory fell into the patient. The piece was retrieved and insert was replaced. The procedure continued and the replacement harmonic curved shears insert also broke and fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda relating to this event: 2955842-2010-00436. | The harmonic curved insert accessory was returned and evaluated. Engineering found the insert to be returned with the clamp arm detached from the distal end. The shaft features that mates with the clamp arm are bent backwards. No other damage was found. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1934509 | 1905066 | 1767075 | 2955842201000508 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-29 | 2010-11-22 | No | Yes | Health Professional | 40016907 | No | No | 2010-10-29 | No | Yes | Unkown | 2010
40016907 | 2010-10-15 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued and the replacement harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the third harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the fourth harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient, replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch reports were sent to the fda in relation to this event: 2955842-2010-00492, 2955842-2010-00506, 2955842-2010-00507. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with the clamp arm detached from the distal end. The shaft feature that mates with the clamp arm pins is bent backwards. The insert showed evidence of overloading and or hyperextension of the clamp arm resulting in the breakage. No other damage was found. On (b)(6), 2010, (b)(6) stated that 4 harmonic curved shear (hcs) inserts broke during the myomectomy case on (b)(6) 2010. She stated that the surgeon was torquing the jaws of the hcs insert to tear tissue away from the myoma and using a hooking motion on the bottom which was causing the hcs inserts to break. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1934510 | 1905067 | 1802063 | 2955842201000509 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-28 | 2010-11-22 | No | No | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-28 | Device Not Retu | Yes | Initial | 2010
40016907 | 2010-10-15 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued and the replacement harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the third harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the fourth harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient, replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch reports were sent to the fda in relation to this event: 2955842-2010-00492, 2955842-2010-00506, 2955842-2010-00509. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with the clamp arm detached from the distal end. The shaft feature that mates with the clamp arm pins is bent backwards. The insert showed evidence of overloading and or hyperextension of the clamp arm resulting in the breakage. No other damage was found. On (b)(6), 2010 (b)(6) stated that 4 harmonic curved shear (hcs) inserts broke during the myomectomy case on (b)(6) 2010. She stated that the surgeon was torquing the jaws of the hcs insert to tear tissue away from the myoma and using a hooking motion on the bottom which was causing the hcs inserts to break. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1934511 | 1905068 | 1802063 | 2955842201000507 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-28 | 2010-11-22 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-28 | Device Not Retu | Yes | Initial | 2010
40016907 | 2010-10-15 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued and the replacement harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the third harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the fourth harmonic curved shears insert; however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient, replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch reports were sent to the fda in relation to this event: 2955842-2010-00492 2955842-2010-00509 2955842-2010-00507. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with the clamp arm detached from the distal end. The shaft feature that mates with the clamp arm pins are bent backwards. The insert showed evidence of overloading and or hyperextension of the clamp arm resulting in the breakage. No other damage was found. On (b)(6) 2010, (b)(6) stated that 4 harmonic curved shear (hcs) inserts broke during the myomectomy case on (b)(6) 2010. She stated that the surgeon was torquing the jaws of the hcs insert to tear tissue away from the myoma and using a hooking motion on the bottom which was causing the hcs inserts to break. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1934502 | 1905082 | 1802063 | 2955842201000506 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-28 | 2010-11-22 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-28 | Device Not Retu | Yes | Initial | 2010
IS3000 60P4 | 2010-11-02 | Injury | Other | It was reported that during a da vinci si sacrocolpopexy with hysterectomy procedure, while the surgeon was dissecting the i. P. Ligaments and round ligaments, the patients ureter was found to be injured. A urologist was called into the case to perform a ureterotomy and successful anastomosis of the ureter. The procedure was delayed approximately 45 minutes for the ureter injury repair, however the planned surgical procedure was completed. | Based on the information provided it was determined that the da vinci si surgical system worked as intended, with no system malfunction having had occurred, and the system did not cause or contribute to the patient's injury. The compression injury experienced by the patient was believed to have been caused by the complex patient's pathology. On (b)(4), 2010 the surgeon stated that the patient is recovering well and has since the stent, placed for the ureter injury, removed. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1934627 | 1905176 | 1802172 | 2955842201000510 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-11-02 | 2010-11-22 | Yes | No | Health Professional | IS3000 60P4 | No | No | 2010-11-02 | Device Not Retu | No | Reuse | 2010
42022703 | 2010-11-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a piece of the pk dissecting forceps instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument grips exhibit an offset of approximately. 015 inches at the tips when closed. The yaw pulley was found broken and conductor wire is broken and detached from the connection at the grip. The yaw pulley is missing a. 160 inch by. 060 inch piece opposite the conductor break allowing the grip to move within the pulley and have a larger range of motion in yaw. Per the case notes, the instrument fragment was retrieved from the patient. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1934584 | 1905192 | 1802188 | 2955842201000511 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-11-08 | 2010-11-22 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2010-11-08 | Device Not Retu | No | Reuse | 2010
PS2000 | 2010-04-13 | Malfunction |  | Internal wire within davinci s surgical robot system broke during the surgical procedure but was able to be retrieved easily from patient and was not a retained object. |  | DAVINCI S SURGICAL SYSTEM | DAVINCI S SURGICAL SYSTEM | 1944140 | 1915122 | 1812118 | 1915122 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-22 | 2010-11-22 | No | Yes | Physician | PS2000 | Yes | No Answer Provided |  |  |  |  | 2010
40016907 | 2010-10-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a piece from the harmonic curved shears insert, installed on the harmonic curved shears instrument, broke off and fell into the patient. The fragment was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with a missing clamp arm. The clamp arm connection joint was damaged and signs of over-extension of the clamp arm were observed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1932886 | 1903773 | 1802063 | 2955842201000503 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-19 | 2010-11-20 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-19 | Device Not Retu | Yes | Initial | 2010
40018304 | 2010-10-19 | Malfunction |  | Intuitive surgical, inc. Was advised of this event by the user facility on (b)(6) 2010 and on (b)(6) 2010, medwatch (b)(4) was received reporting the following event: it was reported that during a da vinci s prostatectomy procedure, the surgeon observed a plastic part of the tip of the permanent cautery hook instrument broke off. The pieces were retrieved and the procedure was successfully completed. There was no harm, adverse outcome or injury to the patient as a result of the instrument malfunction. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one ear of distal clevis to have broken off at its base. The clevis ear and the conductor cap were returned detached from the instrument. The extruded boss feature on yaw pulley hub is intact, but bent/deformed. Engineering has concluded that the breakage may be due to overloading the tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1932887 | 1903774 | 1800770 | 2955842201000502 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-20 | 2010-11-20 | No | Yes | Health Professional | 40018304 | Device Returned To Manufacturer | No | 2010-10-20 | Device Not Retu | No | Reuse | 2010
42018405 | 2010-10-22 | Malfunction |  | It was reported that during a da vinci si thymectomy procedure, a piece of the plastic wheel on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor cap is missing from the distal end; however, the cap was not returned with instrument. The conductor wire is exposed between the distal clevis and yaw pulley as a result of the damage. The yaw pulley extruded boss feature that mates with the cap is broken off, which allowed the cap to slip out. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1933007 | 1903894 | 1800890 | 2955842201000504 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-10-22 | 2010-11-20 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2010-10-22 | Device Not Retu | No | Reuse | 2010
50028106 | 2010-11-19 | Injury | Required Intervention | The bovie cautery unit cover had a defect in it. When the cautery was activated it arced through the defect causing a laceration of the patient's ileac vein. The vein was repaired. The cover device was inspected prior to use and no tears or punctures were noted, however two additional events that did not result in patient harm occurred the day before this event with the same device. One looked like a small tear or puncture, and the other looked like a small slit in the material. The packaging was not retained, however, the lot # from the storage area for the same product was c10222. ======================health professional's impression======================defect in cautery cover caused patient injury when cautery device activated. |  | DAVINCI TIP COVER ACCESSORY | DAVINCI TIP COVER ACCESSORY | 1944143 | 1915124 | 1812120 | 1915124 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-19 | 2010-11-19 | Yes | No | Physician | 50028106 | Yes | No Answer Provided |  |  |  |  | 2010
IS2000 A51 P3 | 2010-10-18 | Other | Other | It was reported that during a da vinci s sigmoid colectomy procedure, the site experienced a battery fault causing the system to switch to battery back up. The system also could not be connected to the patient side cart (psc) and a red fiber cable error was displayed. Prior to calling isi for technical support, the surgeon made the decision to convert the surgery to traditional open techniques to complete the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the reported issues experienced by the customer were associated with the system power cord not being property seated into the system and a dirty red fiber optic cable. The power cord provides power to the system. The red system cable connects the surgeon console and the patient side cart and passes video, audio, and data during system operations. The system was repaired by properly seating the power cord and cleaning the fiber cable. As of november 16, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1930931 | 1901883 | 1798879 | 2955842201000501 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-18 | 2010-11-17 | Yes | Yes | Health Professional | IS2000 A51 P3 | Yes | No | 2010-10-18 | No | No | Reuse | 2010
42019408 | 2010-10-18 | Malfunction |  | It was reported that while stitching the vaginal cuff during a da vinci s hysterectomy procedure, the surgeon noticed that a piece of the mega needle driver instrument was missing when trying to hold the needle. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | (b)(4) the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the carbide insert had broken off one of the grips. The broken piece was returned with the instrument. No other damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1928457 | 1898735 | 1795731 | 2955842201000499 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-19 | 2010-11-15 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2010-10-19 | Device Not Retu | No | Reuse | 2010
420183 | 2010-08-23 | Malfunction |  | On october 20, 2010, intuitive surgical received a copy of medwatch (b)(4) with the following information: title: xxxxx event desc: during pediatric robotic laparoscopic nissen fundoplication procedure, the davinci permanent cautery hook sleeve cracked and a section broke off. The piece was retrieved and the instrument removed from the field. The hook had been used eight times. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | On november 8, 2010, the initial reporter informed intuitive surgical that the affected instrument could not be located. If located, the instrument will be returned to intuitive surgical for evaluation. Since the instrument has not been returned for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1928476 | 1898741 | 1775182 | 2955842201000500 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-10-20 | 2010-11-15 | No | Yes | Health Professional | 420183 | No | No | 2010-10-20 | No | No | Reuse | 2010
400169 | 2010-11-10 | Malfunction |  | End piece of device broke off. |  | DA VINCI HARMONIC CURVED SHEARS INSERT | ENDOWRIST INSTRUMENT FOR DAVIN | 1939787 | 1910900 | 1807896 | 1910900 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-15 | 2010-11-15 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2010
400169 | 2010-11-10 | Malfunction |  | End piece broke off and was retrieved. |  | DAVINCI HARMONIC CURVED SHEARS INSERT | INSTRUMENT ULTRASONIC SCALPEL | 1939789 | 1910922 | 1807918 | 1910922 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-15 | 2010-11-15 | No | Yes | Physician | 400169 | Yes | No Answer Provided |  |  |  |  | 2010
42018405 | 2010-08-24 | Malfunction |  | On october 15, 2010, maude event report (b)(4) was received: pt seen in operating room by dr for frozen d&c, robotic lah and, soon after the beginning of the procedure, the cautery spatula tip inside the pt's abdomen broke off. A search for all the broken pieces resumed and were retrieved by dr, surgery proceeded as planned. No injury to the patient was noted. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1927569 | 1897732 | 1794728 | 2955842201000488 | 1 | NAY | Manufacturer | OtherUser facility | Nurse | Initial | 2010-10-15 | 2010-11-12 | No | Yes | Health Professional | 42018405 | No | Yes | 2010-10-15 | No | No | Unkown | 2010
42017909 | 2010-07-30 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The fragment was removed and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- (b)(4) the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found a piece of the left scissor blade to be broken off from the instrument; the broken piece was returned. The blade fractured at the cross section of the grip cable crimp resulting in half of the cable crimp to be exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1927574 | 1897735 | 1794731 | 2955842201000491 | 1 | NAY | Manufacturer | DistributorOther | Physician | InitialFollowup | 2010-11-16 | 2010-11-12 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-11-16 | Device Not Retu | No |  | 2010
40016907 | 2010-07-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the harmonic curved shears insert would not hold on to the harmonic curved shears instrument and fell into the patient. The insert was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to have a broken curved blade and clamp arm. The broken surface of the blade is flat with a small chip, evidence that external force to the blade tip was applied. The clamp arm joint showed evidence of over extension. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1927576 | 1897737 | 1794733 | 2955842201000490 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-13 | 2010-11-12 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-13 | Device Not Retu | Yes | Initial | 2010
40016907 | 2010-10-11 | Malfunction |  | It was reported that during a da vinci liver resection procedure, the harmonic curved shears instrument, with the harmonic curved shears insert installed, was making a clicking noise. At this time it was noted that the foam tip fell into the patient. The fragment was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with damage to the teflon pad, referred to by the customer as the foam tip. The teflon pad and curved blade were found to exhibit burn marks. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1927573 | 1897757 | 1794753 | 2955842201000489 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-14 | 2010-11-12 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-14 | Device Not Retu | Yes | Initial | 2010
40016907 | 2010-10-15 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s myomectomy procedure, the harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- It was reported that during a da vinci s myomectomy procedure, the harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued and the replacement harmonic curved shears insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the third harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient and replaced. The procedure continued using the fourth harmonic curved shears insert, however, the replacement insert broke and fell into the patient. The insert was retrieved from the patient, replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following med watch reports were sent to the fda in relation to this event: 2955842-2010-00506 2955842-2010-00509 2955842-2010-00507. | Manufacturer Narrative: --/-- The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to be returned with the clamp arm detached from the distal end. The shaft feature that mates with the clamp arm pins at the distal end are bent. The insert showed evidence of overloading and or hyperextension of the clamp arm resulting in the breakage. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2010, (b)(6) stated that 4 harmonic curved shear (hcs) inserts broke during the myomectomy case on (b)(6) 2010. She stated that the surgeon was torquing the jaws of the hcs insert to tear tissue away from the myoma and using a hooking motion on the bottom which was causing the hcs inserts to break. | HARMONIC CURVED SHEARS INSERT | Initial | 1927554 | 1897758 | 1802063 | 2955842201000492 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2010-11-18 | 2010-11-12 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-11-18 | Device Not Retu | Yes |  | 2010
IS1200 A43P9 | 2010-10-14 | Other | Other | It was reported that prior to using the da vinci surgical system to perform any tasks for a prostatectomy procedure, the surgical staff experienced non-recoverable system error code #207. The system was restarted multiple times, however, system error code #207 continued to recur. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure and reschedule it to a later date. No patient harm was reported. | Conclusions - the investigation conducted by the field service engineer concluded that system error code #207 was associated with a multiple servo driver (msd) printed circuit assembly (pca) board. The system contains a msd pca for each system arm, which provides the drive current to the servo motors associated with each degree of freedom of a system arm. The system was repaired by replacing the affected msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #207 occurs when a voltage fault is detected on the main board power of a msd pca. Upon determining this condition, the safety systems put the da vinci in a "non-recoverable safe state. " the msd pca was returned to isi for failure analysis investigation. Engineering discovered that an axis was out of calibration, thus generating the system error codes experienced by the customer. As of november 10, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1948032 | 1902709 | 1799705 | 2955842201000496 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-14 | 2010-11-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2010-10-14 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-10-13 | Other | Other | It was reported that during the beginning of a da vinci s prostatectomy procedure, the patient side card (psc) would not power on. The site attempted to troubleshoot the issue by checking cable connections, applying emergency power off (epo) to the system; however, the issue persisted. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to perform the planned procedure. No patient harm was reported. | Conclusions - the investigation conducted by field service engineering (fse) found the customer reported complaint to be associated with the patient -side power distribution board (ppd). The ppd distributes power and signals to components in the patient side cart (psc). The system was repaired by replacing the affected ppd. As of november 11, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1947545 | 1902710 | 1799706 | 2955842201000498 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2010-10-13 | 2010-11-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-10-13 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2010-10-15 | Other | Other | It was reported that prior to using the da vinci surgical system to perform any surgical tasks for a prostatectomy procedure, the surgical staff experienced multiple occurrences of system error code #21008. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No patient harm was reported. | Conclusions - the investigation conducted by the field service engineer concluded that system error code #21008 was associated with the right master tool manipulator (mtmr) gimbal. The mtm gimbal is a subsection of the complete mtm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #21008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a "recoverable safe state. " the mtmr gimbal was returned to isi for evaluation. Engineering discovered that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of november 10, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1948024 | 1902722 | 1799718 | 2955842201000497 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-15 | 2010-11-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2010-10-15 | Yes | No | Reuse | 2010
420183 | 2010-11-01 | Malfunction | Other | Davinci robot permanent cautery hook was in use for the 10th time. Half of the plastic housing broke off in the patient. It is yellow, and the doctor was unable to find and remove the small yellow part that came off. Dates of use: (b)(6) 2010. Diagnosis or reason for use: total hysterectomy. |  | INTUITIVE SURGICAL | PERMANENT CAUTERY HOOK | 1936610 | 1902912 | 1799908 | MW5018220 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-11-08 | 2010-11-12 | No | Yes | Health Professional | 420183 | Yes | Yes |  |  |  |  | 2010
42017206 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tips on the maryland bipolar forceps instrument were observed to be offset. No pt harm, adverse outcome, or injury to the pt was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of grips. There is a. 015" offset at the tips. The bent grip has a slight separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering also observed charring on a yaw pulley. The side of one yaw pulley exhibits a char mark and has localized melting adjacent to the outer grip surface. The other yaw pulley does not exhibit any charring. While the evidence is inconclusive, engineering concluded that charring observed on the instrument indicates that unintended arcing may have occurred. The instrument passed electrical continuity testing. No other damage was found. On (b)(6) 2010, it was reported by the initial reporter that an arcing event during the surgical procedure did not occur. | MARYLAND BIPOLAR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1948375 | 1903074 | 1800070 | 2955842201000494 | 1 | NAY | Manufacturer | User facilityOther | Nurse | Initial | 2010-10-13 | 2010-11-12 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2010-10-13 | Yes | No | Reuse | 2010
42011704 | 2010-10-13 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure, the console surgeon noticed that the "beak" of the 5mm needle driver instrument wouldn't close all the way. The pt side surgeon removed the instrument to reseat it and when the instrument was re-inserted through the trocar, one side of the instrument "beak" fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one needle driver grip (which the customer referred to as the "beak") detached from the distal end. The grip had fractured at the junction between the grip arm and hub where there is a 90 degree angle. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1948392 | 1903075 | 1800071 | 2955842201000493 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-13 | 2010-11-12 | No | Yes | Health Professional | 42011704 | Yes | Yes | 2010-10-13 | Yes | No | Reuse | 2010
42017909 |  | Malfunction |  | It was reported that prior to using the monopolar curved scissors instrument in a da vinci si hysterectomy procedure, it was noticed that the insulation was bent at the location where the joints meet. No pt harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute an fda reportable event, however, engineering discovered evidence of arcing during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube. The entire tube extension is circumferentially broken at the base of the axial keys and as a result, the tube extension can be rotated 360 degrees. No pieces are missing. Engineering also observed char marks on the outer surface of the distal end of the main tube. The instrument was disassembled to find additional char marks on the hypotubes and on the inner surface of the metal reinforcement ring. The char marks leads engineering to conclude that the instrument may have arced during a previous surgical procedure and the arcing most likely occurred after the tube extension breakage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1948399 | 1903083 | 1800079 | 2955842201000495 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-10-13 | 2010-11-12 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-10-13 | Yes | No | Reuse | 2010
| 2010-11-05 | Malfunction |  | Intuitive robotic maryland bipolar forceps had a frayed wire from the gear by the tip. |  | MARYLAND BIPOLAR FORCEPS | ENDOWRIST BIPOLAR CAUTERY INST | 1936164 | 1906879 | 1803875 | 1906879 | 1 | NAY | User Facility |  |  | Initial | 2010-11-12 | 2010-11-12 | No | Yes | UNKNOWN |  |  | No Answer Provided |  |  |  |  | 2010
| 2010-11-08 | Malfunction |  | Intuitive surgical robotic monopolar scissor (endowrist) instrument did not cauterize when used attached to robot. Cautery cord used and functioned properly for another instrument. Scissor tip changed and still did not cauterize. |  | HOTSHEARS MONOPOLAR CURVED SCISSOR | ENDOWRIST MONOPOLAR CAUTERY IN | 1936165 | 1906880 | 1803876 | 1906880 | 1 | NAY | User Facility |  |  | Initial | 2010-11-12 | 2010-11-12 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-10-10 | Other | Other | It was reported that during the beginning of a da vinci si surgical procedure, the site experienced error code #23017 on the endoscopic camera manipulator (ecm). No surgical tasks had been performed at this time; however, the pt was under anesthesia when the surgeon decided to perform the surgery using traditional open techniques. The pt was reported as recovering well with no post operative complications. | Conclusions: the investigation conducted by field service engineering found system error code #23017 and #23008 in the system error logs. It was determined that these errors were associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state. " system error code #23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specific tolerance for agreement. Upon determining this condition, the safety systems put davinci in a "recoverable safe state. " the ecm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and replaced the motor encoder to repair the ecm. As of november 08, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1960624 | 1912866 | 1809862 | 2955842201000487 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-10-10 | 2010-11-09 | Yes | No | Health Professional | IS3000 A60P4 | Yes | No | 2010-10-10 | Yes | No | Reuse | 2010
| 2010-10-18 | Malfunction |  | Surgeon was stitching vaginal cuff and did not see fragment break off, but noticed piece missing when he tried to hold needle. It was part of the jaw. |  | ENDO WRIST NEEDLE DRIVER | MEGA NEEDLE DRIVER | 1936722 | 1907637 | 1804633 | 1907637 | 1 | GDF | User Facility |  | RISK MANAGER | Initial | 2010-11-05 | 2010-11-08 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2010-10-07 | Other | Other | It was reported that during the beginning of a da vinci s prostatectomy procedure, the surgical staff experienced system error code 23017 on a endoscopic camera manipulator(ecm) arm. The site recycled power to the system and when powering back on, system error 25511 was experienced. The site undocked the ecm, power system down and exercised the ecm axis. The system powered up and homed successfully. The customer attempted to dock the ecm and the system immediately faulted with system error code 23017. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found system error code 23017 and 25511 to be associated with a endoscopic camera manipulator (ecm) arm. Endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. Error code 25511 occurs on startup when the system runs a diagnostic test on the current control loop (which regulates the electric current driving motors) and determined it to be out of specification. When this diagnostic fault is triggered, the system transitions to a non-recoverable safe state. The ecm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the ecm to have a defective motor/encoder assembly which was replaced to repair the ecm. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1922171 | 1891761 | 1787377 | 2955842201000482 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-07 | 2010-11-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-10-07 | Device Not Retu | No | Reuse | 2010
|  | Malfunction |  | It was reported that approximately (b)(6) ago, during a da vinci hysterectomy procedure, the pad on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was immediately retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1922283 | 1891816 | 1788812 | 2955842201000485 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-10-06 | 2010-11-05 | No | Yes | Health Professional |  | No | Yes | 2010-10-06 | No | No | Unkown | 2010
|  | Malfunction |  | It was reported that approximately (b)(6) ago, during a da vinci hysterectomy procedure, the pad on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was immediately retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1922286 | 1891819 | 1788815 | 2955842201000484 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-10-06 | 2010-11-05 | No | Yes | Health Professional |  | No | Yes | 2010-10-06 | Device Not Retu | No | Unkown | 2010
42018304 | 2010-10-07 | Malfunction |  | It was reported that during a da vinci s adrenalectomy procedure, a piece of plastic from the distal end of the permanent cautery hook instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a broken distal clevis pin, thus causing the disc to fall off. The clevis and pulleys do not have an arc track or exhibit damage. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1922265 | 1891831 | 1788827 | 2955842201000483 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-10-07 | 2010-11-05 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2010-10-07 | Device Not Retu | No | Reuse | 2010
|  | Malfunction |  | It was reported that approximately two years ago, during a da vinci hysterectomy procedure, the pad on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was immediately retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1922309 | 1891835 | 1788831 | 2955842201000486 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-10-06 | 2010-11-05 | No | Yes | Health Professional |  | No | Yes | 2010-10-06 | No | No | Unkown | 2010
IS3000 A60P4 | 2010-08-26 | Injury | Other | It was reported that prior to starting a da vinci si hysterectomy procedure, the site received a vision system error. The isi representative on site began to troubleshoot and found the left eye on the surgeon side console was not working. The surgeon and or staff were notified of the system status; however, the surgeon decided to perform the case with one eye working in 2d vision. During the procedure the patient's ureteral was injured and a urologist was called in to perform the repair and place a stent. At this time the case was converted to traditional open surgical techniques to repair the patient's ureteral and complete the planned procedure. The patient was required to stay in the hospital an additional 24-48 hours due to the open incision as well as the stent removal. On (b)(4) 2010, intuitive surgical received maude event report (b)(4) for this event. | The investigation conducted by field service engineering found the camera cable had a bad left eye channel. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon side console. The system was repaired by replacing the defective camera cable. As of november 4, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1922371 | 1891889 | 1788885 | 2955842201000481 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-08-26 | 2010-11-05 | Yes | Yes | Health Professional | IS3000 A60P4 | Yes | No | 2010-08-26 | Device Not Retu | No | Reuse | 2010
400169 | 2010-11-03 | Malfunction |  | Could not get shears to open. |  | HARMONIC CURVED SHEARS INSERT DA VINCI | ENDOWRIST INSTRUMENT FOR DAVIN | 1932198 | 1903244 | 1800240 | 1903244 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-05 | 2010-11-05 | No | Yes | Invalid Data | 400169 | Yes | No Answer Provided |  |  |  |  | 2010
420184 | 2010-08-27 | Malfunction |  | Surgeon noticed a piece of metal inside patient's abdomen. Further investigation noted that part of the robotic permanent cautery spatula instrument was cracked and missing a piece. The piece was removed from the abdomen and taken off the field. ======================health professional's impression======================unknown. |  | PERMANENT CAUTERY SPATULA | ENDOWRIST INSTRUMENT FOR DAVIN | 1935395 | 1906183 | 1803179 | 1906183 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-05 | 2010-11-05 | No | Yes | Physician | 420184 |  | No Answer Provided |  |  |  |  | 2010
400180 | 2010-10-06 | Other | Other | It was reported that approximately 2 to 3 hours into a da vinci si hysterectomy procedure, during node dissection, arcing was observed coming from the tip of the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The distal end of the instrument was lying on the patient's bowel at the time of the arcing, resulting in a burn to the patient's bowel. The surgeon repaired the burn with stitches and no additional complications were reported. The planned surgical procedure was completed and no adverse outcome was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory is returned or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1920980 | 1891341 | 1788337 | 2955842201000480 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-10-06 | 2010-11-04 | Yes | Yes | Health Professional | 400180 | No | No | 2010-10-06 | No | Yes | Unkown | 2010
DAVINCI S HD | 2010-10-20 | Injury | Required Intervention | Pt arrived to operating room for davinci assisted hysterectomy. The procedure was started and the robot was docked. Early into the procedure, the robot had a non-recoverable fault. Intuitive surgical technical support was called. The fault was not able to be recovered by technical support. The robotic portion of the procedure had to be aborted, and the rest of the case had to be done laparoscopically without the robot. The pt did not have any direct harm. |  |  | DAVINCI ROBOT | 1926994 | 1894027 | 1791023 | MW5018056 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-10-26 | 2010-11-04 | No | Yes | Health Professional | DAVINCI S HD | Yes | Yes |  |  |  |  | 2010
| 2010-10-25 | Malfunction |  | The davinci tip cover accessory had holes in it. This was discovered prior to use, and the lot was pulled from the shelf and returned to the manufacturer for credit. Had it been used, the potential existed for there to be arcing. |  | DAVINCI TIP COVER ACCESSORY | TIP COVER ACCESSORY | 1924656 | 1894884 | 1791856 | 1894884 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-11-04 | 2010-11-04 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2010-10-04 | Other | Other | It was reported that during the beginning of a da vinci s prostatectomy procedure, the surgical staff experienced system error code 23017 on a patient side manipulator (psm) arm. The site restarted the system multiple times without success. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found system error code 23017 to be associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the psm to have a defective motor/encoder assembly which was replaced to repair the psm. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1920035 | 1890381 | 1787377 | 2955842201000479 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-10-04 | 2010-11-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-09-19 | Device Not Retu | No | Reuse | 2010
400180 | 2010-08-30 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1918871 | 1888565 | 1785561 | 2955842201000476 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-10-05 | 2010-11-02 | No | Yes | Health Professional | 400180 | No | No | 2010-10-05 | No | Yes | Unkown | 2010
IS1200 A43P9 | 2010-10-06 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site experienced a system error code 20008. The site attempted to restart the system but the system would not power back on. With the assistance of an isi technical support engineer (tse), the site emergency powered off (epo) the system and restarted, exercised the effected patient side manipulator (psm) arm, reseated the psm sterile adapter, and power cycled the system, however, the system error code 20008 continued to occur. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering found that system error code 20008 experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The 20008, system error code appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering evaluation found that the potentiometer malfunctioned, thus generating the system errors experienced by the customer. As of (b)(4) 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1918872 | 1888582 | 1785578 | 2955842201000477 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-06 | 2010-11-02 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-10-06 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-10-06 | Other | Other | It was reported that during a da vinci si hysterectomy procedure the site was unable to get the bipolar instrument to cauterize. The isi representative onsite stated that the electrical surgical unit (esu) was responding to the surgeon side console foot pedal, however the customer stated that the instrument would not fire. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) was unable to replicate the customer reported complaint. The da vinci si system was tested with the bipolar instruments, esu, and cables used during the procedure and no trouble was found. All devices responding when cauterization was activated appropriately. As of november 2, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1918855 | 1888586 | 1785582 | 2955842201000478 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-06 | 2010-11-02 | Yes | No | Health Professional | IS3000 A60P4 | Yes | No | 2010-10-06 | Device Not Retu | No | Reuse | 2010
| 2010-09-21 | Malfunction |  | Pt had da vinci assisted supracervical hysterectomy with right salpingo-oophorectomy. Halfway through the cautery of the cervix, there was a sudden smoking from the monopolar cautery cord where it attached to the instrument outside of the pt's body. Immediately all cauterization was stopped and the instrument and the cord were replaced. There was issue with any thermal damage to the pt's skin or tissues internally. The smoke was limited to the cord where it attached to the instrument. |  | DA VINCI MONOPOLAR CURVED SCISSORS | DA VINCI MONOPOLAR CURVED SCIS | 1927099 | 1894080 | 1791076 | MW5018072 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-11-02 | 2010-11-02 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2010
IS3000 A60P5 | 2010-09-30 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff was unable to get the pk dissecting forceps instrument to work. The instrument, cables, and electrosurgical unit were swapped, however, the issue continued to recur. The site indicated that there were no system errors or alarms generated from this issue. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | An intuitive surgical field service engineer (fse) went onsite and evaluated the system. The fse conducted mechanical and performance testing and found the system to be fully functional, without defect, and performing to specifications. The root cause of the customer reported failure mode cannot be determined. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1916753 | 1886212 | 1783208 | 2955842201000469 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-09-30 | 2010-10-29 | Yes | Yes | Health Professional | IS3000 A60P5 | Yes | No | 2010-09-30 | Device Not Retu | No | Reuse | 2010
40018001 |  | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, a hole was observed in the 8mm monopolar curved scissors (mcs) tip cover accessory. The customer reported malfunction does not in itself constitute an fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the mcs tip cover accessory. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating that an arcing event occurred. The hole size is approximately. 020 and the hole center is located. 195 above the silicone to sleeve interface. The hole appears to have a small tear on one edge. The silicone also exhibits a. 125 long mechanical tear through the entire wall thickness in the general vicinity of the hole. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1916819 | 1886226 | 1783222 | 2955842201000471 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-10-07 | 2010-10-29 | No | Yes | Health Professional | 40018001 | Device Returned To Manufacturer | No | 2010-10-07 | Device Not Retu | Yes | Unkown | 2010
IS3000 A60P5 | 2010-10-07 | Other | Other | Event Description: --/-- On (b)(4) 2011, additional information was provided to intuitive surgical by the patient's legal representative. The patient is alleging that significant remnants of the thyroid tissue were left behind during the da vinci si thyroidectomy procedure. As a result, the patient is now experiencing additional illnesses for which treatment is ongoing and the full residual effects are unknown at this time. | Manufacturer Narrative: --/-- The investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the surgeon side console (ssc) power supply. The ssc power supply converts the ac voltage from the wall outlet to dc voltage for use by the system. The system was repaired by replacing the affected power supply. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital.It was reported that during a da vinci si thyroidectomy procedure, the surgeon console lost power. The site turned the system off and upon restart a surgeon console not connected message was displayed. A technical support engineer attempted to troubleshoot the issue via telephone and had the site check the cable connections, cycle the breaker switches on the surgeon console, and emergency power off the system. These actions were unable to resolve the issue. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2010, additional information was provided by the surgeon who performed the surgical procedure. The surgeon indicated that two post operative examinations were performed on the patient and the examinations revealed that the patient was recovering well. The surgeon also mentioned that antibiotics were prescribed during one of the examinations, however, the surgeon would not provide any details as to the reason why. At a later date, the patient informed the surgeon that they had developed an infection and was being treated by another physician for the infection. The surgeon requested that the patient see them for treatment of the infection, however, the patient did not follow-up with the surgeon. The surgeon has no additional information concerning the patient's current condition. --/-- Additional Manufacturer Narrative: --/-- DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL SYSTEM | 1916821 | 1886228 | 1783224 | 2955842201000470 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 10072010 | 2010-10-29 |  | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 8242011 |  | No | Reuse |  | 0
42014304 | 2010-09-24 | Malfunction |  | It was reported that during a da vinci s gynecological procedure a piece of the 5mm maryland dissector instrument tip broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one instrument grip has a. 120 x. 075 piece broken off at the tip. The tip fractured at the root of one of the teeth and the outer surface of the grip exhibits light scratches. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1916817 | 1886252 | 1783248 | 2955842201000470 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-04 | 2010-10-29 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2010-10-04 | Device Not Retu | No | Reuse | 2010
42018405 | 2010-10-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgeon observed that the permanent cautery spatula (pcs) instrument tip was stuck inside the prograsp instrument. The surgeon used gentle pressure to separate the two instruments and then noticed that the insulating tip of the pcs instrument had broken off into multiple pieces inside of the patient's abdomen. The pcs instrument was immediately removed and the broken pieces were retrieved. The planned surgical procedure was completed with a replacement instrument of the same type and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the ceramic sleeve at the grip tip had completely broken off into several pieces. The broken pieces were returned with the instrument. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instrument. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1916820 | 1886253 | 1783249 | 2955842201000472 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-10-07 | 2010-10-29 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2010-10-07 | Device Not Retu | No | Reuse | 2010
42018405 | 2010-09-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed a foreign object within the surgical field. The object was removed and inspected. After inspection, the object was determined to be the yellow insulation covering from the permanent cautery spatula instrument. The procedure was successfully completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a broken distal clevis ear on one side. No arc trac was found on the clevis and pulleys indicating energy did not unintentionally escape. Electrical continuity test passed. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1916988 | 1886285 | 1783281 | 2955842201000474 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-30 | 2010-10-29 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | Yes | 2010-09-30 | Device Not Retu | No | Reuse | 2010
42018304 | 2010-08-31 | Malfunction |  | It was reported that during common bile duct exploration procedure, while dissecting gallbladder tissue the surgeon observed tissue sticking on the tip bottom of the permanent cautery hook instrument. The surgeon requested that the surgical staff remove the instrument and upon inspection of the instrument, it was noted by the rn that the yellow part on the instrument could conduct electricity the planned surgical procedure was completed with an instrument of the same type and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the yaw pulley, distal clevis and proximal clevis exhibit charring on the same side. The yaw pulley extruded boss feature that retains the conductor cap is broken off and the conductor cap is missing, thus exposing the conductor wire. The wire exhibits damage to the insulation and the bare wire is exposed. Engineering concluded that the damage to the instrument most likely contributed to unintended arcing. The instrument passed electrical continuity testing. No other damage was found. Upon follow up with the initial reporter, it has been confirmed that during the surgical procedure, no components from the permanent cautery hook instrument fell into the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1916989 | 1886286 | 1783282 | 2955842201000475 | 1 | NAY | Manufacturer | User facilityDistributor | Other | Initial | 2010-10-01 | 2010-10-29 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-10-01 | Device Not Retu | No | Reuse | 2010
420172 07 | 2010-09-27 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the jaws on the maryland bipolar forceps instrument melted. No other information was provided. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip base exhibits damage related to arcing. The yaw pulleys have melted material and the conductor wire cover also exhibits damage, however, there is no damage to the wire. The clevis and grip tips do not have an arc trac. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1914792 | 1884253 | 1781249 | 2955842201000463 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-27 | 2010-10-27 | No | Yes | Health Professional | 420172 07 | Device Returned To Manufacturer | No | 2010-09-27 | Device Not Retu | No | Reuse | 2010
400169 | 2010-09-30 | Malfunction |  | It was reported that during a da vinci thyroidectomy procedure, the harmonic curved shears insert broke and a piece fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Multiple attempts have been made to obtain the accessory back however at the time of this report the accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1914895 | 1884256 | 1781285 | 2955842201000467 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-30 | 2010-10-27 | No | Yes | Health Professional | 400169 | No | Yes | 2010-09-30 | No | Yes | Unkown | 2010
42017206 | 2010-09-27 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the jaws on the maryland bipolar forceps instrument melted. No other information was provided. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | (b)(4) the instrument was returned and evaluated. Per the customer reported complaint, engineering observed melting at the grip base and yaw pulleys. One conductor wire broke off from the pulley. The clevis and cable do not exhibit damage. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1914797 | 1884273 | 1781249 | 2955842201000464 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-27 | 2010-10-27 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-09-27 | Device Not Retu | No | Reuse | 2010
42018905 | 2010-01-12 | Malfunction |  | On (b)(6) 2010, medwatch uf report (b)(4) was received: after a da vinci double fenestrated instrument was removed from the pt (robotic assisted lap assisted vaginal hysterectomy, bilateral salpingo-oophorectomy), the instrument tip broke off and was not detected by x-ray. The surgeon opened, tissue morcellated, no identifiable part found in wound or morcellated tissue. The area was irrigated well and searched - no foreign objects were detected in the wound area. | Multiple attempts have been made to obtain the instrument or a photograph of the instrument; however, at the time of this report, the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1914887 | 1884284 | 1781280 | 2955842201000465 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2010-09-28 | 2010-10-27 | No | Yes | Health Professional | 42018905 | No | No | 2010-09-28 | No | No | Unkown | 2010
400169 | 2010-09-30 | Malfunction |  | It was reported that during a da vinci thyroidectomy procedure, the harmonic curved shears insert broke and a piece fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Multiple attempts have been made to obtain the accessory back however, at the time of this report the accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1914894 | 1884289 | 1781285 | 2955842201000466 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-30 | 2010-10-27 | No | Yes | Health Professional | 400169 | No | Yes | 2010-09-30 | No | Yes | Unkown | 2010
IS2000 A51P8 | 2010-09-28 | Other | Other | It was reported that while preparing for a da vinci s prostatectomy procedure, the site was unable to power on the system. With the assistance of an isi field service engineer (fse) and an isi technical support engineer, the site emergency powered off (epo) the system, rebooted the psc and cleaned the fiber optic cable; however, the issue persisted. The patient was under anesthesia when the surgeon decided to abort the planned procedure and to reschedule later on the same day. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the red system cable. The red system cable connects the surgeon console and the patient side cart and passes video, audio, and data during system operations. The system was repaired by replacing the affected cable. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1914995 | 1884346 | 1781342 | 2955842201000468 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-28 | 2010-10-27 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-09-28 | Device Not Retu | No | Reuse | 2010
IS3000 A60P5 | 2010-09-28 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the customer experienced system error code 31030 and the fiber cable led turned red. With the assistance of an isi technical support engineer (tse) the site replaced the fiber cable and restarted the system, however, the system error code 31030 continued to occur. The tse had the customer disable the affected patient side manipulator (psm) arm and at this time system error code 23 began to appear. The patient was under anesthesia for approximately 1. 5 hours when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes experienced by the customer were associated with a set up joint (suj) and embedded serializer set up joint (essj). The suj is an unpowered linkage which allows the surgical team to position the active patient side manipulator (psm) arm and supports the weight of the psm. The essj is the printed circuit assembly (pca) inside a suj that monitors the sensors for each of the four set-up joints and transmits information between the psm and the rest of the system. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 31030 is reported when the vision cart powers up and the patient side cart (psc) or surgeon side console (ssc) fails to restart. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The suj was returned to isi for failure analysis investigation. Engineering was immediately able to replicate the error 23. In the process of isolating the issue, a patient side manipulator (psm) arm and a remote arm controller (rac) were also replaced. Engineering did not find any issues with the psm and rac. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1911473 | 1880689 | 1777685 | 2955842201000458 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-28 | 2010-10-25 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-09-28 | Device Not Retu | No | Reuse | 2010
IS3000 A60P5 | 2010-10-01 | Other | Other | It was reported that while beginning a da vinci si prostatectomy procedure, the customer experienced a non recoverable system error 25553. With the assistance of an isi technical support engineer (tse), the customer power cycled the system. The system did not error upon restart, however, the surgeon decided to abort the planned procedure. The patient was under anesthesia and the port incisions were made when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes experienced by the customer were associated with the endoscopic camera manipulator (ecm) arm, a master tool manipulator (mtm) arm, a set up joint (suj), and two remote arm controllers (rac). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The mtm refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The setup joint is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected ecm, mtm, suj, and racs. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25553 is reported when the da vinci safety system determines a hardware fault reaction logic occurred on a local rac. System error code 23009 is reported when the da vinci safety systems determined that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The components were returned to isi for failure analysis investigation. Engineering was able to replicate the error 23009 on the ecm. No trouble was found on the mtm, two racs or the suj. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1911438 | 1880690 | 1777685 | 2955842201000459 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-10-01 | 2010-10-25 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2010-10-01 | Device Not Retu | No | Reuse | 2010
IS2000 |  | Death | Death | Event Description: --/-- The surgeon did not specify what type of cancer this patient was diagnosed with; however, did advise that the cancer was neither a t3 or t4. This information is being provided to clarify information provided in a follow up report to the fda on (b)(6) 2011. --/-- Additional Event Description: --/-- On (b)(6), 2011, the initial reporter provided additional information related to this event. He advised that the patient underwent a da vinci s partial laryngectomy procedure and experienced bleeding from the laryngeal artery 7 days post op. The patient had aspirated blood and a secondary surgery was performed to successfully control the bleeding. The patient was intubated and expired 1 to 2 days later due to a chest infection from the aspiration. The initials reporter advised that the patient's death was unrelated to the da vinci s surgical system. --/-- Additional Event Description: --/-- While gathering information for a study, the surgeon was made aware of 3 events that occurred 4 to 7 days post a da vinci s tors procedure where the patient expired secondary to complication related to their cancer. Several attempts have been made to gather additional information without success. For these similar events, please see mdr's 2955842-2010-00460 and 2955842-2010-00462. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The surgeon has indicated that the patient's death was unrelated to the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. As of october 22, 2010, no similar instances of this event have been reported to intuitive surgical. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | Reuse | 1912699 | 1882150 | 1779146 | 2955842201000461 | 1 | NAY | Manufacturer | ForeignHealth Professional | Physician | FollowupFollowup | 2010-09-26 | 2010-10-25 | Yes | No | Health Professional | IS2000 | No | Yes | 2011-04-01 | Device Not Retu | No |  | 2010
IS2000 |  | Death | Death | Event Description: --/-- While gathering information for a study, the surgeon was made aware of 3 events that occurred 4 to 7 days post a da vinci s tors procedure where the patient expired secondary to complication related to their cancer. Several attempts have been made to gather additional information without success. For these similar events, please see mdr's 2955842-2010-00460 and 2955842-2010-00461. --/-- Additional Event Description: --/-- On (b)(6) 2011, the initial reporter provided additional information related to this event. He advised that the patient underwent a da vinci s partial laryngectomy procedure and experienced bleeding from the laryngeal artery 5 days post op. The patient had aspirated blood and a secondary surgery was performed to successfully control the bleeding. The patient was entubated and expired 1 to 2 days later due to a chest infection from the aspiration. The initial reporter advised that the patient's death was unrelated to the da vinci s surgical system. | Manufacturer Narrative: --/-- The surgeon has indicated that the patient's death was unrelated to the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. As of october 22, 2010, no similar instances of this event have been reported to intuitive surgical. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | Reuse | 1912700 | 1882151 | 1779147 | 2955842201000462 | 1 | NAY | Manufacturer | ForeignHealth Professional | Physician | InitialFollowup | 2010-09-26 | 2010-10-25 | Yes | No | Health Professional | IS2000 | No | Yes | 2011-04-01 | Device Not Retu | No |  | 2010
IS2000 |  | Death | Death | Event Description: --/-- While gathering information for a study, the surgeon was made aware of 3 events that occurred 4 to 7 days post a da vinci s tors procedure where the patient expired secondary to complication related to their cancer. Several attempts have been made to gather additional information without success. For these similar events, please see mdr's 2955842-2010-00461 and 2955842-2010-00462. --/-- Additional Event Description: --/-- On (b)(6) 2011, the initial reporter provided additional information related to this event. He advised that the patient had a history of pulmonary and cardiac problem to surgery and refused a tracheotomy post da vinci s partial laryngectomy procedure. The patient expired approximatley 1 month post op secondary to a pulmonary infection due to the aspirations. The initial reporter advised that the patient's death was unrelated to the da vinci s surgical system. | Manufacturer Narrative: --/-- The surgeon has indicated that the patient's death was unrelated to the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. As of october 22, 2010, no similar instances of this event have been reported to intuitive surgical. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | Reuse | 1912703 | 1882154 | 1779150 | 2955842201000460 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2010-09-26 | 2010-10-25 | Yes | No | Health Professional | IS2000 | No | Yes | 2011-04-01 | Device Not Retu | No |  | 2010
IS2000 | 2010-09-13 | Death | Death | Event Description: --/-- It was reported that during a da vinci s gastrectomy procedure performed on (b)(6) 2010, the surgeon noticed that the patient's pancreas was damaged. The surgeon repaired the damaged pancreas and a drain was placed around the pancreas. The planned surgical procedure was completed with no system malfunctions. Post-operatively, the patient was diagnosed with pancreatic fistula and his blood hepatic enzyme level was high. On the evening of (b)(6) 2010, a ct scan was performed and it was discovered that the upper intestinal membrane artery was blocked and bowel ischemia was found around the perfused area of the upper intestinal membrane artery. The patient underwent another surgical procedure and a drain was again placed around the pancreas. A gastrostomy tube and duodenal fistula tube were also placed. While in icu, the patient's metabolic acidosis did not improve and his bowel color was getting worse. On (b)(6) 2010, the patient's abdomen was re-opened and his 4m small bowel and right colon were resected. Necrotizing fasciitis occurred after the surgery. The blood poisoning could not be controlled and the patient passed away on (b)(6) 2010. --/-- Additional Event Description: --/-- It was reported that during a da vinci s gastrectomy procedure performed on (b)(6) 2010, the surgeon noticed that the patient's pancreas was damaged. The surgeon repaired the damaged pancreas and the planned surgical procedure was completed with no system malfunctions. Post-operatively, there was occlusion of the patient's superior mesenteric artery and necrosis of the colon occurred. On (b)(6) 2010, the patient underwent another surgical procedure to resect his bowel. Necrotizing fasciitis occurred after the surgery and the patient passed away on (b)(6) 2010, due to multi-organ failure. | Manufacturer Narrative: --/-- The hospital has indicated that the patient death was unrelated to the da vinci s surgical system. Based on the information provided, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred. As of october 19, 2010, no similar instances of this event have been reported to intuitive surgical. A follow-up medwatch report will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- (b)(4). | DA VINCI S SURGICAL SYSTEM | Reuse | 1910944 | 1880162 | 1777158 | 2955842201000452 | 1 | NAY | Manufacturer |  | Other | FollowupInitial | 2010-09-22 | 2010-10-22 | Yes | No | Health Professional | IS2000 | No |  | 2010-10-26 | Device Not Retu | No |  | 2010
IS2000 A51P8 | 2010-09-24 | Injury | Required Intervention | It was reported that during a da vinci s nephrectomy procedure, the surgeon experienced visualization issues related to the patient's anatomy and as a result, the patient's aorta was inadvertently damaged and the surgeon made the decision to convert to open surgical techniques to complete the planned procedure. | Based on the information provided by the attending surgeon, it was determined that the da vinci s surgical system did not cause or contribute to the patient's injury. The hospital has continued to use the system to perform surgery on patients. As of (b)(6) 2010, no adverse events have been experienced with the site's da vinci s surgical system and no similar instances of this event have been reported to isi. An intuitive surgical field service engineer (fse) went onsite and evaluated the system. The fse conducted mechanical and performance testing and found the system to be fully functional, without defect, and performing to specifications. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1910915 | 1880163 | 1777159 | 2955842201000453 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-24 | 2010-10-22 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2010-09-24 | Device Not Retu | No | Reuse | 2010
42017206 | 2010-09-29 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the plastic portion near the tip of the maryland bipolar forceps instrument was burnt when the bipolar energy was first activated. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed an arc track and melted materials at the grip base on the yaw pulleys, indicating that an arcing event occurred. Both conductor wires remained inside the yaw pulley and the grip cables were frayed. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1910951 | 1880165 | 1777161 | 2955842201000454 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2010-09-29 | 2010-10-22 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-09-29 | Device Not Retu | No | Reuse | 2010
42018304 | 2010-09-30 | Malfunction |  | It was reported that while dissecting the myoma from the uterus during a da vinci si myomectomy procedure it was noticed that a piece of the plastic at the distal tip of the permanent cautery hook instrument was missing. The surgical staff inspected the patient's abdomen and pelvis, however, they were unable to locate the missing piece. The procedure was continued and after dissection and repair of the uterus was completed, a second inspection was done with irrigation. The missing piece was located and successfully removed from the patient without incident. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one distal clevis ear of the instrument is broken off at its base. The broken piece is approximately. 300 x. 235 and is located on the same side as the conductor wire and cap. The yaw pulley hub extruded boss feature that holds the cap is broken off and the conductor cap is missing. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1910954 | 1880169 | 1777165 | 2955842201000455 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-09-30 | 2010-10-22 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-09-30 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-09-23 | Malfunction |  | It was reported that toward the end of a da vinci s prostatectomy procedure, and while performing lymph node dissection, one of the blades from the monopolar curved scissors instrument broke off and fell into the patient. The piece was retrieved and the instrument was replaced to successfully complete the procedure with no injury to the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a section of the left blade broken off. The broken piece was not returned. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1911174 | 1880407 | 1777403 | 2955842201000456 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-23 | 2010-10-22 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-09-23 | Device Not Retu | No | Reuse | 2010
40016907 | 2010-09-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The insert accessory was retrieved and the procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found the clamp arm detached from the distal end. The shaft features that mate with the clamp arm pins are bent backwards. Engineering concluded that breakage of the harmonic curved shears (hcs) insert accessory was most likely caused by mishandling or misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1911159 | 1880421 | 1777417 | 2955842201000457 | 1 | LFL | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-23 | 2010-10-22 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-09-23 | No | No | Reuse | 2010
IS2000 A51P8 | 2010-09-21 | Other | Other | It was reported that during a da vinci surgical procedure, the surgical staff experienced a loss of vision on the right eye of the high resolution stereo viewer (hrsv) located on the surgeon side console. The site noticed that the connectors to the camera head appeared to be broken. The patient was under anesthesia when the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the camera head. The system was repaired by replacing the affected camera head. The camera head was returned to isi for evaluation. Engineering evaluation revealed that internal connectors were found to be loose or damaged. The unit was repaired by replacing the connectors. As of october 20, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1910239 | 1879163 | 1776159 | 2955842201000450 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-21 | 2010-10-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-09-21 | Device Not Retu | Yes | Reuse | 2010
IS1200 A43P9 | 2010-09-22 | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure, the surgical staff experienced recurring system error message excessive force. With the assistance of an isi clinical sales representative (crs) and technical support engineer (tse) the site undocked the system from the patient and restarted the system. At this time, the site experienced system error code 212. The site then powered down the system, reset the breaker and applied emergency power off (epo) several times with no success. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error message excessive force and system error code 212 were associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The excess force system message has an associated system error code, 10025. System error code 10025 appears when the da vinci safety system detects a failure to enable amps. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The engineering analysis of the returned mtmr confirmed that the issue was caused by a defective harness. The mtmr was repaired by replacing the harness. As of may 20, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1910241 | 1879169 | 1776165 | 2955842201000451 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-09-22 | 2010-10-21 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-09-22 | Device Not Retu | No | Reuse | 2010
42030901 | 2010-09-20 | Malfunction |  | It was reported that while sewing the vaginal cuff closed during a da vinci si sacrocolpopexy procedure the metal wire threads from the wrist of the mega suturecut needle driver instrument frayed and snapped. The incident was not observed by the surgical staff and the instrument damage was noticed after the breakage had already occurred. Part of the metal threads fell inside of the patient and some were removed, however, the surgical staff is unsure if all fragments were retrieved. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close cable broken near the proximal pulleys and proximal clevis cable hole. As a result, the idler pulley spins freely and the cable segment with the frayed strands sticks out at the wrist. The proximal clevis cable hole exhibits wear on one edge. The other cables at the wrist are not damaged. No other damage was found. Intuitive surgical 8mm endowrist instruments for use on the da vinci si surgical system have been tested for material biocompatibility. The materials in these instruments have been tested for biocompatibility in accordance with (b)(4) for their intended use (blood path, indirect (4. 2. 3. A) and limited exposure (4. 3. A)). | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1907691 | 1876690 | 1773686 | 2955842201000446 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-21 | 2010-10-20 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2010-09-21 | Device Not Retu | No | Reuse | 2010
42022703 | 2010-09-20 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, after removal of the pk dissecting forceps instrument from the patient, it was observed that tip of the instrument was melted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal clevis is melted and has a light char mark at the grip base. The clevis and pulleys do not exhibit any damage and the conductor wires are present. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1908894 | 1877965 | 1774961 | 2955842201000447 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-21 | 2010-10-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2010-09-21 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-09-16 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while taking down the patient's uterus, arcing to the patient's uterus from the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument was observed. The tip cover accessory was replaced to successfully complete the procedure. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1908907 | 1877968 | 1774964 | 2955842201000449 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-21 | 2010-10-20 | No | No | Health Professional | 40018010 | No | Yes | 2010-09-21 | No | Yes | Unkown | 2010
40018302 | 2010-08-24 | Malfunction |  | It was reported that during a da vinci ureteral re-implantation procedure, the surgeon observed arcing and smoke coming from the wrist of permanent cautery hook instrument. The instrument was removed and the surgical procedure was completed with an instrument of the same type. The site also observed that during the surgical procedure, the patient had a small incision in the bladder. The incision was repaired with absorbable sutures. The site was unable to determine how the incision was introduced to the patient's bladder. There was no harm, adverse outcome or injury to the patient as a result of the instrument malfunction. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the base of the distal yaw pulley is melted and has char marks, indicating that an arcing event occurred. The electrical wire is also exposed. The instrument passed electrical continuity testing. No other damage was found. Based on the information provided by the site, it has been determined that the damage to the patient's bladder was unrelated to the arcing event, as arcing related injuries would exhibit burn and char marks. Per the initial reporter, there was no patient harm, adverse outcome or injury as a result of the arcing event. The patient was released from the hospital on (b)(6) 2010 and has not returned to the hospital due to any post operative complications. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1908877 | 1878001 | 1774997 | 2955842201000448 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-21 | 2010-10-20 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | Yes | 2010-09-21 | Device Not Retu | No | Reuse | 2010
400180 | 2010-09-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci hysterectomy procedure arcing was observed through the side of the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was immediately removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The tip cover accessory was returned and evaluated. The silicone has a. 055 long tear through the entire wall thickness, located approximately. 270 above the interface with the sleeve. The silicone does not have melting around the tear. The outer surface of the silicone also exhibits a couple of scratches/gouges on the opposite side of the tear, located approximately. 045 above the silicone to sleeve interface. No evidence of arcing was observed. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip cover accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument and/or tip cover accessory are returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | Unkown | 1905450 | 1874374 | 1771370 | 2955842201000444 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2010-09-22 | 2010-10-19 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | Yes | 2010-10-28 | Device Not Retu | Yes |  | 2010
IS2000 A51P8 | 2010-09-19 | Other | Other | It was reported that prior to starting a da vinci s hysterectomy procedure, the surgical staff experienced no vision in the right side monitor of the surgeon's console. The patient was under anesthesia when the surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with the transformer located within one of the camera control units (ccu). The ccus control the quality of the video images and produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the transformer of the affected ccu. As of october 19, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1905436 | 1874383 | 1771378 | 2955842201000443 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-09-19 | 2010-10-19 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-09-19 | Device Not Retu | No | Reuse | 2010
420184 | 2010-08-10 | Malfunction |  | On september 20, 2010, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda with the following information: title: xxxxx. Event desc: while a robotic prostatectomy was in progress, the permanent cautery spatula was in use. The tip broke off in the abdomen and had to be retrieved. Device usage problem: other. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned (post-evaluation) or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1905437 | 1874384 | 1771379 | 2955842201000445 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-09-20 | 2010-10-19 | No | Yes | Health Professional | 420184 | No | No | 2010-09-20 | No | No | Unkown | 2010
REF 420209 | 2010-10-06 | Malfunction |  | Surgeon suturing vaginal cervical cuff using davinci robot. The suture cut needle driver suturing device suddenly displayed frayed metal wires in articulating joint of needle driver. Intuitive rep examined tool after incident. ======================health professional's impression======================does not know. |  | DAVINCI S SUTURECUT NEEDLE DRIVER | ENDOWRIST NEEDLE HOLDER | 1912775 | 1882180 | 1779176 | 1882180 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-10-18 | 2010-10-18 | No | Yes | Physician | REF 420209 | Yes | No Answer Provided |  |  |  |  | 2010
42017909 | 2010-07-07 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the nurse found the tip cover accessory installed on the monopolar curved scissors (mcs) instrument to be torn and burned. The mcs instrument was removed and replaced with an instrument of the same type and a new tip cover was installed to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed an arc trac at the distal end of the instrument's tube extension. The arc trac covers both the tube extension and the proximal clevis and is approximately 0. 12 inches by 0. 05 inches in length. No other damage was found. The tip cover accessory used in the procedure was not returned for evaluation. Therefore, the root cause of the customer reported failure mode cannot be determined. If the accessory is returned for evaluation or if additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1902740 | 1871295 | 1768385 | 2955842201000438 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-09-20 | 2010-10-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2010
42018304 | 2010-08-30 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the plastic part on the permanent cautery hook instrument broke off and fell into the patient. The broken piece was retrieved. The planned surgical procedure was completed and no patient harm adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument tip did not have any broken parts. The instrument also passed electrical continuity testing. Engineering also observed that the edge of the proximal clevis is melted and has char marks, indicating that an arcing event occurred. Disassembly of the instrument found that the electrical wire insulation located at the proximal end of the main tube exhibits damage. Engineering concluded that damage to the electrical wire insulation mostly likely caused the instrument to arc. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1902900 | 1871423 | 1768467 | 2955842201000439 | 1 | NAY | Manufacturer | User facilityDistributor | Other | Initial | 2010-09-17 | 2010-10-15 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-09-17 | Device Not Retu | No | Reuse | 2010
42018305 | 2010-08-03 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the plastic portion of the permanent cautery hook instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that instrument distal clevis was broken with a piece missing. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1902908 | 1871425 | 1768469 | 2955842201000440 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-09-17 | 2010-10-15 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2010-09-17 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-08-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, one tip of the monopolar curved scissors instrument broke and fell into the patient. The piece was retrieved and the procedure was successfully completed using a replacement instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a piece of the left blade had broken off and is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1903205 | 1871716 | 1768712 | 2955842201000441 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-09-17 | 2010-10-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-09-17 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-09-16 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the site experienced a patient cart running on battery message displayed on their monitor. With the assistance of an isi technical support engineer, the site verified that the electrical cables were properly plugged into the wall socket and the breaker was in the correct position. The site also removed instruments from the system, applied emergency power off to the system and turned off the breaker on patient side cart before restarting the system. They system message continued. The site re-installed the instruments but could not activate the instruments from surgeon side console. The monitor displayed the message no battery backup. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the surgeon side console (ssc) power supply. The ssc power supply converts the ac voltage from the wall outlet to dc voltage for use by the system. The system was repaired by replacing the affected power supply. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1903208 | 1871744 | 1768753 | 2955842201000442 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-16 | 2010-10-15 | Yes | No | Health Professional | IS3000 A60P4 | Yes | No | 2010-09-16 | No | No | Reuse | 2010
42022702 |  | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, it was observed that the tip of the pk dissecting forceps instrument was bent. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one of the grips was bent and is approximately 0. 25 offset at the tip. While the evidence is inconclusive, engineering concluded that damage to the grip was most likely due to misuse/mishandling. Engineering also observed that the yaw pulley on the instrument exhibits slight charring on the outer surface near the distal clevis, indicating that an arcing event occurred. No other damage was found. On (b)(6), 2010, the initial reporter reported that there was no report by the surgical staff indicating that an arcing event during the surgical procedure occurred. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1901637 | 1869816 | 1767041 | 2955842201000434 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-09-14 | 2010-10-14 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-09-14 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-09-22 | Injury | Other | It was reported that approximately 2. 5 hours during a da vinci si cystectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument and made contact with the patient's bowel which required repair via stitch. The procedure was reported to be completed successfully without further issue. The patient was reported to be recovering as expected and did not sustain any lasting injury or adverse outcome. | The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, however during an (b)(6), 2010 conversation with robotics coordinator (b)(6) stated that intra operative collisions likely occurred during the procedure and was the preliminary cause of the event however this has not been confirmed. If the accessory is returned for evaluation or if more information is obtained a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1901688 | 1869880 | 1767055 | 2955842201000437 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-23 | 2010-10-14 | Yes | Yes | Health Professional | 40018010 | No | No | 2010-09-23 | No | Yes | Unkown | 2010
400180 | 2010-09-15 | Injury | Other | It was reported that approximately 30 minutes during a da vinci s pulmonary lobectomy procedure the site witnessed an arc of energy escape from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The arced energy struck the patient's pulmonary vein and due to excessive bleeding, the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. | The tip cover accessory was disposed of by the hospital due to the patient being hiv positive and the tip cover accessory being a one time use device that cannot be re-sterilized. The isi representative present for the procedure followed up with the surgeon and confirmed that the patient had no related complications is recovering as expected. The surgeon stated that due to the patient's pre-existing health issues the case would have likely been converted to open to complete. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1901656 | 1869882 | 1767068 | 2955842201000435 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-09-15 | 2010-10-14 | Yes | Yes | Health Professional | 400180 | No | No | 2010-09-15 | No | Yes | Unkown | 2010
40016907 | 2010-09-10 | Malfunction |  | Event Description: --/-- It was reported that approximately 30 minutes during a da vinci si hysterectomy procedure the harmonic curved shears insert accessory fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The harmonic curved insert accessory was returned and evaluated. Engineering confirmed the customer reported complaint and found the insert to have a broken clamp arm, with the clamp arm detached from the distal end. The shaft features that mate with the clamp arm are bent and exhibit scratch marks. The clamp arm joint showed evidence of overloading. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory is returned for evaluation a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERTACCESSORY | Initial | 1901665 | 1869892 | 1767075 | 2955842201000436 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2010-10-19 | 2010-10-14 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-10-19 | Device Not Retu | Yes |  | 2010
IS2000 A51P8 | 2010-09-14 | Other | Other | It was reported that while starting a da vinci s hysterectomy procedure, the site was unable to align their endoscope. With the assistance of an isi technical support engineer, the site exchanged the scope, verified the camera controller unit (ccu) settings, re-seated the scope sterile adapter and re-aligned the scope. At this time, the image appeared to be in alignment; however, the site stated that the camera head sterile adapter appeared to be crooked. The patient was under anesthesia for approximately 40 minutes when the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the reported scope mis-alignment issue experienced by the customer was associated with the sterile adapter having been incorrectly seated. The camera arm sterile adapter provides the point of connection for the endoscope to the camera arm. The system was repaired by having the site re-seat the camera arm sterile adapter. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1897946 | 1865628 | 1763498 | 2955842201000433 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-14 | 2010-10-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-09-14 | No | No | Reuse | 2010
IS1200 A43P9 | 2010-09-17 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced recurring system error code 20013. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that system error code 20013 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1898193 | 1865844 | 1763631 | 2955842201000431 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-09-17 | 2010-10-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-09-17 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-09-20 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced system error code 21013 and was unable to home the system. The site exercised the right master tool manipulator, however, system error code 21013 continued to recur. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer determined that system error code 21013 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 21013 appears when the davinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts davinci in a recoverable safe state. The mtmr was returned to isi for failure analysis investigation. Engineering was unable to replicated the customer reported failure, however, an axis potentiometer and motor encoder assembly were replaced as a precaution. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1898194 | 1865845 | 1763632 | 2955842201000430 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-09-20 | 2010-10-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2010-09-20 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-09-16 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, there was no text, icons or images displayed on the right side monitor of the surgeon's console. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. The hrsv was returned to isi for evaluation. Engineering discovered that the right side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1898197 | 1865855 | 1763634 | 2955842201000432 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2010-09-16 | 2010-10-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2010-09-16 | Device Not Retu | No | Unkown | 2010
42017909 | 2010-09-07 | Malfunction |  | It was reported that during a da vinci s surgical procedure, a part of a blade from the monopolar curved scissors instrument broke and fell into the patient. The piece was retrieved and the procedure was successfully completed. | The instrument was returned with the detached blade and evaluated. Per the customer reported complaint, engineering observed that the blade fractured at the cross section of the grip cable crimp. The cable crimp has slipped out from crimp pocket and the fractured plane of the blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1894142 | 1861380 | 1759952 | 2955842201000427 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-07 | 2010-10-08 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-09-07 | Device Not Retu | No | Initial | 2010
40016907 | 2010-09-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the harmonic curved shears insert broke off and fell into the patient. The planned surgical procedure was completed and there was no report of patient harm, adverse outcome or injury. No other details were provided. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. In addition, several attempts have been made to obtain additional information concerning the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1895370 | 1862608 | 1760805 | 2955842201000429 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-08 | 2010-10-08 | No | Yes | Health Professional | 40016907 | No | No | 2010-09-08 | No | Yes | Unkown | 2010
40016907 | 2010-09-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the harmonic curved shears insert broke off and fell into the patient. The planned surgical procedure was completed and there was no report of patient harm, adverse outcome or injury. No other details were provided. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. In addition, several attempts have been made to obtain additional information concerning the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1895404 | 1862628 | 1760805 | 2955842201000428 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-08 | 2010-10-08 | No | Yes | Health Professional | 40016907 | No | No | 2010-09-08 | No | Yes | Unkown | 2010
400169 | 2010-09-08 | Malfunction |  | It was reported that during a da vinci s thymectomy procedure, the tip of the instrument broke off and fell into the patient. The broken piece was retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(6) 2010, isi received medwatch (b)(4) with the following event description: during a robotic thymectomy procedure, while using an intuitive surgical harmonic curved shears insert in the da vinci robot-s, one portion of the shear snapped off into the thoracic cavity, rendering the equipment unusable. Md was able to retrieve 2 small pieces from the shears. The equipment was removed from the table, reported to the intuitive rep, central sterile supply, and a chest x-ray was taken upon completion; no objects seen in patient by chest films. Pt appears to not have any harm from the event. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided in the site's (b)(4), the broken pieces from the insert accessory were successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1893449 | 1860753 | 1759381 | 2955842201000426 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-09 | 2010-10-07 | No | Yes | Health Professional | 400169 | No | No | 2010-09-09 | No | Yes | Unkown | 2010
420207 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while mobilizing the patient's 975g uterus from side to side, the tenaculum forceps instrument broke and 7 or 8 small shreds of black plastic from the instruments sheath fell into the patient's abdomen. The broken pieces were retrieved laparoscopically and the planned procedure was completed vaginally. It was also reported that the instrument had only been in use for 30 minutes prior to the breakage and that a instrument wire used to control the movement had also broke. The surgeon indicated that the instrument broke due of the excess size of the patients uterus. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the information provided by the initial reporter, the broken instrument pieces were successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1893477 | 1860818 | 1759406 | 2955842201000425 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-09-07 | 2010-10-07 | No | Yes | Health Professional | 420207 |  | Yes | 2010-09-07 | No | No | Unkown | 2010
420183 | 2010-08-23 | Malfunction |  | During pediatric robotic laparoscopic nissen fundoplication procedure, the davinci permanent cautery hook sleeve cracked and a section broke off. The piece was retrieved and the instrument removed from the field. The hook had been used eight times. |  | PERMANENT CAUTERY HOOK | ENDOWRIST MONOPOLAR CAUTERY IN | 1909046 | 1878186 | 1775182 | 1878186 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-10-05 | 2010-10-06 | No | Yes | Physician | 420183 | No | No Answer Provided |  |  |  |  | 2010
42018404 | 2010-09-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the pulley disc on the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1885368 | 1850522 | 1750172 | 2955842201000416 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-01 | 2010-10-01 | No | Yes | Health Professional | 42018404 | No | No | 2010-09-01 | No | No | Unkown | 2010
IS3000 A60P4 | 2010-09-03 | Death | Death | It was reported that post a da vinci si sacrocolpopexy procedure, the patient returned to the hospital and was found to to have ileus. Several procedures were performed; however, the patient expired. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | The hospital has been contacted several times to obtain additional information; however, as of the date of this report, the site has not provided any additional information. No malfunctions of the da vinci s surgical system were reported to have occurred during the procedure, and the hospital has continued to use the system to perform surgery on patients. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1885529 | 1850577 | 1750245 | 2955842201000419 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-09-03 | 2010-10-01 | Yes | No | Health Professional | IS3000 A60P4 | No | No | 2010-09-03 | No | No | Reuse | 2010
IS3000 A60P4 |  | Death | Death | It was reported that approximately one week post a da vinci si hysterectomy procedure, the patient returned to the hospital and was found to have sepsis. The patient had a do not resuscitate (dnr) order on file and expired one week later. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | The hospital has been contacted on several occassions to obtain additional information; however, complete event details and medical records have not been provided to intuitive surgical. No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1885560 | 1850589 | 1750268 | 2955842201000420 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-03 | 2010-10-01 | Yes | No | Health Professional | IS3000 A60P4 | No | No | 2010-09-03 | No | No | Reuse | 2010
IS2000 A51P8 | 2010-09-06 | Other | Other | It was reported that during the beginning of a da vinci s surgical procedure, system error code 23002 occurred several times. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the surgical staff emergency power-off the system. This action resolved the issue for a few minutes and then the system would no longer "home. " the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 23002 was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to the right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23002 appears when the da vinci s safety system determines that the position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer) were beyond mechanical limits. Upon determining this condition, the safety systems put davinci in a recoverable safe state. The mtml was returned to isi for evaluation. Engineering discovered that the hall effect sensor had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1885541 | 1850593 | 1750272 | 2955842201000417 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-09-06 | 2010-10-01 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-09-06 | Device Not Retu | No | Reuse | 2010
40016907 | 2010-09-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The insert accessory was retrieved and the surgeon proceeded with the procedure using a replacement insert accessory. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found a piece of the curved blade to be broken off at the distal end. The blade fractured near the base of the clamp arm and the broken piece was not returned. The fracture surface has some ridges, it is not flat. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1885578 | 1850633 | 1750288 | 2955842201000421 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-03 | 2010-10-01 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2010-09-03 | Device Not Retu | Yes | Unkown | 2010
IS3000 A60P4 |  | Death | Death | It was reported that approximately 10 days post a da vinci si surgical procedure, the patient returned to the hospital with unknown symptoms and expired. Reportedly, the patient's demise was unrelated to the da vinci si surgical system. | The hospital has been contacted on several occassions to obtain additional information; however, complete event details and medical records have not been provided to intuitive surgical. No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1885546 | 1850640 | 1750292 | 2955842201000422 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-03 | 2010-10-01 | Yes | No | Health Professional | IS3000 A60P4 | No | No | 2010-09-03 | No | No | Reuse | 2010
42017909 | 2010-09-07 | Malfunction | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, one of the blades from the monopolar curved scissors instrument broke off and fell into the patient. The procedure was extended 3. 5 hours due to the time spent by the surgical staff looking for the broken blade. A gel port was added and a port incision was extended in order to inspect the patient's bowel and remove the broken blade. The planned surgical procedure was completed. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned (post-evaluation) or if additional information is received. On (b)(6) 2010, the initial reporter informed intuitive surgical that no post-surgical complications were experienced by the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1885582 | 1850641 | 1750298 | 2955842201000418 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-09-07 | 2010-10-01 | Yes | Yes | Health Professional | 42017909 | No | Yes | 2010-09-07 | No | No | Unkown | 2010
40017909 | 2009-07-21 | Malfunction |  | It was reported that after a da vinci si hysterectomy procedure, the surgical staff had observed that the monopolar curved scissors instrument was missing one blade. The component was located in the patient's pelvis area during x-ray and the surgeon was able to retrieve the missing piece through open surgery. No patient harm, adverse outcome or injury was reported. | The instrument and broken blade were returned and evaluated. Per the customer reported complaint, the broken blade had fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit damage at the tip except general wear of the blade. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1885595 | 1850648 | 1750303 | 2955842201000423 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-09-02 | 2010-10-01 | No | Yes | Health Professional | 40017909 | Device Returned To Manufacturer | No | 2010-09-02 | Device Not Retu | No | Reuse | 2010
420008 | 2010-09-29 | Malfunction |  | The staff reported that the obturator broke when placed in the trocar on two different patients, with two different surgeons, in two different rooms for two different surgeries on the same day. It broke near the same spot in both cases. The staff were able to retrieve the item. They reported this as a recognized defect for this item. They did not feel that the break was related to surgical technique. Similar breakage has occurred in 5 other obturators. ======================health professional's impression======================there appears to be a common area of stress and/or weakness in the obturator. Staff feels that the plastic is not holding up to repeated steam sterilizations. Reprocessing has been completed according to manufacturer's recommendations. We are planning a simulation to see if we can re-produce the defect to evaluate product and technique. ======================manufacturer response per site reporter======================the rep thought it might be related to the way we process the equipment. He was not onsite, nor had he seen the actual defect. Manufacturer is coming to our facility to evaluate breakage. ======================health professional's impression======================there appears to be a common area of stress and/or weakness in the obturator. We are planning a simulation to see if we can re-produce the defect to evaluate product and technique. ======================manufacturer response per site reporter======================the rep thought it might be related to the way we process the equipment. He was not onsite, nor had he seen the actual defect. |  | CANNULA OBTURATOR | SYSTEMSURGICALCOMPUTER CONTROL | 1903776 | 1872629 | 1769625 | 1872629 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-10-01 | 2010-10-01 | No | Yes | Physician | 420008 | Yes | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-09-02 | Other |  | It was reported that while setting up the da vinci si system for a prostatectomy procedure, the left eye in the surgeon side console was providing a yellow image. The patient had been administered anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that the image issue experienced by the customer was associated with the system camera controller and cable. The camera controller produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon side console. The system was repaired by replacing the affected camera controller and camera cable. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1881377 | 1846670 | 1746864 | 2955842201000409 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-09-02 | 2010-09-28 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-09-02 | Device Not Retu | No | Reuse | 2010
420184 | 2010-06-22 | Malfunction |  | On (b)(6) 2010, user facility medwatch report (b)(4) was received: while the surgeon was using the da vinci maryland bipolar forceps, it stopped working and broke while in use inside the patient. It was removed from the field and inspected. It was found to have pieces broken off of it. The patient was later found to have three broken plastic pieces in the abdominal area which were removed. | On (b)(6) 2010, (b)(6) was contacted regarding the user facility medwatch report (b)(4) received stating maryland bipolar forceps instrument (part number 420172 and lot number m10090731-18) had broken and fell inside of a patient during a da vinci hysterectomy procedure. It was relayed to (b)(6) that isi had received this part on (b)(4) on (b)(4) 2010 for a different reported event. Additionally, failure analysis of this instrument found that it was not broken or missing any parts. (b)(6) stated she would follow up with operating room staff to ensure the correct instrument had been reported and returned. On (b)(4) 2010, a follow up call to (b)(6) was made to request confirmation that the maryland bipolar forceps instrument reported on user facility medwatch report (b)(4) was correct. (b)(6) stated she had not heard back from her operating room staff but will try again to obtain information. On (b)(6) 2010, (b)(6) contacted isi to state that the nurse present for the hysterectomy procedure on (b)(6) 2010 confirmed that a different instrument, a permanent cautery spatula (part number 420184), was the instrument that broke and that the instrument had since been discarded. (b)(6) stated she had no further information to provide isi. Due to the permanent cautery spatula instrument having been discarded by the site, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1881380 | 1846672 | 1698268 | 2955842201000410 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-09-03 | 2010-09-28 | No | Yes | Health Professional | 420184 | No | No | 2010-09-03 | No | No | Unkown | 2010
42017909 | 2010-09-08 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure one of the jaws of the monopolar curved scissors instrument broke and fell into the patient. The jaw was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The monopolar curved scissors instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned or if additional information is received a follow up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1881383 | 1846678 | 1746867 | 2955842201000412 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-08 | 2010-09-28 | No | Yes | Health Professional | 42017909 | No | No | 2010-09-08 | No | No | Unkown | 2010
IS2000 A51P8 | 2010-09-09 | Other | Other | It was reported that prior to starting a da vinci s mitral valve repair procedure, the site lost vision on the left video channel. With the assistance of an isi technical support engineer (tse), the site replaced the camera cable; however, the left vision loss continued to occur. The patient had been under anesthesia for approximately one hour when the surgeon decided to complete the case using traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the image issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of september 28 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1881382 | 1846691 | 1746834 | 2955842201000411 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2010-09-09 | 2010-09-28 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2010-09-09 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-09-08 | Other | Other | It was reported that while beginning a da vinci s thoracic procedure, the site experienced problems with the focus controller. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the procedure and reschedule for a later date. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by the isi field service engineer found that the focus controller cable was not fully connected to the camera. The focus controller is designed to adjust the endoscope lens focus to sharpen the surgical image. The focus controller cable connects the focus controller to the camera. The system was repaired by re-seating the focus controller cable to the camera. As of (b)(4) 2010, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1881421 | 1846710 | 1746897 | 2955842201000413 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-09-08 | 2010-09-28 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-09-08 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-09-13 | Other | Other | It was reported that while starting a da vinci s surgical procedure, the surgeon placed his head into the surgeon side console and experienced a black image in the left eye. With the assistance of an isi technical support engineer and the clinical sales representative onsite, the site replaced the camera head and camera cable, however, the black image in the left eye continued to occur. The patient was under anesthesia and the port incisions were made when the surgeon decided to complete the planned procedure using traditional open surgical techniques. No adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the image issue experienced by the customer was associated with the camera controller unit (ccu). The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by replacing the affected camera controller unit. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1881448 | 1846717 | 1746902 | 2955842201000414 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-09-13 | 2010-09-28 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-09-13 | Device Not Retu | No | Reuse | 2010
42011704 | 2010-09-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure the 5mm needle driver instrument broke and a piece fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse event, or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The site returned the 5mm needle driver instrument. Per the customer reported complaint, engineering found one grip to be detached. The grip was fractured at the base, where the grip arm meets the grip hub. The outer surface of the grip does not exhibit any obvious damage. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The 5mm needle driver instrument has not been returned for evaluation. The root cause of the customer reported malfunction cannot be determined. If the instrument is returned or if additional information is received, a follow up mdr will be submitted. | 5MM NEEDLE DRIVER INSTRUMENT | Reuse | 1881471 | 1846767 | 1746928 | 2955842201000415 | 1 | NAY | Manufacturer | DistributorOther | Other | FollowupInitial | 2010-10-06 | 2010-09-28 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-10-06 | Device Not Retu | No |  | 2010
42017805 | 2010-08-23 | Malfunction |  | It was reported that during a da vinci s mitral valve repair procedure, during insertion of the curved scissors instrument into the trocar, a blade from the curved scissors instrument broke off and fell on the patient's lung. The broken piece was immediately retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or complications were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece from the left scissor blade missing. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The blade edge of the intact blade exhibits chipping. | CURVED SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1880047 | 1845721 | 1746018 | 2955842201000408 | 1 | NAY | Manufacturer | User facilityDistributor | Physician | Initial | 2010-08-31 | 2010-09-27 | No | Yes | Health Professional | 42017805 | Device Returned To Manufacturer | Yes | 2010-08-31 | Device Not Retu | No | Reuse | 2010
420183 | 2010-08-25 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the permanent cautery hook instrument broke and pieces fell into the patient. Not all pieces were believed to have been retrieved. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned (post-evaluation) or if additional information is received. Intuitive surgical 8mm endowrist instruments for use on the da vinci s surgical system have been tested for material biocompatibility. The materials in these instruments have been tested for biocompatibility in accordance with (b)(4) for their intended use (blood path, indirect (4. 2. 3. A) and limited exposure (4. 3. A)). | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1877571 | 1844137 | 1744788 | 2955842201000402 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-08-25 | 2010-09-24 | No | Yes | Health Professional | 420183 | No | No | 2010-08-25 | No | No | Unkown | 2010
40017208 | 2010-08-26 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, while dissecting tissue, a piece from the tip of the maryland bipolar forceps instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one grip detached and broken off. The grip fractured within the yaw pulley, at the cross section of where the grip is swaged to retain the conductor wire. Installation of the broken piece that backs into the yaw pulley shows that there is a bend near the fracture surface, as the tips have an offset. Engineering concluded that such damage to the instrument is indicative of overloading at the instrument tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1877911 | 1844329 | 1744921 | 2955842201000403 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-27 | 2010-09-24 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2010-08-27 | Device Not Retu | No | Reuse | 2010
40018007 | 2010-08-23 | Malfunction |  | It was reported that during a da vinci nephrectomy procedure, arcing was observed coming from the tip of the monopolar curved scissors (mcs) instrument with an mcs tip cover accessory installed. The mcs tip cover accessory was replaced to complete the planned procedure. No patient harm was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory is returned for evaluation a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1878094 | 1844407 | 1744987 | 2955842201000404 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-08-27 | 2010-09-24 | No | Yes | Health Professional | 40018007 | No | Yes | 2010-08-27 | Device Not Retu | Yes | Unkown | 2010
42022703 | 2010-08-20 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the pk dissecting forceps instrument broke off and fell into the patient. Despite requests for additional information, no additional information has been reported by the site. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a broken yaw pulley on the opposite side of the conductor wire cover. The yaw pulley was chipped and the conductor wire broke off the crimp. No damage was found on the clevis. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1878079 | 1844410 | 1744990 | 2955842201000407 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-25 | 2010-09-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2010-08-25 | No | No | Reuse | 2010
42017909 | 2010-08-26 | Malfunction |  | It was reported that one hour into a da vinci s hysterectomy procedure, the surgical staff noticed the tip of the monopolar curved scissors instrument broke off and fell into the patient. The surgeon was able to retrieve the tip and change the instrument without causing the patient any harm. The planned surgical procedure was completed with an instrument of the same type. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a broken scissor blade. The tip of the right blade was chipped off, approximately 0. 09 x 0. 05 in size. No bending in the blades was found. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1878097 | 1844411 | 1744991 | 2955842201000406 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-26 | 2010-09-24 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-08-26 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-08-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a piece from the left blade broke off and is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1878080 | 1844412 | 1744992 | 2955842201000405 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-26 | 2010-09-24 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-08-26 | Device Not Retu | No | Reuse | 2010
IS2000 A51 P8 | 2010-08-23 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced system error code 23008. With the assistance of an isi technical support engineer, the site turned off the system and back drove the affected master tool manipulator two times, however, upon restarting the system, the fault recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found that system error code 23008 experienced by the site was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a nonrecoverable safe state. The mtmr was returned to isi for failure analysis investigation. Engineering evaluation found the axis 4 motor and encoder were defective and likely generating the system error code experienced by the customer. As of (b)(4) 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1870668 | 1838830 | 1740181 | 2955842201000401 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-08-23 | 2010-09-20 | Yes | Yes | Health Professional | IS2000 A51 P8 | Device Returned To Manufacturer | Yes | 2010-08-23 | Device Not Retu | No | Reuse | 2010
PS3000 | 2010-08-26 | Injury | Hospitalization | During robotic assisted laparoscopic hysterectomy, 3-d capability stopped working. Surgeon opted to proceed robotically in 2-d mode. Ureteral injury converted to open procedure. |  | DA VINCI SI HD | ROBOTIC SURGICAL SYSTEM | 1892656 | 1852915 | 1788885 | MW5017637 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-09-17 | 2010-09-20 | Yes | Yes | Health Professional | PS3000 |  | Yes |  |  |  |  | 2010
420189 | 2010-01-12 | Malfunction |  | After the da vinci double fenestrated instrument was removed from the pt (robotic assisted lap assisted vaginal hysterectomy, bilateral salpingo-oophorectomy), the instrument tip broke off and was not detected by x-ray. The surgeon opened, tissue morcellated, no identifiable part found in wound or morcellated tissue. The area was irrigated well and searched - no foreign objects were detected in the wound area. ====================== health professional's impression======================the instrument tip broke off of the double fenestrated grasper. |  | ENDOWRIST | GRASPER FENESTRATED | 1889244 | 1855009 | 1754351 | 1855009 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-09-20 | 2010-09-20 | No | Yes | Invalid Data | 420189 | Yes | No Answer Provided |  |  |  |  | 2010
42020702 | 2010-07-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while grasping the patient's enlarged uterus, the teneculum forceps instrument broke and fragments fell into the patient. The fragments were successfully retreived and the surgical staff continued with the procedure after some delay. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | TENECULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1867989 | 1837550 | 1739083 | 2955842201000398 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2010-08-17 | 2010-09-16 | No | Yes | Health Professional | 42020702 | No | No | 2010-08-17 | Device Not Retu | No | Unkown | 2010
40016906 | 2010-08-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The insert accessory was retrieved and the procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument accessory is returned for evaluation a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1867991 | 1837552 | 1739086 | 2955842201000399 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-17 | 2010-09-16 | No | Yes | Health Professional | 40016906 | No | No | 2010-08-17 | No | Yes | Unkown | 2010
40016906 | 2010-08-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The insert accessory was retrieved and the procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found one insert with a broken curved blade and the other was found to have a broken clamp arm. One side of the hinge features that mate with clamp arm pins is severely bent backwards. Damage is consistent with hyperextension of clamp arm. Engineering concluded that breakage of the harmonic curved shears (hcs) insert accessory was most likely caused by mishandling or misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1867992 | 1837553 | 1739086 | 2955842201000400 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-17 | 2010-09-16 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | No | 2010-08-17 | Device Not Retu | Yes | Initial | 2010
42018404 | 2010-07-05 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s lobectomy procedure, the surgeon found that part of the plastic at the wrist of the permanent cautery spatula had broken off and fell into the patient's body. The fragment was retrieved and the procedure was successfully completed using a replacement instrument of the same type. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one distal clevis ear broken off at its base. The size of the broken piece is approximately. 285 x. 235 and was not returned. The ear had broken off on the opposite side of the conductor wire and cap. The broken ear piece was not returned, however, the conductor cap was returned taped to the housing. Additionally engineer observed a missing conductor cap and broken ceramic sleeve. The conductor cap is missing as a result of the clevis breakage. The yaw pulley extruded boss feature that holds cap in place is broken off. The distal end of ceramic sleeve has a. 090 x. 090 piece broken off. Spatula is bent slightly where it transitions from a cylindrical to a flat shape. Engineer has concluded that the clevis breakage and observed damage may be due to excessive loading at the tip of the instrument. No other damage found. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 1860710 | 1832256 | 1734376 | 2955842201000395 | 1 | NAY | Manufacturer | ForeignDistributor | Other | InitialFollowup | 2010-08-13 | 2010-09-13 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2011-02-16 | Device Not Retu | No |  | 2010
42020504 | 2010-08-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was removing the patient's uterus through laparotomy incision, the console surgeon observed a piece of the fenestrated bipolar instrument in the patient. The piece was removed with no harm to the patient. It was reported that the patient had large fibroids in their uterus causing the surgeon to have to remove via laparotomy instead of morcellating as per usual technique. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one broken grip tip at the yaw pulley. The broken point is inside yaw pulley near the conductor cable connection point with clevis. With grip tip installed, engineering observed the tip to have a slight bend. No other damage found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1860705 | 1832257 | 1734397 | 2955842201000396 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-08-11 | 2010-09-13 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2010
400180 | 2010-07-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure the surgeon was cutting some fascia when the surgeon felt that the movement of the monpolar curved scissors (mcs) instrument wire wasn't normal. The nurse removed the mcs instrument from the patient and observed that the mcs tip cover accessory installed on the mcs instrument was burnt and had a hole. The instrument was replaced and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument and tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument and/or tip cover accessory are returned (post-evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument and tip cover accessory were returned for evaluation. The tip cover has various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicative of arcing. The multiple holes indicate multiple arcing events occurred. The hole sizes range from. 010 to. 040 with varying shapes (round and oblong) and are located from. 025 to. 335 above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. Visual inspection of the instrument shows char marks on the distal and proximal clevis. The instrument does not have any burrs or unusually sharp edges in the area of the char marks that would contribute to arcing. With the tip cover installed on the instrument, the position of the holes line up with the char marks on the instrument. | 8MM MCS TIP COVER ACCESSORY | Unkown | 1861030 | 1832491 | 1734543 | 2955842201000397 | 1 | NAY | Manufacturer | ForeignDistributor | Other | InitialFollowup | 2010-08-13 | 2010-09-13 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | No | 2010-09-17 | Device Not Retu | Yes |  | 2010
42009308 | 2010-08-09 | Malfunction |  | It was reported that during a da vinci s gynecological procedure a piece of wire was found in the patient. The wire was immediately removed and the prograsp forceps instrument that was in use at the time was also removed for inspection. The surgical staff evaluated the instrument however they did not see any damage and noted that there was no system issues during the procedure. The planned surgical procedure was completed and no patient complication, injury or adverse outcome was reported. | The instrument has not been returned for evaluation, however, an isi field engineer onsite evaluated the instrument and found the instrument to be fully functional with no physical damage. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1859587 | 1831273 | 1733552 | 2955842201000394 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-08-16 | 2010-09-10 | No | Yes | Health Professional | 42009308 | No | No | 2010-08-16 | No | No | Reuse | 2010
42018305 | 2010-08-12 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while removing a large myoma, the side of the permanent cautery hook instrument fractured and a piece fell into the patient. The piece was removed and the planned surgical procedure was completed. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Engineering found one distal clevis ear to be broken off at it's base. The yaw pulley hub extruded boss feature that retains the conductor cap is broken and the conductor cap is nearly completely severed off at the yaw pulley. The broken distal clevis ear and detached conductor cap were returned with the instrument. Electrical continuity failed. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1859581 | 1831274 | 1733553 | 2955842201000393 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-08-12 | 2010-09-10 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | Yes | 2010-08-12 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-08-09 | Other | Other | It was reported that while preparing for a da vinci si hysterectomy procedure, the surgeon side console (ssc) experienced left eye video loss in the high resolution stereo viewer (hsrv). With the assistance of an isi technical support engineer and an isi da vinci specialist, the site powered down the system and re-seated all camera head cables; however, the issue persisted. The core video inputs were switched in an attempt to isolate the issue. The site replaced the camera head and camera cable; however, this issue continued. Emergency power off (epo) was applied and cables were swapped at the rear of the dual camera control unit and the video loss then moved to the right video channel. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the camera head, camera cable and the double camera controller (doco). The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the doco on the vision cart. The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected camera head, camera cable and deal camera control unit. The components were returned to isi for evaluation. Engineering evaluation revealed that the doco had experienced a burnt fuse thereby causing the vision issue experienced by the site. As of september 1, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1857507 | 1829625 | 1732069 | 2955842201000392 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-08-09 | 2010-09-09 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-08-09 | Device Not Retu | No | Reuse | 2010
| 2010-09-08 | Malfunction |  | During a robotic thymectomy procedure, while using an intuitive surgical harmonic curved shears insert in the davinci robot-s, one portion of the shear snapped off into the thoracic cavity, rendering the equipment unusable. Md was able to retrieve 2 small pieces of the shears. The equipment was removed from the table, reported to the intuitive rep, central sterile supply, and a chest x-ray was taken upon completion; no objects seen in patient by chest films. Pt appears to not have any harm from the event. ====================== manufacturer response for harmonic curved shears as per site reporter======================have asked vendor to share info if they choose to analyze the item. |  | HARMONIC | CURVED SHEARS | 1870160 | 1838650 | 1740046 | 1838650 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-09-09 | 2010-09-09 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2010
| 2010-08-10 | Malfunction |  | While a robotic prostatectomy was in progress, the permanent cautery spatula was in use. The tip broke off in the abdomen and had to be retrieved. |  | ENDOWRIST | MONOPOLAR CAUTERY | 1858836 | 1830666 | 1732960 | 1830666 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-08-24 | 2010-09-08 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-08-13 | Other | Other | It was reported that during the anastomosis portion of a da vinci si prostatectomy procedure, the customer experienced system error code 252. The system was restarted, however, system error code 90 appeared at restart. The surgical staff verified that all connections were secure on the vision cart (and that all external video connections were removed) and emergency powered off the system three times. System error code 90 continued to appear during each restart. The site then replaced the camera and camera cable and restarted the system, however, system error code 90 appeared again during restart. The surgeon decided to convert to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes 90 and 252 were associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 90 appears when a system board reads a voltage value out of the normally acceptable range. System error code 252 is a sympathetic error and appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the safety system puts da vinci si in a non-recoverable safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1854218 | 1827213 | 1729953 | 2955842201000390 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-08-13 | 2010-09-07 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-08-13 | Device Not Retu | No | Reuse | 2010
420002 | 2010-08-05 | Other | Other | It was reported that after a da vinci s hysterectomy procedure was successfully completed, the patient developed blisters on the skin in their abdomen area due to an allergic reaction to the metal composition of the instrument cannula accessory. The console surgeon was aware of the patient's allergy to metal, however, a decision was made to proceed with the surgical procedure. Steroids were used to treat the patient's blisters. | On (b)(6) 2010, the console surgeon reported that the patient was doing well, however, the console surgeon would not provide any other patient information. Per the information provided by the initial reporter, it was determined that the da vinci s surgical system worked as intended and no system malfunctions occurred during the surgical procedure. The hospital has continued to use the system to perform surgery on patients. As of september 1, 2010, no adverse events have been experienced with the site's da vinci s surgical system and no similar instances of this event have been reported to isi. | 8MM INSTRUMENT CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 1854220 | 1827216 | 1729956 | 2955842201000391 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-08-09 | 2010-09-07 | Yes | No | Health Professional | 420002 | No | No | 2010-08-09 | No | No | Unkown | 2010
42019408 | 2010-08-05 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip of the mega needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a broken carbide insert at the grip tip. The grip tip itself is not bent, and the broken carbide piece is approximately 0. 11 x 0. 13 in size. No other damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1852872 | 1826143 | 1728987 | 2955842201000388 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-08-05 | 2010-09-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2010-08-05 | Device Not Retu | No | Reuse | 2010
42020504 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip on the fenestrated bipolar forceps instrument cracked near the cable. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument grips are not cracked or damaged. Functional testing revealed that the grips open/close properly when the input disks are manually rotated. Engineering also observed that the distal end components exhibit charring and have an arc track, indicating that an arcing event occurred. One yaw pulley cover exhibits charing/melting on its outer edge near the grip base and there is an arc track at the yaw pulley cover that extends down the side of one yaw pulley. The side of the yaw pulley has charring with material removal adjacent to the grip cable. There is a slight separation between this yaw pulley and its cover at the glue joint. While the evidence is inconclusive, engineering concluded that damage to the instrument was most likely the result of arcing in which unintended current flowed to the yaw pulley. The instrument passed electrical continuity testing. On (b)(6), 2010, isi followed up with (b)(6) in materials management and she reported that an arcing event during the planned surgical procedure was not observed. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1852856 | 1826144 | 1728988 | 2955842201000389 | 1 | NAY | Manufacturer | User facility | Health Professional | Initial | 2010-08-10 | 2010-09-03 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | Yes | 2010-08-10 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-08-05 | Other | Other | It was reported that towards the end of a da vinci s myomectomy procedure, the nurse noticed that one of the monopolar curved scissors instrument tips had broken off and was missing. The procedure was delayed so that an x-ray of the patient could be performed. The missing instrument tip was located (via x-ray) at the posterior side of the uterus and stuck on the rectum, causing minor bleeding on the rectum. The broken tip was retrieved from the patient cavity and the planned surgical procedure was completed. No additional patient harm, adverse outcome, nor injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off near the distal clevis pin. The breakage surface is uneven and the cable crimp is exposed due to the break. The detached blade was returned with the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1850577 | 1824174 | 1727178 | 2955842201000384 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-08-05 | 2010-09-02 | Yes | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-08-05 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-07-29 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the patient received a bowel injury secondary to arcing from the tip of the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. The surgeon stitched the affected area and successfully completed the procedure with no further issues. No additional patient harm, adverse outcome or injury has been reported. | The tip cover accessory was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The mcs instrument used during the time of the event was returned and evaluated. Engineering found no evidence of arcing and no physical damage was observed on the instrument. A follow-up mdr will be submitted if the accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1849206 | 1823086 | 1726283 | 2955842201000386 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-08-02 | 2010-09-01 | No | Yes | Health Professional | 40018010 | No | No | 2010-08-02 | No | Yes | Unkown | 2010
40018302 | 2010-07-23 | Malfunction |  | It was reported that during a da vinci nephrectomy procedure, a plastic slice from the top of the permanent cautery hook instrument broke off and fell into the patient's abdomen. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the conductor cover broken off of the distal clevis. There is no evidence of arcing or damage to the wire and cable. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1849359 | 1823168 | 1726331 | 2955842201000387 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | PHARMACIST | Initial | 2010-08-02 | 2010-09-01 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | Yes | 2010-08-02 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-06-28 | Malfunction |  | It was reported that approximately 1 hour during a da vinci si hysterectomy procedure, burning could be seen through the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The tip cover was removed and replaced and the procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1846473 | 1821054 | 1724500 | 2955842201000385 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-08-02 | 2010-08-31 | No | Yes | Health Professional | 40018010 | No | No | 2010-08-02 | No | Yes | Unkown | 2010
42017910 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors (mcs) instrument was dull. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as a latex cut test revealed that the scissors cut cleanly through. 006 latex. The blades do not exhibit damage. Engineering also observed charring on the distal end components. There is a char mark on the proximal clevis, near the window that reveals the proximal clevis seal. The clevis seal has a. 030 x. 030 section directly below the char mark that exhibits material removal. Engineering concluded that while the evidence is inconclusive, damage to the instrument is indicative that an arcing event occurred. The clevis does exhibit any burrs or sharp edges near the damaged section. The instrument passed electrical continuity testing. No other damage was found. Although the mcs tip cover accessory used in conjunction with the monopolar curved scissors (mcs) instrument was not returned for evaluation, engineering found that the location of the damage on the proximal clevis roughly aligns where the tip cover accessory silicone meets the ultem sleeve. On (b)(4) 2010, isi followed up with rn, (b)(6) in the operatin room department and she reported that arcing during the surgical procedure was not observed and no patient harm, adverse outcome or injury occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1843452 | 1819007 | 1722758 | 2955842201000382 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-07-29 | 2010-08-27 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2010-07-29 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-07-28 | Other | Other | It was reported that during a da vinci s gynecological procedure, the customer experienced a nonrecoverable system error code 25. With the assistance of an isi technical support engineer, the site power cycled the system, cycled the breakers and emergency power off (epo) was applied; however, the attempt to resolve the issue was unsuccessful as the error reoccurred at boot up. Patient side manipulator (psm) arm 4 was disabled and the system successfully powered on and homed with the three remaining psm arms. The system powered back up and homed with no issues. The site indicated that the 4th arm was not used for this procedure and would continue the procedure using 3 arms. The customer called back during the procedure to report the occurrence of a system error code 25310 followed by a gear icon. With the assistance of the tse, the site power cycled the system and experienced system error code 814. The site then power cycled the system, cycled the breakers and emergency power off (epo) was applied; however, non recoverable system error code 810 appeared. The patient side cart (psc) was power cycled and epo'd several more times with no change. The patient had been under anesthesia and port incisions had been placed when the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system errors were associated with fluid contamination on the printed circuit assembly boards causing them to malfunction. The pca boards performs numerous tasks related to the function, control and hardware fault monitoring for a specific arms. The system was repaired by replacing the affected pca boards. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1843626 | 1819116 | 1722838 | 2955842201000383 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-28 | 2010-08-27 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-07-28 | Device Not Retu | No | Reuse | 2010
| 2010-08-11 | Injury | Required Intervention | A (b)(6) female admitted with diagnosis of fibroid uterus to undergo hysterectomy via da vinci system. Intra-op report obtained of problems with 3 harmonic shear inserts. First one broke at the tip during surgery, piece retrieved. Second one was working fine at first then quit. Third one was tested outside the pt and broke during testing. Note: above reported to mr (b)(6) with intuitive surgical, (b)(4) assigned. |  | HARMONIC CURVED SHEARS | SHEAR INSERTS | 1846708 | 1816096 | 1720105 | MW5017207 | 1 | NAY | Voluntary |  | Nurse | Initial | 2010-08-18 | 2010-08-24 | Yes | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2010
IS2000 A51P8 | 2010-08-04 | Other | Other | It was reported that during a da vinci s hysterectomy procedure a patient side manipulator (psm) arm was not following correctly. The surgical staff installed a different instrument on the psm and reseated the instrument sterile adapter, however, the issue continued to recur. An isi technical support engineer was contacted, who requested that the surgeon use another psm arm. The surgeon declined and made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation was conducted by an isi field service engineer (fse), who was unable to replicate the patient side manipulator (psm) issue experienced by the customer. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The fse test drove the system at the surgical site and the found the system to be within specification and working as intended. The system error logs were also reviewed and no errors associated with the customer reported failure mode were observed. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1836450 | 1811584 | 1715868 | 2955842201000380 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-08-04 | 2010-08-23 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-08-04 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-07-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while bi-valving the uterus for vaginal removal, one of the grips of the monopolar curved scissors instrument broke and fell into the patient. The broken grip was immediately retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the left grip to be broken off and missing. The blade fractured at the cross section of the grip cable crimp causing half of the cable crimp to be exposed. The fractured plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1834896 | 1810272 | 1714802 | 2955842201000375 | 1 | NAY | Manufacturer |  | Other | Initial | 2010-07-26 | 2010-08-20 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-07-29 | Other | Other | It was reported that while preparing to start a da vinci s epicardial ablation procedure, upon powering up the system, the site experienced system error code 3. With the assistance of an isi technical support engineer the site power cycled the system, however upon restart the patient side cart battery leds were red. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the customer reported complaint was related to the patient side power distribution (ppd) pcs board. The patient side power distribution board is a printed circuit (pca) board located on the patient side cart that divides the electrical power feed into subsidiary circuits while providing a protective circuit breaker for each circuit. The system was repaired by replacing the ppd. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1834898 | 1810273 | 1714803 | 2955842201000377 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-07-29 | 2010-08-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2010-07-29 | Device Not Retu | No | Reuse | 2010
42018404 | 2010-07-23 | Malfunction |  | It was reported that during a da vinci si gynecological surgical procedure, the tip of the permanent cautery spatula instrument was burned up. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the grip base to be burned and melted at the yaw pulley to conductor wire interface. The conductor wire cap is melted and damage was found on the cable and yaw pulley. The proximal clevis was found to be broken at the main tube interface. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1834899 | 1810300 | 1714791 | 2955842201000376 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-27 | 2010-08-20 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2010-07-27 | Device Not Retu | No | Reuse | 2010
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the maryland bipolar instrument was stiff and hard to move. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument passed recognition and engagement testing. Additional functional testing revealed that the instrument moved intuitively with a full range of motion in all directions and the grips opened and closed properly. There was no excessive friction in any axis of motion. Engineering also observed charring on the yaw pulley and tube damage. The outer surface of the yaw pulley cover exhibits a char mark and localized melting/thermal damage, indicating that an arcing event occurred. There is no charring between the grips or the yaw pulley. The char mark position may indicate unintended current flowed to the yaw pulley cover. The distal end of the main tube has a 1. 8 long section with light material removal and has a rough surface finish. The damaged area is axially aligned with the tube. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1834978 | 1810366 | 1714815 | 2955842201000379 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-21 | 2010-08-20 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-07-21 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-07-26 | Malfunction |  | It was a reported that during a da vinci si surgical procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the tip of one scissor blade is broken off. The broken blade is approximately 0. 160 x 0. 420. The fracture plane of the broken blade is not straight and the blade has a chip on the edge below the midpoint. The intact blade does not exhibit any obvious cracking. No other damage was found. | MONOPLOAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1834979 | 1810367 | 1714816 | 2955842201000378 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-26 | 2010-08-20 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-07-26 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-07-30 | Other | Other | It was reported that during a da vinci sacrocolpopexy procedure the surgical staff experienced system error code 20013. The system was rebooted multiple times, however, error code 20013 continued to recur. The surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 20013 experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts davinci in a recoverable safe state. The ecm was returned to isi for failure analysis investigation, however, engineering was unable to replicate the customer reported failure mode. Based on the system error logs, an axis potentiometer and motor/encoder will be replaced as a precaution. As of (b)(6), 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1835100 | 1810514 | 1714929 | 2955842201000381 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2010-07-30 | 2010-08-20 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-07-30 | Device Not Retu | No | Reuse | 2010
420205 | 2010-07-14 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, while performing a vaginal cuff closure and using the fenestrated bipolar forceps instrument, unintended cautery occurred. The planned surgical procedure was completed and no patient complications were reported. | The fenestrated bipolar forceps instrument was not returned for failure analysis, however, isi field service engineering investigation concluded that the issue experienced by the customer was associated with a non isi supplied cautery cable that connects the instrument to the electrical surgical unit (esu). When the surgeon moved the patient side manipulator (psm), the cautery cable also moved, thus causing unintended cautery activation of the instrument. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. Replacement of the cautery cable resolved the issue experienced by the customer. An additional investigation conducted by engineering included a review of the system and tool error logs which did not indicate any system faults or errors that would have caused or contributed to the issue experienced by the site. On (b)(4) 2010, isi field service engineering follow-up with the site revealed that there were no recurrences of the issue. As of (b)(4) 2010, the site has continued to use the da vinci si system with no reported adverse events. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1828935 | 1802299 | 1707512 | 2955842201000366 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-07-14 | 2010-08-17 | Yes | Yes | Health Professional | 420205 | No | No | 2010-07-14 | No | No | Unkown | 2010
IS2000 A51P8 | 2010-07-21 | Other | Other | It was reported that during set up for a da vinci s surgical procedure the surgeon experienced a foggy image through both eyes of the high resolution stereo viewer simultaneously. The site informed an isi technical support engineer (tse) that they attempted to used 4 different endoscopes; however, the foggy image persisted. They site then attempted to re-calibrate the scopes and checked proper settings. The site tested their scope warmer but was unwilling to perform additional troubleshooting. The patient was under anesthesia and port incisions had been make before the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the camera head. The system was repaired by replacing the affected camera head. The camera heads were returned to isi for evaluation. Engineering evaluation revealed that the camera head was dropped; thereby causing the stage to be damaged. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1828106 | 1801372 | 1706780 | 2955842201000370 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-21 | 2010-08-16 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-07-21 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-06-28 | Malfunction |  | It was reported that during a da vinci s nephrectomy procedure while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, unintended arcing was seen. The tip cover was immediately removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering found that there is a hole burned through the silicone of the tip cover accessory. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The size of the hole is approximately 0. 008 inches in diameter and the tip cover exhibits multiple tears on the outer surface. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1828126 | 1801383 | 1706761 | 2955842201000368 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-07-19 | 2010-08-16 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-07-19 | Device Not Retu | Yes | Initial | 2010
IS2000 A51P8 | 2010-07-21 | Other | Other | It was reported that during a da vinci s urological procedure, the site had difficulty with the system powering on. Prior to calling isi technical support, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by a field service engineer (fse) confirmed the customer reported problem was associated with the surgeon side console's power cord being unplugged. The site was made aware of the issue and the power cord was plugged in to repair the reported problem. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1828180 | 1801448 | 1706816 | 2955842201000371 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-07-21 | 2010-08-16 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2010-07-21 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-07-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, unintended arcing was seen. The tip cover was immediately removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering found that there is a hole burned through the silicone of the tip cover accessory. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The size of the hole is approximately 0. 009 inches in diameter and the tip cover exhibits multiple tears on the outer surface. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1827326 | 1799755 | 1706761 | 2955842201000369 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-07-19 | 2010-08-13 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-07-19 | Device Not Retu | Yes | Initial | 2010
42009308 | 2010-07-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a wire from the prograsp forceps instrument fell into the patient. The wire was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm customer reported compliant as the instrument was found to be intact with no broken or frayed cables or wires. The instrument was placed on an in house test system and was found to move intuitively in all directions. The instrument was fully function with no damage found. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1827391 | 1799803 | 1705337 | 2955842201000367 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2010-07-16 | 2010-08-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2010-07-16 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-07-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors (mcs) tip cover accessory, installed on a mcs instrument, nicked the patient causing an injury to the small bowel that was successfully repaired with stitches. The tip cover was removed and replaced and the planned surgical procedure was completed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burned though the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole diameter is roughly 0. 020 and is located 0. 200 above the silicone to sleeve interface. The tip cover also has various mechanical tears in the general vicinity of the holes. The instruments and accessories user manual specifically states: - inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. - do not use another instrument to clean the tip cover accessory. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1827418 | 1799836 | 1705361 | 2955842201000374 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-07-22 | 2010-08-13 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2010-07-22 | Device Not Retu | Yes | Initial | 2010
400180 | 2010-07-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, it was observed that the plastic piece on the monopolar curved scissors (mcs) instrument was burnt. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the proximal clevis (metal, not plastic) exhibits a char mark, with the silicone proximal clevis seal exhibiting material removal. Visual inspection shows a char mark on the proximal clevis, directly above the window that reveals the proximal clevis seal. The clevis seal has a. 085 x. 025 section directly below the char mark that exhibits material removal. Engineering concluded that while the evidence is inconclusive, damage to the instrument indicates that an arcing event occurred. The clevis does not have any burrs or sharp edges near the damaged section. The instrument passed electrical continuity testing. No other damage was found. Although the mcs tip cover accessory used in conjunction with the monopolar curved scissors instrument was not returned for evaluation, engineering found that the location of the damage on the returned instrument's proximal clevis roughly aligns where the tip cover accessory silicone and ultem sleeve meet. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1827382 | 1799862 | 1705367 | 2955842201000373 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-15 | 2010-08-13 | No | Yes | Health Professional | 400180 | No | No | 2010-07-15 | Device Not Retu | Yes | Initial | 2010
IS2000 A51P8 | 2010-07-12 | Injury | HospitalizationLife Threatening | It was reported that the patient underwent a successful and uneventful da vinci s right colectomy procedure. Two days post procedure, on (b)(6), 2010, the patient was taken back to surgery for internal bleeding at the hepatic flexure. The bleeding was stopped and the patient was brought to icu. After recovery the patient was discharged home and is reported as doing well. | During additional investigation, the console surgeon reported that the patient is recovering well with no further complications and has indicated that the cause of the bleeding has been undetermined. It is also unknown if the injury was caused by inital port placement. He went on to state that this type of event could also occur in traditional laproscopic techniques. The site had advised that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the patient injury. As of (b)(4), 2010, no adverse events or system malfunctions have been experienced with the site's da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1827425 | 1799863 | 1705368 | 2955842201000372 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-07-15 | 2010-08-13 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-07-15 | No | No | Reuse | 2010
42018304 | 2010-07-13 | Malfunction |  | It was reported that 45 minutes into a da vinci s myomectomy procedure while changing instruments, the plastic housing at the distal end of the permanent cautery hook instrument's shaft became dislodged. The instrument was difficult to remove and as a result the trocar was removed with the instrument attached. At this time the site noticed that a fragment from the instrument had fallen into the patient. The fragment was successfully retrieved. This event caused a 20 minute delay; however, the planned surgical procedure was completed and no patient complications, harm, or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to be broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result. Both the proximal clevis and the main tube are missing pieces. The clevis is missing roughly half of the section that mates with the main tube interface wall. Engineering has concluded that the damage is likely due to overloading at the tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1826650 | 1798315 | 1704019 | 2955842201000364 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-07-13 | 2010-08-12 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2010-07-13 | Device Not Retu | No | Reuse | 2010
42018404 | 2010-07-13 | Malfunction |  | It was reported that 45 minutes into a da vinci s myomectomy procedure while changing instruments, the plastic housing at the distal end of the permanent cautery spatula instrument's shaft became dislodged. The instrument was difficult to remove and as a result the trocar was removed with the instrument attached. At this time the site noticed that a fragment from the instrument had fallen into the patient. The fragment was successfully retrieved. This event caused a 20 minute delay; however, the planned surgical procedure was completed and no patient complications, harm, or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to be broken at the proximal clevis to main tube interface. The clevis is dislodged from main tube as a result. Both the proximal clevis and main tube are missing pieces. The clevis is missing roughly half of the section that mates with the main tube interface wall. The clevis ear exhibits scratches/gouges above the breakage. Engineering has determined that the breakage is likely due to overloading at the tip. No other damage found. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1826651 | 1798316 | 1704020 | 2955842201000365 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-07-13 | 2010-08-12 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | Yes | 2010-07-13 | Device Not Retu | No | Reuse | 2010
| 2010-08-09 | Malfunction |  | While the surgeon was operating from the surgeon's console, arm # 1 of the davinci patient console came down on the suction irrigator and caused irrigation to shoot out for approximately 4-5 seconds until the surgical scrub could remove the irrigator from beneath the arm. Some of the irrigation ran into the base of the patient cart and at the same time the robot began alarming unrecoverable fault, turned red and the camera arm seized up. Davinci technical support immediately called and after extensive trouble shooting and several hard shut downs, the robot could not be restarted. The robot was undocked and the surgeon converted to a laparoscopic supracervical hysterectomy. There was no injury to patient during the incident and the patient remained stable. |  | DA VINCI | SURGICAL ROBOT | 1841109 | 1816536 | 1720541 | 1816536 | 1 | NAY | User Facility |  |  | Initial | 2010-08-12 | 2010-08-12 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
420006 | 2010-08-11 | Malfunction |  | The da vinci large needle driver cable broke during the procedure. The device was removed from the patient and the sterile field. There was no evidence of patient injury. This is a multi-use item, and it was the first time this device had been used. ====================== manufacturer response ======================the device was given to the business office to contact the rep and have the device replaced. |  | DA VINCI S | LARGE NEEDLE DRIVER | 1841110 | 1816537 | 1720542 | 1816537 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-08-12 | 2010-08-12 | No | Yes | Physician | 420006 | No | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-07-03 | Death | Death | It was reported that approximately 1 to 2 hours during a da vinci si colectomy / abdominoperineal resection procedure, the patient began to abnormally bleed. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. Two days post the da vinci si procedure, the patient expired. No additional information was provided. | The hospital was contacted to obtain additional information, however, despite several attempts the site has decided to not provide any information related to this event. On (b)(6) 2010, an isi clinical sales representative contacted dr. (b)(6), the operating surgeon, to obtain additional information however the surgeon did not respond. On august 10, 2010, an isi clinical sales representative stated she had left several messages with dr. (b)(6) however no response had been received. On august 10, 2010, an isi clinical sales manager stated in a conversation he had with dr. (b)(6) regarding this event, dr. (b)(6) stated that the patient had many pre-existing contributing health issues but did not wish to provide further details. No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. As of (b)(6) 2010, no adverse events have been experienced with the site's system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1824137 | 1795608 | 1701507 | 2955842201000360 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-07-12 | 2010-08-11 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2010-07-12 | No | No | Reuse | 2010
42017909 | 2010-07-10 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, a jaw of the monopolar curved scissors instrument broke and fell into the patient. The piece was retrieved and the procedure was successfully completed using a replacement instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one of the jaws was broken. The instrument was found with the left blade broken at the grip base. Both cable crimps are still in place. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1824166 | 1795641 | 1701516 | 2955842201000361 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-07-12 | 2010-08-11 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-07-12 | Device Not Retu | No | Initial | 2010
IS3000 A60P4 | 2010-07-12 | Other | Other | It was reported that during the beginning of a da vinci si prostatectomy procedure and while the system was being docked to the patient, the site experienced system error code 907 and the led on all patient side manipulator (psm) arms were illuminated red except for the third arm. All arms could be clutched expect the third arm. The site attempted to troubleshoot the problem by rebooting the system without success. With the assistance of an isi technical support engineer (tse), the site powered down the system, reconnected connection cables, and powered on the system; however, the system error continued and system error code 31030 was now displayed on the monitor. The tse had the site exercise arm 3 without success and suggested that the site could continue with the procedure using two of the three arms. At this time, the surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes 907 and 31030 were associated with the system's embedded serializer for setup joints (essj). The essj gathers position information and sends this information as serialized data to the remote arm controller (rac). The system was repaired by replacing the affected essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 907 and 31030 indicates a communication problem between the rac and essj. Upon determining this condition, the safety system puts da vinci si in a non-recoverable safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1824540 | 1796023 | 1701894 | 2955842201000362 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-07-12 | 2010-08-11 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-07-12 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-03-01 | Other |  | On (b)(6) 2010, it was reported that during a da vinci prostatectomy procedure, the staff at (b)(6) experienced multiple system error code 20013. With the assistance of an isi technical support engineer, the site reviewed the system error logs and confirmed the system error code. The site then undocked the system from the patient and turned off the system. The site unplugged a system cable and inspected the pins before reconnecting. The system was rebooted and they were able to successfully proceed with the case. On july 12, 2010, intuitive surgical received a legal complaint stating that the patient involved with the above event had suffered an injury. On july 12, 2010, the clinical quality coordinator at the site was contacted for further information and reported that during the procedure, the surgical staff experienced a non-recoverable system error and converted the case to traditional open surgical techniques to successfully complete the procedure with no patient harm or complications. The patient was discharged 2 days later and is not believed to have returned due to post operative complications. The hospital was unaware of any adverse events related to this procedure and as a result, has not reported any adverse events to the fda. | Based on the initial complaint received from the site, an investigation was conducted by field service engineering which found system error code 20013 in the system error logs as well as system error code 21013. The system error codes could not be replicated; however, the fse determined that the system error was likely associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error codes 20013 and 21013 appear when the davinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts davinci in a recoverable safe state. The ecm was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, an axis potentiometer assembly was replaced. As of the date of this report, no additional information related to the patient injury was known by the hospital and has not been provided by the patient's legal representation. If additional information becomes available, a follow-up mdr will be submitted. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1824746 | 1796227 | 1702084 | 2955842201000363 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-07-12 | 2010-08-11 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-07-12 | Device Not Retu | No | Reuse | 2010
IS2000 A51P7 | 2009-06-23 | Injury | Other | On (b)(6) 2010, isi was presented with legal documentation regarding the litigation between the plaintiff and eleven parties including intuitive surgical. The documentation states that on or about (b)(6) 2009, through (b)(6) 2009, a patient who underwent a gyn procedure at the university hospital of (b)(6) was seriously and permanently injured. No other details related to this event were reported. | Several attempts have been made to obtain additional information at the site. On july 14, 2010, (b)(4), was called and a voicemail was left requesting additional information regarding this event. On august 5, 2010, (b)(4), a second voicemail message was left to request additional information regarding this event. As of the date of this report, isi has not received additional information from this site related to this event. If additional information becomes available, a follow-up medwatch report will be submitted. As of august 5, 2010, the site has continued to use the system with no reported adverse events. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1817847 | 1787665 | 1694361 | 2955842201000355 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-07-07 | 2010-08-06 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2010-07-07 | No | No | Reuse | 2010
42018305 | 2010-06-29 | Malfunction |  | It was reported that post a da vinci si surgical procedure the cautery tip of the permanent cautery hook instrument was melted. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed evidence of arcing on the distal clevis conductor cable. The conductor cable cover was missing and the conductor cable cover was damaged. Additionally, the grip base showed signs of arcing. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1818027 | 1788060 | 1694773 | 2955842201000357 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-07-12 | 2010-08-06 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2010-07-12 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-07-09 | Other |  | It was reported that during a da vinci si surgical procedure, the customer experienced a system error. The patient was under anesthesia when the surgeon decided to abort the planned procedure and reschedule for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error experienced by the customer was associated with a patient side manipulator (psm) setup-joint axis. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The setup-joint (suj) is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The system was repaired by replacing the affected suj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The suj was returned to isi for failure analysis investigation. Engineering was able to replicate the customer reported system error when the suj axis was moved to the top range of motion. The suj was repaired by replacing the internal electrical cable harness. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1818010 | 1788101 | 1694793 | 2955842201000356 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-07-09 | 2010-08-06 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-07-09 | Device Not Retu | No | Reuse | 2010
40018302 |  | Malfunction |  | It was reported that during a da vinci surgical procedure, the wire in the jaws of the permanent cautery hook (pch) instrument was loose. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the conductor cap cover at the distal clevis missing, thus causing the conductor wire to be loose. Engineering also observed evidence of arcing near the silicon connection base, indicating that an arcing event occurred. There was no damage to the conductor wire insulation and the instrument passed electrical continuity testing. No other damage was found. On(b)(6) 2010, isi followed up with (b)(4) (initial reporter) in (b)(4) and she verified that no piece from the permanent cautery hook instrument fell into the patient. She also reported that arcing during the surgical procedure was not observed and no harm, adverse outcome or injury to the patient occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1818106 | 1788184 | 1694869 | 2955842201000358 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2010-07-07 | 2010-08-06 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | No | 2010-07-07 | Device Not Retu | No | Reuse | 2010
42017907 | 2010-05-18 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, a scissor tip on the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed no patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as both of the scissor blades were found to be intact and undamaged. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Initial | 1818217 | 1788221 | 1694902 | 2955842201000359 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2010-07-09 | 2010-08-06 | No | Yes | Health Professional | 42017907 | Device Returned To Manufacturer | Yes | 2010-08-06 | Device Not Retu | No |  | 2010
| 2010-07-29 | Malfunction |  | The da vinci surgical robot was tested prior to the procedure and there were no issues identified. It was docked and instruments were attached. The surgeon sat at the console but the arms would not work. The machine was rebooted and redocked and the error still occurred. After several attempts to fix the problem the surgery was switched to laparoscopic total abdominal hysterectomy. The technician from intuitive surgical, inc. Found a 20008 error reading and replaced arm 1. Preventive maintenance was also completed at this time. The robot is currently in use and functioning properly. |  | DA VINCI | SURGICAL ROBOT | 1824822 | 1796284 | 1702158 | 1796284 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-08-06 | 2010-08-06 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
IS3000 A60P4 | 2010-07-06 | Other |  | It was reported that prior to starting a da vinci si surgical procedure, the site experienced system error 31030. The patient was under anesthesia when the site decided to abort the planned procedure and reschedule for another date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering found system error code 31030, system error code 31065, and system error code 23 in the system error log. Engineering concluded that these system error codes were associated with a patient side manipulator (psm) remote arm controller (rac); however, the system error code was not reproducable. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The engineer reconfigured multiple psm and rac locations; however, the system error code 23 was not reproducable. Additionally, the engineer noticed that the patient side cart (psc) emergency power off (epo) breaker has been pressed, likely causing the system error codes 31030 and 31065. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the psm. System error code 31030 is reported when the system fails to restart due to the patient side cart not being reset when the other system console has been reset. System error code 31065 is reported when the subsystem was connected after the system was turned on. It is believed that the site tried to restart the system multiple times causing system error codes 31030 and 31065. In the event of these communication issues, the system records the fault and transitions to a safe state. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1816126 | 1784975 | 1694793 | 2955842201000352 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-07-06 | 2010-08-05 | Yes | Yes | Health Professional | IS3000 A60P4 | Yes | No | 2010-07-06 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-07-06 | Other | Other | It was reported that towards the end of a da vinci s prostatectomy procedure, the surgical staff experienced non-recoverable system error code # 264. An isi technical support engineer (tse) attempted to troubleshoot the issue via telephone and had the site power down, reset the breaker, emergency power off and power back on the system, however, error code # 264 immediately returned. The system error logs were reviewed by the tse and in addition to error code # 264, error code # 290 was also observed within the logs. The tse had the site power down the system again, perform an emergency power off and breaker reset, had the surgeon move the right master controller to various positions, however, error code # 264 recurred upon power up. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error codes #264 and #290 were associated with the multiple servo driver (msd) printed circuit assembly (pca) board. The system contains two msd pcas which provide drive current to the servo motors associated with each degree of freedom of a system arm. The system was repaired by replacing the affected msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #290 occurs when the system detects that the msd board did not perform required computations during a servo cycle. Upon determining this condition, the safety systems put the da vinci in a "recoverable safe state. " system error code #264 is a sympathetic fault providing additional info to the field service engineer. The msd pca was returned to isi for failure analysis investigation. Engineering discovered that one of the pins was bent, thus generating the system error codes experienced by the customer. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1850649 | 1795129 | 1701078 | 2955842201000354 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-07-06 | 2010-08-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-07-07 | Yes | No | Reuse | 2010
FORCE TRIAD | 2010-07-19 | Malfunction |  | Monopolar cut still engaged after surgeon took foot off pedal. |  | FORCETRIAD | ELECTROSURGICAL GENERATOR | 1831945 | 1806748 | 1711544 | 1806748 | 1 | GEI | User Facility |  |  | Initial | 2010-08-05 | 2010-08-05 | No | Yes | Physician | FORCE TRIAD | Yes | No Answer Provided |  |  |  |  | 2010
FORCE TRIAD | 2010-07-19 | Malfunction |  | Monopolar cut still engaged after surgeon took foot off pedal. |  | FORCETRIAD | ELECTROSURGICAL GENERATOR | 1831945 | 1806748 | 1711544 | 1806748 | 1 | GEI | User Facility |  |  | Initial | 2010-08-05 | 2010-08-05 | No | Yes | Physician | FORCE TRIAD | Yes | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2010-07-05 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced a poor 3d image and non intuitive movement of the camera arm. With the assistance of an isi technical support engineer, the site tried a replacement endoscope and recalibrated the 3d image; however, the poor image continued to persist. The surgeon did not want to continue to troubleshoot and did not want to perform the procedure in 2d vision and decided to convert the planned procedure to traditional open surgical techniques. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field engineering discovered that the left camera input cable was connected to the right camera controller unit (ccu) and the right camera input cable was connected to the left ccu, thus causing the customer reported vision issue. The camera cables connect the camera to the ccu. The ccu then calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by reconnecting the camera cables into the correct ccu input ports. The isi field engineer made the surgeon aware of the issue and provided training on proper installation od the camera cables. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1814227 | 1781816 | 1689648 | 2955842201000351 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-07-05 | 2010-08-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-07-05 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-07-09 | Other | Other | It was reported that system error code 23023 occurred while draping the patient for a da vinci s hysterectomy procedure. After the system was power cycled, system error code 1 occurred and the system would not complete power on. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that system error code 23023 was associated with a patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 23023 appears when the power supply voltage on the cannula sensor is not as expected, indicating that the cannula sensor signal cannot be trusted. System error code 1 appears when the safety system determines that a power supply voltage was out of range. Upon determining these conditions, the safety system puts da vinci s in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure mode, however, the main harness assembly was replaced as a precaution. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1814397 | 1782247 | 1689912 | 2955842201000349 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-07-09 | 2010-08-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-07-09 | Device Not Retu | No | Reuse | 2010
42022703 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, it was observed that the plastic on the wrist is damaged. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire cover at the yaw pulley exhibits damage. The distal clevis located at the damaged conductor cover has an arcing track and exhibits melting, indicating that an arcing event occurred. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1814519 | 1782334 | 1689980 | 2955842201000353 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-07-06 | 2010-08-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2010-07-06 | Device Not Retu | No | Reuse | 2010
420172 | 2010-06-22 | Malfunction |  | While the surgeon was using the davinci (b)(4) bipolar forceps it stopped working and broke while in use inside the patient. It was removed from the field and inspected. It was found to have pieces broken off of it. The patient was later found to have three broken plastic pieces in the abdominal area which where removed. ====================== health professional's impression======================device malfunction/failure. |  | ENDOWRIST | FORCEPS BIPOLAR MARYLAND | 1846542 | 1821152 | 1724596 | 1821152 | 1 | NAY | User Facility |  |  | Initial | 2010-08-03 | 2010-08-03 | No | Yes | Physician | 420172 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
| 2010-06-28 | Malfunction |  | A physician was using scissors and cauterizing close to the bladder area and noted a burned area on the tip cover. The incident occurred at approximately one hour into the surgical procedure. The physician stopped the procedure. The sheath/tip cover was changed. Areas including the patient's bladder, ureters and bowel were inspected. No injury was noted. The surgical procedure continued using shears with a new tip cover. ====================== health professional's impression======================potential tip cover failure due to design issues. This type of event has occurred before. However, the hospital has not had any incidents since (b) (6) 2008. |  | MICROTEK | TIP COVER ACCESSORY | 1830594 | 1804859 | 1709879 | 1804859 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-08-02 | 2010-08-02 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
| 2010-07-12 | Malfunction |  | Davinci needle tip driver frayed. |  | DAVINCI MEGA SUTURECUT NEEDLE DRIVER | MEGA SUTURECUT NEEDLE DRIVER | 1816255 | 1783207 | 1690701 | MW5016937 | 1 | NAY | Voluntary |  | Nurse | Initial | 2010-07-29 | 2010-07-30 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2010
IS2000 A51P8 | 2010-07-09 | Other | Other | It was reported that during a da vinci s surgical procedure, non-recoverable system error code #297 occurred. The surgical staff powered down the system, performed an emergency power off, and reset the breaker multiple times, however, error code #297 continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #297 was associated with the embedded serializer setup joint (essj) printed circuit assembly. The essj monitors the primary and secondary joint position sensors (potentiometers) for each of the joints of the setup joint arm. The setup joint arm is the passive portion of the patient side arm that aids in correctly positioning the active, motorized patient side arms. The system was repaired by replacing the affected essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #297 indicates that the system detected that an electronic component was reporting an incorrect configuration. System error code #297 then puts the da vinci s system into a non-recoverable "safe state. " the essj was returned to isi for evaluation. Engineering discovered that the ball grid array component of the essj had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872817 | 1791902 | 1698216 | 2955842201000350 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2010-07-09 | 2010-07-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-07-09 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-07-01 | Other | Other | It was reported that prior to using the system to perform any surgical tasks for a da vinci s prostatectomy procedure, it was observed that the right eye image viewed from the surgeon's console was blue. The patient was under anesthesia and no port placements had been made when the surgeon made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the systems vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. ". | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1871381 | 1791906 | 1698240 | 2955842201000347 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2010-07-01 | 2010-07-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-07-01 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-07-08 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgical staff experienced recurring recoverable system error code #23000. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the surgical staff exercise the left master tool manipulator, restart, and emergency power-off the system. These actions were performed multiple times, however, system error code #23000 continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23000 was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to the right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23000 appears when the da vinci s safety system determines that the secondary joint position sensor indicates that the joint has traveled beyond the allowed physical range of the joint. Upon determining this condition, the safety system puts da vinci s in a "recoverable safe state". The mtml was returned to isi for evaluation. Engineering discovered that the hard-stop had failed due to a pushed in pin, thus generating the system error code experienced by the customer. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872461 | 1791921 | 1698222 | 2955842201000348 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-07-08 | 2010-07-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-07-08 | Yes | No | Reuse | 2010
IS3000 A60P4 |  | Injury | Other | It was reported that post a successful da vinci si prostatectomy procedure, the patient experienced unspecified nerve damage. | On june 29, 2010 the site's robotics coordinator (b)(4) stated that an anesthesiologist (dr. (b)(6)) notified her that she was positioning patients incorrectly for robotic assisted procedures, resulting in multiple patients experiencing post procedure nerve damage. Ms (b)(4) stated that she has been positioning patients in the same manner and has not had any previous problems. The isi clinical sales representative believes that ms (b)(4) is positioning patient's in an appropriate manner based on what she has observed. - on july 13, 2010, ms (b)(4) stated that she has not had any further follow up from the dr (b)(6) after multiple requests to him for more information. The isi clinical sales representative has also attempted to reach dr (b)(6) without success. - on july 20, 2010, ms (b)(4) requested that isi follow up with dr (b)(6) directly as she has no further information. - on july 22, 2010, an urgent message was left for dr (b)(6) requesting more information about these events. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. As of july 23, 2010, no further occurrences of these events have been reported by this site. Based on the information provided by the site it was determined that the da vinci si surgical system did not malfunction in a way that would have led or contributed to the patient's injury. Please also see related manufacturer medwatch reports: 2955842-2010-00340 2955842-2010-00342 2955842-2010-00343. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1808992 | 1775169 | 1683910 | 2955842201000341 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-29 | 2010-07-29 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2010-06-29 | No | No | Reuse | 2010
| 2010-07-27 | No Answer Provided |  | Mega suture cut hand piece would not open and close while on the robotic arm. When we pulled it out and manually checked the device it would work. The device would not work on the robotic arm. The device was on use 5 of 10. ====================== health professional's impression======================staff could not identify a reason for the failure. |  | ENDOWRIST | MEGA SUTURECUT NEEDLE DRIVER | 1812198 | 1778964 | 1687464 | 1778964 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-07-29 | 2010-07-29 | No |  | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
40016907 | 2010-06-29 | Malfunction |  | It was reported that during a da vinci s assisted cystectomy procedure, the tip of the harmonic curved shears insert accessory broke off and fell into the patient. The location of the broken piece was identified via x-ray and the broken piece was successfully retrieved during the open surgical procedure portion of the planned case. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was broken off and one side of the hinge feature that accepts the clamp arm pin is bent. The evidence is not conclusive, however, the damage is consistent with hyperextension of the clamp arm. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do no use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1874294 | 1791622 | 1697981 | 2955842201000346 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-29 | 2010-07-29 | No | Yes | Health Professional | 40016907 | Yes | Yes | 2010-06-29 | Yes | Yes | Initial | 2010
IS1200 A43P9 | 2010-06-29 | Other |  | It was reported that during the beginning of a da vinci surgical procedure, the customer experienced system error codes #20005 and #20003. The patient had been administered anesthesia and ports had been placed when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes #20003 and 22005 were associated with the patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error codes #20003 and #22005 appear when the da vinci safety system determines that the measured angular position of one or more robotic joints on the specified manipulator deviated from the commanded position by more than a specified tolerance. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state. " the psm was returned to isi for failure analysis investigation. Engineering evaluation found that the cable tensioning of the psm to be out of specification, thus generating the system errors experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872836 | 1791898 | 1698213 | 2955842201000345 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2010-06-29 | 2010-07-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-06-29 | Yes | No | Reuse | 2010
42018404 | 2010-06-22 | Malfunction |  | On june 30, 2010, user facility medwatch (b)(4) was received with the following event description. "event desc: permanent cautery spatula was inserted through a trocar sleeve in the normal fashion, into chest cavity during robotic assisted coronary artery bypass graft surgery. When the device was retracted back into the camera view, it was noted that the tip of the cautery spatula was broken off the device. X-ray was taken, then c-arm was brought into the room. The device was detected on both types of x-ray. Operation resumed as planned and the piece was retracted and saved. The total timing of added search was 80 minutes. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The instrument was returned and evaluated. Per the customer reported complaint engineering observed the instrument's spatula had broken off at the distal clevis tip. There are no signs of charing or arcing. The spatula metal rod broke off at the middle circular hole cut out. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1871827 | 1791973 | 1698268 | 2955842201000344 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-06-30 | 2010-07-29 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2010-06-23 | Yes | No | Reuse | 2010
IS3000 A60P4 |  | Injury | Other | It was reported that post a successful da vinci si prostatectomy procedure, the patient experienced unspecified nerve damage. | On june 29, 2010, the site's robotics coordinator (b)(4) stated that an anesthesiologist (dr. (b)(6) notified her that she was positioning patients incorrectly for robotic assisted procedures, resulting in multiple patients experiencing post procedure nerve damage. Ms. (b)(4) stated that she has been positioning patients in the same manner and has not had any previous problems. The isi clinical sales representative believes that ms. (b)(4) is positioning patient's in an appropriate manner based on what she has observed. - on july 13, 2010, ms. (b)(4) stated that she has not had any further follow up from the dr. (b)(6) after multiple requests to him for more information. The isi clinical sales representative has also attempted to reach dr. (b)(6) without success. - on july 20, 2010, ms. (b)(4) requested that isi follow up with dr. (b)(6) directly as she has no further information. - on july 22, 2010, an urgent message was left for dr. (b)(6) requesting more information about these events. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. As of july 23, 2010, no further occurrences of these events have been reported by this site. Based on the information provided by the site it was determined that the da vinci si surgical system did not malfunction in a way that would have led or contributed to the patient's injury. Please also see related manufacturer medwatch reports: 2955842-2010-00341 2955842-2010-00342 2955842-2010-00343. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1808208 | 1773407 | 1682250 | 2955842201000340 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-29 | 2010-07-28 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2010-06-29 | No | No | Reuse | 2010
IS3000 A60P4 |  | Injury | Other | It was reported that post a successful da vinci si prostatectomy procedure the patient experienced unspecified nerve damage. | On (b)(4) 2010 the site's robotics coordinator ((b)(4)) stated that an anesthesiologist (dr (b)(6)) notified her that she was positioning patients incorrectly for robotic assisted procedures, resulting in multiple patients experiencing post procedure nerve damage. (b)(4) stated that she has been positioning patients in the same manner and has not had any previous problems. The isi clinical sales representative believes that (b)(4) is positioning patient's in an appropriate manner based on what she has observed. On (b)(6) 2010, (b)(4) stated that she has not had any further follow up from the dr (b)(6) after multiple requests to him for more information. The isi clinical sales representative has also attempted to reach dr (b)(6) without success. On (b)(6) 2010, (b)(4) requested that isi follow up with dr (b)(6) directly as she has no further information. On (b)(6) 2010, an urgent message was left for dr (b)(6) requesting more information about these events. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. As of (b)(6) 2010, no further occurrences of these events have been reported by this site. Based on the information provided by the site it was determined that the da vinci si surgical system did not malfunction in a way that would have led or contributed to the patient's injury. Please also see related manufacturer medwatch reports: 2955842-2010-00340 2955842-2010-00341 2955842-2010-00342. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1808228 | 1773427 | 1682272 | 2955842201000343 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-29 | 2010-07-28 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2010-06-29 | No | No | Reuse | 2010
IS3000 A60P4 |  | Injury | Other | It was reported that post a successful da vinci si prostatectomy procedure, the patient experienced unspecified nerve damage. | On june 29, 2010 the site's robotics coordinator (b)(4) stated that an anesthesiologist (b)(6) notified her that she was positioning patients incorrectly for robotic assisted procedures, resulting in multiple patients experiencing post procedure nerve damage. Ms. (b)(4) stated that she has been positioning patients in the same manner and has not had any previous problems. The isi clinical sales representative believes that ms. (b)(4) is positioning patient's in an appropriate manner based on what she has observed. - on july 13, 2010, ms. (b)(4) stated that she has not had any further follow up from the dr. (b)(6) after multiple requests to him for more information. The isi clinical sales representative has also attempted to reach dr. (b)(6) without success. - on july 20, 2010, ms. (b)(4) requested that isi follow up with dr. (b)(6) directly as she has no further information. - on july 22, 2010, an urgent message was left for dr. (b)(6) requesting more information about these events. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. As of july 23, 2010, no further occurrences of these events have been reported by this site. Based on the information provided by the site it was determined that the da vinci si surgical system did not malfunction in a way that would have led or contributed to the patient's injury. Please also see related manufacturer medwatch reports: 2955842-2010-00340 2955842-2010-00341 2955842-2010-00343. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1808229 | 1773428 | 1682273 | 2955842201000342 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-29 | 2010-07-28 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2010-06-29 | No | No | Reuse | 2010
400169 | 2010-03-18 | Malfunction |  | On (b)(6), 2010, intuitive surgical, inc. Received a maude event report (foi) for voluntary report (b)(4) from the fda. The event details are provided below: during a surgical procedure the da vinci harmonic instrument fell apart. The end jaw piece was found in the pt and removed. No additional information was provided. | Based on the limited information provide in the maude event report, intuitive surgical is unable to provide additional information concerning the reported event. If additional information becomes available, a follow-up mdr will be submitted. A review of our complaint database was performed; however, there was no similar event found related to this report. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1806051 | 1770709 | 1679721 | 2955842201000330 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-06-23 | 2010-07-27 | No | Yes | Health Professional | 400169 | No | No | 2010-06-23 | No | Yes | Unkown | 2010
42011704 | 2010-06-23 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, while suturing the patient's ueterous, the tip of the 5mm needle driver instrument broke and fell into the patient. The broken piece was retrieved and the surgeon proceeded with the procedure using a replacement 5mm needle driver instrument. After some time of use, the tip of the replacement instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. Mfr report 2955842-2010-00329 was submitted concerning the second event. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument has a broken jaw tip. The cable is not frayed or broken. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1806079 | 1770763 | 1679735 | 2955842201000328 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-23 | 2010-07-27 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-06-23 | Device Not Retu | No | Reuse | 2010
42011704 | 2010-06-24 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s abdominal cyst excision procedure, the tip of the 5mm needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the jaw grip broken. The cable is not broken or frayed. Removal of the housing revealed that the clamping pulley has slight traces of chemical residue. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- (b)(4). (b)(6). The event details are provided below: robotic needle holder tip broke off inside the body cavity when surgeon was suturing. Broken piece retrieved from body cavity with a new needle driver replacement. Broken needle driver given to resource nurse, (b)(6). | 5MM NEEDLE DRIVER INSTRUMENT | Reuse | 1806066 | 1770779 | 1679791 | 2955842201000331 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-07-01 | 2010-07-27 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-07-27 | Device Not Retu | No |  | 2010
42018404 | 2010-05-18 | Malfunction |  | It was reported that during a da vinci s surgical procedure the permanent cautery spatula instrument would only cauterize in short sessions. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed that the conductor wire was damaged on the outer wire sleeve cover, which likely caused the cautery issue experienced by the customer. Engineering also found burnt and melted materials at the yaw pulley and conductor wire cover, indicating that an arcing event occurred. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1806085 | 1770796 | 1679801 | 2955842201000333 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-06-23 | 2010-07-27 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2010-06-23 | Device Not Retu | No | Reuse | 2010
42011704 | 2010-06-23 | Malfunction |  | It was reported that the 5mm needle driver instrument was used as a replacement to complete a da vinci s myomectomy procedure. After some time of use, the tip of the instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. Med watch report 2955842-2010-00328 was submitted concerning the first event. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument has a broken jaw tip. The cable is not frayed or broken. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1806092 | 1770805 | 1679735 | 2955842201000329 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-23 | 2010-07-27 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2010-06-23 | Device Not Retu | No | Reuse | 2010
400169 | 2010-06-08 | Malfunction |  | On (b)(4) 2010, intuitive surgical, inc. Received uf/importer report (b)(4)from the fda. The event details are provided below: during the laparoscopy part of the robotic case the instrument being used was the harmonic. The harmonic insert piece of the instrument bent and broke. The piece was obtained whole and removed from patient. Dr was aware. The product was obtained and given to materials management. No harm to the patient. No additional information was provided. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1806098 | 1770807 | 1679812 | 2955842201000334 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-06-23 | 2010-07-27 | No | Yes | Health Professional | 400169 | No | No | 2010-06-23 | No | Yes | Unkown | 2010
IS2000 A51P8 | 2010-06-23 | Other | Other | It was reported that during a da vinci s cardiac procedure, the customer experienced a non-recoverable system error code #1. The customer was unable to clear the system error code after restarting the system multiple times. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm or complications were reported. | During the investigation conducted by the field service engineer (fse), a review of the system error log was performed and the occurrence of system error code #1 was confirmed. The fse also observed multiple occurrences of system error code #23003 in the system error log. The fse concluded that the system error code #1 and system error code #23003 were associated with a pt side manipulator (psm) arm. The psm is an instrument arm located on the pt side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarms (system generated fault codes) functioned as designed and there was no injury to the pt. System error code #1 appears when the da vinci s safety system determines that a power supply voltage was out of range. Upon determining this condition, the safety system puts da vinci s in a "non-recoverable safe state". System error code #23003 appears when the da vinci s safety system determines that the arm did not follow the test trajectory within a specified margin. The psm was returned to isi for eval, however, engineering was unable to replicate the customer reported failure mode. As a precaution, the embedded serializer pt manipulator (espm), an axis potentiometer assembly, and an axis motor/encoder were replaced. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1878496 | 1773805 | 1682633 | 2955842201000336 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-06-24 | 2010-07-23 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-06-24 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-06-25 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, a screw was becoming loose on the endoscopic camera manipulator (ecm) and an audible scratching sound could be heard. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by isi field service engineering discovered that the loose screw issue experienced by the customer was associated with the two loose screws in the ecm. The system was repaired by adjusting the screws. As of 7/21/2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1878509 | 1773806 | 1682634 | 2955842201000338 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2010-06-25 | 2010-07-23 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2010-06-25 | Yes | No | Reuse | 2010
IS3000 A60P4 | 2010-06-23 | No Answer Provided | Other | It was reported that during a da vinci si prostatectomy procedure, the site was unable to swap the pt side manipulator's (psm's). With the assistance of a technical support engineer, the site pressed e-stop and fault override; however, the psm's still would not swap. The site verified that the psm's light-emitting diode (led) was illuminate blue indicating the psm was ready for surgeon control; however, the surgeon made the decision to undock the pt from the system and abort the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with a damaged footswitch cable. The system was repaired by replacing the affected footswitch cable. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1820099 | 1773816 | 1682663 | 2955842201000337 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-06-23 | 2010-07-23 | Yes | Yes | Health Professional | IS3000 A60P4 | Yes | No | 2010-06-23 | Yes | No | Reuse | 2010
40018010 | 2010-06-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a spark was observed coming from a hole in the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The planned surgical procedure was completed with a replacement mcs instrument tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory and/or mcs instrument are returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1881231 | 1776286 | 1684951 | 2955842201000332 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-30 | 2010-07-23 | No | Yes | Health Professional | 40018010 | Yes | No | 2010-06-30 | Device Not Retu | Yes | Unkown | 2010
42022702 | 2010-05-20 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure while removing tissue for biopsy, one side of the pk dissecting forceps instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1881165 | 1776302 | 1684976 | 2955842201000335 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2010-06-24 | 2010-07-23 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer |  | 2010-09-07 | Device Not Retu | No | Reuse | 2010
42000605 | 2010-05-07 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a "screw" from the large needle driver instrument fell into the patient's abdomen. The piece was removed and the procedure was successfully completed with a replacement instrument of the same type. No patient complications, harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have a broken distal clevis pin, and missing distal clevis idler pulleys. The cable tension was affected although instrument still was able to grip and hold during live system test. No other damage was found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1881208 | 1776303 | 1684977 | 2955842201000339 | 1 | NAY | Manufacturer | User facilityForeign | UNKNOWN | Initial | 2010-06-23 | 2010-07-23 | No | Yes | Health Professional | 42000605 | Yes | No | 2010-06-23 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-06-10 | Injury | Other | It was reported that post a successful da vinci s prostatectomy procedure, the patient experienced blindness. The patient has been seen by a nuero optomologist who confirmed that the blindness was permanent. | No malfunction of the da vinci s system was alleged. Several attempts to obtain additional information related to this event have been made to the surgeon, risk management and surgical management at the site; however, no response has been received. The hospital has continued to use the system to perform surgery on patients. As of (b)(6), 2010, no adverse events have been experienced with the site's da vinci s surgical system and no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1803117 | 1768152 | 1677281 | 2955842201000327 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-22 | 2010-07-22 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-06-22 | No | No | Reuse | 2010
IS2000 A51P8 | 2010-06-21 | Death | Death | It was reported that post a successful da vinci s prostatectomy procedure, the patient went into cardiac arrest and expired. Reportedly, the patient's demise was unrelated to the da vinci s surgical system. | The hospital was contacted to obtain additional information, however, despite several attempts the site has decided to not provide any information related to this event. No malfunctions of the da vinci s surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1801430 | 1766264 | 1675641 | 2955842201000326 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-22 | 2010-07-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-06-22 | No | No | Reuse | 2010
42018304 | 2010-06-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the tip of the permanent cautery hook (pch) instrument was broken. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. | The instrument was returned and evaluated. Engineering observed the entire ceramic sleeve to be missing from the distal end. Engineering determined that the damage was likely due to mishandling. Engineering also observed the yaw pulley to exhibit localized charring/melting in the area where the ceramic sleeve use to be installed. This damage would not likely occur if the sleeve were still intact, as ceramic sleeve insulates this area during cautery activation. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796276 | 1759998 | 1669691 | 2955842201000323 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-16 | 2010-07-16 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-06-16 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-06-17 | Malfunction |  | It was reported that 2. 5 hours into a da vinci si prostatectomy procedure, one blade from the monopolar curved scissors (mcs) instrument fell off. The piece was located by x-ray and removed. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned with the scissor blade detached. Per the customer reported complaint, engineering observed one scissor blade to be broke off with the blade fracture at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit damage at the tip. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796279 | 1760005 | 1669701 | 2955842201000322 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-17 | 2010-07-16 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-06-17 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-06-17 | Injury | HospitalizationLife Threatening | It was reported that during a da vinci s hysterectomy procedure, the patient experienced a higher than normal blood loss requiring a blood transfusion. The patient was reported as being stable when the surgical staff made the decision to continue with the procedure without incident. One day post operation, the patient was observed to have internal bleeding and was brought in the operating room where it was determined that her iliac artery had been nicked. The surgical staff controlled the bleeding by packing the affected area and keeping the patient over the weekend. Two days later the bleeding was found to be under control and the packing was removed. No additional patient injury or adverse event has been reported. The surgeon has reported that the patient is doing well with no further issues. | Based on the information provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the patient injury. It is undetermined if use error was a contributing factor to this event as per the surgeon, it is possible that he unknowingly nicked the patient's illiac artery during the procedure. On (b)(6) 2010, the surgeon reported that the patient is recovering well with no further complications. As of july 16, 2010, no adverse events or system malfunctions have been experienced with the site's da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1796400 | 1760065 | 1669736 | 2955842201000319 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-21 | 2010-07-16 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-06-21 | No | No | Reuse | 2010
42017909 | 2010-05-05 | Malfunction |  | It was reported that during a da vinci s surgical procedure, while coagulating tissue, the grip on the monopolar curved scissors (mcs) instrument broke and a piece from the instrument fell into the patient. The broken piece was not retrieved from the patient. The planned surgical procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796636 | 1760234 | 1669919 | 2955842201000324 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2010-05-15 | 2010-07-16 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-07-26 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796631 | 1760250 | 1669915 | 2955842201000318 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-07-16 | No | Yes | Health Professional | 42017909 | No | Yes | 2010-05-10 | No | No | Unkown | 2010
42017909 | 2010-03-23 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the hospital, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796649 | 1760251 | 1669916 | 2955842201000317 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-07-16 | No | Yes | Health Professional | 42017909 | No | Yes | 2010-05-10 | No | No | Unkown | 2010
42017909 | 2010-03-23 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip of the monopolar curved shears (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the hospital, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796638 | 1760253 | 1669916 | 2955842201000321 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-07-16 | No | Yes | Health Professional | 42017909 | No | Yes | 2010-05-10 | No | No | Unkown | 2010
42017909 | 2010-04-06 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip of one of the blades on the monopolar curved shears (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the hospital, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1796639 | 1760254 | 1669938 | 2955842201000320 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-07-16 | No | Yes | Health Professional | 42017909 | No | Yes | 2010-05-10 | No | No | Unkown | 2010
IS2000 A51P4 |  | Injury | Disability | It was reported that during a da vinci s prostatectomy procedure performed on (b)(6) 2008, the surgeon experienced interference in the left eye image viewed from the surgeon's console. The surgeon switched the image from 3-dimensional (3-d) to 2-dimensional (2-d) to proceed with the procedure. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. The customer-reported-event does not in itself constitute a reportable event, however, on june 17, 2010, isi received a legal summons and complaint filed by the patient, alleging that due to the amount of time added by completing the surgery in 2-d (the surgery lasted over 8 hours) the patient sustained a severe and permanent right leg injury, specifically right calf compartment syndrome. The complaint further alleges that the patient has undergone numerous and extensive medical procedures, including surgery and therapy. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The camera cable was returned to the original equipment manufacturer (oem) for evaluation. The oem observed that the left channel of the camera cable had a loose connection, thus causing the vision issue experienced by the customer. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1796669 | 1760256 | 1669939 | 2955842201000325 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-06-17 | 2010-07-16 | Yes | Yes | Health Professional | IS2000 A51P4 | Device Returned To Manufacturer | No | 2010-06-17 | Device Not Retu | No | Reuse | 2010
42017909 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while dissecting tissue, one side of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retreived and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade is broken off the scissor grip. The cable crimps are still in place on the yaw pulley. Removal of the housing revealed that the clamping pulleys have slight traces of chemical residue. No other damage was found. Per the information provided by the distributor, the broken piece was successfully retreived from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1795419 | 1758920 | 1668723 | 2955842201000316 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-06-15 | 2010-07-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-06-15 | Device Not Retu | No | Reuse | 2010
DA VINCI 1S200 | 2010-06-23 | Injury | Hospitalization | Pt scheduled for robotic laparoscopic prostatectomy. After the procedure began and the robot was docked to the trocars, the third arm was not being recognized from the robotic console. Sales rep was present for case and tried troubleshooting and speaking with company rep. All attempts were made to resolve the issue. The surgeon spoke with the pt's family and a decision was made to abort the case and reschedule for a later time. The surgical services dept. Director contacted the company and the equipment was serviced the same day of the incident. The equipment was repaired and cleared for use. |  | DA VINCI | ROBOTIC SYSTEM | 1804112 | 1765026 | 1674422 | MW5016760 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-07-15 | 2010-07-15 | No | Yes | Health Professional | DA VINCI 1S200 | Yes | Yes |  |  |  |  | 2010
DA VINCI 1S200 | 2010-06-23 | Injury | Hospitalization | Pt scheduled for robotic laparoscopic prostatectomy. After the procedure began and the robot was docked to the trocars, the third arm was not being recognized from the robotic console. Sales rep was present for case and tried troubleshooting and speaking with company rep. All attempts were made to resolve the issue. The surgeon spoke with the pt's family and a decision was made to abort the case and reschedule for a later time. The surgical services dept. Director contacted the company and the equipment was serviced the same day of the incident. The equipment was repaired and cleared for use. |  | DA VINCI | ROBOTIC SYSTEM | 1804112 | 1765026 | 1674422 | MW5016760 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-07-15 | 2010-07-15 | No | Yes | Health Professional | DA VINCI 1S200 | Yes | Yes |  |  |  |  | 2010
42017206 | 2010-06-15 | Malfunction |  | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that during a da vinci si gynecological lysis of adhesions, while the surgeon was moving instruments, the wrist of the maryland bipolar forceps instrument was inadvertently forced along the ribbed portion of the cannula. The surgical staff noticed that a cable from the maryland bipolar forceps instrument had broken and three pieces fell into the patient. The fragments were immediately retrieved from the patient and the instrument was replaced with another instrument of the same type. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. As of july 13, 2010 there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering found the pitch up cable is broken near the distal clevis hub and proximal clevis cable hole. The cable segment that contains the crimp is still installed in the clevis and the broken strands stick out at the wrist. The clevis does not exhibit any wear marks or damage and other cables at the wrist are not damaged. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 1792521 | 1755495 | 1665664 | 2955842201000314 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | InitialFollowup | 2010-06-15 | 2010-07-13 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-06-15 | Device Not Retu | No |  | 2010
42017908 | 2010-06-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci hysterectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The piece was retrieved and the procedure was completed. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering found the right blade to be broken off and missing with a 0. 39 inch by 0. 13 inch piece missing. No bending of the remaining blade was found and no other damage was found. Per the case notes the missing instrument components which fell inside of a patient was retrieved. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1792526 | 1755514 | 1665678 | 2955842201000315 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2010-09-07 | 2010-07-13 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-06-16 | Device Not Retu | No |  | 2010
IS3000 A60P3 | 2010-04-08 | Other | Other | It was reported that during a da vinci si prostatectomy procedure the site experienced a white image in the left eye view of the surgeon console. The site swapped the camera cable ends in an attempt to the resolve the issue, however, the issue moved to the right side image. Prior to contacting an isi technical support engineer for assistance the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering discovered that the vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci si vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1792506 | 1755524 | 1665673 | 2955842201000178 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-04-08 | 2010-07-13 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-04-08 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-06-15 | Other | Other | It was reported that while preparing the da vinci system for a prostatectomy procedure, the site experienced system error code 21003 and the system would not pass the homing test upon start up. The patient was under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. | The investigation conducted by field service engineering determined that system error code 21003 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. Upon further inspection, a plastic obstruction was found to be blocking the psm's movement. The system was repaired by removing the obstruction. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. During system startup, each arm is commanded move through a range of motion in order to verify functionality. System error code 21003 appears when the da vinci safety system determines that the arm did not perform the commanded motions during startup within a specified tolerance. Upon determining these conditions, the safety system puts da vinci in a recoverable safe state. As of july 8, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1790508 | 1753173 | 1663520 | 2955842201000308 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-06-15 | 2010-07-12 | Yes | Yes | Health Professional | IS1200 A43P9 | No | No | 2010-06-15 | Device Not Retu | No | Reuse | 2010
| 2010-07-08 | Malfunction |  | The device was not used on patient during procedure; it was noted to have a missing blade. ====================== health professional's impression======================no adverse event. |  | HARMONIC | CURVED SHEARS INSERT | 1801470 | 1766289 | 1675658 | 1766289 | 1 | LFL | User Facility |  | RISK MANAGER | Initial | 2010-07-12 | 2010-07-12 | No | Yes | Invalid Data |  | No | No Answer Provided |  |  |  |  | 2010
42018304 | 2010-06-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, one of the jaws on the permanent cautery hook (pch) instrument broke off. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. | The customer reported failure mode as the reported instrument is a hook instrument and does not have jaws. The hook is intact and does not exhibit damage. Engineering also observed that the conductor wire is broken at the distal end, near the yaw pulley interface. The yaw pulley and inner surface of the distal clevis ear exhibits charring, indicating that an arcing event occurred. The conductor cap is missing as is one of the idler pulleys on the distal clevis. The yaw pulley extruded boss feature on the same side as conductor cap is broken off. Breakage of the boss feature allowed the cap to slip out. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1788843 | 1751475 | 1661910 | 2955842201000304 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-17 | 2010-07-10 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2010-06-17 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-06-10 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while dissecting tissue, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient's abdomen. The broken piece was retrieved and the planned procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with one scissor blade broken off and is missing. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The tip of the intact blade does not exhibit any damage. The instrument passed electrical continuity testing. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1788844 | 1751476 | 1661911 | 2955842201000305 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-10 | 2010-07-10 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-06-10 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-06-10 | Death | Death | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, an unknown event occurred requiring the site to use a deliberator on the patient. After this event, the site wanted to continue with the scheduled procedure; however, the system would not power back on. A technical support engineer was called to troubleshoot the problem and the power breaker on the surgeon side consoled (ssc) was found to require cycling. At this time the system immediately powered up and the site confirmed that the system is working properly. Shortly after the surgeon console powered up, it was reported that the patient's aorta was nicked and the surgeon decided to convert the procedure to traditional open surgical techniques due to patient complications. It was reported that the patient expired shortly after the conversion. --/-- Additional Event Description: --/-- On (b)(6), 2010 (b)(6), risk manager at (b)(6) hospital, submitted a letter after reviewing medwatch manufacturer report number 2955842-2010-00306 stating: at no time during the laparoscopic procedure did the da vinci si surgeon side console experience a loss of power while in use nor was the aorta nicked as indicated in the medwatch report. Although a call was placed by operating room personnel to an intuitive surgical technical support field engineer on (b)(6), this call was placed after the procedure concluded. It was noted during a check of the equipment after the procedure that the surgical side console would not power on. It is my understanding that during the troubleshooting conversation between the operating room personnel and the technical support field engineer that it was reported to the technical support field engineer that a defibrillator had been plugged into the same electrical wall socket as the surgical side console and that the defibrillator had been activated. Please be adviced that the defibrillator was not activated during the da vinci procedure. | Manufacturer Narrative: --/-- Based on the information provided, it was determined that the da vinci si surgical system did not cause or contribute to the patient's demise. Multiple requests for additional information, including the surgical report and patient's autopsy report, have been made however at the time of this report no additional information has been provided by the hospital to intuitive surgical. The hospital has continued to use the system to perform surgery on patients. As of july 8, 2010, no adverse events have been experienced with the site's da vinci si surgical system and no similar instances of this event has been reported to isi. --/-- Additional Manufacturer Narrative: --/-- As stated by (b)(6) the initial event reported by intuitive surgical on july 9, 2010 via medwatch manufacturer report number 2955842-2010-00306, was inaccurate. (b)(6) stated that she is unable to provide intuitive surgical any additional information about this case and if the fda requires additional information that intuitive surgical refer them directly to her. | DA VINCI SI SURGICAL SYSTEM | Reuse | 1788845 | 1751478 | 1661912 | 2955842201000306 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | InitialFollowup | 2010-08-27 | 2010-07-10 | Yes | Yes | Health Professional | IS3000 A60P4 | No | No | 2010-08-20 | Device Not Retu | No |  | 2010
42022702 | 2010-04-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a plastic piece of the pk dissector forceps instrument wrist broke and fell into the patient. The plastic piece was retrieved and the planned surgical procedure was completed with a replacement instrument. No patient harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal clevis to be chipped and the conductor cable broken off of the grip jaw. The instrument housing was removed and chemical residue was found on the clamping pulleys. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1788846 | 1751479 | 1661913 | 2955842201000303 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2010-06-11 | 2010-07-10 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | Yes | 2010-06-11 | Device Not Retu | No | Reuse | 2010
IS3000 A60P4 | 2010-06-14 | Other |  | It was reported that prior to starting a da vinci si prostatectomy procedure, the site experienced a system error code 31030 when starting up the system. With the assistance of a technical support engineer, the site power cycled the system and reset the cycle breakers on the vision side cart and the patient side cart, however the system error 31030 reoccurred when the system was powered on. The patient was under anesthesia for approximately one hour when the surgical staff made the decision to abort the planned surgical procedure and reschedule to a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering found system error code 31030 and system error code 23 in the system logs and determined they were associated with an embedded serializer set up joint (essj) on the setup joint (suj), which is an unpowered configurable arm which supports the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The essj is a printed circuit assembly(pca) inside the setup joint that monitors the sensors in the suj and relays information between the ecm and controlling processor. The system was repaired by replacing the affected essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by the software to denote communication faults in the low voltage differential signal carrying information about the ecm. System error code 31030 occurs when the patient side cart fails to restart when the core attempts to restart. In the event of these communication issues, the system records the fault and transitions to a safe state. The essj was returned to isi for failure analysis. Engineering was able to immediately replicate the system error 23 when wiggling the essj cable harness. The essj cables and internal electrical harness were replaced to resolve the system errors. As of july 8, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURIGCAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1788887 | 1751480 | 1661914 | 2955842201000307 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-14 | 2010-07-10 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-06-14 | Device Not Retu | No | Reuse | 2010
40016907 | 2010-06-09 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The hcs insert was retrieved and the planned surgical procedure was completed using a replacement insert accessory. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the clamp arm missing from the distal end. One side of the features that mates with the clamp arm pins are bent slightly. The curved blade also exhibits various scratch marks on its outer surface. Evidence not conclusive, but damage may be due to hyperextension of the clamp arm or similar kind of mishandling. No other damage found. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1788888 | 1751481 | 1661915 | 2955842201000311 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-10 | 2010-07-10 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-06-10 | No | Yes | Initial | 2010
420179 | 2010-04-15 | Malfunction |  | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the nurse observed that the mcs tip cover accessory installed on the monopolor curved scissors (mcs) instrument was torn. Additional examination performed by the nurse after the instrument was removed, revealed that both the tip cover accessory and mcs instrument exhibited burn marks. The planned procedure was completed with a replacement instrument and no patient harm, adverse outcome or injury was reported. The following med watch report was also sent to the fda regarding the mcs tip cover accessory used in conjunction with the mcs instrument: 2955842-2010-00310. | Manufacturer Narrative: --/-- The monopolar curved scissors (mcs) instrument and mcs tip cover accessory were returned and evaluated. Per the customer reported complaint, engineering observed that the mcs instrument with the tip cover accessory installed has an arc trac through the tip cover accessory at the tube extension and main tube connection point, indicating that an arcing event occurred. The mcs instrument also exhibits charing and melted material. No other damage was found. Refer to follow-up mdr (001) 2955842-2010-00310 regarding the evaluation results of the returned mcs tip cover accessory. --/-- Additional Manufacturer Narrative: --/-- The mcs instrument and tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument and tip cover accessory are returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1788895 | 1751488 | 1661922 | 2955842201000309 | 1 | NAY | Manufacturer | DistributorOther | Nurse | InitialFollowup | 2010-06-11 | 2010-07-10 | No | Yes | Health Professional | 420179 | Device Returned To Manufacturer | Yes | 2010-07-28 | Device Not Retu | No |  | 2010
40018010 | 2010-04-15 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the nurse observed that the mcs tip cover accessory installed on the monopolor curved scissors (mcs) instrument was torn. Additional examination performed by the nurse after the instrument was removed, revealed that both the tip cover accessory and mcs instrument exhibited burn marks. The planned procedure was completed with a replacement instrument and no patient harm, adverse outcome or injury was reported. The following medwatch report was also sent to the fda regarding the mcs instrument used in conjunction with the tip cover accessory: 2955842-2010-00309. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover was returned installed on the monopolar curved scissors (mcs) instrument. The tip cover has a. 225 x. 190 hole burned through the ultem sleeve and silicone, indicating that an arcing event occurred. The top of the hole is located along the silicone to sleeve interface. The sleeve has an axially oriented trail of charring and a couple scratch marks below the hole. The silicone part of the tip cover has a couple of mechanical tears on the opposite side of the hole. The location of the hole on the tip cover aligns perfectly with char mark on mcs instrument. No other damage was found. Please refer to follow-up mdr (001) 2955842-2010-00309 regarding the evaluation results of the returned mcs instrument. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | Unkown | 1788896 | 1751490 | 1661923 | 2955842201000310 | 1 | NAY | Manufacturer | DistributorOther | Nurse | InitialFollowup | 2010-06-11 | 2010-07-10 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-07-29 | Device Not Retu | Yes |  | 2010
IS2000 A51P8 | 2010-06-16 | Other | Other | It was reported that during a da vinci s gynecologic procedure, the left eye image viewed from the surgeon's console was black. The surgical staff replaced the camera cable, however, the issue remained. The patient was under anesthesia for approximately 2 hours when the surgeon made the decision to convert to open surgical techniques to complete the planned procedure due to the vision issue and the patient's enlarged uterus. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the initial vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. It was later discovered that the site had inadvertently replaced the defective camera cable with another defective camera cable and as a result, the surgical staff continued to have vision issues. Further investigation revealed that the camera cable was known to be defective, however, it was not properly segregated by the site. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1789335 | 1751948 | 1662360 | 2955842201000312 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-06-16 | 2010-07-10 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-06-16 | Device Not Retu | No | Reuse | 2010
40020504 | 2010-06-16 | Malfunction |  | On (b)(4) 2010, intuitive surgical received a copy of medwatch uf/importer report (b)(4) from the fda, with the following information: event desc: during laparoscopic prostatectomy, md looked at the left bipolar forceps and could only find 1 arm of the bipolar forceps. The surgeon looked around inside pt's abdomen and found the other half of the arm and removed it. Intuitive rep was in the room at the time and stated that there was no such recorded episode. Nothing unusual had occurred prior to this with regard to the davinci. Manufacturer response for fenestrated bipolar forceps, intuitive surgical awaiting formal request for return of device. (b)(4). Laparoscopic prostatectomy device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned (post-evaluation) or if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1789341 | 1751958 | 1662366 | 2955842201000313 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-06-30 | 2010-07-10 | No | Yes | Health Professional | 40020504 | No | No | 2010-06-30 | No | No | Unkown | 2010
IS1200 A43P9 | 2010-06-15 | Other |  | It was reported that during a da vinci prostatectomy procedure, the customer experienced repeated occurrences of system error, which then became non-recoverable. An isi technical support engineering attempted to troubleshoot the issue via telephone and had the surgical staff emergency power off and restart the system twice. The site was able to re-home the system and then error recurred when the right master tool manipulator (mtmr) was moved. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a non-recoverable safe state. The mtmr was returned to isi for failure analysis investigation. Engineering evaluation found that the harness conductor had malfunctioned, thus generating the system error experienced by the customer. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1785336 | 1747231 | 1658099 | 2955842201000301 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-06-15 | 2010-07-06 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-06-15 | Device Not Retu | No | Reuse | 2010
420006 | 2010-06-10 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a fragment of the large needle driver instrument fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1785346 | 1747235 | 1658103 | 2955842201000302 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-06-10 | 2010-07-06 | No | Yes | Health Professional | 420006 | No | Yes | 2010-06-10 | No | No | Unkown | 2010
IS2000 A51P8 | 2010-06-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgical staff experienced multiple occurrences of recoverable system error code 23013. The site power cycled the system, however, the same error code continued to appear. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the surgeon exercise the right master tool manipulator (mtmr). The site was able to use the system for approximately 2 minutes when the system faulted again with error code 23013. The patient had been under anesthesia for approximately 1. 5 hours and the system had been docked to the patient for approximately 30 minutes when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23013 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23013 appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The mtmr was returned to isi for evaluation. Engineering determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of july 6, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1785299 | 1747263 | 1658123 | 2955842201000300 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-06-15 | 2010-07-06 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-06-15 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-06-01 | Death | Death | Event Description: --/-- It was reported that the patient underwent a successful da vinci s transoral robotic surgery on (b)(6), 2010 to remove a tumor at the base of the patient's tongue. The physician made the decision to keep the patient over the weekend and on (b)(6), 2010, the patient was seen by a resident who observed him vomiting blood. The patient required intubation; however, prior to completing intubation, the patient expired. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- Based on the information provided by the console surgeon, it was determined that the da vinci s surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the patient's demise. The console surgeon advised that an autopsy was performed on (b)(6), 2010, which found the primary cause of death to be a gi bleed caused by a stressed induced ulcer. The surgeon reported that the patient had been on a type of cleansing diet prior to the surgery that may have contributed to this event. No additional information was provided. The hospital has continued to use the system to perform surgery on patients. As of (b)(6), 2010, no adverse events have been experienced with the site's da vinci s surgical system and no similar instances of this event have been reported to isi. --/-- Additional Manufacturer Narrative: --/-- Product code has been removed per fda request. | DA VINCI S SURGICAL SYSTEM | Reuse | 1783350 | 1744642 | 1655658 | 2955842201000291 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | FollowupInitial | 2010-06-01 | 2010-07-02 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2010-06-29 | Device Not Retu | No |  | 2010
40017908 | 2010-05-06 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The procedure was converted to traditional laparoscopic techniques in order to retrieve the broken piece. The procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the left blade was returned detached from the instrument. The blade was found to have fractured at the cross section of the grip cable crimp resulting in half of the cable crimp becoming exposed. The fractured plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1783905 | 1745261 | 1656232 | 2955842201000297 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-06-08 | 2010-07-02 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2010-06-08 | Device Not Retu | No | Reuse | 2010
400169 | 2010-06-07 | Other | Other | It was reported that the harmonic curved shears (hcs) insert accessory installed on the hcs instrument was used as replacement to complete a da vinci s hysterectomy procedure. After some time of use, the insert accessory broke and a piece of the insert accessory fell into the patient. The broken piece was not retrieved. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda regarding the first event: 2955842-2010-00295. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2010, rn, (b)(6) reported that during the procedure, the surgical staff observed that the surgeon grasped another endowrist instrument also being used during the procedure with the insert accessory, thus causing the insert accessory to break. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1783933 | 1745293 | 1656256 | 2955842201000294 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-07 | 2010-07-02 | Yes | Yes | Health Professional | 400169 | No | No | 2010-06-07 | No | Yes | Unkown | 2010
42020703 | 2010-06-03 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, after 30 minutes of use, the proximal clevis on the tenaculum forceps instrument broke, thus preventing removal of the instrument from the trocar. When the surgical staff attempted to remove the instrument, pieces from the instrument fell into the patient's abdomen. The broken pieces were retrieved and the planned procedure was completed. Simultaneous removal of the tenaculum forceps instrument and trocar was required. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to be broken at the proximal clevis to the main tube interface. The clevis is dislodged from main tube as a result. Both the proximal clevis and main tube are missing pieces. Per the case notes, all of the broken pieces were successfully removed from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1783935 | 1745297 | 1656253 | 2955842201000292 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-06-03 | 2010-07-02 | No | Yes | Health Professional | 42020703 | Device Returned To Manufacturer | No | 2010-06-03 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci prostatectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the distributor, the broken piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the left scissor blade broken off. Both of the cable crimps remain on the yaw pulleys. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1783951 | 1745299 | 1656255 | 2955842201000296 | 1 | NAY | Manufacturer | DistributorOther | Other | InitialFollowup | 2010-06-08 | 2010-07-02 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-07-09 | Device Not Retu | No |  | 2010
40016907 | 2010-06-07 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, that the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The insert accessory was retrieved and the surgeon proceeded with the procedure using a replacement insert accessory. After some time of use, a piece from the replacement insert accessory broke off and fell into the patient. The broken piece from the replacement insert accessory was not retrieved. No patient harm, adverse outcome or injury was reported. The following med watch report was sent to the fda regarding the second event: 2955842-2010-00294. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering found the hinge feature at the distal end of the inner tube which mates with the clamp arm pins, to be bent. As a result the clamp becomes loose or disengaged. The clamp arm was returned with the insert, however no parts are missing. The teflon pad remains affixed to the arm. The curved blade presents gouge and scratch marks on different sections from the tip to the base. Engineering concluded that breakage of the harmonic curved shears (hcs) insert accessory was most likely caused by intra-operative instrument collisions. No other damage was found. On (b)(6), 2010, rn, (b)(6) reported that during the procedure the surgical staff observed that the surgeon grasped another endowrist instrument also being used during the procedure with the insert accessory, thus causing the insert accessory to break. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1783936 | 1745300 | 1656256 | 2955842201000295 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-06-07 | 2010-07-02 | Yes | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-06-07 | Device Not Retu | Yes | Unkown | 2010
40018010 | 2010-05-10 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, while preforming dissection, the surgeon observed arcing to adjacent tissue from the monopolar curved scissors (mcs) with the mcs tip cover accessory installed. Examination of the tip cover accessory by the surgical staff revealed that the tip cover had two holes. The planned surgical procedure was completed with a replacement tip cover. No patient complications or adverse outcome was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the tip cover has multiple holes approximately roughly. 020 in diameter burned through the silicone. The silicone exhibits localized melting around the holes, indicative that multiple arcing events occurred. The holes are located. 200 and. 300 above the silicone to sleeve interface. The tip cover also has a mechanical tear in the silicone near one of the holes. The endowrist instruments instructions for use (ifu) specifically state: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4), 2010 user facility med watch report (b)(4)was received indicating that the patient was in stable condition post operatively and no complications were reported during the follow up examination. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1783972 | 1745424 | 1656371 | 2955842201000298 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-03 | 2010-07-02 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-06-03 | Device Not Retu | No | Unkown | 2010
40017908 | 2010-05-03 | Malfunction |  | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that during a da vinci hysterectomy procedure, while performing a pelvic lymphadenectomy a blade on the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade was returned detached from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the physician, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1783975 | 1745442 | 1656398 | 2955842201000299 | 1 | NAY | Manufacturer | User facilityOther | Physician | InitialFollowup | 2010-06-08 | 2010-07-02 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | Yes | 2010-07-06 | Device Not Retu | No |  | 2010
40016907 | 2010-05-17 | Malfunction |  | On (b)(4) 2010, isi received user facility medwatch (b)(4) reporting the following. Event desc: the harmonic da vinci scissor grasper: the tip of the instrument broke off during robotic surgery. The tip was retrieved. (b)(4). Robotic assisted left nephrectomy device usage problem: device failed. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was separated from the insert and the hinge feature at the distal end of the inner tube that mates with the clamp arm pins are bent. No parts are missing. The teflon pad remains affixed to arm, but presents some damage around the midpoint where small pieces were removed. The curved blade presents gouge and scratch marks on different sections from the tip to the base. Evidence suggests the insert was misused/mishandled; however, this is not conclusive. No other damage found. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1782998 | 1744326 | 1655332 | 2955842201000288 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-06-01 | 2010-07-01 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-06-01 | Device Not Retu | Yes | Initial | 2010
400077 | 2010-05-18 | Malfunction |  | It was reported that during a da vinci s surgical procedure, pieces of the plastic from the 8mm cannula seal accessory broke off of the seal and fell into the patient. The pieces were retrieved to complete the planned procedure. No patient harm, adverse outcome or injury was reported. The site reports that this has occurred in one additional case. See mdr's 2955842-2010-00290. | The cannula seal accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument or accessory are returned for evaluation, a follow-up mdr will be submitted. | 8MM CANNULA SEAL ACCESSORY | ENDOSCOPIC ACCESSORY | 1782994 | 1744342 | 1655366 | 2955842201000289 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-02 | 2010-07-01 | No | Yes | Health Professional | 400077 | No | Yes | 2010-06-02 | No | Yes | Unkown | 2010
400077 | 2010-06-01 | Malfunction |  | It was reported that during a da vinci s surgical procedure, pieces of the plastic from the 8mm cannula seal accessory broke off of the seal and fell into the patient. The pieces were retrieved to complete the planned procedure. No patient harm, adverse outcome or injury was reported. The site reports that this has occurred in one additional case. See mdr's 2955842-2010-00289. | The cannula seal accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument or accessory are returned for evaluation, a follow-up mdr will be submitted. | 8MM CANNULA SEAL ACCESSORY | ENDOSCOPIC ACCESSORY | 1782955 | 1744344 | 1655368 | 2955842201000290 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-02 | 2010-07-01 | No | Yes | Health Professional | 400077 | No | Yes | 2010-06-02 | No | Yes | Unkown | 2010
42017909 | 2010-05-26 | Malfunction |  | It was reported that during a da vinci si colectomy procedure while working on the pt's gallbladder, the scissors on the monopolar curved scissors (mcs) instrument broke and fell into the pt. The planned surgical procedure was completed, however, the broken scissors were not retrieved from the pt. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if add'l info is received. On (b)(6) 2010, the distributor (initial reporter) reported that no add'l procedure was required to locate the broken piece that was not retrieved from the pt. The pt's hospitalization was not prolonged and the pt has not returned to the hospital with any post operative complications. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1806307 | 1763511 | 1673001 | 2955842201000293 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-06-01 | 2010-07-01 | Yes | Yes | Health Professional | 42017909 | No | No | 2010-06-01 | Device Not Retu | No | Unkown | 2010
42018304 | 2010-05-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the yellow housing on the tip of the permenant cautery hook instrument broke and fell into the patient. The surgeon was unable to retrieve the broken piece. The planned surgical procedure was completed and no patient complications, harm, injury or adverse outcome was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one ear of the distal clevis to be broken off and missing. The missing piece measures approximately. 281 inches by. 240 inches. The conductor wire cap was also found to be missing. Electrical continuity passed. Multiple attemps to obtain further information from the site regarding the patient status has been made however at the time of this report the site has not responded. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1780433 | 1741223 | 1652547 | 2955842201000287 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-06-02 | 2010-06-29 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | Yes | 2010-06-02 | Device Not Retu | No | Reuse | 2010
| 2010-06-24 | Malfunction |  | According to mfr, potts scissors are re-usable ten times. Following the cleaning after the first use, 2-8mm "o" rings were found with instrument. No longer usable, requiring operating room to order new ones. |  | DA VINCI | POTTS SCISSORS | 1783268 | 1743026 | 1654183 | MW5016543 | 1 | LRW | Voluntary |  | Nurse | Initial | 2010-06-25 | 2010-06-28 | No | Yes | Service Personnel |  | Yes | Yes |  |  |  |  | 2010
ISI200 A43P9 | 2010-05-27 | Other | Other | It was reported that during a da vinci surgical procedure, one of the instrument arms became stuck and the surgical staff could not get it to move when pressing the clutch button. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Conclusions: the investigation conducted by the isi field service engineer concluded that the system issue experienced by the customer was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. The ria pca performs numerous tasks related to the function, control and hardware fault monitoring for a specific arm. There is a ria pca assigned to each system arm. The system was repaired by replacing the affected ria pca. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872363 | 1743037 | 1654193 | 2955842201000285 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-05-27 | 2010-06-25 | Yes | Yes | Health Professional | ISI200 A43P9 | Yes | No | 2010-05-27 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-05-27 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced system error codes #23008 and #23013. The site undocked the arm and restarted the system; however, the system error codes returned after 5 minutes of use. Prior to calling an isi technical support engineer, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Conclusions: an investigation was performed by a field service engineer (fse). The system error logs were reviewed and system error codes 23008 and 23013 were confirmed to have occurred. The fse also observed system error code #23004. The fse determined that the system error codes were associated with the system's psm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error codes #23008 and #23013 appear when the da vinci s safety system determines that the angular position (#23008) or rate of change in angular position (#23013) of one or more robotic joints on the specified manipulator, as measured by that joint primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. The system error code #23004 appears when the da vinci s safety system detects that the measured rate of change in joint velocity differs from the expected value by more than a specified tolerance. This error is intended to detect rare failures in the internal memory or processor that controls the robot motors. It can be triggered without any system deficiencies when instruments with a lot of friction are moved rapidly or when two arms collide outside the patient. Upon determining these conditions, the safety system put davinci in a "recoverable safe state. ". | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872359 | 1743039 | 1654195 | 2955842201000286 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2010-05-27 | 2010-06-25 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-05-27 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2010-05-26 | Other | Other | It was reported that during a da vinci surgical procedure, the site experienced system error code #20008. An isi technical support engineer (tse) was contacted via telephone, who had the site reboot the system. The system was restarted with no issues and the planned procedure was successfully completed. The tse was contacted after the procedure, who informed the surgical staff that the system should not be used for another procedure as the system error code would likely return and become nonrecoverable. The site decided to proceed with their next scheduled da vinci surgical procedure and prior to using the system to perform any tasks, the site again experienced system error code #20008. The tse attempted to troubleshoot the issue and had the surgical staff power cycle and emergency power off the system multiple times, however, system error code #20008 continued to recur during homing. The patient had been under anesthesia for approximately 1 hour and 15 minutes when the surgeon made the decision to abort the planned procedure and reschedule it to a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer (fse) concluded that system error code #20008 was associated with an input output distribution (iod) / remote interface adapter (ria) cable. The fse found a recessed pin at the connection end of the cable, preventing the cable from staying seated. The iod/ria cable is a bundle of wires which transmits power and sensor information between the surgeon side cart and the ria board associated with a particular slave arm. The system was repaired by replacing the affected ria cable. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #20008 appears when the da vinci safety system determines the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's, primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a "nonrecoverable safe state". As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1872170 | 1743065 | 1654239 | 2955842201000278 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-05-26 | 2010-06-25 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-05-26 | Yes | No | Reuse | 2010
40017909 | 2010-06-03 | Malfunction | Required Intervention | It was reported that after a da vinci surgical procedure was completed, the surgeon noticed that a piece had broken off of the monopolar curved scissors instrument used during the case. A radiography was performed on the pt and the broken instrument piece was observed to be inside of the pt. An add'l laparoscopic procedure was performed to remove the broken piece from the pt cavity. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was missing from the instrument. The blade had fractured at the cross section of the grip cable crimp, exposing half of the cable crimp. The fracture plane of the blade is nearly straight. The detached blade was returned with the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1878529 | 1762759 | 1672249 | 2955842201000283 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2010-06-03 | 2010-06-25 | Yes | Yes | Health Professional | 40017909 | Yes | No | 2010-06-03 | Yes | No | Reuse | 2010
42017206 | 2010-05-27 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, electrical current was not transmitting from the tip of the maryland bipolar forceps instrument. The site was unable to use the instrument and the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported problem. The instrument failed electrical continuity testing and was found to have damage to the tip of the conductor wires. The housing was removed to find one wire detached from the bipolar pin. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1878524 | 1762760 | 1672250 | 2955842201000284 | 1 | NAY | Manufacturer | ForeignDistributor | UNKNOWN | Initial | 2010-05-27 | 2010-06-25 | Yes | Yes | Health Professional | 42017206 | Yes | No | 2010-05-27 | Yes | No | Reuse | 2010
42022702 | 2010-05-27 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon was sealing a vessel when the tip of the pk dissecting forceps instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. The instrument had been in use for one hour when the breakage occurred. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-eval) or if additional info is received. Per the info provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1806136 | 1762782 | 1672252 | 2955842201000282 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-05-27 | 2010-06-25 | No | Yes | Health Professional | 42022702 | No |  | 2010-05-27 | Device Not Retu | No | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic curved shears (hcs) insert installed on the hcs instrument broke and fell into the patient. The insert was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | Conclusion: the insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps. " per the information provided by the physician, the insert accessory was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1792374 | 1740420 | 1651827 | 2955842201000279 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2010-05-25 | 2010-06-24 | No | Yes | Health Professional | 400169 | No | Yes | 2010-05-25 | Device Not Retu | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that the harmonic curved shears (hcs) insert accessory installed on the hcs instrument was used as replacement to complete a da vinci s hysterectomy procedure. After some time of use, the insert accessory broke and fell into the patient. The insert accessory was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch report was sent to the fda regarding the first event: 2955842-2010-00276. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. There are no components in the harmonic curved shears instruments or inserts that meet the definition of medical sharps". Per the information provided by the physician, the broken pieces were successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1863745 | 1740431 | 1651816 | 2955842201000277 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-05-25 | 2010-06-24 | No | Yes | Health Professional | 400169 | No | Yes | 2010-05-25 | Device Not Retu | Yes | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The hcs insert was retrieved and the surgeon proceeded with the procedure using a replacement insert accessory. After some time of use, the replacement insert accessory broke and fell into the patient. The replacement insert accessory was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The following medwatch report #2955842-2010-00277 was sent to the fda regarding the second event. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the information provided by the physician, the broken pieces were successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1863749 | 1740432 | 1651837 | 2955842201000276 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-05-25 | 2010-06-24 | No | Yes | Health Professional | 400169 | No | Yes | 2010-05-25 | Device Not Retu | Yes | Initial | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The hcs insert was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps. " per the information provided by the physician, the insert accessory was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1797985 | 1740433 | 1651838 | 2955842201000280 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-05-25 | 2010-06-24 | No | Yes | Health Professional | 400169 | No | Yes | 2010-05-25 | Device Not Retu | No | Unkown | 2010
400169 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the harmonic curved shears (hcs) insert accessory installed on the hcs instrument broke and fell into the patient. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | Conclusions: the insert accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert accessory is returned (post engineering evaluation) or if additional information is received. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps. " per the information provided by the physician, the insert accessory was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1792530 | 1740435 | 1651840 | 2955842201000281 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2010-05-25 | 2010-06-24 | No | Yes | Health Professional | 400169 | No | Yes | 2010-05-25 | Device Not Retu | Yes | Unkown | 2010
| 2010-06-16 | Malfunction |  | The shaft of the scissors became flexible, but not totally broken in half. The disposable sheath bent as it was being introduced into the port. The scissors were removed from the field and replaced with new scissors. There was no harm to the patient. |  | ENDOWRIST | CURVED SHEARS | 1785130 | 1746925 | 1657821 | 1746925 | 1 | NAY | User Facility |  |  | Initial | 2010-06-24 | 2010-06-24 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2010
42017908 | 2010-05-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the shaft of the monopolar curved scissors instrument broke at the orange/black junction. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Electrical continuity passed. In addition, engineering also observed the tube extension to present char marks and localized melting on a section of the fractured surface. The instrument was disassembled to find char marks on the hypotubes and the inner surface of the reinforcement metal ring. The evidence suggests an arcing event has occurred. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1863734 | 1740413 | 1651820 | 2955842201000275 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-24 | 2010-06-23 | No | Yes | Health Professional | 42017908 | Yes | No | 2010-05-24 | Yes | No | Reuse | 2010
DAVINCI S | 2010-06-22 | Malfunction |  | Permanent cautery spatula was inserted through a trocar sleeve in the normal fashion, into chest cavity during robotic assisted coronary artery bypass graft surgery. When the device was retracted back into the camera view, it was noted that the tip of the cautery spatula was broken off the device. X-ray was taken, then c-arm was brought into the room. The device was detected on both types of x-ray. Operation resumed as planned and the piece was retracted and saved. The total timing of added search was 80 minutes. |  | INTUITIVE SURGICAL INC | PERMANENT CAUTERY SPATULA | 1773943 | 1733910 | 1645736 | 1733910 | 1 | NAY | User Facility |  | Nurse | Initial | 2010-06-22 | 2010-06-22 | No | Yes | Physician | DAVINCI S | Yes | Yes |  |  |  |  | 2010
IS3000 A60P4 |  | Injury | Other | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that after a da vinci si hysterectomy procedure was performed on (b) (6), 2010, the pt claimed to have neuropathy. No system malfunction occurred and the planned procedure was successfully completed. | Manufacturer Narrative: --/-- On june 23, 2010, and july 2, 2010, additional information was provided by dr. (b) (6), the console surgeon who performed the da si hysterectomy procedure when the event occurred. The patient was properly positioned in deep trendelenburg, with thick padding under the patient's body and buttocks. The procedure was prolonged as it was dr. (b) (6) first hysterectomy procedure using the system and the uterine manipulator was ineffective. The planned surgical procedure was completed and from dr. (b) (6) point of view, the patient recovered seamlessly from the case. Post-operatively, the patient initially complained of not being able to bear weight on her left lower extremity. A neurology in-hospital consultation was ordered and the patient recovered her ability to walk within hours after the consultation. One week later, the patient complained of one episode of pain, which increased in intensity over the next week. The patient saw the same neurologist as an outpatient and began taking neurotin. The patient requested a second opinion, which was immediately ordered and the patient had a second opinion four days later. An mri and emg were ordered and both were negative. The patient was referred to a specialist who diagnosed the patient's symptoms as superficial nerve neuropathy and increased the patient's neurontin dose. The patient returned to work ten weeks post-op. --/-- Additional Manufacturer Narrative: --/-- On june 8, 2010, an email was sent to the console surgeon requesting additional info regarding the event. On june 14, 2010, a follow-up email was sent to the console surgeon. On june 17, 2010, a voice message was left for the hospital's director of pt services requesting additional info regarding the event. As of june 18, 2010, no response has been received from the console surgeon nor the hospital's director of pt services. Since no devices have been returned for eval and no additional info has been received, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional info is received. | DA VINCI SI SURGICAL SYSTEM | Reuse | 1777688 | 1735761 | 1647484 | 2955842201000270 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | InitialFollowup | 2010-05-21 | 2010-06-18 | Yes | No | Health Professional | IS3000 A60P4 | No |  | 2010-06-23 | Device Not Retu | No |  | 2010
42017909 | 2010-05-19 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument had been in use for 1 hour and 30 minutes when "one jaw of the scissor broke off and fell into the pt. " the broken piece was retrieved and the planned surgical procedure was completed with a 45 minute delay. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1785583 | 1735889 | 1647627 | 2955842201000267 | 1 | NAY | Manufacturer | Distributor | NOT APPLICABLE | Initial | 2010-05-19 | 2010-06-18 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer |  | 2010-06-21 | Device Not Retu | No | Reuse | 2010
42017206 | 2010-05-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the (b)(4) bipolar forceps instrument jaws did not align. No pt harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode and found one of the instrument grips to be bent, causing side to side misalignment of the grips. There is an offset of approx. 081 inches at the tips and the bent grip has separation of the yaw pulley and yaw pulley cover at the glue joint. The instrument passed electrical continuity testing. Engineering also observed that the instrument presents char marks on the yaw pulley near the distal clevis ear and the conductor wire jacket has char marks and localized melting around the same area. The conductor wire has exposed cables. Evidence suggests that an arcing event has occurred. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1849071 | 1735890 | 1647628 | 2955842201000269 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2010-05-25 | 2010-06-18 | No | Yes | Health Professional | 42017206 | Yes | No | 2010-05-25 | Yes | No | Reuse | 2010
40017908 | 2010-05-11 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, one of the blades on the monopolar curved scissors (mcs) instrument broke off and fell into the pt. The broken blade was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported failure mode, engineering observed that the left blade was returned detached from the instrument. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The blades and distal clevis exhibit various small scratches in the axial direction. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1786402 | 1735891 | 1647629 | 2955842201000263 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | PHARMACIST | Initial | 2010-05-20 | 2010-06-18 | No | Yes | Health Professional | 40017908 | Yes | Yes | 2010-05-20 | Yes | No | Reuse | 2010
42017909 | 2010-05-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the blade broke off of the monopolar curved scissors instrument and fell into the pt. The blade was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is broken at the distal idler and the broken piece was returned with the instrument. The right blade was not bent or damaged. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1786370 | 1735892 | 1647630 | 2955842201000268 | 1 | NAY | Manufacturer | Health Professional | Nurse | Initial | 2010-05-28 | 2010-06-18 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2010-05-28 | Yes | No | Reuse | 2010
40016907 | 2010-03-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure the harmonic curved shears insert remained "blocked" into the pt's stomach and a different insert broke into the pt's stomach. The piece was retrieved and the planned procedure was completed with a replacement insert. No pt harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed the insert to be returned with the insert clamp broken. Microscope imaging showed damage to the shaft and clamp base that indicate possible over bending and hyperextending of the clamp. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1785555 | 1735910 | 1647637 | 2955842201000273 | 1 | NAY | Manufacturer | User facilityDistributor | BIOMEDICAL ENGINEER | Initial | 2010-05-19 | 2010-06-18 | No | Yes | Health Professional | 40016907 | Yes | No | 2010-05-19 | Yes | Yes | Initial | 2010
42017909 | 2010-05-19 | Malfunction |  | It was reported that 45 minutes into a da vinci si hysterectomy procedure, one blade from the monopolar curved scissors (mcs) instrument fell off. It took approximately 20 minutes to locate the piece before successfully removing it from the pt. The planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the left scissor blade to be broken. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane is nearly straight. Both blades present indentations around the midpoint of the sharp edge. Electrical continuity passed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1785564 | 1735911 | 1647638 | 2955842201000271 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-05-19 | 2010-06-18 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-05-19 | Yes | No | Reuse | 2010
40016907 | 2010-03-24 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the harmonic curved shears insert remained "blocked" in the pt's stomach and broke while retrieving it causing a piece to fall into the pt. The piece was retrieved and the planned procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed the insert to be returned with the insert clamp broken. Microscope imaging showed damage to the shaft and clamp base that indicate possible over bending and hyperextending of the clamp. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1786387 | 1735913 | 1647640 | 2955842201000272 | 1 | NAY | Manufacturer | User facilityDistributor | BIOMEDICAL ENGINEER | Initial | 2010-05-19 | 2010-06-18 | No | Yes | Health Professional | 40016907 | Yes | No | 2010-05-19 | Yes | Yes | Initial | 2010
2955842 | 2010-05-21 | Malfunction |  | It was reported that during sterilization of instruments, the hose clamp came off the da vinci sonicpro cleaning system resulting in hot water coming out from the unit and flooding the surrounding area. A staff member attempted to turn off the unit and was burned by hot water. No further info was provided. | The investigation conducted by the steris field service tech concluded that the customer reported problem was caused by a loose clamp that is used to hold the hose against the incoming water pressure. To correct the issue, the steris tech retightened that clamp and installed another clamp on top of it. As of 06/18/2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SONICPRO CLEANING SYSTEM | NONE | 1786406 | 1735914 | 1647641 | 2955842201000274 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-05-21 | 2010-06-18 | No | Yes | Health Professional | 2955842 | Yes | No | 2010-05-19 | Yes | No | Reuse | 2010
| 2009-09-18 | Injury | Required Intervention | On may 19, 2010, a copy of medwatch report # (b)(4) was received from the fda, with the following event description. "volun 3-apr-2010: my obgyn dr (b) (6) performed a robotic assisted hysterectomy. This resulted in injury to both ureters. The right ureter had to be reimplanted, and the left ureter was stented. Dr (b) (6) was using a medical device called the 'da vinci'. It was supposed to be the state of the art medical device in gynecological surgeries. The hosp i was in (b) (6) in (b) (6). The hospitals that use the da vinci are supposed to have a privileging or credentialing system to determine the requirements prior to performing robotic surgeries. I do not believe that my doctor had been trained to use this medical device, that is why my injuries occurred. Also, i have been unable to get in touch with anyone at (b) (6) hosp who can tell me if this particular hosp has set up a privileging system in the matter of this medical device. I have not been able to find out if the hosp has followed the regulations that must be followed, in which a report has to be filled out when a pt is injured. I will continue to ask questions from the hosp about what and why i was injured on (b) (6) 2009. ". | As the event location and physician's name were not provided on voluntary medwatch report # (b)(4), isi contacted the fda division of postmarket surveillance on may 20, 2010, to request additional info regarding this event. On may 21, 2010, fda employee (b)(4) informed isi that no additional info could be provided and the initial reporter requested that no further info be release. A search of our complaint database was performed and a preexisting complaint record could not be matched to this event. To the best of our knowledge, this event was not reported to isi. Therefore, it is unk which da vinci surgical system and associated instruments and accessories were used during the time of the event. As a result, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional info is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1791962 | 1750496 | 1661007 | 2955842201000266 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-05-19 | 2010-06-18 | Yes | Yes | Health Professional |  | No | No | 2010-05-19 | Device Not Retu | No | Unkown | 2010
IS3000 A60P3 | 2010-05-13 | Death | Death | Event Description: --/-- It was reported that after the robotic portion of a colon resection was successfully performed, the patient suffered a heart attack while the surgeons were closing the patient using standard medical procedure. The patient was unable to be resuscitated and expired on the table. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- On (b) (6) 2010, the console surgeon performing the procedure advised that the system performed as intended with no issues and that the patient's death was not related in any way to the da vinci si surgical system. The surgeon went on to state that the patient had a known pre-existing condition that was directly related to this event; however, the exact condition was not disclosed to isi. He also mentioned the patient had suffered a heart attack during a previous surgical procedure performed several years ago and the risks for this colon resection were known in advance. --/-- Additional Manufacturer Narrative: --/-- Product code has been removed per fda request. | DA VINCI SI SURGICAL SYSTEM | Reuse | 1769384 | 1727307 | 1639587 | 2955842201000265 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2010-05-18 | 2010-06-17 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2010-06-29 | Device Not Retu | No |  | 2010
IS2000 A51P8 | 2010-05-18 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced an issue with the patient side manipulator arm (psm) not recognizing the cannula. With the assistance of an isi technical support engineer (tse), the site attempted to troubleshoot the issue by trying several cannulae and verifying the system set up was correct; however, the issue persisted at this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the issue experienced by the customer was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for evaluation. Engineering discovered that the cannula mount was damaged. The proximal pin inside the assembly cannula mount holding the cannula sensors in place was missing causing the engagement issue experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1769386 | 1727309 | 1639589 | 2955842201000264 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-18 | 2010-06-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-05-18 | Device Not Retu | No | Reuse | 2010
| 2010-06-16 | Malfunction |  | During laparoscopic prostatectomy, md looked at the left bipolar forceps and could only find 1 arm of the bipolar forceps. The surgeon looked around inside pt's abdomen and found the other half of the arm and removed it. Intuitive rep was in the room at the time and stated that there was no such recorded episode. Nothing unusual had occurred prior to this with regard to the davinci. ====================== manufacturer response for fenestrated bipolar forceps, intuitive surgical======================awaiting formal request for return of device. |  | ENDOWRIST | FORCEPS FENESTRATED BIPOLAR | 1770065 | 1728727 | 1640804 | 1728727 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-06-17 | 2010-06-17 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2010-05-18 | Other | Other | It was reported that prior to using the da vinci s surgical system to perform any tasks, the site experienced non-recoverable system error code 212. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code 212 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety system puts da vinci s in a non-recoverable safe state. The mtmr was returned to isi for evaluation. Engineering was able to replicate the customer reported failure mode and it was determined that an axis motor/encoder had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT ACCESSOR | 1766048 | 1723687 | 1636220 | 2955842201000255 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-05-18 | 2010-06-15 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-05-18 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure there was a collision between the instruments and part of the monopolar curved scissors instrument fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse event or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off of the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The detached blade was also returned. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1766056 | 1723702 | 1636235 | 2955842201000260 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-05-19 | 2010-06-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-19 | Device Not Retu |  | Reuse | 2010
42017909 | 2010-04-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the surgeon tried to close the monopolar curved scissors instrument, but the tip didn't move and one of the tips fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the right scissor blade was broken off of the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. Per the information provided by the initial reporter, the broken blade was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1766057 | 1723705 | 1636284 | 2955842201000259 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-05-19 | 2010-06-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-19 | Device Not Retu |  | Reuse | 2010
42017909 | 2010-04-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the surgeon tried to close the monopolar curved scissors instrument, but the tip didn't move and one of the tips fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off of the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The detached blade was also returned. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1766058 | 1723709 | 1636284 | 2955842201000258 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-05-19 | 2010-06-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-19 | Device Not Retu |  | Reuse | 2010
42017909 | 2010-05-11 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the monopolar curved scissors instrument stopped working. The instrument was removed and the surgical nurse noticed that one of the instrument tips was missing. The missing tip was found inside of the patient and retrieved. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off of the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The detached blade was returned with the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1766074 | 1723750 | 1636235 | 2955842201000262 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-05-19 | 2010-06-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-19 | Device Not Retu |  | Reuse | 2010
42017909 | 2010-04-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure there was a collision between the instruments and one of the monopolar curved scissors instrument tips broke and fell into the patient. The broken piece was retrieved and the planned procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off of the instrument. The blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. Per the information provided by the initial reporter, the broken blade was successfully retrieved from the patient. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1766075 | 1723752 | 1636284 | 2955842201000261 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-05-19 | 2010-06-15 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-19 | Device Not Retu |  | Reuse | 2010
IS1200 A43P9 | 2010-05-17 | Other | Other | It was reported that 20 minutes into a da vinci prostatectomy procedure, an error message was received indicating that an instrument arm may be stowed incorrectly. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the surgical staff undock all of the system arms, power off the surgeon console and reseat a system cable. The system was restarted and the same error message occurred after homing was completed. The surgeon made the decision to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted field service engineering concluded that the system error message experienced by the customer was associated with a setup joint (suj) arm z-axis string pot. The suj is an unpowered arm located on the patient side cart that operates as a configurable support for the powered instrument arm. The reported message occurs when the reading from the vertical sliding joint on the suj is consistent with the arm being stowed (not ready for surgery). The system was repaired by replacing the affected z-axis string pot. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1764991 | 1722501 | 1635188 | 2955842201000256 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-05-17 | 2010-06-14 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-05-17 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure the tip of the monopolar curved scissors instrument fired and pulled away from the shaft. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension is dislodged from the main tube and the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. Engineering also observed that the tube extension has a section below the breakage that exhibits charring/melting, indicating that an arcing event occurred. The instrument was disassembled to find the hypotubes exhibiting darkened marks where the tube extension broke. The inner surface of metal tube reinforcement ring and tube extension have dark marks in the same location as the external charring. It was determined that the arcing event like occurred after the breakage of the tube extension. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1764992 | 1722502 | 1635189 | 2955842201000257 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-17 | 2010-06-14 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-17 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the blade of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1761661 | 1720023 | 1632952 | 2955842201000248 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-14 | 2010-06-11 | No | Yes | Health Professional | 42017909 | No | No | 2010-05-14 | No | No | Unkown | 2010
IS1200 A43P9 | 2010-05-14 | Other | Other | It was reported that during system start-up for da vinci surgical procedure, the surgical staff received a message notifying them to check the yellow cable. An isi technical support engineer attempted to troubleshoot the issue via telephone and had the surgical staff power down, emergency power off (epo) the system, and check for bent pins on both ends of the yellow cable. All of the pins were straight. The system was restarted and the same message appeared. The same actions were performed, except that the yellow cable was swapped with another system cable after the epo. The system was restarted again, with the same result. The patient was under anesthesia when the surgical staff made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | Investigation conducted by field service engineering concluded that the error message was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. An ria pca is assigned to each system arm and performs numerous functions related to the function, control and hardware fault monitoring for its assigned arm. It was determined that the yellow cable connected to the affected ria pca may have been hot plugged while the surgeon side console was turned on, subsequently damaging the ria pca. The system was repaired by replacing the affected ria pca. The ria pca was returned to isi for failure analysis evaluation. Engineering discovered a damaged capacitor, thus generating the system error message experienced by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1761872 | 1720068 | 1625676 | 2955842201000250 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-05-14 | 2010-06-11 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No |  | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-07 | Other | Other | It was reported that during a da vinci s nephrectomy procedure, the patient sustained a superficial burn to the bowel due to arcing from the monopolar curved scissors (mcs) instrument with mcs tip cover installed. Repair of the affected area was not required as there was no significant damage. The planned surgical procedure was completed with a replacement instrument and tip cover of the same type and there were no post operative complications experienced by the patient. | The monopolar curved scissors instrument and tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1762047 | 1720135 | 1633020 | 2955842201000251 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2010-05-12 | 2010-06-11 | Yes | Yes | Health Professional | 42017909 | No | Yes | 2010-05-12 | No | No | Unkown | 2010
40018010 | 2010-05-11 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while working on the bladder neck, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The procedure was completed with a replacement mcs instrument tip cover accessory and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found two holes burned through the silicone. The silicone exhibits localized melting around the holes, indicating that there were multiple occurrences of arcing. The hole sizes range from 0. 020 to 0. 030 with varying shapes (round and oblong). The holes are located adjacent to each other, 0. 350 and 0. 390 above the silicone to sleeve interface. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1762053 | 1720136 | 1633021 | 2955842201000252 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-12 | 2010-06-11 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-05-12 | Device Not Retu | Yes | Unkown | 2010
IS3000 A60P4 | 2010-05-14 | Other | Other | It was reported that during a da vinci si surgical procedure, the customer experienced system error code 31030 and the touchpad located on the surgeon side console (ssc) was black with no illumination. With the assistance of an isi technical support engineer (tse), the site attempted to troubleshoot the issue by cycling power to the system multiple times and powering on the ssc in stand alone mode; however, the system error continued to recure. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer confirmed that system error code 31030 was associated with the touchpad. The touchpad is located in the middle of the surgeon console armrest and provides the means for selecting various system functions. The system was repaired by replacing the affected touchpad. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 31030 appears when the da vinci si safety system determines that a subcomponent of the system did not successfully complete its startup process. Upon determining this condition, the safety system puts da vinci si in a recoverable safe state. As of june 10, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1762062 | 1720138 | 1633023 | 2955842201000249 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-14 | 2010-06-11 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2010-05-14 | Device Not Retu | No | Reuse | 2010
42020702 | 2010-05-13 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the surgical staff had difficulty removing the tenaculum forceps instrument from the instrument arm. When attempting to remove the instrument, the instrument broke apart and shreds of fiberoptics fell into the patient. A piece of the instrument proximal clevis was protruding, so the surgeon had to manually break off the piece in order to pull the instrument out of the cannula. Some pieces were retrieved, however, some fiberoptic pieces remained inside of the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument is broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result of the damage and the proximal clevis and main tube are missing pieces. The clevis is missing the entire section below the ears and half of the main tube interface wall is broken off. (b)(4). | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1762096 | 1720173 | 1633057 | 2955842201000253 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-05-13 | 2010-06-11 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | No | 2010-05-13 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2009-08-23 | Death | Death | Event Description: --/-- It was reported that a patient passed away on (b) (6) 2009, eleven days after undergoing a urologic procedure with the da vinci s surgical system. | Manufacturer Narrative: --/-- On june 11, 2010, a message requesting additional information regarding the event was left at the office of the console surgeon. A follow-up mdr will be submitted if additional information is received. --/-- Additional Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- Product code has been removed per fda request. | DA VINCI S SURGICAL SYSTEM | Reuse | 1763241 | 1721314 | 1634129 | 2955842201000254 | 1 | NAY | Manufacturer | Company Representative | Health Professional | FollowupFollowup | 2009-09-01 | 2010-06-11 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2010-06-29 | Device Not Retu | No |  | 2010
42018404 | 2010-02-01 | Malfunction |  | On may 11, 2010, medwatch uf/importer report (b)(4) was received from the fda, with the following event description: the insulation on the tip came off, but was recovered. Pieces taken to spd (sterile processing department) for cleaning. Detached pieces were lost during cleaning. This event did not involve an electrophysiology procedure or an attempted electrophysiology procedure. The original intended procedure was robotic myomectomy chromopertubation. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1758833 | 1716835 | 1629917 | 2955842201000247 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2010-05-11 | 2010-06-10 | No | Yes | Health Professional | 42018404 | No | Yes | 2010-05-11 | No | No | Unkown | 2010
| 2010-06-08 | Malfunction |  | During the laparoscopy part of the robotic case the instrument being used was the harmonic. The harmonic insert piece of the instrument bent and broke. The piece was obtained whole and removed from patient. Dr. Was aware. The product was obtained and given to materials management. No harm to the patient. |  | HARMONIC | ULTRASONIC ENERGY INSTRUMENT | 1765266 | 1723025 | 1635712 | 1723025 | 1 | NAY | User Facility |  |  | Initial | 2010-06-10 | 2010-06-10 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the wrist of the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. At the time of the event the surgeon was attempting to go around a large fundal fibroid with the wrist of the instrument at an acute angle. The tip cover was replaced and the procedure continued, however, arcing was observed each time the acute angle was made with the mcs instrument affecting two additional tip covers. The mcs instrument was checked for burrs by the surgical staff and was found to be within specification. The case was successfully completed. No patient harm or adverse outcome was reported. Please also see mdr 2955842-2010-00244 and 2955842-2010-00246. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was returned with two holes burned though the silicone. The silicone exhibits localized melting around the holes which indicates that multiple arcing event occurred. The silicone also has a tear in the axial direction at the distal tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1757436 | 1715267 | 1611367 | 2955842201000245 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-06-09 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-05-10 | Device Not Retu | Yes | Unkown | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the wrist of the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. At the time of the event the surgeon was attempting to go around a large fundal fibroid with the wrist of the instrument at an acute angle. The tip cover was replaced and the procedure continued, however, arcing was observed each time the acute angle was made with the mcs instrument affecting two additional tip covers. The mcs instrument was checked for burrs by the surgical staff and was found to be within specification. The case was successfully completed. No patient harm or adverse outcome was reported. Please also see mdr 2955842-2010-00244 and 2955842-2010-00245. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was returned with a holes burned though the silicone. The silicone exhibits localized melting around the hole which indicates that an arcing event occurred. There is a dark material embedded in the silicone located next to the hole. Roughly 90 degrees from the hole, the silicone has a small tear in the axial direction. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1757437 | 1715268 | 1611367 | 2955842201000246 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-06-09 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-05-10 | Device Not Retu | Yes | Initial | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the wrist of the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. At the time of the event the surgeon was attempting to go around a large fundal fibroid with the wrist of the instrument at an acute angle. The tip cover was replaced and the procedure continued, however, arcing was observed each time the acute angle was made with the mcs instrument affecting two additional tip covers. The mcs instrument was checked for burrs by the surgical staff and was found to be within specification. The case was successfully completed. No patient harm or adverse outcome was reported. Please also see mdr 2955842-2010-00245 and 2955842-2010-00246. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, the tip cover was returned with various mechanical tears and various holes burned though the silicone. The silicone exhibits localized melting around the holes which is an indication that multiple arcing events occurred. One of the holes has a tear adjacent to one side. The tip cover also has various mechanical tears in the general vicinity of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1757453 | 1715284 | 1611367 | 2955842201000244 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-05-10 | 2010-06-09 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-05-10 | Device Not Retu | Yes | Initial | 2010
IS1200 A43P9 | 2010-05-11 | Other | Other | It was reported that while starting a da vinci mitral valve repair procedure, the site had difficulty getting two of the patient side manipulator (psm) arms to respond. The patient was under anesthesia and the port incisions were made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by a field service engineer (fse) found no system error codes in the system log. After speaking further with the surgical staff, the fse concluded that the reported issue experienced by the customer was likely associated with rapid exchanges between the patient side manipulator (psm) arms, which caused the da vinci surgical system to lock up. System verification was performed and the da vinci surgical system was found to be fully operation and performing to specification. The fse conducted additional training with the surgical staff to recreate the issue and show them how to correct the issue if it should occur in future procedures. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1753866 | 1712047 | 1625676 | 2955842201000240 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-11 | 2010-06-07 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2010-05-11 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-05-11 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, when the surgeon activated the maryland bipolar forceps instrument, slight energy activation could be seen at the monopolar curved scissor (mcs) instrument tips resulting in a superficial burn to the patient's pelvic fascia. With the assistance of an isi technical support engineer (tse), the site exchanged the instrument energy cords, checked the patient grounding pad and exchanged the mcs instrument for another instrument of the same type, however, the issue continued to occur. The planned surgical procedure was completed and no further patient harm was reported. Additionally, the patient did not require surgical repair due to the superficial burn. | Further investigation found that the site was using an incompatible electrocautery surgical unit (esu) with the da vinci si surgical system. The site replaced the unapproved esu with a compatible esu as indicated in the da vinci si surgical system instructions for use (ifu). No malfunctions of the da vinci si surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. As of june 3, 2010, no similar instances of this event have been reported at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1753870 | 1712049 | 1625678 | 2955842201000239 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-11 | 2010-06-07 | Yes | No | Health Professional | IS3000 A60P3 | Yes | No | 2010-05-11 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-26 | Malfunction |  | It was reported that during a da vinci s surgical procedure the tip of the monopolar curved scissor instrument was found to be bent. No instrument pieces fell into the patient and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument tube extension to be dislodged from the main tube. The entire tube extension wall is circumferentially broken at the base of the axial keys as a result and the tube can be rotated 360 degrees. No pieces are missing. The tube extension has a section below the breakage that exhibits charring and melting. The metal tube reinforcement ring has a dark mark directly above the charring. The instrument was disassembled to find the hypotube also exhibit darkened marks where the tube extension broke. Evidence indicates an arcing event occurred, likely after the breakage of the tube extension. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1753878 | 1712062 | 1625690 | 2955842201000241 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-05-13 | 2010-06-07 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-13 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-10 | Malfunction |  | It was reported that 10 minutes into a da vinci si partial nephrectomy procedure and while performing a blunt dissection of the ureter towards the hilum, one jaw/blade from the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken blade was retrieved and the planned surgical procedure was completed with another instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The intact blade does not exhibit damage at the tip. Electrical continuity passed. No other damage was observed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1753897 | 1712069 | 1625694 | 2955842201000242 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-10 | 2010-06-07 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-10 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-05-12 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the tip of the monopolar curved scissors instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one scissor blade was broken off. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is straight. The detached blade was returned with the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1754673 | 1712466 | 1626001 | 2955842201000243 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-12 | 2010-06-07 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-12 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-05-10 | Other | Other | It was reported that during a da vinci gynecological procedure, the customer experienced system error code 20008. The site undocked the system from the patient and restarted the system. The left master tool manipulator (mtm) was not free to home and with the assistance of an isi technical support engineer (tse), the site exercised the left master. This did not resolve the problem as system error 20008 fault returned. The site then powered down the system, reset the breaker, pressed emergency power off, and exercised axis 2 of the left mtm for approximately 30 seconds. Once the system was restarted, the system continued to experience issues homing and system error code 20008 returned. At this time, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer confirmed that system error code 20008 occurred. Engineering also found system error code 21008 present in the system error logs. Engineering determined that these system errors were associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtml gimbal. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error codes 20008 and 21008 appear when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The mtml gimbal was returned to isi for evaluation. Engineering discovered that a screws had broken, thus generating the system error code experienced by the customer and preventing the system from homing. As of june 3, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1753219 | 1711231 | 1624915 | 2955842201000237 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2010-05-10 | 2010-06-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2010-05-10 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-02 | Injury | Required InterventionHospitalization | On may 4, 2010, a wall street journal (wsj) article titled "surgical robot examined in injuries" was published. The article included details concerning pt injuries and a pt death at (b) (6) hospital. None of the allegations of pt injuries in the wsj article were reported to isi, however, on (b) (6) 2010, upon follow up with the hospital, (b) (6), the vp of community relations, confirmed one pt injury. Reportedly, after a da vinci s hysterectomy procedure, the pt did not have urine output. On (b) (6) 2009, a cystoscopy with ureteroscopy and exploratory laparoscopy procedure was performed which revealed that the pt had bilateral ureteral transection. Bilateral ureteral reimplantations and a bilateral ureteral stent placement was required to repair the affected area. In addition, (b) (6) stated that (b) (6) has no record of the other pt injuries/pt death related to the da vinci s surgical system as indicated in the wsj article. Damage to the pt's ureter prolonged the pt's hospitalization and the pt has returned to the hospital due to post operative complications. The hospital was unable to provide additional info concerning the pt's status. | An additional investigation with the site's risk management department was performed. It was reported that the damage to the pt's ureter was unrelated to the da vinci s system as there was no malfunction of the system during the surgical procedure. They also advised that they would be unable to provide the surgeon's opinion as to the cause of the pt injury due to peer review protection. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1761221 | 1715883 | 1629136 | 2955842201000238 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-05 | 2010-06-04 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-05-05 | Device Not Retu | No | Reuse | 2010
400180 VER 10 | 2010-05-10 | Malfunction |  | Patient with adenocarcinoma of prostate underwent a davinci prostatectomy with extended pelvic lymph node dissection. During the procedure, the surgeon noted arcing from the cautery to adjacent tissue when he used the robotic scissors. Scissors tip cover was changed and it was discovered that there were 2 holes in the scissor tip cover. Scrub rn stated the cover inserted smoothly when it was placed on the scissor. Surgery continued without known complications. Patient had a high jackson-pratt drainage, rapid return of bowel function, & minimal pain. His hemoglobin was stable postoperatively. Jackson-pratt drainage decreased and the patient was sent home with jackson-pratt in place. Patient was seen the following week for jackson-pratt drain removal. ====================== health professional's impression======================unknown cause of arcing====================== manufacturer response for insulating tip cover for robotic scissors, intuitive tip cover accessory======================requested that the device be returned for investigation. |  | TIP COVER ACCESSORY | INSULATING TIP COVER SCISSORS | 1753785 | 1711965 | 1625601 | 1711965 | 1 | NAY | User Facility |  |  | Initial | 2010-06-03 | 2010-06-03 | No | Yes | Physician | 400180 VER 10 | Yes | No Answer Provided |  |  |  |  | 2010
40018010 | 2010-05-05 | Other | Other | It was reported that approx 20 minutes into a da vinci si hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, arcing to the pt's uterus was observed. No pt harm, or adverse outcome was reported. | The monopolar curved scissors instrument and mcs tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional info is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1769735 | 1725677 | 1621794 | 2955842201000236 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-05 | 2010-06-03 | Yes | No | Health Professional | 40018010 | No | No | 2010-05-05 | Device Not Retu | No | Unkown | 2010
IS2000 A51P8 | 2010-05-04 | Other | Other | It was reported that during a da vinci s surgical procedure, the site had difficulty with instrument engagement on a pt side manipulator (psm) arm. The site tried multiple instruments and reseated the sterile adapter, however, the intermittent instrument engagement continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Investigation conducted by a field service engineer (fse) confirmed the customer reported problem as he had difficulty engaging an instrument on one of the pt side manipulator (psm) arms. The fse also found system error code #23014 in the system log which was associated with a second patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing both affected psms. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #23014 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". Both psms was returned to isi for failure analysis investigation. Engineering confirmed the customer reported instrument engagement problem on one psm and replaced multiple axis' wrist assembly shafts as well as multiple axis' potentiometers. Engineering found the second psm associated with the system error code #23014, faulted immediately and replaced the axis 6 motor. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1775619 | 1725678 | 1638096 | 2955842201000235 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-05-04 | 2010-06-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-05-04 | Yes | No | Reuse | 2010
40018010 | 2010-05-04 | Malfunction |  | On (b)(4) 2010, medwatch uf/importer report (b)(4) was received from (b)(6) hospitals, with the following information: event desc: we had two incidents during the same surgery with the davinci s tip cover accessory. This tip is used on the da vinci monopolar scissors to cover the shaft, exposing only the scissor tips. This is to protect the patient from any possible tissue burn during surgery. Two tip covers from the same lot melted during robotic procedure. A third cover from a different lot was used without further incident. These issues were initially reported to isi on (b)(4) 2010. The event occurred during a da vinci s hysterectomy procedure and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2010-00232 was submitted for the first occurrence of arcing. | The tip cover accessory was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The mcs instrument used during the time of the event was returned and evaluated. Engineering found no evidence of arcing and no physical damage was observed on the instrument. A follow-up mdr will be submitted if the accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750837 | 1707866 | 1621794 | 2955842201000233 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | Device Not Retu | Yes | Unkown | 2010
40018010 | 2010-05-04 | Malfunction |  | On (b)(6) 2010, medwatch uf/importer report (b)(4) was received from (b)(6) hospitals, with the following information: event desc: we had two incidents during the same surgery with the davinci s tip cover accessory. This tip is used on the da vinci monopolar scissors to cover the shaft, exposing only the scissor tips. This is to protect the patient from any possible tissue burn during surgery. Two tip covers from the same lot melted during robotic procedure. A third cover from a different lot was used without further incident. These issues were initially reported to isi on (b)(4) 2010. The event occurred during a da vinci s hysterectomy procedure and no patient harm, adverse outcome or injury was reported. Medwatch mfr report 2955842-2010-00233 was submitted for the second occurrence of arcing. | The tip cover accessory was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The mcs instrument used during the time of the event was returned and evaluated. Engineering found no evidence of arcing and no physical damage was observed on the instrument. A follow-up mdr will be submitted if the accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750839 | 1707868 | 1621794 | 2955842201000232 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | Device Not Retu | Yes | Unkown | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure a burn hole was observed on the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. The planned procedure was successfully completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The mcs instrument used during the time of the event was returned and evaluated. Engineering found no evidence of arcing and no physical damage was observed on the instrument. A follow-up mdr will be submitted if the accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750840 | 1707869 | 1621791 | 2955842201000234 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On may 7, 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00229 & 2955842-2010-00230 were submitted for the first and second occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750841 | 1707870 | 1621794 | 2955842201000231 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(4) 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00229 & 2955842-2010-00231 were submitted for the first and third occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750842 | 1707871 | 1621794 | 2955842201000230 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(4) 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00230 & 2955842-2010-00231 were submitted for the second and third occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750843 | 1707872 | 1621794 | 2955842201000229 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-04-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(4) 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00226 & 2955842-2010-00227 were submitted for the first and second occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750844 | 1707873 | 1621797 | 2955842201000228 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-04-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(4) 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00226 & 2955842-2010-00228 were submitted for the first and third occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750845 | 1707874 | 1621797 | 2955842201000227 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
40018010 | 2010-04-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover was replaced and arcing was observed coming from a hole on the replacement tip cover. Another tip cover was installed and arcing was noticed coming from the third tip cover accessory used. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. On (b)(6) 2010, the site indicated that the three affected tip covers were discarded. Medwatch mfr reports 2955842-2010-00227 and 2955842-2010-00228 were submitted for the second and third occurrences of arcing. | Since the tip cover accessory was discarded by the site, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1750846 | 1707875 | 1621797 | 2955842201000226 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-05-04 | 2010-06-02 | No | Yes | Health Professional | 40018010 | No | No | 2010-05-04 | No | Yes | Unkown | 2010
IS2000 A51P8 | 2010-05-10 | Other | Other | It was reported that during the beginning of a da vinci s surgical procedure, the site experienced double vision with the image displayed in the surgeon console. The site tried multiple endoscopes, however, the vision issue continued to recur. The patient was under anesthesia the port incisions had been made when the surgical staff made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the isi field service engineer concluded that the vision issue experienced by the customer was associated with the camera, which the site reportedly dropped prior starting the surgical procedure. The camera produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. A camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The camera was returned to the original equipment manufacturer (oem) for evaluation. The oem observed that the camera stage had been damaged from an impact, thus resulting in the vision issue experienced by the customer. As of may 26, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1748891 | 1705169 | 1619234 | 2955842201000223 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-10 | 2010-06-01 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-05-10 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade had fractured at the cross section of the grip cable crimp and detached from the instrument. Half of the cable crimp is exposed and the fracture plane of the blade is nearly straight. The broken left blade was returned with the instrument. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1748901 | 1705170 | 1619235 | 2955842201000224 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-07 | 2010-06-01 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-07 | Device Not Retu | No | Reuse | 2010
420227 | 2010-02-04 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the patient's sigmoid colon was cauterized as it was being retracted with the pk dissecting forceps instrument. The surgeon's intent was to apply cautery energy to a monopolar curved scissors (mcs) instrument also being used during the procedure; however, he erroneously applied cautery energy to the pk dissecting forceps instrument. The affected area was over-sewn to repair damage to the patient's sigmoid colon. The planned surgical procedure was completed and there were no post operative complications experienced by the patient. | The pk dissecting forceps instrument was not returned for evaluation; however, based on the information provided, it was determined that the da vinci si surgical system, instruments or accessories did not malfunction. It was determined that the surgeon inadvertently applied cautery energy to the incorrect instrument. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1769925 | 1725168 | 1637603 | 2955842201000225 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2010-05-03 | 2010-06-01 | Yes | No | Health Professional | 420227 | No | Yes | 2010-05-03 | Device Not Retu | No | Unkown | 2010
40016907 | 2010-04-29 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the harmonic curved shears insert fell into the patient. The piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed the insert to be returned with the blade detached from the distal end. The shaft features that mates with the clamp arm pins are bent backwards which is indicative of a mechanical overload. The curved blade has scratch marks on the side that makes contact with the clamp arms' teflon pad. The distal end of shaft has various scratches which have removed the black paint. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1745876 | 1702518 | 1616671 | 2955842201000218 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-04-29 | 2010-05-28 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-04-29 | Device Not Retu | Yes | Initial | 2010
42017909 | 2010-05-03 | Malfunction |  | It was reported that during a da vinci si pyeloplasty procedure, one side of the monopolar curved scissors (mcs) instrument blade broke off and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The intact blade does not exhibit damage at tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1745879 | 1702519 | 1616672 | 2955842201000219 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-03 | 2010-05-28 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-03 | Device Not Retu | No | Reuse | 2010
42017206 | 2010-01-22 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument was not recognized by the system. The planned procedure was successfully completed with another instrument and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Engineering placed the instrument on an in-house system and recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions and the pogo pins do not stick and are not contaminated. The customer reported failure mode could not be replicated. Engineering also observed a char mark on one yaw pulley cover, indicating that an arcing event occurred. The charring is located at a corner of the grip base. There is also a slight separation at the glue joint between the yaw pulley and the yaw pulley cover. No charring is present on the other grips' yaw pulley. No other damage was found. | MAYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1745881 | 1702520 | 1616673 | 2955842201000220 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-05-03 | 2010-05-28 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-05-03 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-04-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the staff experienced non-recoverable system error code 809 and noticed smoke and sparks near the patient cart (pc). The surgical staff attempted to restart the system two times, however, the pc would not shut down. An isi technical support engineering (tse) was onsite and the emergency power off (epo) button was pressed, however, the pc would not turn off. The tse then opened the pc cover and turned the battery switch off directly. The pc was undocked from the patient and the surgeon decided to convert to traditional laparoscopic surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that system error code 809 was associated with the patient back plane (pbp) printed circuit assembly (pca) board. Also during the investigation, damage was observed on the cable connecting the pbp pca to the axillary power board (apb) pca. The system was repaired by replacing the pbp pca and damaged cable. The apb pca and pc battery box module were also replaced as a precaution. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The pbp pca board was returned to isi with significant charring and heat damage around and underneath the connectors. Engineering observed signs of liquid intrusion on the board, which likely caused a short circuit across multiple pins and resulted in a high current being discharged from the battery. The high current then overheated the electronics and caused the smoke experienced by the customer. The short circuit also prevented the pc from powering down using the power button. Unplugging the pc from the wall did not shut the pc down since the pc was running on battery power. The short circuit also bypassed the emergency power off (epo) button, leading to a small amount of current continuing to flow through the system which kept some of the leds on the pc lit. The current was limited to approximately 90 ma by resettable fuses located on the apb pca, which were triggered shortly after the current exceeded 90 ma. Both the apb and pbp pcas worked as designed and limited the excess current generated. It was determined that some of the physical evidence is consistent with the presence of a very brief flame, however, any flame which may have occurred would have been non-sustainable, brief, and snuffed out, as the pcb pca and connector are flame-retardant, consistent with ul standards. On (b)(4) 2010, isi provided the site the investigation results and recommended that they avoid using excessive amounts of cleaning fluid when cleaning the system. Additionally, care should also be taken to avoid spilling fluid on the system as seepage of fluid into the pc base can cause damage to the system's internal electronics. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1745911 | 1702526 | 1616679 | 2955842201000222 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-04-01 | 2010-05-28 | No | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-04-01 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-04-28 | Death | Death | It was reported that a pt who underwent a da vinci s hysterectomy procedure on (b) (6), 2010, passed away on (b) (6), 2010. The attending surgeon indicated that the surgical procedure went well and there were no sys malfunctions. Postoperatively, the pt's abdominal pain was minimal and her vital signs were stable, however, she had a severe migraine headache the morning of her discharge. The pt declined medication for her headache and was discharged on (b) (6), 2010. Later that evening, the pt complained of fatigue and became unconscious. The pt expired the following day. On (b) (6), 2010, the surgeon also indicated that the autopsy was not yet complete and isi will be forwarded a copy of the autopsy report when it becomes available. | Based on the info provided by the attending surgeon, it was determined that the da vinci s surgical system worked as intended, no sys malfunction occurred and the sys did not cause or contribute to the pt's demise. The hosp has continued to use the sys to perform surgery on pts. As of (b) (6), 2010, no adverse events have been experienced with the site's da vinci s surgical sys and no similar instances of this event have been reported to isi. A follow-up mdr will be submitted when the pt autopsy report is rec'd from the site/surgeon. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1768289 | 1724830 | 1637281 | 2955842201000217 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2010-04-28 | 2010-05-28 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-04-28 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-24 | Death | Death | It was reported that a pt who underwent a da vinci s tors procedure on (b) (6), 2010, expired on (b) (6), 2010. On may 4, 2010, the surgeon who performed the surgical procedure informed isi that there were no acute complications with the case and the pt was discharged 4-5 days after the procedure was performed. The surgeon also mentioned that follow-up was conducted with the pt approx one week after the operation and the pt was found to be doing well. Shortly after, the pt began experiencing post-operative bleeding and passed away. The surgeon indicated that he does not believe that the robotic procedure was the direct cause of this unfortunate event, but rather this is the type of complication that it associated with these type of surgical procedures. On may 20, 2010, the surgeon indicated that they do not have access to the pt's death certificate or the hospital records from the institution that the pt was taken to on (b) (6), 2010. Additionally, the pt's family informed the surgeon via telephone that they did not seek an autopsy. | Based on the info provided by the attending surgeon on may 4 & 20, 2010, it was determined that the da vinci s surgical system worked as intended, no sys malfunction occurred and the sys did not cause or contribute to the pt's demise. The hosp has continued to use the sys to perform surgery on pts. As of may 27, 2010, no adverse events or sys malfunctions have been experienced with the site's da vinci s surgical sys and no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1768254 | 1724831 | 1637282 | 2955842201000216 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2010-04-29 | 2010-05-28 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-04-29 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-30 | Malfunction |  | It was reported that during a da vinci si pelvic lymphadenectomy procedure, one of the monopolar curved scissors instrument "jaws" broke off and fell into the pt. The broken piece was retrieved and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade has broken off at the distal end and was not returned. The blade fractured at the cross section of the grip cable crimp. The cable crimp has slipped out of the crimp pocket. The fracture plane of the blade is nearly straight. No other damage was found. Per the info provided by the initial reporter, the broken piece was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1775557 | 1726862 | 1639210 | 2955842201000221 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-05-04 | 2010-05-28 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-05-04 | Yes | No | Initial | 2010
| 2010-04-21 | Malfunction |  | It was reported that during a da vinci si surgical procedure, shavings from the intuitive surgical instrument were observed to have come off of the instrument and fall into the pt. The surgeon attempted to remove all the debris. No pt harm, adverse outcome or injury was reported. No additional info was provided. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. |  | ENDOSCOPIC ELECTROSURGICAL | 1831479 | 1713390 | 1626882 | 2955842201000208 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-04-21 | 2010-05-27 | No | Yes | Health Professional |  | No | No | 2010-04-21 | Device Not Retu | No | Unkown | 2010
42017909 | 2010-04-19 | Malfunction |  | It was reported that during a da vinci si surgical nephrectomy procedure, part of the monopolar curved scissors (mcs) instrument blade broke off and fell into the pt. The broken piece was removed and the surgical procedure was completed with an instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one scissor blade is broken off and missing. The blade fractured at the cross section of the cable crimp. The fracture plane is straight. The intact blade does not exhibit damage at the tip. The instrument passed electrical continuity testing. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1831513 | 1713409 | 1626898 | 2955842201000209 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-19 | 2010-05-27 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2010-04-19 | Yes | No | Reuse | 2010
40017908 | 2010-04-05 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure a blade from the monopolar curved scissors instrument broke off and left into the pt's pelvis. The broken piece was retrieved and the planned surgical procedure was completed after a delay. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1831505 | 1713410 | 1626899 | 2955842201000210 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-05 | 2010-05-27 | No | Yes | Health Professional | 40017908 | No | Yes | 2010-04-05 | Device Not Retu | No | Unkown | 2010
42020906 | 2010-03-30 | Malfunction |  | It was reported that during a da vinci s gastric bypass procedure, a piece of the pulley on the suture cut needle driver instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed. There was no pt harm, adverse outcome or injury reported. | The suturecut needle driver instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. However, engineering review of photographic images provided by the customer revealed that the suturecut needle driver instrument appears to have a broken grip cable. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if additional info is received. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1833464 | 1714426 | 1586941 | 2955842201000207 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2010-05-17 | 2010-05-27 | No | Yes | Health Professional | 42020906 | No | No | 2010-05-17 | Device Not Retu | No | Unkown | 2010
42001503 | 2010-04-22 | Malfunction |  | It was reported that during system takedown after a da vinci s surgical procedure was successfully completed, the surgical staff observed a hole in one of the instrument arm drapes. The hole was located near the release button of the instrument sterile adapter. | The drape was not returned to isi, however, engineering evaluation was performed via photographs taken of the affected drape. Per the customer reported complaint, engineering observed that the plastic film covering the 1 x 1 square opening in the sterile adapter was missing. It cannot be determined if the plastic film went missing prior to draping, while in use during the surgical procedure, or during system takedown. | DA VINCI S INSTRUMENT ARM DRAPE | ENDOSCOPIC ACCESSORY | 1743604 | 1699892 | 1614244 | 2955842201000214 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-04-26 | 2010-05-26 | No | Yes | Service Personnel | 42001503 | No | No | 2010-04-26 | No | Yes | Unkown | 2010
42018404 | 2010-04-27 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the permanent cautery spatula (pcs) instrument was in a bent position, preventing removal of the instrument from the cannula by the surgical staff. Simultaneous removal of the pcs instrument and cannula was required. The planned surgical procedure was completed and there was no harm, adverse outcome or injury to the pt. Examination of the instrument by the surgical staff revealed that a piece from the pcs instrument tip was broken off and missing. The surgical staff was unable to locate the missing piece and it is unk if the missing piece fell into the pt. | The permanent cautery hook instrument was not returned for evaluation, therefore the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional info is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1775592 | 1726859 | 1639207 | 2955842201000215 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-27 | 2010-05-26 | No | Yes | Health Professional | 42018404 | No | No | 2010-04-27 | Device Not Retu | No | Unkown | 2010
| 2010-03-18 | Malfunction |  | During a surgical procedure the davinci harmonic instrument fell apart. The end jaw piece was found in the pt and removed. |  | HARMONIC CURVED SHEARS INSERT | HARMONIC CURVED SHEARS INSERT | 1751462 | 1702799 | 1616914 | MW5016127 | 1 | GEI | Voluntary |  | Other | Initial | 2010-05-25 | 2010-05-25 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2010
400180 | 2010-04-30 | Malfunction |  | It was reported that during a da vinci s surgical procedure arcing was observed when using the monopolar curved scissors (mcs) instrument, resulting in a small hole on the mcs tip cover accessory installed on the instrument. The planned procedure was successfully completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory used during the time of the event was not returned, however, the instrument was. Engineering inspected the instrument and found no evidence of arcing. No char marks were observed on the metal components at the distal end and the instrument did not have any burrs/damage on any components that would contribute to arcing. No physical damage was found on the instrument. Since the tip cover accessory was not returned for evaluation, the root cause of the customer reported failure mode cannot be determined. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1740621 | 1696771 | 1611367 | 2955842201000213 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-30 | 2010-05-24 | No | Yes | Health Professional | 400180 | No | No | 2010-04-30 | No | No | Unkown | 2010
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the tips of the (b)(4) bipolar forceps instrument stopped moving. The planned procedure was successfully completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Engineering placed the instrument on an in-house system and recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. The customer reported failure mode could not be replicated. 48 - engineering also observed a char mark on one yaw pulley cover, indicating that an arcing event occurred. The charring is located at a corner of the grip base. There is also a slight separation at the glue joint between the yaw pulley and the yaw pulley cover. No charring is present on the other grips' yaw pulley. No other damage was found. | MAYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1740622 | 1696772 | 1611368 | 2955842201000212 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-04-28 | 2010-05-24 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-04-28 | Device Not Retu | No | Reuse | 2010
42020504 | 2010-04-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, one of the fenestrated bipolar forceps instrument tips broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one instrument grip broken off at its base, just above the top surface of the yaw pulley. The fracture surface is not quite straight and the grip appears to have a twist just below the fracture surface. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1740623 | 1696773 | 1611369 | 2955842201000211 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-26 | 2010-05-24 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No | 2010-04-26 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed after a delay of 2 or 3 minutes and with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the left grip blade to be detached from the instrument. The broken blade had fractured at the cross section of the cable crimp causing half of the cable crimp to become exposed. The fractured plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1738965 | 1695810 | 1610561 | 2955842201000205 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-04-28 | 2010-05-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-04-28 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-22 | Malfunction |  | It was reported that during a da vinci s nephrectomy procedure a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed after a delay of 2 or 3 minutes and with a replacement instrument of the same type. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the left grip blade to be detached from the instrument. The broken blade had fractured at the cross section of the cable crimp causing half of the cable crimp to become exposed. The fractured plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1738968 | 1695811 | 1610562 | 2955842201000206 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-04-22 | 2010-05-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-04-28 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-04-29 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced multiple occurrences of system error code 20008 and system error code 20013. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 20008 and 20013 were associated with the left master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 is generated when the rate of change in position 'an error' is detected as measured by that joint's primary control sensor (encoder) differs from the measurement of the secondary sensor (potentiometer) by more than a specified tolerance for agreement. System error code 20008 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a recoverable safe state. The engineering analysis of the returned mtml confirmed that the issue was caused by a potentiometer on the arm. The arm was repaired by replacing the potentiometer. As of may 20, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1738972 | 1695812 | 1610563 | 2955842201000204 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-04-30 | 2010-05-21 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-04-30 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a jaw from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is fractured and the broken piece was not returned with the instrument. The broken blade had fractured at the cross section of the cable crimp resulting in the cable crimp being exposed and the fracture plane is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1738973 | 1695813 | 1610564 | 2955842201000192 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-23 | 2010-05-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-04-23 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-04-20 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the image was flickering in both eyes of the surgeon console. The camera head and cable were replaced in an attempt to resolve the issue, however, the flicker continued to recur. An isi technical support engineering (tse) was then contacted, who provided troubleshooting instructions via telephone. The tse requested that the staff ensure that the camera cable connector is fully seated, shutdown and reboot the system, cycle the circuit breaker on the back of the vision cart while the system was off, and reseat the vision cable. After all of these actions were performed, the flickering remained. The patient had been under anesthesia for approximately 1 hour and 15 minutes and was docked to the system when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The doco was returned to isi for evaluation and engineering determined that the flickering issue experienced by the customer was most likely caused by a poor conducting connection between the camera and camera cable. As may 20, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1738360 | 1695317 | 1610119 | 2955842201000201 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-20 | 2010-05-20 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-04-20 | Device Not Retu | No | Reuse | 2010
42017908 | 2010-04-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, one side of the monopolar curved scissors (mcs) instrument blade broke off and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. Both blades present indentations around the midpoint of the sharp edge. Electrical continuity passed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1738379 | 1695319 | 1610121 | 2955842201000203 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-04-20 | 2010-05-20 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-04-20 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-04-21 | Other | Other | It was reported that during a da vinci surgical procedure, the customer experienced a non-recoverable system error code 20013. With the assistance of a technical support engineer, the site pressed the silent alarm button; however, it was non responsive. The site then undocked the system, turned off power, and reseated cables. The system was restarted and was homed successfully. Shortly after, system error code 20013 returned. The site pressed the e-stop button and then pressed the fault override button to clear the system error code. At this time, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering determined that system error code 20013 was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a non-recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of may 19, 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1738380 | 1695320 | 1610122 | 2955842201000202 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-21 | 2010-05-20 | Yes | No | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-04-21 | Device Not Retu | No | Reuse | 2010
| 2010-05-17 | Malfunction |  | The harmonic da vinci scissor grasper: the tip of the instrument broke off during robotic surgery. The tip was retrieved. |  | DA VINCI | SHEARS CURVED HARMONIC | 1739533 | 1696004 | 1610701 | 1696004 | 1 | NAY | User Facility |  | Nurse | Initial | 2010-05-20 | 2010-05-20 | No | Yes | Physician |  | Yes | Yes |  |  |  |  | 2010
IS3000 A60P3 | 2010-04-19 | Other | Other | It was reported that after a da vinci si hysterectomy procedure was performed, the patient had a reddened area on their left shoulder where the shoulder pads were used. The patient was also complaining of intense shoulder pain. | On (b)(4) 2010, a voice message was left for the initial reporter requesting additional information regarding the event. On (b)(4) 2010, another voice message was left for the initial reporter requesting additional information regarding the event. On may 18, 2010, another voice message was left for the initial reporter requesting additional information regarding the event. As of may 19, 2010, no response has been received from the initial reporter. Since no devices have been returned for evaluation and no additional information has been received from the initial reporter, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1737172 | 1694303 | 1609236 | 2955842201000200 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-19 | 2010-05-19 | Yes | No | Health Professional | IS3000 A60P3 | No | Yes | 2010-04-19 | No | No | Reuse | 2010
| 2010-05-04 | Malfunction |  | We had two incidents during the same surgery with the davinci s tip cover accessory. This tip is used on the da vinci monopolar scissors to cover the shaft, exposing only the scissor tips. This is to protect the patient from any possible tissue burn during surgery. Two tip covers from the same lot melted during robotic procedure. A third cover from a different lot was used without further incident. ====================== health professional's impression======================bad lot====================== manufacturer response for tip cover accessory, da vinci s======================requested return. |  | DA VINCI S | ACCESSORY TIP COVER | 1738701 | 1695569 | 1610349 | 1695569 | 1 | NAY | User Facility |  |  | Initial | 2010-05-07 | 2010-05-18 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
| 2010-05-04 | Malfunction |  | We had two incidents during the same surgery with the davinci s tip cover accessory. This tip is used on the da vinci monopolar scissors to cover the shaft, exposing only the scissor tips. This is to protect the patient from any possible tissue burn during surgery. Two tip covers from the same lot melted during robotic procedure. A third cover from a different lot was used without further incident. ====================== health professional's impression======================bad lot====================== manufacturer response for tip cover accessory, da vinci s======================requested return. |  | DA VINCI S | ACCESSORY TIP COVER | 1738701 | 1695569 | 1610349 | 1695569 | 1 | NAY | User Facility |  |  | Initial | 2010-05-07 | 2010-05-18 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
IS3000 A60P3 | 2010-04-20 | Other | Other | It was reported that during system setup for a da vinci si surgical procedure, a hospital employee felt a little shock when assisting with draping of the camera. The hospital employee had the camera head in her right hand and the shock occurred when she grabbed the camera cable and the silver handles on the camera with her left hand. No additional harm was reported. | Investigation was conducted by a field service engineer (fse). The fse evaluated the system by performing an electrical safety test and physical inspection. The fse was unable to recreate the customer reported failure mode. As a precaution, the system camera and camera cable were replaced. As of may 13, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1733905 | 1691467 | 1606673 | 2955842201000198 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-20 | 2010-05-17 | No | Yes | Health Professional | IS3000 A60P3 | Yes | No | 2010-04-20 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the shaft of the monopolar curved scissors instrument broke at the distal end. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension dislodged from the main tube. The main tube is intact, however, the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Engineering also observed charring on various distal end components. The tube extension has a section located directly below the fracture surface that exhibits charring/localized melting, indicating that an arcing event likely occurred. The instrument was disassembled and char marks were found on the main tube, hypotubes, and inner surface of the metal reinforcement ring. The evidence is not conclusive, however, the breakage of the tube extension likely created a path for energy to escape. No other damage was found. | MONPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1733926 | 1691472 | 1606678 | 2955842201000199 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-04-22 | 2010-05-17 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-04-22 | Device Not Retu | No | Reuse | 2010
420143 | 2010-04-02 | Malfunction |  | On (b)(6), 2010, isi received uf/importer report (b)(4) with the following event description: (b)(6) old female underwent a robotic-assisted left hemi-throidectomy on (b)(6) 2010. The plastic like sheath coating approximately one centimeter by 0. 5cm sheard off the maryland dissector. A piece was retrieved from the patient's incision. A piece was retrieved from the patient's incision. The patient's surgical site was irrigated by the physician. | The instrument used during the surgical procedure has not been returned for evaluation. A follow up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6), 2010, (b)(6) in risk management reported that it is unknown what surgical task was being performed when the plastic like sheath coating sheared off inside of the patient. Significant irrigation of the affected area was required to remove the broken piece and lengthen the surgical procedure. The site believes all of the pieces were retrieved. The patient was discharged from the hospital and to the best of (b)(6) knowledge, the patient has not experienced any post operative complications that required further hospitalization. | MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1732637 | 1690248 | 1605524 | 2955842201000194 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-04-15 | 2010-05-16 | No | Yes | Health Professional | 420143 | No | No | 2010-04-15 | No | No | Unkown | 2010
42017206 | 2010-04-08 | Malfunction |  | It was reported that the tips of the maryland bipolar forceps instrument would not open during a da vinci s surgical procedure. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument passed recognition and engagement testing. Additional functional testing revealed that the instrument moved intuitively with a full range of motion in all directions and the grip opened and closed properly. Engineering also observed charring of the insulation material, indicating that an arcing event occurred. One yaw pulley cover exhibits a char mark at one corner of the grip base. The other yaw pulley and yaw pulley cover are undamaged. The instrument passed electrical continuity testing. No other damage was observed. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1733136 | 1690752 | 1606024 | 2955842201000193 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-19 | 2010-05-16 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-04-19 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-15 | Malfunction |  | Event Description: --/-- It was reported that three hours into da vinci s total nephrectomy procedure and while the surgeon was using the monopolar curved scissors (mcs) instrument to move tissue, the tip of the instrument broke off and fell into the pt. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The intact blade does not exhibit damage at the tip. Electrical continuity passed. No other damage was observed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1745723 | 1694805 | 1609700 | 2955842201000195 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2010-04-15 | 2010-05-14 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-05-07 | Device Not Retu | No |  | 2010
400180 | 2010-04-15 | Malfunction |  | It was reported that 15 minutes into a da vinci si prostatectomy procedure, arcing was observed at the tip of the monopolar curved scissors (mcs) instrument with mcs tip cover accessory installed. The tip cover was removed and replaced with another mcs tip cover accessory; however, arcing was also observed coming from the tip of the instrument with tip cover installed. The second tip cover was replaced to successfully complete planned procedure. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. Also see mdr# 2955842-2010-00197 for the second tip cover related to this event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1745785 | 1694807 | 1609702 | 2955842201000196 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-04-15 | 2010-05-14 | No | Yes | Health Professional | 400180 | No | No | 2010-04-15 | Device Not Retu | Yes | Initial | 2010
400180 | 2010-04-15 | Malfunction |  | It was reported that 15 minutes into a da vinci si prostatectomy procedure, arcing was observed at the tip of the monopolar curved scissors (mcs) instrument with mcs tip cover accessory installed. The tip cover was removed and replaced with another mcs tip cover accessory; however, arcing was also observed coming from the tip of the instrument with tip cover installed. The second tip cover was replaced to successfully complete planned procedure. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval a follow-up mdr will be submitted. Also see mdr# 2955842-2010-00196 for the first tip cover related to this event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1745731 | 1694810 | 1609702 | 2955842201000197 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-04-15 | 2010-05-14 | No | Yes | Health Professional | 400180 | No | No | 2010-04-15 | Device Not Retu | Yes | Initial | 2010
IS3000 A60P3 | 2010-04-12 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgeon noticed a burn mark on the patient's abdomen, located at one of the port incisions where a cannula accessory was used. The planned surgical procedure was completed with no additional patient harm. | On (b)(6) 2010, a voice message was left for the initial reporter requesting additional information regarding the event. On (b)(6) 2010, another voice message was left for the initial reporter requesting additional information regarding the event and that the instrument(s) and cannula used at the port location where the burn mark was noticed be returned to isi for evaluation. As of may 13, 2010, no response has been received from the initial reporter. Since no devices have been returned for evaluation and no additional information has been received from the initial reporter, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1729227 | 1686674 | 1602057 | 2955842201000191 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-13 | 2010-05-13 | Yes | No | Health Professional | IS3000 A60P3 | No | Yes | 2010-04-13 | No | No | Reuse | 2010
400180 | 2010-04-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The planned surgical procedure was completed with no patient harm, adverse outcome or injury reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1720462 | 1678075 | 1593850 | 2955842201000190 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-13 | 2010-05-10 | No | Yes | Health Professional | 400180 | No | Yes | 2010-04-13 | No | Yes | Unkown | 2010
IS2000 A41P8 | 2010-04-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the vision was not ideal when using the system's 12mm schoelly 0 degree endoscope. After the surgical procedure was successfully completed, the endoscope was taken to the biomedical department and the biomedical technician observed that a lens was missing. It cannot be determined if the lens went missing prior to inserting the endoscope into the patient or while it was inserted in the patient during the surgical procedure. It was also reported that the site had been using sterrad 100nx in their endoscope cleaning and sterilization process. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment manufacturer (oem) for evaluation. Per the customer reported complaint, the oem confirmed that an endoscope lens was missing. Damaged glue joints were observed, which resulted in fluid invasion and subsequent damage to the optical components. Sterility is the responsibility of the person/institution performing the sterilization. Isi's reprocessing instructions for cleaning, disinfection, and sterilization information for reusable instruments, accessories and endoscopes used with da vinci, da vinci s and da vinci si systems lists isi's recommendations for methods of sterilization and the parameters. Sterrad 100nx used by the site is not recommended by isi and has not been validated for efficacy (i. E. Process achieves expected cleaning/sterilization results) and compatibility (i. E. Process doesn't damage the product). | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1718323 | 1676902 | 1592798 | 2955842201000183 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-04-14 | 2010-05-07 | No | Yes | Health Professional | IS2000 A41P8 | Device Returned To Manufacturer | No | 2010-04-14 | Device Not Retu | No | Unkown | 2010
42017909 | 2010-04-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure the tip of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one scissor blade was broken off. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is straight. The detached blade was returned with the instrument. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1718332 | 1676904 | 1592799 | 2955842201000184 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-04-09 | 2010-05-07 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-28 | Device Not Retu | No |  | 2010
40018402 | 2010-04-07 | Malfunction |  | It was reported that during a da vinci left internal mammary artery takedown procedure, the surgical staff observed that a plastic part of the permanent cautery spatula instrument was broken at the tip after it was inserted into the patient. While trying to remove the instrument, the broken part fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1718339 | 1676906 | 1592800 | 2955842201000185 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2010-04-09 | 2010-05-07 | No | Yes | Health Professional | 40018402 | No | Yes | 2010-04-09 | No | No | Unkown | 2010
40016907 | 2010-04-12 | Malfunction |  | It was reported that during a da vinci s heller myotomy procedure, the blade of the harmonic curved shears (hcs) insert was dull, would not close, arced and broke off and fell into the patient's abdominal cavity. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert was returned and evaluated. Per the customer reported complaint, the clamp arm was returned detached from the insert. The clamp arm itself is not damaged, however, the shaft features at the distal end that mate with the clamp arm pins are bent. The direction of bending is consistent with hyperextension of the clamp arm. While the evidence is inconclusive, the damage is most likely due to mishandling and misuse. The curved blade which is not designed for mechanical cutting is not dull and exhibits no damage. Additionally, energy modality as reported by the customer(harmonic vibrations) of the generator does not produce electrical arcing, however, there is a component at the distal end of the insert that is broken off. There are no components in the harmonic curved shears insert that meet the definition of medical sharps. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1718562 | 1677102 | 1592968 | 2955842201000186 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-21 | 2010-05-07 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-04-21 | Device Not Retu | Yes | Initial | 2010
IS1200 | 2007-07-05 | Injury | Required Intervention | On 04/08/2010, isi received voluntary medwatch (b)(4) from the fda department of health and human services reporting the following: volun 01-feb-2010: my name is (b)(6). I am the personal rep of the estate of (b)(6), on (b)(6) 2007, my mother had a davinci robot assisted laparoscopic hysterectomy at (b)(6) medical center in (b)(6). It is my understanding that during this event, dr (b)(6) lost visualization and the instrument on the robot cut the right common iliac artery leading to hemorrhage, cardiopulonary arrest and near exsaguination. The info i have been provided with indicates that this may have been reported to the fda as a maude adverse event associated with the davinci surgical robot. I have not been able to locate such a report during my online search. Physician: (b)(6) md. Facility: (b)(6) medical center. No additional information has been provided, including information related to the fields in (b)(6). | The site and physician's name were not provided on the initial voluntary medwatch (b)(4) and on 04/09/2010 (b)(4) at the fda was contacted. She advised that the initial reporter requested to remain anonymous. The site and physician's name also could not be released. A database search was conducted and a prexisting complaint record could not be matched to this reported event. To the best of our knowledge, intuitive surgical was not advised of this event from this site and as a result, the affected da vinci surgical system and/or instruments and accessories have not been evaluated. The root cause of the customer reported failure mode cannot be determined. If additional information is provided a follow-up mdr will be submitted. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1718571 | 1677109 | 1573517 | 2955842201000188 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-04-08 | 2010-05-07 | Yes | Yes | Health Professional | IS1200 |  | No | 2010-04-08 | No |  | Unkown | 2010
420179 | 2010-03-29 | Injury | Required Intervention | On april 14, 2010, isi rec'd medwatch uf/importer report (b) (4) with the following event description: "a pt was undergoing laparoscopic total hysterectomy in the main operating room using robotic equipment. The surgeon was exchanging the devices to allow for cautery of a small bleeder. The iliac artery was nicked in error. A vascular surgeon was called and arrived quickly to repair the vessel. The pt was transferred to the intensive care unit for observation. The pt is doing well. " note: during the case at the very end with insertion of robotic laparoscopic scissors through the right robotic sleeve, there was injury to the right external iliac artery. This resulted in laparotomy and repair by vascular surgery. On april 15, 2010, isi rec'd amended medwatch uf/importer report (b) (4) with the following add'l info: "the surgical assistant physician is a seasoned surgeon but was not trained in robotic surgery. The estimated blood loss for the pt was 3 liters. ". | The instrument used during the surgical procedure has not been returned for eval. However, based on the info provided in the customer's amended medwatch report (b) (4), damage to the pt's iliac artery was the result of an use error by the surgeon assisting at the pt side cart (psc). The psc is the operative component of the da vinci s system, and its primary function is to support the instrument arms and camera arm. The pt cart operator works in the sterile field, assisting the surgeon console operator by exchanging instruments and endoscopes, and by performing other pt-side activities. On may 7, 2010, (b) (6) an rn and operating room team leader reported that to the best of her knowledge, the pt has not returned to the hosp due to any post operative complications. A review of our training database revealed that on (b) (6) 2010, the pt side surgeon assistant completed isi's pt side assist training lab. As of (b) (6) 2010, there have been no reported recurrences of the issue at this hosp. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1736109 | 1692421 | 1568809 | 2955842201000189 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-03-31 | 2010-05-07 | Yes | No | Health Professional | 420179 | No | No | 2010-03-31 | Device Not Retu | No | Unkown | 2010
42022702 |  | Malfunction |  | It was reported that there is a "burn" on the side of the pk dissecting forceps instrument. The "burn" was not observed during a surgical procedure. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed charring in an area approx. 045" x. 048" on one side, above the distal clevis's ear and the conductor's cap. These findings indicate that an arcing event occurred. Arcing most likely occurred due to leakage of energy through the insulating material, which can occur as a result of overloading at the tip. Engineering also observed that the grip on the same side is slightly bent causing an offset of approx. 020" between both grips. Engineering concluded that the damage was most likely the result of overloading at the tip. The instrument passed electrical continuity testing. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1744331 | 1694050 | 1609010 | 2955842201000187 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-03 | 2010-05-07 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-03 | Yes | No | Reuse | 2010
| 2010-02-01 | Malfunction |  | The insulation on the tip came off, but was recovered. Pieces taken to spd(sterile processing department) for cleaning. Detached pieces were lost during cleaning. |  | ENDOWRIST | PERMANENT CAUTERY SPATULA | 1720465 | 1678076 | 1593851 | 1678076 | 1 | NAY | User Facility |  |  | Initial | 2010-05-05 | 2010-05-05 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2010
40018010 | 2010-04-05 | Malfunction |  | It was reported that during a da vinci s surgical procedure, holes were observed to be burned into the monopolar curved scissors (mcs) tip cover accessory that was installed on a mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. No additional information was provided. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1714638 | 1673286 | 1589478 | 2955842201000180 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-04-05 | 2010-05-04 | No | Yes | Health Professional | 40018010 | No | No | 2010-04-05 | No | Yes | Unkown | 2010
IS3000 A60P3 | 2010-04-06 | Other | Other | It was reported that while preparing for a da vinci si surgical procedure, the surgeon side console (ssc) would not power up. The site attempted to troubleshoot the issue by verifying proper power cord connection and confirming power from the wall's electrical outlet. With the assistance of an isi technical support engineer (tse), the site cycled power using the circuit breaker on the ssc and checked the fiber led to verify that it was off. The patient was under anesthesia for approximately 30 minutes, when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the surgeon side console (ssc) power supply. The ssc power supply converts the ac voltage from the wall outlet to dc voltage for use by the system. The system was repaired by replacing the affected power supply. (b)(4). | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1714639 | 1673287 | 1589479 | 2955842201000181 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-06 | 2010-05-04 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-04-06 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-04-05 | Malfunction |  | It was reported that while the surgeon was dissecting tissue during a da vinci s hysterectomy procedure, one tip of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up medwatch report will be submitted if the instrument is returned (post-evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CUVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1714734 | 1673345 | 1589522 | 2955842201000182 | 1 | LRW | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-04-05 | 2010-05-04 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-05-05 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-03-31 | Malfunction |  | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that during a da vinci s nephrectomy procedure, while resecting tissue using the monopolar curved scissors (mcs) instrument, arcing from the tip of the mcs instrument occurred, and the tissue specimen that was being removed from the patient was burned. The planned procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The electrical surgical unit (esu) setting ifu information reference in medwatch manufacturer report number 2955842-2010-00174 submitted on april 30, 2010 was stated incorrectly. The isi instrument and accessory user manual states the valley lab force fx esu, which the site was using during this case, has a maximum setting requirement of 38 watts on the coag fulgurate setting when used with the mcs instrument and tip cover accessory. The site reported using a setting of 50 watts on the coag fulgurate setting during this procedure, which was above the recommended maximum setting. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering found that the instrument tube extension was dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. There are no missing instrument pieces and the instrument passed electrical continuity testing. Inspection of the instrument did not reveal evidence that an arcing event occurred. Disassembly of the instrument revealed that the hypotubes and conductor wire at the distal end of the instrument are not damaged. No other damage was found. On april 2, 2010, isi contacted (b)(6) a scrub nurse assisting during the procedure. (b)(6) indicated that the tip of the mcs instrument where the orange and black meet appeared to be misaligned and had a gap where energy escaped from. She also reported that the electrical surgical unit (esu) setting was 50 watts fulgurate. (b)(6) was advised that isi's ifu states that the maximum setting for fulgurate energy is 70 watts. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1711807 | 1671251 | 1587636 | 2955842201000174 | 1 | GEI | Manufacturer | User facility | Other | InitialFollowup | 2010-10-14 | 2010-04-30 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-04-01 | Device Not Retu | No |  | 2010
40016907 | 2010-03-12 | Malfunction |  | It was reported that during a da vinci s thyroidectomy procedure the tips of the harmonic curved shears insert accessory broke off and fell into the patient. The piece was retrieved and the planned surgical procedure was completed with a replacement insert. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was missing from the distal end and one side of the hinge feature that accepts the clamp arm pin is bent. The evidence is not conclusive, however, it appears that the damage was most likely due to hyperextension of the clamp arm or other mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1711820 | 1671258 | 1587640 | 2955842201000175 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-07 | 2010-04-30 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-04-07 | Device Not Retu | Yes | Initial | 2010
40016907 | 2010-03-30 | No Answer Provided |  | It was reported that during a da vinci s left nephroureterectomy procedure, the tip of the harmonic curved shears (hcs) insert accessory fell into the patient. The piece was retrieved and the planned surgical procedure was completed with a replacement insert. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was missing from the insert and the hinge feature at the distal end of the inner tube that mates with the clamp arm pins are bent. Engineering also observed that the insert shaft is bent around the midpoint, forming an angle of approximately 110 degrees. The shaft's black coating finish is flaking off at the bending location, exposing the metal shaft. No other damage was found. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1711821 | 1671259 | 1587641 | 2955842201000176 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-09 | 2010-04-30 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | No | 2010-04-09 | Device Not Retu | Yes | Initial | 2010
IS2000 A51P8 | 2010-04-05 | Other | Other | It was reported that while preparing to start a da vinci s prostatectomy procedure, the site stated that the light source was very dim. With the assistance of an isi technical support engineer, the site replaced the light cord; however, the light was still dim. The site did not have a replacement illuminator lamp module and tried a secondary light source, however, it did not produce enough light. The patient was under anesthesia and the port incisions were made when the surgeon made the decision to abort the planned surgical procedure. | The investigation conducted concluded that the vision issue experienced by the customer was associated with the illuminator lamp module. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator lamp module. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1711835 | 1671268 | 1587649 | 2955842201000179 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-05 | 2010-04-30 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2010-04-05 | No | No | Reuse | 2010
40016907 | 2010-03-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure the tip of the harmonic curved shears insert fell into the patient. The piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. On (b)(6) 2010, medwatch uf / importer report with patient identifier (b)(6) was received: jaw of harmonic scalpel broke while being used intraoperatively. Piece was retrieved by surgeon without incident. No harm to this patient. See intuitive (b)(4). | The insert was returned and evaluated. Per the customer reported complaint, engineering observed the insert has a. 360 inch long piece of the curved blade detached from the distal end. The blade broke off near where the blade transitions from straight to curved. The fractured surface is slightly jagged and there are scratch marks on the curved blade below the fractured surface. The clamp arm is intact and undamaged. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 1711836 | 1671269 | 1587650 | 2955842201000177 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-04-09 | 2010-04-30 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2010-04-09 | Device Not Retu | Yes | Initial | 2010
42020906 | 2010-03-30 | Malfunction |  | It was reported that during a da vinci s gastric bypass procedure, a piece of the pulley on the suturecut needle driver instrument broke off and fell into the patient. An x-ray performed, did not reveal the location of the broken piece. The planned surgical procedure was completed and there was no patient harm, adverse outcome or injury reported. | The suturecut needle driver instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The isi representative present during the procedure observed that the surgeon was working laterally under port 1 and the wrist of the suturecut needle driver instrument was positioned at a sharp angle and was pulled up into the cannula when the instrument breakage occured. It was determined by the surgeon and the isi representative that breakage of the suturecut needle driver instrument was most likely caused by the instrument wrist scrapping against the cannula. Approximately one week post op, the surgeon reported to the isi representative that the patient is doing well. As of (b)(6), 2010, there have been no reported recurrences of this event at this hospital. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1710726 | 1670461 | 1586941 | 2955842201000173 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-30 | 2010-04-29 | No | Yes | Health Professional | 42020906 | No | No | 2010-03-30 | No | No | Unkown | 2010
42022702 | 2010-03-29 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the pk dissecting forceps instrument was observed to be burnt. The planned surgical procedure was completed with no report of patient harm, adverse outome or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one yaw pulley cover exhibits charring at the corner of the grip base, indicating that an arcing event occurred. There is separation between the yaw pulley and yaw pulley cover at the glue joint. The other yaw pulley does not exhibit charring. The instrument passed electrical continuity testing. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1709383 | 1669268 | 1585972 | 2955842201000172 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-29 | 2010-04-28 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-29 | Device Not Retu | No | Reuse | 2010
42022702 | 2010-03-29 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure, the tip of the pk dissecting forceps instrument was observed to be burnt. The planned surgical procedure was completed with no report of patient harm, adverse outcome or injury. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one yaw pulley exhibits a char mark on its side, adjacent to a grip open cable, indicating that an arcing event occurred. No charring is exhibited on other parts of both yaw pulleys or the grip base. The instrument passed electrical continuity testing. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- | PK DISSECTING FORCEPS INSTRUMENT | Initial | 1709384 | 1669269 | 1585972 | 2955842201000171 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-03-29 | 2010-04-28 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-29 | Device Not Retu | No |  | 2010
42022702 |  | Malfunction |  | Event Description: --/-- It was reported that the tips of the pk dissecting forceps instrument were observed to be out of alignment. No surgical procedure was being performed at the time this issue was discovered. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering observed that one grip is bent, causing side to side misalignment of the grips. There is a 0. 020 offset at the tips. The bent grip has separation of the yaw pulley and pulley cover at the glue joint, indicative of overloading at the tip. Engineering also observed that both yaw pulley covers exhibit charring at the grip base, indicating that an arcing event occurred. The yaw pulley cover with separation from the yaw pulley at the glue joint and exhibits severe charring. The instrument passed electrical continuity testing. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering observed that one grip is bent, causing side to side misalignment of the grips. There is a 0. 020 offset at the tips. The bent grip has separation of the yaw pulley and pulley cover at the glue joint, indicative of overloading at the tip. Engineering also observed that both yaw pulley covers exhibit charring at the grip base, indicating that an arcing event occurred. The yaw pulley cover with separation from the yaw pulley at the glue joint and exhibits severe charring. The instrument passed electrical continuity testing. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 1709386 | 1669270 | 1585973 | 2955842201000170 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2010-03-29 | 2010-04-28 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-29 | Device Not Retu | No |  | 2010
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument were out of alignment. No surgical procedure was being performed at the time this issue was discovered. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event may have occurred during a previous surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed that both grips are properly aligned. The instrument passed recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument passed electrical continuity testing. The customer reported failure mode cannot be duplicated. Engineering also observed that one yaw pulley cover presents a char mark on the outside of the grip base indicating that an arcing event occurred. Some separation can be seen between the yaw pulley and yaw pulley cover at the glue joint. There was no damage observed to the other yaw pulley. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1709387 | 1669271 | 1585974 | 2955842201000169 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-29 | 2010-04-28 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-29 | Device Not Retu | No | Reuse | 2010
42017908 | 2010-03-17 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a fragment from the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the cable crimp and half of the crimp is exposed. The fracture plane is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1706589 | 1667175 | 1584080 | 2955842201000168 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-03-29 | 2010-04-26 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-03-29 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-04-01 | Other | Other | It was reported that while beginning a da vinci s prostatectomy procedure, the surgeon was experiencing difficulties focusing the image from the surgeon side console (ssc). With the assistance of an isi technical support engineer, the site attempted to focus from the vision cart and replace the camera cable; however, they were still unable to focus. The site was asked to use a back up camera head but the site did not have one. The patient was under anesthesia when the surgeon decided to abort the scheduled procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that the focus issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci s system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The camera head was returned and evaluated by the original equipment manufacturer (oem). The oem found the camera's stage was loose, causing the customer reported focus vision. The camera was repaired by readjusting and repairing the stage. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1706592 | 1667177 | 1583303 | 2955842201000167 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-04-01 | 2010-04-26 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-04-01 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-26 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the surgeon observed a reddish tint and horizontal lines in the right eye of the high resolution stereo viewer (hrsv) when the camera head cable was moved. The patient was under anesthesia for approximately 45 minutes when the surgeon made the decision to abort and reschedule the planned surgical procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of april 22, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1705527 | 1666281 | 1583295 | 2955842201000163 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-26 | 2010-04-25 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-03-26 | Device Not Retu | No | Reuse | 2010
42022702 | 2010-01-07 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the pk dissecting forceps instrument stopped cauterizing. The instrument was removed and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. Although the customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence of arcing. No patient harm was reported, however, we assessed the evaluation results and conservatively decided to report our findings as it cannot be determined whether or not the arcing occurred during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip, making cautery at grips non-functional. The instrument failed electrical continuity testing. Engineering also observed that the yaw pulley is charred and broken, indicating that an arcing event occurred. The bottom section of the yaw pulley exhibits a char mark on the same side as the wire breakage. The exposed bare wire most likely created a path for arcing. The same yaw pulley is missing a. 165 x. 060 piece opposite of the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. It is indeterminable as to how the piece of yaw pulley broke. No other damage was found. Per the case notes, there is no indication from the site that the missing instrument component fell inside of a patient during a surgical procedure. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1705531 | 1666285 | 1583299 | 2955842201000152 | 1 | NAY | Manufacturer | OtherHealth Professional | Other | Initial | 2010-03-15 | 2010-04-25 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-25 | Device Not Retu | No | Reuse | 2010
42017206 | 2010-01-15 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the (b)(4) bipolar forceps instrument stopped cauterizing. The instrument was removed and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a reportable event, however, during internal evaluation of the instrument, engineering discovered evidence of arcing. No patient harm was reported, however, we assessed the evaluation results and conservatively decided to report our findings as it cannot be determined whether or not the arcing occurred during the surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip, thus preventing cautery from functioning at the grips. The instrument failed electrical continuity testing. Engineering also observed that both yaw pulleys exhibit severe charring with material removal and localized melting at the grip base, indicating that an arcing event occurred. Engineering concluded that damage to the yaw pulleys most likely occurred prior to breakage of the conductor wire. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1705533 | 1666288 | 1583301 | 2955842201000151 | 1 | NAY | Manufacturer | OtherDistributor | Other | Initial | 2010-03-15 | 2010-04-25 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2010-03-25 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-03-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a piece of the monopolar curved scissors instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade missing from the instrument. The missing blade had fractured at the cross section of the cable crimp and the fracture plane is straight. An indentation was also observed near the tip of the intact blade. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1705534 | 1666289 | 1583302 | 2955842201000165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-30 | 2010-04-25 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-03-30 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-03-31 | Other |  | It was reported that during a da vinci prostatectomy procedure, the surgical staff experienced multiple occurrences of system error code 212 which would not clear when pressing the fault override button. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The mtmr was returned to isi for failure analysis investigation. Engineering evaluation found that the harness conductor had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1705535 | 1666290 | 1583303 | 2955842201000164 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-31 | 2010-04-25 | Yes | No | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-03-31 | Device Not Retu | No | Reuse | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
IS 1200 | 2010-04-21 | Injury | Required InterventionHospitalization | Urologist called to report problems with robotic equipment The urologist had to convert the robotic prostatectomy case to open when the system stopped working The next week the surgeon had two robotic prostatectomies During the first case there were 10 alarms with the error 20013 that said press fault override button Despite this problem the first case went as planned without any other issues During the second case the same problem with the fault alarm occurred three times on the third time the fault did not override The manufacturer was contacted using the davinci help number We were instructed to undock the robot completely and power down I was instructed to check some connections on the console which were fine Then we turned on the robot and started again We had another fault that we could not override so i called the company again had the surgeon press the emergency stop button and then fault override This was to release the instruments grasps on the tissue At that time there was a bleeder that hadnt fully gotten under control because the robot was frozen There were three more faults before we had another irrecoverable fault The surgeon decided to completely undock the robot and convert to an open case A sales rep came to do a system check on the robot and it should have been ready to perform today After speaking with three representatives from the davinci company i was told it was most likely a component sensor in the arm that was not working correctly We were also having trouble with the brightness of the light but we changed the light cord and the picture was normal There was some problem with our haronic as well but we opened a new instrument and it continued to work The surgeon complained that one of the eyes was blinking on and off Manufacturer technician is servicing the equipment shortly health professionals impressionequipment problems 1 Lighting was 12 the intensity 2 One eye was blinking offon 3 Problem with green arm surgeon feels that the machine keeps getting bandaid fixes poor equipment maintainance The case yesterday the red arm would not work light bulb changed but still blinding on and off Our facility contracts directly with manufacturer for maintainance with service agreements manufacturer response for robot da vincimanufacturer was on site for maintainance im waiting for a report Our faciity contracts witht the manufacturer for them to provide the maintainance on this piece of equipment |  | DA VINCI | ROBOT | 1709705 | 1669451 | 1586126 | 1669451 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-04-22 | 2010-04-23 | Yes | Yes | Physician | IS 1200 | Yes | No Answer Provided |  |  |  |  | 2010
42017908 | 2010-03-25 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure and while dissecting tissue, one blade from the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Conclusion: the instrument was returned and evaluated. Per the customer, reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the cable crimp and half of the crimp is exposed. The fracture plane is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1714477 | 1671119 | 1587527 | 2955842201000166 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-25 | 2010-04-23 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2010-03-25 | Yes | No | Reuse | 2010
40017909 | 2010-03-22 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure one blade of the monopolar curved scissors instrument broke off and fell into the patient's pelvis. The instrument was removed and the procedure was delayed approximately 20-30 minutes as the surgical staff was locating and retrieving the broken piece. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1702281 | 1663854 | 1581191 | 2955842201000160 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-22 | 2010-04-21 | No | Yes | Health Professional | 40017909 | No | No | 2010-03-22 | No | No | Unkown | 2010
42017909 | 2010-03-24 | Malfunction |  | It was reported that during a da vinci hysterectomy, a metal piece of the monopolar curved scissors instrument was observed on the video screen to have broken off. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the grip cable crimp and the fracture plane is nearly straight. Half of the cable crimp is exposed. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1702749 | 1664187 | 1581530 | 2955842201000161 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-24 | 2010-04-21 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-03-24 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-24 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the customer experienced system error code 212 and an instrument engagement issue which prevented full function of the instrument installed on the affected patient side manipulator (psm). With the assistance of an isi clinical sales representative, the site was able to resolve the system error code 212. The site attempted to resolve the instrument engagement issue with the affected psm by replacing the drape, reseating the sterile adapter and instrument; however, the instrument engagement issue remained. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure before contacting an isi technical support engineer for troubleshooting. No patient harm was reported. | The investigation conducted by the technical support engineer confirmed that site experienced multiple system error code 212. An additional investigation performed by the field service engineer concluded that system error code 212 was associated with the left master tool manipulator. The instrument engagement issue could not be duplicated and no known reoccurrences have been reported. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a non-recoverable safe state. The mtml was returned to isi for failure analysis investigation. Engineering evaluation was not able to duplicate the customer reported failure, however, the motor encoder assembly was replaced as a precautionary measure. As of april 20, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1702754 | 1664201 | 1581550 | 2955842201000162 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-24 | 2010-04-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-03-24 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-03-26 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors instrument broke and a fragment fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint the left blade was returned detached from the instrument. Engineering observed the blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1701372 | 1663093 | 1580505 | 2955842201000155 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-26 | 2010-04-20 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-03-26 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-03-18 | Injury | Other | It was reported that post a da vinci myomectomy procedure, the patient experienced numbness in her hands and fingers and sharp pain and tingling in her legs. Per the surgeon that performed the operation, the myomectomy procedure took longer than anticipated, approximately 5 hours, with the patient in the trendelenburg position. Per the operative report the patient tolerated the entire procedure well with good hemostasis and minimal blood loss. | During a follow up appointment on (b) (6) 2010 the patient no longer had hand or finger numbness however, she was experiencing sharp pain and tingling in her legs. On (b) (6) 2010, the patient was seen by her physician who believes the patient may be experiencing mild paresthesia as she continued to feel sharp pains in her left leg. She was referred to a neurologist for follow up treatment. Based on the information provided, it was determined that the da vinci surgical system likely did not malfunction in a way that would cause nor contribute to the patient injury, however, the patient's positioning and the extended procedure duration may have contributed to the event. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1701378 | 1663094 | 1580506 | 2955842201000156 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-03-24 | 2010-04-20 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2010-03-24 | No | No | Reuse | 2010
420179 | 2010-03-01 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the surgeon accidentally cut the patient's ureter while using the monopolar curved scissors (mcs) instrument. The surgeon placed a precautionary stitch in the patient's ureter. The planned surgical procedure was completed and no additional patient harm or adverse outcome was reported. | On (b) (6) 2010 the patient returned the to hospital and required a ureteral stricture to be stented. Per the isi clinical sales representative, the surgeon that performed the da vinci si hysterectomy procedure stated that the system, instruments and/or accessories in no way malfunctioned during the procedure. The mcs instrument has not been returned for evaluation and therefore a root cause cannot be determined. If additional information is received or if the instrument is returned, a follow-up mdr will be submitted. As of april 20, 2010, no adverse events have been experienced with the site's system. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1701395 | 1663099 | 1580510 | 2955842201000159 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-23 | 2010-04-20 | Yes | No | Health Professional | 420179 | No | No | 2010-03-23 | No | No | Unkown | 2010
400180 | 2010-03-01 | Injury | Other | It was reported that 3 weeks post a da vinci si hysterectomy procedure the patient experienced pain and visited the emergency room. It was found that the patient had a urethral stricture which required a stent. | Per the surgeon that performed the da vinci si hysterectomy procedure, the type of injury the patient suffered correlates to a cautery incident. The surgeon noted that during the procedure she had to change the monopolar curved scissors (mcs) tip cover accessory multiple times due to holes forming. It was determined afterwards that the electrosurgical unit (esu) settings were above the isi recommended setting. The mcs tip cover accessory and mcs instrument have not been returned for evaluation. Therefore root cause of the customer reported failure mode cannot be determined. If the instrument or accessory are returned for evaluation, or if more information is obtained, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1701396 | 1663100 | 1580511 | 2955842201000157 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-23 | 2010-04-20 | Yes | Yes | Health Professional | 400180 | No | No | 2010-03-23 | No | Yes | Unkown | 2010
40018010 | 2010-03-03 | Malfunction |  | It was reported that during a da vinci s surgical procedure the monopolar curved scissors (mcs) tip cover accessory, installed on a mcs instrument, was nicked by a fenestrated bipolar forceps instrument. The tip cover was immediately removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the mcs tip cover accessory. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole size is approximately. 020 in diameter and is located 0. 135 above the silicone to sleeve interface. A small tear was found in the melted area around the hole and mechanical tearing through the silicone was observed approximately 0. 125 above the hole. Another tear was discovered at the distal tip. No other damage was found. The instruments and accessories user manual specifically states: inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. Do not use another instrument to clean the tip cover accessory. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1701425 | 1663110 | 1580516 | 2955842201000158 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-03-24 | 2010-04-20 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2010-03-24 | Device Not Retu | Yes | Initial | 2010
400180 | 2010-03-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on a mcs instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover accessory is returned (post-evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1698954 | 1661189 | 1578763 | 2955842201000153 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-17 | 2010-04-18 | No | Yes | Health Professional | 400180 | No | Yes | 2010-03-17 | No | Yes | Unkown | 2010
40017908 | 2009-09-30 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, one tip of the monopolar curved scissors instrument broke and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument and broken blade were returned and evaluated. Per the customer reported complaint, the broken blade had fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit damage at the tip. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1734018 | 1668535 | 1585299 | 2955842201000154 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-03-19 | 2010-04-16 | No | Yes | Health Professional | 40017908 | Yes | No | 2010-04-09 | Yes | No | Reuse | 2010
42014204 | 2010-02-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s tonsillectomy procedure, the black cover of the shaft of the 5mm monopolar cautery instrument peeled off and fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed black tube insulation damage towards the distal end of the instrument. The damaged section is approximately. 900 long and is located 1. 1 above the snake wrist. The insulation is gouged and exhibits tearing and peeling, resulting in exposure of the metal shaft. The evidence is not conclusive, however, it was determined that the damage was most likely due to mishandling/misuse. The instruments and accessories user manual specifically states: before use, all instruments should be inspected for damage or irregularities. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Always use caution when inserting or removing instruments through the cannula: before inserting the instrument into the sterile adapter and cannula, make sure the wrist is straight and close the end effectors (if applicable). Isi requested additional information from the initial reporter on (b)(4) 2010. To date, no response has been received. --/-- Additional Manufacturer Narrative: --/-- On (b)(6), 2010, additional information was received from the distributor indicating that the broken instrument pieces were successfully retrieved from the patient. | 5MM MONOPOLAR CAUTERY INSTRUMENT | Reuse | 1697761 | 1660020 | 1577700 | 2955842201000150 | 1 | NAY | Manufacturer | ForeignDistributor | Other | InitialFollowup | 2010-03-16 | 2010-04-15 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2010-05-03 | Device Not Retu | No |  | 2010
420179 | 2010-02-26 | Malfunction |  | It was reported that when the vagina was being amputated from the cervix towards the end of a da vinci s hysterectomy procedure a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was not retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The initial reporter indicated that the instrument was disposed of, therefore, will not be returning to isi for evaluation. | Since the instrument was disposed of, evaluation of the device cannot be performed and the root cause of the customer reported failure mode cannot be determined. The materials in the intuitive surgical 8mm and 5mm endowrist instruments for use on the da vinci s surgical system have been tested for material biocompatibility in accordance with iso (b)(4) standards for their intended use (blood path, indirect (4. 2. 3. A) and limited exposure (4. 3. A)). On (b)(6) 2010, a video recording of the surgical procedure was provided to isi. To date, isi has been unable to view the video due to the format it was recorded in. Isi will continue to try a variety of video players and format conversions in an attempt to view the video. If successful, a follow-up mdr will be submitted after the video has been reviewed or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1695250 | 1658184 | 1576013 | 2955842201000149 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-03-14 | 2010-04-13 | No | Yes | Health Professional | 420179 | No | No | 2010-03-14 | No | No | Unkown | 2010
42017909 | 2010-03-16 | Malfunction |  | It was reported that during a da vinci s cystectomy procedure a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left scissor blade was broken off of the instrument. The left blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. The detached blade was not returned with the instrument. No other damage was found. Per the information provided by the initial reporter, the broken blade was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1691782 | 1655227 | 1573324 | 2955842201000142 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-17 | 2010-04-09 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-03-17 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure one side of the tip of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument and detached left scissor blade were returned and evaluated. Per the customer reported complaint, engineering observed that the left blade had fractured at the cross section of the grip cable crimp. The fracture plane of the blade is nearly straight and half of the cable crimp is exposed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1691790 | 1655244 | 1573341 | 2955842201000143 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-03-16 | 2010-04-09 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-03-16 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-18 | Death | Death | Event Description: --/-- It was reported that post a da vinci s cox cryo-maze atrial fibrillation procedure, the patient suffered from hepatorenal failure and disseminated intravascular coagulation (dic). It was reported that the patient expired two days after the procedure. Reportedly, the patient's demise was unrelated to the da vinci s surgical system. --/-- Additional Event Description: --/-- | Manufacturer Narrative: --/-- No malfunctions of the da vinci s surgical system were reported to have occurred during the procedure and the hospital has continued to use the system to perform surgery on patients. As of april 8, 2010, no similar instances of this event have been reported to isi. --/-- Additional Manufacturer Narrative: --/-- Product code has been removed per fda request. | DA VINCI S SURGICAL SYSTEM | Reuse | 1691799 | 1655250 | 1573346 | 2955842201000141 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | InitialFollowup | 2010-03-20 | 2010-04-09 | Yes | No | Health Professional | IS2000 A51P8 |  | Yes | 2010-06-29 | Device Not Retu | No |  | 2010
IS1200 A43P9 | 2010-03-10 | Other | Other | It was reported that approximately 45 minutes into a da vinci prostatectomy procedure, while the surgeon was performing a dissection, the monopolar foot petal was pressed, however, the bipolar energy became activated. The bipolar instrument was holding tissue and created a superficial laceration type burn to the patient's bowel. A general surgeon reviewed the injury and determined that the burn was not through and through and did not require repair, however, the surgeon placed one suture in the patient's bowel as a precaution. The planned surgical procedure was completed with the same equipment and no further instances of the event was reported. | The investigation conducted by field service engineering (fse) consisted of reviewing system error logs and performing functional and electrical tests. The fse's review of the logs did not reveal any system errors and the system functioned as intended with no issues. The root cause of the customer reported failure mode cannot be determined. As of april 8, 2010, the hospital has continued to use the system with no reported recurrences of the reported issue. The instruments used in the case have not been returned for evaluation as of the time of this report. If additional information is received and / or if the instruments are returned and evaluated, a follow-up mdr will be submitted. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1691998 | 1655423 | 1568082 | 2955842201000140 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2010-03-10 | 2010-04-09 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2010-03-10 | Device Not Retu | No | Reuse | 2010
420227 | 2010-03-17 | Malfunction |  | It was reported that approximately 4 hours into a da vinci s sacrocolpopexy with hysterectomy procedure and while holding tissue with the endowrist prograsp instrument, the pk dissecting forceps instrument rubbed against the endowrist prograsp instrument causing shavings from the exterior coat of the instrument sheath to fall into the patient. The point of contact between the two instruments was out of the surgeon's view during the procedure; however, the surgeon observed shavings in the patient while viewing the abdomen at the end of the procedure. The shavings were irrigated and suctioned out of the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have various scratch marks and abrasions with material removed at the distal end. The scratches are not axially aligned with the tube. The area that contains scratches also has some abrasions parallel to the tube axis. There is another area with material removal 180 degrees from the scratches. This damaged section is 1. 5 long and parallel to the tube axis. No other damage found. See mdr 2955842-2010-00148 for report on endowrist prograsp instrument. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1692004 | 1655425 | 1573510 | 2955842201000146 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-17 | 2010-04-09 | No | Yes | Health Professional | 420227 | Device Returned To Manufacturer | Yes | 2010-03-17 | Device Not Retu | No | Reuse | 2010
42009307 | 2010-03-17 | Malfunction |  | It was reported that approximately 4 hours into a da vinci s sacrocolpopexy with hysterectomy procedure and while holding tissue with the endowrist prograsp instrument, the pk dissecting forceps instrument rubbed against the endowrist prograsp instrument causing shavings from the exterior coat of the instrument sheath to fall into the patient. The point of contact between the two instruments was out of the surgeon's view during the procedure; however, the surgeon observed shavings in the patient while viewing the abdomen at the end of the procedure. The shavings were irrigated and suctioned out of the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have various scratch marks and abrasions with material removed at the distal end. The scratches are not axially aligned with the tube. The area that contains scratches also has some abrasions parallel to the tube axis. There is another area with material removal 180 degrees from the scratches. This damaged section is 1. 5 long and parallel to the tube axis. No other damage found. See mdr 2955842-2010-00146 report on the pk dissecting forceps instrument. ,,. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1692007 | 1655426 | 1573511 | 2955842201000148 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-17 | 2010-04-09 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | Yes | 2010-03-17 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-10 | Other | Other | It was reported that while beginning a da vinci s surgical procedure, when docking the endoscopic camera manipulator (ecm) arm, the site was unable to lock or physically hold the ecm cannula mount in place. The patient was under anesthesia and the port placements were made when the surgeon made the decision to abort and reschedule the planned surgical procedure for a later date. No patient harm was reported. | The investigation conducted by field service engineering confirmed that the endoscopic camera manipulator (ecm) cannula mount was defective. The ecm cannula mount is designed to hold the ecm cannula securely in place. The system was repaired by replacing the affected cannula mount. As of april 8, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1692015 | 1655430 | 1573515 | 2955842201000147 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-03-10 | 2010-04-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-03-10 | Device Not Retu | No | Reuse | 2010
42022702 | 2010-03-01 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the insulation at the tip of the pk dissecting forceps instrument was observed to be thin. No additional information was provided. No patient harm, adverse outcome or injury was reported. Although the customer reported malfunction does not in itself constitute a reportable incident, during internal evaluation of the instrument, engineering found evidence of arcing. No patient harm was reported, however, we assessed the evaluation results and conservatively decided to report our findings as it cannot be determined whether or not the arcing occurred during a surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the conductor wire presents exposed wires due to arcing resulting in damage to the insulation. Both yaw pulleys exhibit charring on the same side as the conductor wire. The instrument passed electrical continuity testing. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1692016 | 1655431 | 1573516 | 2955842201000145 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2010-03-10 | 2010-04-09 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2010-03-10 | Device Not Retu | No | Reuse | 2010
| 2007-07-05 | Injury | Life Threatening | On march 12, 2010, isi received maude event report (b) (4) with the following event description: my name is xxx. I am the personal rep of the estate of xxx on xx xx 2007, my mother has a da vinci robot assisted laparoscopic hysterectomy at xxx medical center in xxx. It is my understanding that during this event dr xxx lost visualization and the instrument on the robot cut the right common iliac artery leading to hemorrhage, cardiopulmonary arrest and near exsanguination. The info i have been provided with indicates that this may have been reported to the fda as an maude adverse event associated with the da vinci surgical robot. I have not been able to locate such a report during my online search. Physician xxx. Medical facility xxx. No additional information was provided. | Due to limited information provided in the maude event report (b) (4), intuitive surgical is unable to provide additional information concerning the reported event. If additional information becomes available, a follow up mdr will be submitted. A review of our complaint database was performed; however, there was no similar event found related to this report. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1692021 | 1655432 | 1573517 | 2955842201000144 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-03-12 | 2010-04-09 | Yes | Yes | Health Professional |  | No | No | 2010-03-12 | No | No | Unkown | 2010
400169 | 2010-03-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the jaws of the harmonic curved shears (hcs) insert fell apart and the thicker portion of the grip fell inside of the patient. The broken piece was retrieved and the planned procedure was completed via traditional open surgical techniques. The initial reporter indicated that the case conversion was due to the patient's anatomy and unrelated to the hcs insert malfunction. | The hcs insert has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert is returned (post-evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 1690565 | 1654056 | 1572284 | 2955842201000138 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-09 | 2010-04-08 | No | Yes | Health Professional | 400169 | No | No | 2010-03-09 | Device Not Retu | Yes | Unkown | 2010
IS1200 A43P1 | 2009-03-30 | Other | Other | On (b)(6) 2010, a pt notified intuitive surgical in response to a request for their pt story. She advised that during a da vinci hysterectomy with sacrocolpopexy procedure, the case was converted to open surgical techniques resulting in an increased hospital stay, bladder trauma lasting one week post op and intermittent kidney pain. On (b)(6) and (b)(6) of 2010, the console surgeon provided info and operative reports related to this event. The info provided indicated that during 2 hours of the sacrocolpopexy portion of the procedure and while attempting to dissect the pt's sacral area and peritoneum, the surgical staff experienced intermittent difficulty obtaining a full range of motion with the pt side manipulator (psm). Multiple experienced surgical technicians were used in an attempt to troubleshoot the issue by repositioning the pt side cart (psc), the center post and the psms however, they were unable to gain a full range of motion in the left psm. Without contacting an isi technical support engineer for assistance, the surgeon made the decision to convert to a minilaparotomy to complete the planned procedure. | During an investigation with the console surgeon, he advised that post procedure, the staff found the pt side manipulators (psm) to be rotated 180 degrees. He also advised that there were no issues with the system and that the positioning of the psms caused the limited range of motion. An additional investigation conducted by engineering included a review of the system error logs which did not indicate any system faults that would have contributed to a limited range of motion in the psms. The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. Based on the info provided by the console surgeon and eval the system error logs, engineering concluded that the issue experienced by the customer was associated with incorrect positioning of the psms by the site. The site corrected the positioning of the arms to correct the reported issue. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1718156 | 1657827 | 1575669 | 2955842201000139 | 1 | NAY | Manufacturer | Other | NOT APPLICABLE | Initial | 2010-03-09 | 2010-04-08 | Yes | No | Health Professional | IS1200 A43P1 | No | No | 2010-03-09 | No | No | Reuse | 2010
IS1200 | 2007-12-14 | Injury | HospitalizationLife ThreateningRequired Intervention | Event Description: --/-- I was subject to the davinci surgical robot which you have approved. In their adverse report to you, the mfr reported there was no adverse effect to the pt. Yes there was. Really too many to list, but the robot broke down three times during the surgery and was sent back to the mfr of the product. I subsequently, within 6 months, had six hospitalizations after the surgery, so someone is lying to you. Either the doctor or the mfr. The following is there report to you: intuitive surgical, inc. Da vinci surgical system endoscopic instrument control system back to search results. Model number is1200 a4. 3p9. Event date (b) (6) 2007. Event type: other pt outcome other; mfr narrative: the investigation conducted by isi field service engineering concluded that system error codes #20009 and #21009 were experienced by the customer and not system error code #20008. This system error code was associated with a pt side manipulator -psm-. The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error codes #20009 and #21009 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor -encoder- and the secondary sensor -potentiometer. The system alarm -system generated fault code- functioned as designed and there was no injury to the pt. Upon determining this condition, the safety sys put da vinci in a "recoverable safe state". Based on the sys error logs two motor encoders were replaced. As of (b) (6), 2007, there have been no reported recurrences of the issue at this hosp. Event description it was reported that during a da vinci procedure, the customer experienced multiple occurrences of sys error code #20008. Initially, the site was able to override the error and continue with the case but the error recurred and could not be overridden. No pt harm, adverse outcome, or injury was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. Search alerts/recalls new search ¦ submit an adverse event report brand da vinci surgical system type of device endoscopic instrument control system mfr -selection d- intuitive surgical, inc. (b) (4). Device event key (b) (4) mdr report key (b) (4). Event key (b) (4). Report number 2955842-2008-00017. Device sequence number 1 product code nay report source mfr source type user facility reporter occupation unk remedial action repair type of report initial report date: 01/07/2008. One device was involved in the event. One pt was involved in the event. Date fda rec'd 01/07/2008. Is this an adverse event report? yes. Is this a product problem report? yes, device operator service personnel device model number is1200 a4. 3p9. Was device available for eval? device returned to mfr. Date returned to mfr: 12/21/2007. Is the reporter a health professional? no. Was the report sent to mfr? no. Date mfr rec'd: 12/14/2007. Was device evaluated by mfr? yes. Date device manufactured: 09/01/2000. Is the device single use? no. Is this a reprocessed and reused single-use device? no. Is the device an implant? no. Is this an explanted device? type of device usage reuse pt treatment data. Date received: 01/07/2008. Pt sequence number: 1. Treatment date: da vinci surgical sys instrument and accessories. Hospitalized 7 times after treatment with device. --/-- Additional Event Description: --/-- I was subject to the davinci surgical robot which you have approved. The mfr reported there was no adverse effect to the pt. Yes there was. Really too many to list but the robot broke down three times during the surgery and was sent back to the mfr of the product. I subsequently, within 6 months, had six hospitalizations after the surgery. Either the doctor or the mfr. Intuitive surgical inc. Da vinci surgical system endoscopic instrument control system. Back to search results: model # is1200. (b) (6). Event date: (b) (6) 2007. Event type: other. Pt outcome: mfr narrative: the investigation conducted by isi field service engineering concluded that system error codes #20009 and #21009 were experienced by the customer and not system error code #20008. This system error code was associated with a pt side manipulator -psm-. The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. System error codes #20009 and #21009 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor -encoder- and the secondary sensor -potentiometer-. The system alarm -system generated fault code- functioned as designed and there was no injury to the pt. Upon determining this condition, the safety system put da vinci in a "recoverable safe state". Based on the system error logs, two motor encoders were replaced. As of jan 07, 2007, there have been no reported recurrences of the issue at this hosp. Event description: it was reported that during a da vinci procedure, the customer experienced multiple occurrences of system error code #20008. Initially, the site was able to override the error and continue with the case, but the error recurred and could not be overridden. No pt harm, adverse outcome, or injury was reported. The procedure was converted to traditional open surgical techniques to finish the planned surgery. Search alerts/recalls. New search / submit an adverse event report. Brand name: da vinci surgical system. Type of device: endoscopic instrument control system. Mfr - intuitive surgical inc. (b) (4). (b) (4). Report# 2955842-2008-00017. Device sequence #1. Product code: nay. Report source, mfr, source type, user facility. Reporter occupation: unk. Remedial action: repair. Type of report initial report date: 01/07/2008. One device was involved in the event. One pt was involved in the event. Date fda received: 01/07/2008. Is this an adverse event report? yes. Is this a product problem report? yes. Device operator: service personnel. Device model #: is1200 (b) (6). Was device available for eval? device returned to mfr. Date returned to mfr: 12/21/2007. Is the reporter a health professional? no. Was the report sent to the mfr? no. Date mfr received: 12/14/2007. Was device evaluated by mfr? yes. Date device manufactured: 09/01/2000. Is the device single use? no. Is this a reprocessed and reused single-use device? no. Is the device an implant? no. Is this an explanted device? type of device usage: reuse. Pt treatment data: date received: 01/07/2008. Pt sequence #1 treatment. Treatment date: 1, da vinci surgical system instruments and accessories. | Manufacturer Narrative: --/-- | DAVINCI SURGICAL ROBOT |  | 983202 | 1660340 | 936901 | MW5015533 | 1 | NAY | Voluntary |  | Patient | FollowupInitial | 2010-04-07 | 2010-04-07 | Yes | Yes | Lay UserPatient | IS1200 | Device Returned To Manufacturer | No |  |  |  |  | 2010
420184 | 2010-03-10 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure and upon removing the permanent cautery spatula instrument from use, a plastic piece from the distal end of the instrument was found to be missing. The piece was not observed to have fallen in the patient nor observed to be in the patient. The site indicated that the instrument had been working prior to its removal and stated that the procedure was converted to traditional open surgical techniques due to the patient's anatomy. As of april 1, 2010, there have been no reports of patient harm, adverse outcome or injury. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1685452 | 1649448 | 1568066 | 2955842201000137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-03-11 | 2010-04-02 | No | Yes | Health Professional | 420184 | No | No | 2010-03-11 | Device Not Retu | No | Unkown | 2010
IS2000 A51P8 | 2010-03-08 | Other | Other | It was reported that while preparing for a da vinci s pyeloplasty procedure, the site experienced system message patient side cart not connected, check red cable and patient side cart power. With the assistance of an isi technical support engineer the site power cycled the system and cleaned and reconnected the patient side cart (pcs) cable, however, the message continued to occur. The patient was under anesthesia when the surgeon decided to abort the planned procedure and complete the procedure using traditional open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the red system cable. The red system cable connects the surgeon console and the patient side cart and passes video, audio, and data during system operations. The system was repaired by replacing the affected cable. As of april 1, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1685478 | 1649464 | 1568081 | 2955842201000136 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2010-03-08 | 2010-04-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-03-08 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-03-09 | Other | Other | It was reported that while preparing to begin a da vinci surgical procedure, the site observed the left eye image switching from green to black. With the assistance of an isi technical support engineer (tse), the site adjusted the camera cable at which time the issue moved to the right eye image. The tse recommended that the site replace the camera cable, however, the site was unable to locate their spare camera cable. The patient was under anesthesia when the surgeon decided to abort the planned procedure. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of april 1, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1685481 | 1649466 | 1568082 | 2955842201000135 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-09 | 2010-04-02 | Yes | Yes | Health Professional | IS1200 A43P9 | No | Yes | 2010-03-09 | No | No | Reuse | 2010
IS2000 A51P8 | 2010-03-12 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the site experienced a green image in both the left and right eye image when the camera cable was moved. The surgeon decided to convert to open surgical techniques to complete the planned procedure due to the camera cable issue and due to having found adhesions within the patient. No patient harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) was unable to replicate the customer reported problem, however, the vision issue experienced by the customer was likely associated with the camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the camera cable. The da vinci s surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1685489 | 1649468 | 1568084 | 2955842201000134 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-12 | 2010-04-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-03-12 | Device Not Retu | No | Reuse | 2010
PS3000 | 2010-04-01 | Injury | Disability | Davinci robot was in the corner in the or and machine was plugged in preparation for the case and the machine made a noise and then smoke came out of the machine. The machine was unplugged and removed from the surgical suite. No harm occurred to anyone. The mfr rep was on site for the case and pulled the battery from the machine. This is an upgraded machine. The lot number supplied is from the upgrade. |  | DAVINCI | NONE | 1691121 | 1652403 | 1570796 | MW5015430 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-04-01 | 2010-04-01 | No | Yes | Health Professional | PS3000 | Device Returned To Manufacturer | Yes |  |  |  |  | 2010
| 2010-03-29 | Injury | Required InterventionHospitalization | A patient was undergoing laparoscopic total hysterectomy in the main operating room using robotic equipment. The surgeon was exchanging the devices to allow for cautery of a small bleeder. The iliac artery was nicked in error. A vascular surgeon was called and arrived quickly to repair the vessel. The patient was transferred to the intensive care unit for observation. The patient is doing well. Note: during the case at the very end with insertion of robotic laparoscopic scissors through the right robotic sleeve, there was injury to the right external iliac artery. This resulted in laparotomy and repair by vascular surgery. The surgical assistant physician is a seasoned surgeon, but was not trained in robotic surgery. The estimated blood loss for the patient was 3 liters. ====================== health professional's impression======================user error, training. |  | HOT SHEARS | INSTRUMENT ENDOWRIST | 1686215 | 1650220 | 1568809 | 1650220 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-03-31 | 2010-03-31 | Yes | No | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
| 2010-03-29 | Injury | Required InterventionHospitalization | A patient was undergoing laparoscopic total hysterectomy in the main operating room using robotic equipment. The surgeon was exchanging the devices to allow for cautery of a small bleeder. The iliac artery was nicked in error. A vascular surgeon was called and arrived quickly to repair the vessel. The patient was transferred to the intensive care unit for observation. The patient is doing well. Note: during the case at the very end with insertion of robotic laparoscopic scissors through the right robotic sleeve, there was injury to the right external iliac artery. This resulted in laparotomy and repair by vascular surgery. The surgical assistant physician is a seasoned surgeon, but was not trained in robotic surgery. The estimated blood loss for the patient was 3 liters. ====================== health professional's impression======================user error, training. |  | HOT SHEARS | INSTRUMENT ENDOWRIST | 1686215 | 1650220 | 1568809 | 1650220 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-03-31 | 2010-03-31 | Yes | No | Physician |  | Yes | No Answer Provided |  |  |  |  | 2010
42018304 | 2010-02-23 | Malfunction |  | It was reported that during a da vinci s partial nephrectomy procedure, the end of the hook detached from the permanent cautery hook instrument and pieces fell into the patient. The pieces were retrieved and the planned surgical procedure was completed with a replacement instrument. The site indicated that the broken pieces were viewed under a magnifying glass to ensure that all parts were retrieved from the patient. No patient harm, adverse outcome or injury was reported. | The instrument and detached piece of the hook were returned and evaluated. Per the customer reported complaint, engineering observed that the hook had fractured within the yaw pulley, at the cross section of the distal most drilled hole. Engineering also observed a broken ceramic sleeve, with a 160" x. 135" piece detached from the distal end. When the detached piece is fit into the yaw pulley, a piece of the sleeve is still missing. No other damage was found. The initial reporter indicated that all broken pieces were successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1707904 | 1648863 | 1567548 | 2955842201000128 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-09 | 2010-03-30 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2010-03-09 | Yes | No | Initial | 2010
40018302 | 2010-03-09 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, a section above the hook of the permanent cautery hook instrument was burning. The instrument was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found arc tracking on the yaw pulley and distal clevis, indicating that an arcing event occurred. Most of the charring is evident on one side of the yaw pulley. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1713379 | 1648849 | 1567535 | 2955842201000130 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-09 | 2010-03-29 | No | Yes | Health Professional | 40018302 | Yes | Yes | 2010-03-09 | Yes | No | Reuse | 2010
420179 | 2010-03-01 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed was melting and arcing was observed. The surgical staff attempted to troubleshoot the issue by replacing the generators and the cable connecting the generator to the instrument. The instrument and tip cover were replaced with another mcs instrument and mcs tip cover accessory, however, the replacement tip cover also melted but no arcing observed. The site was able to successfully use a third mcs instrument and mcs tip cover accessory to complete to planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) concluded that the first issue of tip cover melting and arcing experienced by the customer was associated with a damaged mcs instrument. The fse arrived onsite during the case and observed that the mcs instrument first used in the case was physically damaged at the edge of the orange colored label on the instrument shaft. Additional engineering evaluation was performed per photographs provided by the fse and it was observed that the tube extension is dislodged from the main tube. The entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Evidence is not conclusive, however, it appears that the damage was most likely due to excessive side loading or other type of mishandling. The second mcs instrument used in the case was inspected by the fse and did not appear to have any physical damage, although the tip cover installed on the instrument had a melted spot where the plastic meets the white insulation. The fse was unable to determine if the instrument had malfunctioned and caused the melting of if the customer incorrectly installed the insulin cover on the instrument (not completely covering the orange label). The customer indicated that the latter may have happened, but was uncertain. The fse found no issues with the system. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1713395 | 1648850 | 1567536 | 2955842201000131 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-03-01 | 2010-03-29 | No | Yes | Health Professional | 420179 | No | No | 2010-03-01 | Yes | Yes | Unkown | 2010
42017909 | 2010-03-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the blade of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one scissor blade (left grip) is broken. The blade fractured at the cross section of the cable crimp. The fracture plane is straight. The cable crimp has slipped out of the swageless crimp pocket feature and the grip cable is wedged between the grips. The other blade exhibits indentation-like damage or pitting related to material removal near the tip. The instrument passed electrical continuity testing. No other damage was found. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1711399 | 1648851 | 1567537 | 2955842201000133 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-03-05 | 2010-03-29 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2010-03-05 | Yes | No | Reuse | 2010
42017909 | 2010-03-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed, a "gray and orange plastic piece" fell into the patient. A piece, approximately 2mm in length was not retrieved, however, the "gray" piece was found and removed from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed an oval shaped. 105" x 090" piece missing from the distal end of the tube extension. The missing piece is adjacent to one of the keys that mates with the proximal clevis. There is a scratch mark on the proximal clevis section revealed by the missing piece and various small scratches on the proximal clevis shoulder feature that acts as a hardstop for the tube extension. A tip cover believed to have been used in conjunction with the instrument was also returned with a piece of the sleeve broken off that matches the shape of the missing tube extension piece. Evidence is inconclusive, however, damage to the instrument most likely due to mishandling/misuse. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Manufacturer report # 2955842-2010-00119 was also submitted for the mcs tip cover accessory. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1711615 | 1648853 | 1567537 | 2955842201000120 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-03-05 | 2010-03-29 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2010-03-05 | Yes | No | Reuse | 2010
400180 | 2010-03-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed, a "gray and orange plastic piece" fell into the patient. A piece, approximately 2mm in length was not retrieved, however, the "gray" piece was found and removed from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover was returned with a. 290" x. 270" piece of the ultem sleeve detached. The location and shape of breakage matches the shape of the missing tube extension piece. The sleeve has a couple of gouge marks below the breakage. Most of the pad printing ink has been removed from the sleeve. Conclusions: engineering also observed that the tip cover accessory is damaged as the silicone exhibits various mechanical tears through its thickness in various locations. There is no evidence of arcing. Evidence is inconclusive, however, damage to the tip cover accessory appears most likely due to mishandling/misuse. No other damage was found. Manufacturer report # 2955842-2010-00120 was also submitted for the mcs instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL ACCESSORY | 1711601 | 1648854 | 1567539 | 2955842201000119 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-03-05 | 2010-03-29 | No | Yes | Health Professional | 400180 | Yes | Yes | 2010-03-05 | Yes | Yes | Initial | 2010
40016906 | 2010-03-08 | Malfunction |  | It was reported that during a da vinci surgical procedure, a piece of the harmonic curved shears (hcs) insert fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed the clamp arm was detached from the distal end. One side of the shaft that accepts the clamp arm pins is bent inwards. The curved blade has a couple of scratch marks slightly below the midpoint. Evidence in not conclusive, however, the damage to the insert appears most likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1713305 | 1648855 | 1567540 | 2955842201000132 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-03-08 | 2010-03-29 | No | Yes | Health Professional | 40016906 | Yes | No | 2010-03-08 | Yes | Yes | Initial | 2010
42022702 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the pk dissecting forceps instrument would not grasp. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. See section h10 for additional details. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. The instrument was placed on an in-house test system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in a ll directions and the grips opened and closed properly. The instrument passed a latex pull test, however, one grip is bent causing side to side misalignment of the grips. There is also an offset of approximately. 021" at the tips. Results and conclusions: engineering also observed that the conductor wire has exposed wires due to damage to the insulation as a result of arcing. The damage is located near the yaw pulley exit. The yaw pulley shows localized melting as well. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1711409 | 1648864 | 1567549 | 2955842201000129 | 1 | NAY | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2010-03-11 | 2010-03-29 | No | Yes | Health Professional | 42022702 | Yes | No | 2010-03-11 | Yes | No | Reuse | 2010
IS3000 A60P3 |  | Injury | Other | It was reported that post a da vinci si thyroidectomy procedure, the patient experienced numbness due to radial nerve damage. | On march 18, 2010, the surgeon reported that there was no malfunction of the da vinci si surgical system during the procedure and that the injury to the patient's radial nerve was due to the patient's arm positioning. The surgeon was unable to provide any details concerning the date that the event occurred or other patient details, however, the surgeon stated that the patient has fully recovered and sustained no permanent damage. As of march 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1683478 | 1646666 | 1565552 | 2955842201000122 | 1 | NAY | Manufacturer | Other | Other | Initial | 2010-02-25 | 2010-03-26 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2010-02-25 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2009-09-16 | Injury | Other | It was reported that post a da vinci s thyroidectomy procedure, the patient experienced numbness, due to radial nerve damage. | On march 18, 2010, the surgeon reported that the patient experienced motor weakness in her right hand that prevented her from opening her hand. The surgeon also reported that there was no malfunction of the da vinci s surgical system during the procedure, and that he believes the injury to the patient's radial nerve was due to the patient's arm positioning. In addition, the surgeon reported that after 3 months of out-patient physical therapy the patient has fully recovered and sustained no permanent damage. As of march 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1683467 | 1646669 | 1565555 | 2955842201000121 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-02-07 | 2010-03-26 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2010-02-07 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-02 | Other |  | It was reported that during a da vinci s prostatectomy procedure the site experienced difficulty clutching the endoscopic camera manipulator (ecm) arm. The customer contacted an isi technical support representative (tse) at which time the system began to fault a system error code #23034 and upon restarting the system would not home. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering confirmed the occurrence of system error code #23034. Engineering concluded that system error codes #23034 was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. System error code # 23034 appears when the da vinci s safety systems recognizes a clutch switch on the ecm to be stuck or in a pressed state after a specified amount of time. Upon determining these conditions, the safety systems put da vinci s in a "nonrecoverable safe state". The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The ecm was returned to isi for evaluation. Engineering confirmed the customer reported problem and found the ecm faulted immediately upon instrument clutching. The link-4 switch assembly was replaced to address the issue. As of (b)(6) 2010, there have been no reported recurrences of the issue as this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710352 | 1647073 | 1565939 | 2955842201000123 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-03-02 | 2010-03-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-03-02 | Yes | No | Reuse | 2010
IS3000 A60P3 | 2010-03-03 | Other |  | It was reported that while beginning a da vinci si surgical procedure the site experienced system error code #252. The system was restarted at which time system error code #31030 appeared. With the assistance of an isi technical support engineer the site cycled the power breakers on the vision side cart, emergency powered off the patient side cart and surgeon side console however, the system error code continued to occur. The patient was under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code #31030 and #307 found in the system error logs were associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. System error code 31030 occurs when the system detects that one of the camera controller units in the doco has failed to power on after multiple attempts or has shut down after initially powering up. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710335 | 1647074 | 1565940 | 2955842201000124 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2010-03-03 | 2010-03-26 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | Yes | 2010-03-03 | Yes | No | Reuse | 2010
IS1200 43P9 | 2010-02-04 | Other |  | It was reported that while preparing for a da vinci sacrocolpopexy procedure the camera was dropped. The site was unable to align the endoscope and tried three different endoscopes, however, they were unable to get the vision aligned properly. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found that all three of the sites endoscopes aligned properly with the camera. The fse trained the operating room staff on how to align the camera and endoscope properly as well as troubleshooting techniques. The site ordered a replacement camera as a precautionary measure. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710134 | 1647075 | 1565941 | 2955842201000125 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-04 | 2010-03-26 | Yes | Yes | Health Professional | IS1200 43P9 | No | Yes | 2010-02-04 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-03-11 | Other | Other | It was reported that during a da vinci s surgical procedure the message "trial software, no for human use" appeared after installing instruments onto the system. The system was restarted, however, the same message appeared again after installing an instrument. Multiple instruments were installed, with the same result. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) found that the "patient side cart verification" box was unchecked in the download application. When this box is checked, the download application runs a process which confirms that the installed software matches the released software. If the box is not checked, the confirmation process is not run and the "trial software" message is displayed on the screen when the system is running. It was determined that the box was unchecked when the fse programmed the system the previous day in the process of replacing an instrument arm. The fse was able to repair the system by rechecking the checkbox and reprogramming the system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710380 | 1647077 | 1565943 | 2955842201000126 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-03-11 | 2010-03-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-03-11 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2010-03-09 | Other | Other | It was reported that during the middle of a da vinci prostatectomy procedure the site experienced nonrecoverable system error code #20008. An isi technical support engineer (tse) was contacted, who had the site exercise an axis on one of the system's instrument arms. System error code #20008 would not clear and the tse then had the surgical staff power down, emergency power off the system, reset the console breaker, exercise the axis again, and then power back on the system. The system continued to generate error code #20008. The same actions were repeated, however, the error code continued to recur. A decision was made by the surgeon to abort the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering determined that system error code #20008 was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #20008 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a "nonrecoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure, however, multiple occurrences of the error code were confirmed in the system error logs. Per the data found in the system error logs, two potentiometer assemblies and a motor encoder were replaced. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710369 | 1647078 | 1565944 | 2955842201000127 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-03-09 | 2010-03-26 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-03-09 | Yes | No | Reuse | 2010
40017909 | 2010-03-01 | Malfunction |  | It was reported that while dissecting for a pelvic lymphadenectomy during a da vinci hysterectomy procedure, one of the blades of the monopolar curved scissors (mcs) instrument broke off and fell into the patient's abdomen. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. The initial reporter indicated the mcs instrument would not be returned to isi for evaluation. | Conclusions - since the instrument is not returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1711717 | 1646993 | 1565869 | 2955842201000117 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-03-01 | 2010-03-25 | No | Yes | Health Professional | 40017909 | No | Yes | 2010-03-01 | Device Not Retu | No | Unkown | 2010
42011704 | 2010-02-19 | Malfunction |  | Event Description: --/-- It was reported that while suturing the primary hernia for closure during a da vinci si diaphragmatic hernia repair procedure the tip of the 5mm needle driver instrument fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-eval) or if additional info is received. Per the info provided by the initial rptr, the broken instrument piece was successfully retrieved from the pt. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one needle driver grip was broken off at its base. The size of the broken grip piece is approximately. 225 x. 080. The outer surface of the intact grip has a couple of scratch/gouge marks near the tip. No other damage was found. Additional information can be found. | 5MM NEEDLE DRIVER INSTRUMENT | Unkown | 1701413 | 1646994 | 1565870 | 2955842201000116 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | InitialFollowup | 2010-02-23 | 2010-03-25 | No | Yes | Health Professional | 42011704 | Yes |  | 2010-02-23 | Device Not Retu | No |  | 2010
38003009 | 2010-03-24 | Malfunction |  | Robot was being tested prior to patient being brought into the operating room. The robotic arms didn't move like they should. The over-ride button was pushed, and the same issue occurred. The surgery was cancelled. The robot was rebooted and worked correctly but davinci repair technician advised not to use robot until arm could be replaced. |  | DAVINCI | ROBOT SURGICAL | 1684927 | 1649164 | 1567817 | 1649164 | 1 | NAY | User Facility |  |  | Initial | 2010-03-25 | 2010-03-25 | No | Yes | UNKNOWN | 38003009 | Yes | No Answer Provided |  |  |  |  | 2010
420205 | 2010-02-26 | Malfunction |  | On march 16, 2010, medwatch (b)(4) was received: "after initial inspection by the hospital staff and surgeon, the doctor inserted a fenestrated bipolar forceps instrument, item #420205 used in conjunction with robotic davinci device bearing serial number (b)(4). After some use, a wire was shown to be broken when the instrument was retrieved and a piece of plastic was missing from the end. The missing piece that parted from the forceps is approx. 2mm in diameter. The surgeon looked for this piece in the pt and was not able to find it. According to medical staff, the pt did not suffer any ill affects and did not have a prolonged stay. The pt was discharged as planned. ". | The instrument has not been returned for evaluation, however, the site provided images of the instrument for engineering review. Engineering found the conductor wire to be broken at the yaw pulley exit causing the wire to detach at the grip. The yaw pulley is missing a piece on the same side as the conductor wire breakage. A follow-up mdr will be submitted if additional information is received and/ or if the instrument is physically returned for evaluation. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2618831 | 2586341 | 2483034 | 2955842201000118 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-26 | 2010-03-25 | No | Yes | Health Professional | 420205 | No | No | 2010-02-26 | Device Not Retu | No | Unkown | 2010
42017909 | 2010-02-24 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument, the jaw broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the cable crimp and half of the crimp is exposed, the fracture plane is straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1711360 | 1646903 | 1565785 | 2955842201000115 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-02-24 | 2010-03-23 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-02-24 | Yes | No | Unkown | 2010
| 2010-03-08 | Malfunction | Other | Da vinci balloon trocar tip broke off during removal at end of case. Tip retrieved. |  | DAVINCI TROCAR BALLOON TIP | NONE | 1687322 | 1649756 | 1568364 | MW5015366 | 1 | LXH | Voluntary |  | RISK MANAGER | Initial | 2010-03-19 | 2010-03-23 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2010
| 2010-03-04 | Malfunction | Other | Da vinci balloon trocar tip broke off during removal. Part retrieved. |  | DA VINCI TROCAR BALLOON TIP | NONE | 1687326 | 1649757 | 1568365 | MW5015367 | 1 | LXH | Voluntary |  | RISK MANAGER | Initial | 2010-03-19 | 2010-03-23 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2010
42017909 | 2010-02-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissors instrument came apart and a piece fell into the pt. The piece was retrieved and the planned surgical procedure was successfully completed. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip blade is fractured and the broken piece was returned detached from the instrument. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is nearly straight. Half of the cable crimp is exposed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1709854 | 1646878 | 1565762 | 2955842201000112 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-02-22 | 2010-03-22 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-02-22 | Yes | No | Reuse | 2010
42017909 | 2010-02-23 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors instrument snapped off inside the pt. The piece was retrieved and the planned surgical procedure was successfully completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade was detached from the instrument. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is nearly straight. Half of the cable crimp is exposed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1709841 | 1646879 | 1565763 | 2955842201000111 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-02-23 | 2010-03-22 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-02-23 | Yes | No | Reuse | 2010
42018404 | 2010-02-23 | Malfunction |  | It was reported that during a da vinci s thymectomy procedure, a plastic piece of the permanent cautery spatula instrument came off inside the pt. The piece was retrieved and the planned surgical procedure was successfully completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | Conclusion: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire cap was dislodged from the distal clevis and shows signs of localized melting. The charring can be seen on the yaw pulley, the distal clevis ear, and the conductor wire. Most of the damage on these components are located on the side of the conductor wire. The conductor's insulating jacket melted down exposing wire at the wrist. Electrical continuity failed. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1709831 | 1646880 | 1565764 | 2955842201000108 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2010-02-23 | 2010-03-22 | No | Yes | Health Professional | 42018404 | Yes | No | 2010-02-23 | Yes | No | Reuse | 2010
IS3000 A60P3 | 2010-02-25 | Other | Other | It was reported that during a da vinci si nephrectomy procedure, the site experienced a non recoverable system error fault # 31030. The site restarted the system several times; however, the system error fault continued. With the assistance of an isi technical support engineer (tse), the system error logs were reviewed and error codes #31030 and #281 were found. The tse had the surgical staff power down, emergency power-off (epo) the system, and reset the breaker on the patient side cart (psc) and core. The psc to ssc fiber cables were also swapped. The system generated fault appeared again at start up. The site epo'd the system and reset the breaker again, then restarted the system without success. The procedure had been underway for approximately 1 hour and 30 minutes when the surgeon discussed the surgical options with the family and ultimately converted the planned procedure to open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes #31030 and #281 found in the error logs were associated with a setup joint (suj) arm. The suj is a non-motorized mechanical arm which the surgical team moves to position and support a patient side manipulator (psm) or endoscopic camera manipulator (ecm). It transmits information to and from the manipulator and from its own joint positions to the rac. Errors in this communication at system start-up would induce the system error codes in this case. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The system was repaired by replacing the affected suj. The suj was returned to isi for failure analysis investigation. Engineering confirmed the customer reported complaint as they were able to duplicate the system error codes. As a result, the suj was repaired by replacing several harness cables. As of march 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710291 | 1647068 | 1565934 | 2955842201000110 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2010-02-25 | 2010-03-22 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | No | 2010-02-25 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2010-02-26 | Other |  | It was reported that at the beginning of a da vinci thymectomy procedure, the surgeon experienced vision issues from the left and right eye at the surgeon side console (ssc). The surgeon made the decision to abort the planned procedure and reschedule for a later date. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer was associated with the camera control unit (ccu). The camera control units control the quality of the video images and produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera control unit. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1710185 | 1647069 | 1565935 | 2955842201000109 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-02-26 | 2010-03-22 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2010-02-26 | Yes | No | Reuse | 2010
42017909 | 2010-03-03 | Other | Other | It was reported that during a da vinci s hysterectomy procedure the metal band of the monopolar curved scissors instrument was heating up and burning the patient's tissue. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. During visual inspection, no signs of arcing were observed around the metal reinforcement ring. Functional testing was performed and the instrument properly cauterized at the tips and at no point did energy escape through the main tube or tube extension. The metal ring remained cool to the touch after using the electrocautery for 10 minutes. The instrument was disassembled and no damage was observed on the main tube, tube extension, conductor wire, or cap coupling tube in the area where the metal ring is located. No physical damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1710174 | 1647070 | 1565936 | 2955842201000113 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-03-03 | 2010-03-22 | Yes | Yes | Health Professional | 42017909 | Yes | No | 2010-03-03 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2010-02-01 | Other | Other | On (b)(6) 2010, a copy of medwatch uf/importer report (b)(4) was rec'd from (b)(6) hosp with the following info: pt admitted with left ureteropelvic junction obstruction for attempted robotic pyeloplasty. During the middle of the case, davinci robot lost the video picture. Physician notes state: "the robotic camera began to malfunction. We were not able to appropriately visualize the renal pelvis or the ureter because of the malfunctioning camera. Multiple attempts were made to fix the robotic camera without success. We therefore, decided to just convert to open procedure as we were not able to get the robotic system to work. An incision was made from the level of the twelfth rib towards the midline and carried down to the fascia, which was opened. The pt was awakened and taken to the recovery room in stable condition. " pt was discharged on (b)(6) 2010. Info consent included possibility of converting from laparoscopic to open procedure. The vision issue was initially reported to isi on (b)(6) 2010, however, isi was not made aware at the time of the report that the procedure was completed via traditional open surgical techniques. | Investigation was conducted by an isi field service engineer (fse) who attempted to replicate the vision issue by swapping cables and wiggling the cables throughout the system in relation to the vision path. Multiple power cycles were also performed. The fse was unable to induce the customer reported failure mode. The camera head/cable assemblies were verified and the black/white balance and scope alignment all passed. All connectors and pins were visually inspected and passed. The fse ensured that all boards were properly seated on the surgeon side console and noted that no error logs related to the vision issue were generated. While verifying the connections on the camera controller units and synchronizer, the fse observed that two 75 ohm terminators on the back of the synchronizer were missing. The 75 ohm terminator is a cover placed over an open video connection port to reduce electronic interference. The sys was repaired by installing two 75 ohm terminators on the synchronizer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1687807 | 1649160 | 1540086 | 2955842201000114 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2010-03-01 | 2010-03-22 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2010-03-01 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2010-02-16 | Other | Other | It was reported that during a da vinci s gynecological procedure, the customer experienced non recoverable system error code 23. With the assistance of an isi technical support engineer (tse), the site attempted to power down the system and move the patient side cart to surgeon side cart cable 180 degrees, however, system error 23 continued to occur. After troubleshooting the system error for 20 minutes, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the isi field service engineer concluded that system error code 23 was associated with the patient cart controller, printed circuit assembly (pcc, pca) boards. The pcc, pca is the brains of the patient cart. It receives the fiber optic connection from the surgeon side console and routes communications messages to itself and the four remote arm controllers (rac). It has several other control and monitoring functions for the patient cart system. The system was repaired by replacing the affected pcc, pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The pcc, pca was returned to isi for failure analysis investigation. Engineering found the pcc, pca to have fluid contamination which caused the system error code. As of march 16, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1669837 | 1634907 | 1554864 | 2955842201000107 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-02-16 | 2010-03-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-02-16 | Device Not Retu | No | Reuse | 2010
| 2010-02-23 | Malfunction |  | During a robotic case, the doctor was cleaning the permanent cautery hook with the maryland instrument and the hook broke off inside the patient. |  | ENDOWRIST | INSTRUMENT 8MM CAUTERY | 1675199 | 1640006 | 1559416 | 1640006 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-03-18 | 2010-03-18 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2010-02-15 | Other | Other | It was reported that during a da vinci s surgical procedure, the image through the high resolution stereo viewer (hrsv) became dark after the first few minutes of use. With the assistance of an isi technical support engineer, the site installed a new light guide cable; however, after a short period of time the image became dark again. The system was checked and the correct settings on the camera control unit (ccu) were verified. The light guide cable was found to have signs of charring. The tse recommended that the site install a back up lamp module or illuminator; however, the site made the decision to convert the planned procedure to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer found that the customer reported failure mode was attributed to light guide cables that were damaged due to an illuminator that was reconditioned by a third party. The light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected light guide cable and the illuminator. The site was advised to discontinue the practice of having isi product refurbished by a third party vendor as this could result in product quality issues. As of march 17, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 1668747 | 1633718 | 1553720 | 2955842201000106 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-15 | 2010-03-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2010-02-15 | Device Not Retu | No | Reuse | 2010
42009307 | 2010-02-08 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, fibers from the endowrist prograsp forceps instrument fell into the patient. The fragments were retrieved and the planned surgical procedure was completed. No patient harm, adverse event or complications were reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed the metal proximal clevis has become dislodged from the distal end of the main tube. The main tube interface wall has a section that has collapsed and is missing material. The grips and other metal components at the distal end exhibits scratches and gouges. Engineering concluded that breakage of the instrument is most likely due to overloading of the wrist and grip assembly. No other damage was found. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1668748 | 1633719 | 1553721 | 2955842201000104 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-18 | 2010-03-17 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2010-02-18 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-25 | Other | Other | It was reported that near the end of a da vinci s hysterectomy procedure, while the surgeon was performing the colpotomy using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, the surgeon felt the level of cautery coming from the mcs instrument was not sufficient. The surgeon backed up the camera view and at this time noticed superficial burns to the patient's bowel. A general surgeon was called in to review the patient's bowel and sutures were placed as a precaution. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. The isi representative present for the case inspected the mcs instrument and tip cover accessory after the procedure and noted that the tip cover was installed correctly and there were no visible holes or tears. The electro surgical unit (esu) setting was within isi recommended range. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On march 2, 2010 the isi representative that was present for the case followed up with the surgeon and found that the patient was doing fine and had no adverse effects from the procedure. | MONOPOLAR CURVED SCISORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1668750 | 1633721 | 1553723 | 2955842201000102 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | Initial | 2010-02-25 | 2010-03-17 | Yes | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-03-22 | Device Not Retu | No | Reuse | 2010
42017908 | 2010-03-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The fragment was retrieved and the procedure was completed with a replacement instrument of the same type. No patient harm, adverse event or complications were reported. | The instrument was returned and evaluated. Per the customer reported complaint one scissor blade is broken. The blade fractured at the cross section of the cable crimp. The fracture plane is straight. The intact blade does not exhibit damage at the tip. The instrument passed electrical continuity testing. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1668751 | 1633722 | 1553724 | 2955842201000105 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-03-01 | 2010-03-17 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-03-01 | Device Not Retu | No | Reuse | 2010
IS2000 A51 P8 | 2010-02-24 | Malfunction |  | It was reported that during a da vinci s pyeloplasty procedure, the bifurcated light guide cable began to burn the sterile drape creating a small flame in the operating room. The surgical staff was able to quickly extinguish the flame and the planned surgical procedure was completed with no patient harm reported. | The investigation conducted by field service engineering found the bifurcated light guide cable to have insufficient light transmission, however, no visible damage was found. The bifurcated light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The system was repaired by replacing the affected bifurcated light guide cable. After further investigation, it was discovered that the field service engineer notified the account during a maintenance visit on january 22, 2010, that this cable was defective due to the low light level it produced and should be removed and replaced, however, the account continued to use the affected cable. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1668760 | 1633732 | 1553733 | 2955842201000103 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-02-24 | 2010-03-17 | No | Yes | Health Professional | IS2000 A51 P8 | No | No | 2010-02-24 | Yes | No | Reuse | 2010
42017909 | 2010-02-15 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, one blade from the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken blade was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The cable crimp has slipped out of the swageless crimp pocket feature and the grip cable is wedged between the grips. The intact blade does not exhibit damage at the tip. Electrical continuity passed. No other damage was observed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1665558 | 1630071 | 1550397 | 2955842201000099 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-16 | 2010-03-12 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-16 | Device Not Retu | No | Initial | 2010
IS1200 A43P9 | 2010-02-12 | Other | Other | It was reported that during a da vinci surgical procedure, the site experienced system error 212. The site was unable to clear the system error. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code 212 was associated with a input output distribution board/remote interface adapter cable. The fse found one pin at the connection end of the cable to be recessed. The iod/ria cable is a bundle of wires which transmits power and sensor information between the surgeon side cart and the remote interface adapter (ria) board associated with a particular slave arm. The system was repaired by replacing the affected ria cable. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. As of march 11, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 1665562 | 1630078 | 1550401 | 2955842201000100 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-02-12 | 2010-03-12 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-02-12 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-02-16 | Other | Other | It was reported that while preparing for a da vinci si surgical procedure, the surgeon side console (ssc) lost power. With the assistance of an isi technical support engineer, the site cycled the power breaker on the ssc, checked the power cord connection and plugged in the ssc into multiple power outlets; however, the ssc continued to not power up. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the surgeon side console (ssc) power supply. The ssc power supply converts the ac voltage from the wall outlet to dc voltage for use by the system. The system was repaired by replacing the affected power supply. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1665563 | 1630079 | 1550402 | 2955842201000097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-16 | 2010-03-12 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-02-16 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-23 | Malfunction |  | It was reported that approximately 20 minutes into a da vinci s adrenal structure removal procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient's pelvic cavity. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane not straight. The intact blade does not exhibit damage at tip, however an indentation was observed at the base. Electrical continuity passed. No other damage was observed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1665564 | 1630080 | 1550403 | 2955842201000098 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-23 | 2010-03-12 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-23 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-02-12 | Other | Other | It was reported that during a da vinc s prostatectomy procedure, the customer experienced system error code 23008 while manually moving the endoscopic camera manipulator (ecm). With the assistance of a technical support engineer (tse), the site performed a fault over-ride, however, system error code 23008 recurred. The tse instructed the site to power down the system, reset the patient side cart breaker, emergency power off, and exercise the ecm, however, the issue was not resolved after restarting the system and moving the ecm. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | (b)(4): the system is not available for evaluation. The root cause of the customer reported failure mode cannot be determined, however, on september 6, 2009, during preventative maintenance, an isi field service engineer (fse) experienced several occurrences of system error code 23008 and recommended replacement of the endoscopic camera manipulator (ecm). The site made the decision not to replace the ecm due to budget issues and continues to use the system in it's current condition. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 1665576 | 1630097 | 1550415 | 2955842201000101 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2010-06-02 | 2010-03-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-06-02 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors instrument broke and a piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1664530 | 1629110 | 1549567 | 2955842201000096 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-09 | 2010-03-11 | No | Yes | Health Professional | 42017909 | No | No | 2010-02-09 | No | No | Unkown | 2010
42017604 | 2010-01-21 | Malfunction |  | It was reported that during a da vinci s transoral laryngectomy procedure, the shaft of the bullet nose dissector instrument scratched against the teeth of the patient and fragments from the instrument shaft fell into the patient. The fragments were retrieved. Based on the location of the patient's cancer, the surgeon made the decision to convert the planned procedure to open surgical techniques. No patient harm, adverse event or complications were reported. | The instrument was returned and evaluated. Per the customer reported complaint, the black shaft insulation presents multiple gouge marks within 5 from the interface with the snake wrist that have exposed the metal shaft below. Some insulation pieces are missing. No other damage was found. | BULLET NOSE DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1659587 | 1623308 | 1544218 | 2955842201000090 | 1 | NAY | Manufacturer | ForeignUser facility | BIOMEDICAL ENGINEER | Initial | 2010-02-03 | 2010-03-05 | No | Yes | Health Professional | 42017604 | Device Returned To Manufacturer | Yes | 2010-02-03 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-02-04 | Other | Other | It was reported that prior to starting a da vinci s procedure, the illuminator lamp module blinked several times and then went out. The site attempted to use another illuminator, however, the issue persisted. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm, injury or complications were reported. | The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was associated with the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. As of march 3, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 1659617 | 1623338 | 1544248 | 2955842201000087 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-02-04 | 2010-03-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-02-04 | Device Not Retu | No | Reuse | 2010
420184 | 2010-02-10 | Malfunction | Other | It was reported that during a da vinci si hysterectomy procedure, the site experienced difficulty removing the permanent cautery spatula (pcs) instrument through the cannla accessory. An isi representative attempted to troubleshoot with the site, however, the site was able to remove the instrument on their own. The surgeon made the decision to convert the planned procedure using open surgical techniques due to difficulty with removing soft fibroids imbedded in the patient's uterous robotically. While performing the open procedure, a piece of the pcs instrument was observed inside of the patient. The broken piece was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, however, on (b)(6) 2010, the site stated that the permanent cautery spatula instrument was bent at a 90 degree angle while they attempted to remove and it became jammed in the cannula accessory. A follow up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the fragment was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659618 | 1623339 | 1544249 | 2955842201000094 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-02-10 | 2010-03-05 | No | Yes | Health Professional | 420184 | No | Yes | 2010-02-10 | No | No | Unkown | 2010
IS2000 A51P8 | 2010-02-09 | Other | Other | It was reported that prior to docking the patient side cart to the patient for a da vinci s surgical procedure, the customer experienced system error codes 23020 and 23034. With the assistance of a technical support engineer (tse) the site cycled the endoscopic camera manipulator (ecm) clutch and insertion buttons, however, the system error codes would not clear. The site shut down the system, restarted in standalone mode, cycled the clutch buttons on the ecm and then emergency powered off the system, however, system error code 23034 recurred when the site restarted the system. The patient was under anesthesia and the port placements were made when the surgeon made the decision to abort the planned surgical procedure. No patient harm, injury or complications were reported. | The investigation conducted by field service engineering concluded that system error codes 23020 and 23034 experienced by the customer were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 23020 appears when one of the switches in the manipulator is showing inconsistent signals on it's two switch leads. System error code 23034 appears when the system detects that the ecm switch is stuck. The ecm was returned to isi for failure analysis investigation. Engineering was unable to replicate the system error codes experienced by the customer. It was concluded that system error codes 23020 and 23034 were not replicated because they can occur intermittently or the ecm switch was no longer stuck due to constant cycling. As a precaution, the switch assembly and the flat flex cable were replaced. As of march 4, 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1659622 | 1623344 | 1544253 | 2955842201000088 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-02-09 | 2010-03-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-02-09 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-08 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, one scissor blade (left grip) is broken. The blade fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit damage at the tip. Electrical continuity passed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659623 | 1623345 | 1544254 | 2955842201000093 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-08 | 2010-03-05 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-08 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-04 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed using a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument and broken blade were returned and evaluated. Per the customer reported complaint, the right blade was returned detached from the instrument. The blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed. The fracture plane of blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659624 | 1623346 | 1544255 | 2955842201000092 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-02-11 | 2010-03-05 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-02-11 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-09 | Malfunction |  | It was reported that during a da vinci s surgical procedure the monopolar curved scissors instrument sparked through plastic case when electrocautery was used. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 150 long x. 050 wide char mark on the distal end of the main tube, indicating that an arcing event occurred. The char mark is located directly above the metal tube reinforcement feature. The char mark did not burn through the entire thickness of the main tube wall and the tube does not exhibit axial cracking. Eighty - engineering also observed that the tube extension wall is circumferentially broken at the base of the axial keys and is partially dislodged from the main tube. No pieces are missing. Evidence is not conclusive, but this damage may be due to excessive side loading or other type of mishandling. The instrument was disassembled and char marks were observed on the inner surface of the metal tube reinforcement feature, most likely created as a result of the tube extension breakage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659626 | 1623349 | 1544257 | 2955842201000091 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-09 | 2010-03-05 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-09 | Device Not Retu | No | Reuse | 2010
IS2000 A43P9 | 2010-02-16 | Other | Other | It was reported that at the beginning of a da vinci surgical procedure, the site was unable to move a patient side manipulator (psm) arm and instruments were not being recognized when installed on the arm. An isi technical support engineer (tse) suggested that the customer check the drape and reseat the instrument sterile adapter on the suspect psm, however, the customer stated they were using the psm and would attempt these steps when changing the instrument. The site called back approximately 10 minutes later and stated they were getting nonrecoverable system error code 20008. Upon inspection, the site found broken cables on the suspect psm. At this time, the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer confirmed one patient side manipulator (psm) arm had broken cables and concluded that the system error code 20008 experienced by the customer was associated with this physical damage. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The psm was returned to isi for evaluation. Engineering found the axis 7 cables to be damaged, thus generating the system error code experienced by the customer. As march 4, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1659628 | 1623351 | 1544259 | 2955842201000089 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-16 | 2010-03-05 | Yes | Yes | Health Professional | IS2000 A43P9 | Device Returned To Manufacturer | No | 2010-02-16 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-03 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure and while tissue dissection, the tip of the monopolar curved scissors instrument broke and fell into the patient's abdomen. The surgeon made the decision to convert to traditional open techniques to successfully remove the tip and to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. Per the information provide by the initial reporter, the broken tip was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659630 | 1623354 | 1544261 | 2955842201000095 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-02-03 | 2010-03-05 | Yes | Yes | Health Professional | 42017909 | No | Yes | 2010-02-03 | Device Not Retu | No | Unkown | 2010
42018404 | 2010-02-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si general surgery procedure, a piece of the permanent cautery spatula instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire cap is missing and the boss feature of the yaw pulley that goes through the middle hole of the conductor cap is broken. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 1658595 | 1622509 | 1543492 | 2955842201000086 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2010-02-02 | 2010-03-04 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2010-05-11 | Device Not Retu | No |  | 2010
IS3000 A60P3 | 2010-02-05 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the physician at the surgeon side console intermittently felt a shock in his fingertips from the right master tool manipulator (mtmr). The surgeon switched to the site's second surgeon side console to complete the planned procedure and no further occurrences of shock were reported. No harm to the physician or patient was reported. | Per the customer reported complaint, the investigation conducted by the isi field service engineer found that the system passed electrical safety testing and all values performed to specification, however, as a precaution, the right master tool manipulator (mtmr) was replaced. The master tool manipulator refers to the master controllers which provides the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). An interview with the surgeon conducted by an isi representative on march 2, 2010, found that the surgeon was wearing standard scrubs and shoes during the procedure as well as the same chair that has been previously used in other da vinci cases. He stated that cautery was in use and that the patient was grounded as per the hospital's normal practice. He was unable to confirm that anything had changed in the operating room environment from prior cases. The mtmr was returned to isi for failure analysis investigation, however, the shock experienced by the customer could not be duplicated. As of march 3, 2010, there have been no reported recurrences of the issue at this hospital or by this surgeon. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1658636 | 1622530 | 1543507 | 2955842201000084 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-02-05 | 2010-03-04 | Yes | No | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | Yes | 2010-02-05 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-02-03 | Other | Other | It was reported that during set-up for a da vinci s surgical procedure, there was a collision between one of the instrument arms and the camera arm. After the collision, the instrument arm became unresponsive. The site restarted the system, however, the instrument arm remained unresponsive. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the system issue experienced by the customer was associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. It was determined that the collision between the affected psm and camera arm caused a cable to loosen inside the psm, thus preventing the psm from responding. The fse repaired the loose cable and as of march 3, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1658637 | 1622532 | 1543508 | 2955842201000085 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2010-02-03 | 2010-03-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-02-03 | Device Not Retu | No | Reuse | 2010
400184 | 2010-02-01 | Malfunction |  | On (b)(6) 2010, isi was advised of the following event and on (b)(6) 2010, user facility medwatch with the patient identifier of (b)(6) 2010 was received from the site. It was reported that during a da vinci s hysterectomy procedure and after installation of the permanent cautery spatula instrument, the site experienced an unknown system error that changed the led on the patient side manipulator (psm) arm to yellow. The surgical staff attempted to remove the instrument at a right angle. When straightening the tip of the instrument for removal from the system, the distal portion (plastic part) of the instrument's spatula broke off and fell into the patient's abdomen. The surgeon was working near the patient's sacral and efforts to locate the broken piece were unsuccessful. An x-ray was performed, however, the broken piece was not located. The procedure was completed with an instrument of the same type. To date, there have been no reports of patient harm, injury or complications. | The instrument has not been returned for evaluation; however, the site did provide photographs for review. Per the customer reported complaint, engineering observed the distal clevis ear had broken off at its base. The ear broke on the side opposite the conductor wire, however, the conductor cap is also missing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the user facility medwatch report, the patient did not suffer any ill affects and did not have a prolonged stay. The patient was discharged as planned. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1657673 | 1621491 | 1542540 | 2955842201000083 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-02 | 2010-03-03 | No | Yes | Health Professional | 400184 | No | No | 2010-02-02 | No | No | Unkown | 2010
IS3000 A60P3 | 2010-02-02 | Other | Other | It was reported that during a da vinci si procedure, the site experienced right eye video loss after placing the endoscope inside of the patient. With the assistance of an isi technical support engineer (tse), the site checked the camera cable connections and restarted the system. Restarting the system corrected the vision loss experienced by the customer. The site also reported that a cannula on a patient side manipulator (psm) was not detected. The site was able to resolve the cannula issue after re-docking the psm. The site called back approximately 30 minutes later to report that they experienced system error code 48232 and the message of no video in the right eye. Review of the site's system logs by an isi tse revealed that the site experienced several occurrences of the system error code 48232. The tse instructed the site to power cycle the system and the breakers on the vision side cart, however, the vision issue continued. The tse instructed the site to power cycle the double camera controller (doco), however, there was no change and the system continuously errored. The site called back approximately 15 minutes later to report that they were experiencing intermittent right eye video loss when moving the camera. The site told the tse that they will reconnect the camera cable connection to keep the system in 3-d vision. The patient was under anesthesia for approximately 2 hours and port incisions had been created when the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) determined that the system error code 48232 experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. System error code 48232 occurs when the system detects that one of the camera controller units in the doco has failed to power on after multiple attempts or has shut down after initially powering up. The doco was returned to isi for failure analysis. The customer reported complaint was confirmed as the right ccu of the doco did not function properly. The right ccu was replaced to repair the doco. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1656683 | 1620088 | 1541205 | 2955842201000082 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-02 | 2010-03-02 | Yes | Yes | Health Professional | IS3000 A60P3 | Device Returned To Manufacturer | No | 2010-02-02 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-02-02 | Other |  | It was reported that during a da vinci prostatectomy procedure, the surgical staff was unable to focus the camera head. An isi technical support engineer attempted to troubleshoot the issue and had the site verify that the cable connections were secure at the camera head and controller, however, the vision issue remained. The site also replaced the camera cable without success. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera head was returned to the original equipment manufacturer (oem) for evaluation. The oem discovered that the internal components of the stage had shifted, thus resulting in the vision issue experienced by the customer. The da vinci surgical system user's manual explicitly states that, environmental or equipment failures may cause the da vinci system to be unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1655644 | 1618029 | 1539237 | 2955842201000079 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-02 | 2010-03-01 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-02-02 | Device Not Retu | No | Reuse | 2010
IS3000 A51P8 | 2010-02-04 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the surgeon experienced video loss on the right eye image of the surgeon's console. An isi field service engineer was onsite during the surgical procedure and attempted to resolved the issue by powering down the system and replacing the camera head and camera cable, however, the issue remained. The patient had been under anesthesia for ~1 hour and docked to the system for ~30 minutes when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | During the additional investigation conducted by the field service engineer (fse), the fse determined that the vision issue experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. The doco was returned to isi for evaluation. Engineering discovered that the right ccu had malfunctioned, thus resulting in the right image vision loss experienced by the customer. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1655645 | 1618030 | 1539238 | 2955842201000080 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-02-04 | 2010-03-01 | Yes | Yes | Health Professional | IS3000 A51P8 | Device Returned To Manufacturer | No | 2010-02-04 | Device Not Retu | No | Reuse | 2010
42018304 | 2010-02-03 | Malfunction |  | It was reported that during a da vinci s general surgery procedure the tan piece of the permanent cautery hook (pch) instrument fell into the patient. The piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. The initial reporter indicated that the pch instrument would not be returning to isi for evaluation. | Since the instrument is not returning for evaluation, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1655647 | 1618032 | 1539240 | 2955842201000081 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2010-02-03 | 2010-03-01 | No | Yes | Health Professional | 42018304 | No | No | 2010-02-03 | No | No | Unkown | 2010
42017909 | 2010-02-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument, the jaw broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the cable crimp and the blade presents damage to the sharp edge near the base. The fracture plane is straight and the intact blade does not exhibit damage at the tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654623 | 1617084 | 1538352 | 2955842201000076 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-02-01 | 2010-02-28 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-28 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-01-26 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument, the jaw broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering observed one scissor blade (left grip) to be broken. The blade fractured at the cross section of the cable crimp and the fracture plane is straight. The other blade (right grip) presents damage along the sharp edge. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654626 | 1617088 | 1538354 | 2955842201000077 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2010-01-28 | 2010-02-28 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-28 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-01-19 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure the black cover that covers the wires of the monopolar curved scissors (mcs) instrument broke off, however, nothing fell into the patient. No patient harm, adverse outcome or injury was reported. The customer reported malfunction does not in itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the mcs tip cover accessory, which was installed on the mcs instrument during its return for evaluation. --/-- Additional Event Description: --/-- On (b)(6), 2010, (b)(6), nurse at (b)(6), provided additional information regarding this event. (b)(6) indicated that during a da vinci s prostatectomy procedure the console surgeon observed arcing coming from the monopolar curved scissors (mcs) tip cover accessory that was installed on a mcs instrument. The arcing caused a burn to the patient's bowel and although the surgical staff determined that the burn was not significant, the console surgeon repaired the burn as a precaution. (b)(6) reported that no post-operative complications have been experienced by the patient. (b)(6) also mentioned that the valley lab force fx esu fulgrate setting was set to 50w during the time of the event. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the black cover that covers the wires has broken off and the tube extension is dislodged from the main tube. The entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. The instrument was returned with a tip cover installed. Engineering evaluation of the tip cover observed a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is approximately. 010 in diameter and located. 185 above the silicone to sleeve interface. There are also multiple mechanical tears in the silicone at various locations, including a tear at the distal tip. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instruments and accessories user manual specifically states: - warning: do not exceed the 3kv peak limit. Doing so may result in electrical arcs and alternate site burns. - warning: use the lowest power setting possible for the minimum time necessary to achieve the desired effect. - warning: never increase the power settings without first checking both the active electrode and the patient return electrode and their connections. Use the active electrode or forceps only for the minimum time necessary to achieve the desired surgical effect in order to minimize the possibility of burns. - warning: excessive power levels may result in instrument malfunction and possible patient or user injury. Reduce power setting if any of the following effects are observed: excessive arcing, excessive tissue charring, excessive overheating of end effector (e. G. , end effector glowing red). The instruments and accessories user manual also indicates the corresponding maximum esu power settings to stay below the 3kv limit are as follows for the valleylab force fx esu: coag desiccate mode - max power setting is 120 watts coag fulgrate mode - max power setting is 38 watts coag spray mode - max power setting is 25 watts. | 8MM MCS TIP COVER ACCESSORY | Unkown | 1654628 | 1617090 | 1538356 | 2955842201000073 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2010-02-04 | 2010-02-28 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2010-04-02 | Device Not Retu | Yes |  | 2010
42017908 | 2010-01-04 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a hole was observed on the monopolar curved scissors instrument, resulting in exposed cables. The instrument was removed and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, a hole exposing the cables was observed on the distal end of the instrument sheath. The distal end of the main tube has a. 270 x. 100 hole burnt through it, exposing the hypotubes. The tube has axial cracking on either side of the hole and the tube cracking appears to start at one of the slots that accepts the tube extension keys. The instrument was disassembled and engineering discovered that the hypotubes were charred and the cap coupling tube was partially melted where the hole was created. This damage suggests that an arcing event occurred. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654629 | 1617091 | 1538357 | 2955842201000072 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-02-03 | 2010-02-28 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-02-03 | Device Not Retu | Yes | Reuse | 2010
IS2000 | 2007-08-13 | Death | Death | Event Description: --/-- --/-- Additional Event Description: --/-- It was reported that during a proctored da vinci s prostatectomy procedure, performed on (b) (6) 2007, the case wasn't progressing. The isi representative in attendance during the case asked the hospital's chair of urology to come to the or and make a decision on the procedure. It was decided to convert the case to traditional open surgical techniques. The next day, the isi representative was contacted by the hospital's chief of urology and notified that the patient had passed. The isi representative immediately provided this information to her direct manager at the time; however, it appears that the isi representative's direct manager did not internally escalate the event information via the appropriate channels. The isi representative's direct manager during the time of the event is no longer an isi employee, therefore, it cannot be determined why the information was not escalated appropriately. | Manufacturer Narrative: --/-- Product code has been removed per fda request. --/-- Additional Manufacturer Narrative: --/-- The following information was requested from the proctoring surgeon: patient's medical history, copy of operative report, and patient autopsy report. On february 17, 2010, the protoring surgeon's representative indicated that the information will be forwarded once the records are obtained. A follow-up mdr will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | Reuse | 1654631 | 1617093 | 1538359 | 2955842201000074 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2007-08-14 | 2010-02-28 | Yes | No | Health Professional | IS2000 | No | No | 2010-06-29 | Device Not Retu | No |  | 2010
42017909 | 2010-02-04 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the monopolar curved scissors instrument snapped and fell into the patient's abdomen. The piece was retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the right blade is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654632 | 1617094 | 1538360 | 2955842201000070 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-04 | 2010-02-28 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-04 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-02-02 | Malfunction |  | It was reported that during a da vinci s partial nephrectomy procedure, the tip of the monopolar curved scissors instrument snapped and fell into the patient's abdomen. The broken piece was retrieved and the surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument and broken blade were returned and evaluated. Per the customer reported complaint, the left blade of the scissors is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654633 | 1617095 | 1538360 | 2955842201000071 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-05 | 2010-02-28 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-05 | Device Not Retu | No | Reuse | 2010
42009307 | 2010-02-03 | Malfunction |  | It was reported that approximately 3 hours into a da vinci si prostatectomy procedure, the prograsp forceps instrument malfunctioned. Inspection of the instrument by the surgical staff found that the instrument was disjointed at the junction of the shaft and the endowrist instrumentation, thus requiring simultaneous removal of the cannula and instrument from the patient's abdomen. Shredded fibers from the sheath surrounding the instrument fell into the patient's abdominal cavity. The fibers were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the fragments were successfully retrieved from the patient. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1654636 | 1617099 | 1538363 | 2955842201000069 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-02-03 | 2010-02-28 | No | Yes | Health Professional | 42009307 | No | No | 2010-02-03 | No | No | Reuse | 2010
42017909 | 2010-01-19 | Other | Other | It was reported that after completion of a da vinci s sacrocolpopexy procedure, the surgical staff observed a burn at one of the cannula port incision sites. The burn was located on the skin surface and not internal. The patient did not require an additional procedure and it is unknown how the burn was treated. No additional patient harm, adverse outcome or injury was reported. | The instrument used in the procedure was returned and evaluated. Per the customer reported complaint, the instrument was visually inspected for any signs of arcing and no components exhibited char marks or localized melting. All distal end components are intact and undamaged. Electrical continuity passed. Performance testing was conducted and the instrument moved intuitively in all directions. Latex cut test passed. Instrument is fully functional. No other damage found. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of february 26, 2010, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654637 | 1617102 | 1538364 | 2955842201000078 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-02 | 2010-02-28 | Yes | No | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-02-02 | Device Not Retu | No | Reuse | 2010
VS100 | 2010-02-01 | Malfunction | Other | Pt admitted with left ureteropelvic junction obstruction for attempted robotic pyeloplasty which was converted to open pyeloplasty. During the middle of the case, davinci robot lost the video picture. Physician notes state: ". The robotic camera began to malfunction. We were not able to appropriately visualize the renal pelvis or the ureter because of the malfunctioning camera. Multiple attempts were made to fix the robotic camera without success. We therefore decided to just convert to open procedure as we were not able to get the robotic system to work. An incision was made from the level of the twelfth rib towards the midline and carried down to the fascia, which was opened. The pt was awakened and taken to the recovery room in stable condition. " pt was discharged on (b)(6) 2010. Informed consent included possibility of converting from laparoscopic to open procedure. |  | INTUITIVE DAVINCI ROBOT | DAVINCI ROBOT | 1659065 | 1618900 | 1540086 | 1618900 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2010-02-22 | 2010-02-26 | No | Yes | Health Professional | VS100 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2010
42018404 | 2010-01-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure while dissecting tissue, two pieces of the permanent cautery spatula instrument broke off and fell into the patient. Both pieces were retrieved and the planned surgical procedure was completed with an instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was returned with one of the distal clevis ears missing and the conductor cap detached from the distal end. The clevis ear broke at it's base, on the side opposite of the conductor wire. The size of the missing ear piece is roughly. 285" x. 240". Engineering concluded that the breakage was most likely due to overloading at the tip. The instrument passed electrical continuity testing. No other damage found. Per the information provided by the initial reporter, the broken pieces were successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1737152 | 1627156 | 1547776 | 2955842201000075 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-28 | 2010-02-26 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2010-01-28 | Yes | No | Reuse | 2010
IS3000 A61P8 | 2010-01-26 | Other | Other | It was reported that during a da vinci si surgical procedure, the site experienced system error code 41049 while clutching a patient side manipulator (psm) arm. The system was also faulting with multiple occurrences of system error code 23017. The surgical staff made the decision to not troubleshoot the system errors and aborted the procedure after anesthesia had been administered and ports placed. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes 41049 and 23017 were associated with the remote arm controller (rac) and/or the master tool manipulator (mtm) gimbal. The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected rac and mtm gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The 41049 system error code appears when the da vinci si safety system (supervisor) detected an internal logic error, thus putting the system in a recoverable safe state. System error code 23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The rac and mtm gimbal were returned to isi for evaluation. Engineering determined that the rac drive module was defective, resulting in the customer experienced system error codes. The mtm gimbal was evaluated and found to be within specifications with no defects. As of february 24, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1653554 | 1616224 | 1537600 | 2955842201000068 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-26 | 2010-02-25 | Yes | Yes | Health Professional | IS3000 A61P8 | Device Returned To Manufacturer | No | 2010-01-26 | Device Not Retu | No | Reuse | 2010
PS3000 | 2010-02-16 | Malfunction |  | During a prostatectomy case, the second arm on the davinci faulted according to the physician trying to operate. It would not accept a new instrument. The system was shut down and restarted. Technical support was notified. The arm did regain its abilities but resulted in a delay in completion of the procedure and longer time under anesthesia. There was no harm to the patient. This was the sixth time this system was used. It also failed during its third use. The system is being replaced by the manufacturer. |  | DAVINCI SI | HD SURGICAL SYSTEM | 1657280 | 1620554 | 1541614 | 1620554 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2010-02-25 | 2010-02-25 | No | Yes | Physician | PS3000 | No | No Answer Provided |  |  |  |  | 2010
40018302 | 2010-01-21 | Malfunction |  | It was reported that during a da vinci myomectomy procedure, it was noticed by the surgeon that the permanent cautery hook (pch) instrument was not responding. Inspection of the pch instrument revealed that the "plastic" piece near the hook broke off and fell into the patient's pelvis. The broken piece was retrieved and the surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one ear of the distal clevis is broken off at it's base. The size of the missing piece is approximately. 280" x. 215". The conductor cap is missing as well. The yaw pulley extruded boss feature that mates with the slot on the inner surface of the distal clevis ear is also broken off. The proximal clevis has a hairline crack on the same side as the distal clevis breakage. Engineering concluded that the damage to the instrument was most likely caused by overloading at the instrument tip. No other damage was found. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1704540 | 1627041 | 1547661 | 2955842201000066 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-25 | 2010-02-25 | No | Yes | Health Professional | 40018302 | Yes | Yes | 2010-01-25 | Yes | No | Reuse | 2010
42017909 | 2010-01-28 | Malfunction |  | It was reported that during a da vinci s procedure, the tip of the monopolar curved scissors instrument broke and fell into the pt. The broken piece was retrieved and the planned surgical procedure was successfully completed. There was no pt harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint the left blade is missing from the instrument. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly alright. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1738300 | 1627044 | 1547664 | 2955842201000067 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-01-29 | 2010-02-25 | No | Yes | Health Professional | 42017909 | Yes | No | 2009-01-29 | Yes | No | Reuse | 2010
42017909 | 2010-02-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade is detached and the broken piece was returned with the instrument. The broken blade had fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fractured plane of the blade is nearly straight. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1650082 | 1611429 | 1533214 | 2955842201000064 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-02-01 | 2010-02-22 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-02-01 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-01-28 | Other | Other | It was reported that while preparing the system for a da vinci s surgical procedure, the site experienced horizontal lines and strobing in the vision from the surgeon side console(ssc). With the assistance of an isi technical support engineer, the site re-seated the video cable and changed the camera controller unit settings, however, the image issue continued to occur. The patient had been placed under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule for another date. | The investigation conducted by field service engineering (fse) found that the customer had recently moved the da vinci s system and incorrectly plugged the composite video cable into the camera controller unit (ccu) thus resulting in the vision issues experienced by the customer. The system was repaired by removing the composite video cable. As of february 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1650083 | 1611430 | 1533215 | 2955842201000065 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-28 | 2010-02-22 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2010-01-28 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-01-27 | Injury | Required Intervention | It was reported that approximately 2. 5 hours into a da vinci si hysterectomy procedure, the surgeon began to initiate a tool change to switch instruments. At this time the nurse assistant who was retracting the patient's uterus, dropped the uterus causing the surgeon to lose site of the tennaculum forceps instrument in use. The patient side manipulator (psm) arm that the tennaculum forceps instrument was installed on became unresponsive. The isi representative onsite had the surigcal staff emergency stop the system and perform a fault over-ride, however, the psm continued to be unresponsive. The patient's uterus was moved, and the surgical staff found that the tennaculum forceps instrument was closed on the patient's rectum and they were unable to release the instrument. With the assistance of an isi technical support engineer, the instrument was released and removed from the psm. A general surgeon was requested to attend the procedure and although there were no visible holes, tears, or pinches in the patient's rectum, he placed a few sutures as a precaution. The planned surgical procedure was completed with out any further patient harm, injury or adverse outcome reported. | The isi representative onsite for this procedure stated that the surgeon may have accidentally closed the instrument grips inadvertently at the time the patient's uterus was dropped, causing the instrument to close on the patient's rectum. Since this event, the isi representative has provided additional training to the console surgeon and the supporting surgical staff regarding appropriate instrument tool changes. The patient went home the day after surgery as planned and was placed on stool softer for 1 week as a precaution. On (b) (6) 2010, during the patient's follow up appointment, she reported she was in no pain and had no adverse affects. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1650092 | 1611437 | 1533222 | 2955842201000060 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2010-01-27 | 2010-02-22 | Yes | No | Health Professional | IS3000 A60P3 | Yes | No | 2010-01-27 | Device Not Retu | No | Reuse | 2010
A51P8 IS2000 | 2010-01-26 | Other | Other | It was reported that while beginning a da vinci s prostatectomy procedure, the surgeon was receiving a double vision image from the surgeon side console (ssc). With the assistance of an isi technical support engineer, the site re-aligned their scope several times and exchanged the scope multiple times, however, the double vision continued to occur. The site stated the image looked normal without the scope and when not docked to the patient, however, as soon as they docked the system to the patient and when the camera head was moved, the image lost fusion and gave a double vision appearance. The patient was under anesthesia for approximately 45 minutes and the system had been docked to the patient for approximately 25 minutes when the surgeon decided to abort and reschedule the case for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that the double vision issue experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci s system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The camera head was returned and evaluated by the original equipment manufacturer (oem). The oem found the camera's stage was loose, causing the customer reported double vision. The camera was repaired by readjusting and repairing the stage. As of february 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1650143 | 1611500 | 1533267 | 2955842201000063 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-26 | 2010-02-22 | Yes | Yes | Health Professional | A51P8 IS2000 | Device Returned To Manufacturer | No | 2010-01-26 | Device Not Retu | No | Reuse | 2010
IS3000 A60P3 | 2010-01-21 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the surgeon console's right side image was unusually bright and the colors were off. An isi technical support engineer attempted to troubleshoot the issues over the phone and had the surgical staff reseat the camera cable and then replace the camera head, however, the vision issues remained. The patient was under anesthesia and the port incisions had been made when the surgeon made the decision to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the vision issues experienced by the customer were associated with the system's personality module video acquisition (pmva). The pmva is a printed circuit assembly (pca) which receives the stereo video images from the camera controller and outputs the audio and video for the touchscreen monitor. The system was repaired by replacing the affected pmva. The pmva was returned to isi for evaluation. In-house testing was performed and engineering was unable to replicate the customer reported failure. Additional engineering review confirms that the customer reported symptom is consistent with a loose connection on the pmva that was likely restored during removal from the system and preparation for test. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1737072 | 1622456 | 1543450 | 2955842201000062 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-01-21 | 2010-02-19 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | No | 2010-01-21 | Yes | No | Reuse | 2010
42017909 | 2009-10-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the blade from the tip of the monopolar curved scissors instrument broke while inside of the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported by this account. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was returned detached from the instrument. The blade fracture is located at the cross section of the grip cable crimp. Half of the cable crimp is exposed and the fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643919 | 1604824 | 1526884 | 2955842201000059 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2010-01-22 | 2010-02-18 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-22 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-01-20 | Other | Other | It was reported that 2. 4 hours into a da vinci s hysterectomy procedure, the site experienced double vision issues on the surgeon side console (ssc). The site replaced the endoscope with several other endoscopes with no change. The surgeon made the decision to convert to traditional open techniques to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the technical support engineer concluded that the reported vision issues experienced by the customer were associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The camera head was returned to isi for failure analysis. Engineering evaluation found that the camera head stage was damaged, this resulting in the vision issue. As of february 18, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1643921 | 1604826 | 1526886 | 2955842201000061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-20 | 2010-02-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-01-20 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-01-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument, the bottom tip of the jaw broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The mcs instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643922 | 1604828 | 1526887 | 2955842201000058 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-22 | 2010-02-18 | No | Yes | Health Professional | 42017909 | No | Yes | 2010-01-22 | No | No | Unkown | 2010
400180 | 2010-01-19 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the tip cover accessory installed on the monopolar curved scissors (mcs) instrument developed a hole. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument were returned and evaluated. Engineering observed a 0. 150 inch by 0. 080 inch oblong hole burned through the ultem sleeve and part of the silicone of the tip cover accessory. The center of the hole is located slightly below the silicone to sleeve interface and the ultem and silicone exhibit localized melting and charring indicative of arcing. The tip cover has multiple mechanical tears in different locations in the silicone. The mcs instrument was found to have a char mark on the proximal clevis. The shape and size of the char mark is similar to the hole burned through the tip cover. No other damage was found. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1643924 | 1604831 | 1526889 | 2955842201000057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-19 | 2010-02-18 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | Yes | 2010-01-19 | Device Not Retu | Yes | Initial | 2010
42017909 | 2010-01-19 | Malfunction |  | It was reported that 1. 5 hours into a da vinci s hysterectomy procedure, while dissecting tissue, the surgical staff noticed the tip of the monopolar curved scissors instrument to be missing. It was determined that the missing piece broke off the instrument and a fragment fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported by this account. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was returned detached from the instrument. The blade fracture is located at the cross section of the grip cable crimp. Half of the cable crimp is exposed and the fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643925 | 1604833 | 1526890 | 2955842201000056 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-20 | 2010-02-18 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-20 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-01-21 | Malfunction |  | It was reported that 1. 5 hours into a da vinci s hysterectomy procedure, while dissecting tissue, the surgical staff noticed the tip of the monopolar curved scissors instrument to be missing. It was determined that the missing piece broke off the instrument and a fragment fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported by this account. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was returned detached from the instrument. The blade fracture is located at the cross section of the grip cable crimp. Half of the cable crimp is exposed and the fracture plane of blade is nearly straight. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643926 | 1604834 | 1526890 | 2955842201000055 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-21 | 2010-02-18 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-21 | Device Not Retu | No | Reuse | 2010
420179 | 2010-02-01 | Injury | Required Intervention | A (b) (6) female was admitted to undergo robotic supra cervical hysterectomy. As reported during surgical procedure, surgeon noted scissors not opening. Device extracted from pt to examine, discovering one of the two tips was broken off. Missing part located and retrieved with no apparent injury to pt. Mfr rep (b) (4) notified, issued (b) (4). |  | HOT SHEARS | MONOPOLAR CURVED SCISSORS | 1653925 | 1610422 | 1532250 | MW5014800 | 1 | NAY | Voluntary |  | Nurse | Initial | 2010-02-03 | 2010-02-17 | No | Yes | Health Professional | 420179 | Device Returned To Manufacturer | Yes |  |  |  |  | 2010
420006 | 2010-01-18 | Injury | Required Intervention | A (b) (6) male about to under go robotic assisted laparoscopic prostatectomy, as reported by surgical nurse robotic equipment reviewed prior to start of procedure. Noted two robotic needle drivers to be defective with broken cables. Also defective a third driver, housing came apart. A forth driver was opened and found to be complete for use. No patient injury reported since above was noted prior to start of surgical procedure. Addendum: mfr contacted, response received same day by ms. (b) (6), arrangements made to replace defective robotic parts. (b) (4). |  | ENDOWRIST INSTRUMENT | LARGE NEEDLE DRIVER | 1650857 | 1609392 | 1531380 | MW5014762 | 1 | NAY | Voluntary |  | Nurse | Initial | 2010-02-01 | 2010-02-16 | No | Yes | Health Professional | 420006 | Yes | Yes |  |  |  |  | 2010
38003009 | 2010-02-08 | Malfunction |  | Surgeon was performing a total laparoscopic hysterectomy with left salpingoophorectomy with davinci. As procedure was being completed with suturing there was noted to be a general protection fault code and the robot had to be undocked, re-initialized and then redocked. |  | DAVINCI | ROBOT SURGICAL | 1639626 | 1599916 | 1522175 | 1599916 | 1 | NAY | User Facility |  |  | Initial | 2010-02-15 | 2010-02-15 | No | Yes | Physician | 38003009 | Yes | No Answer Provided |  |  |  |  | 2010
40016907 | 2009-12-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the harmonic curved shears insert accessory, which was inserted into the harmonic curved shears instrument, became unattached and fell into the patient. The insert was retrieved and the planned surgical procedure was completed with no patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory is returned for evaluation a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1638550 | 1598780 | 1521279 | 2955842201000053 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-12-14 | 2010-02-12 | No | Yes | Health Professional | 40016907 | No | No | 2010-12-14 | No | Yes | Unkown | 2010
42020704 | 2010-01-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s colectomy procedure, the tenaculum forceps instrument stopped working. When the surgical staff removed the instrument it was noticed that a plastic piece near the tip was missing. The missing piece was not observed to have fallen in the patient and additional exploration of the patient by the surgical staff could not confirm that the piece fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported complaint cannot be determined. If the instrument is returned and or if additional information is received, a follow up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned for evaluation. Per the customer reported complaint, engineering confirmed the distal end of the instrument is broken. The distal clevis ear on the opposite side of the conductor wire is broken at the base of the clevis ear. The detached ear piece, measuring approximately. 300 x. 215 inches, was not returned. The extruded boss on the yaw pulley hub that mates with the clevis ear is broken off and the conductor cap is missing. Also observed, the instrument is broken at the proximal clevis to main tube interface and the clevis is dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces. Electrical continuity passed. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | Reuse | 1638551 | 1598781 | 1521280 | 2955842201000052 | 1 | NAY | Manufacturer | Company RepresentativeOther | Health Professional | InitialFollowup | 2010-01-14 | 2010-02-12 | No | Yes | Health Professional | 42020704 | Device Returned To Manufacturer | Yes | 2010-02-16 | Device Not Retu | No |  | 2010
40018010 | 2010-01-15 | Other | Other | It was reported that a few hours into a da vinci s prostatectomy procedure, a hole was observed in tip cover accessory, which was installed on the monopolar curved scissors instrument, when stray current coagulated tissue. The site reported that the tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed the tip cover to have various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the hole which is indicative of an arcing event. All tip cover damage is contained within one half of the silicone, between the two parting lines. The hole sizes range from. 010 to 020 with varying shapes (round and oblong). Multiple holes indicate multiple arcing events occured. The tip cover also has various mechanical tears in the general vicinity of the holes. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1638553 | 1598784 | 1521282 | 2955842201000054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-15 | 2010-02-12 | Yes | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-01-15 | No | Yes | Initial | 2010
42022709 | 2010-01-08 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a piece of the pk dissecting forceps broke off and fell into the patient. The customer reported that the blade was removed from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported complaint cannot be determined. If the instrument is returned (post-evaluation) and/or if additional information is received, a follow up mdr will be submitted. There are no components in the pk dissecting forceps instrument that meet the definition of medical sharps. The customer reported blade is believed to have been the forcep grip. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1636427 | 1596895 | 1519670 | 2955842201000050 | 1 | NAY | Manufacturer | User facility | Health Professional | Initial | 2010-01-11 | 2010-02-09 | No | Yes | Health Professional | 42022709 | No | Yes |  | No | No | Unkown | 2010
400180 |  | Other | Other | It was reported that during a da vinci s surgical procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, the patient's tissue was burned. No additional patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation, however, an isi representative onsite a few days after the procedure, inspected the tip cover and did not find any holes or damage. Additionally, the electrosurgical unit (esu) setting used for the procedure was above isi's recommended level. If the accessory is returned for evaluation and / or if additional information is received, a follow up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1636428 | 1596896 | 1519671 | 2955842201000049 | 1 | NAY | Manufacturer | Company Representative | Health Professional | Initial | 2010-01-12 | 2010-02-09 | Yes | Yes | Health Professional | 400180 | No | Yes | 2010-01-12 | No | Yes | Unkown | 2010
42020705 | 2009-12-28 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure a cable on the tenaculum forceps instrument broke and a piece fell into the patient. The piece of cable was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found one pitch cable to be broken at the distal end. The other pitch cable is loose at the instrument wrist, however it was not visibly damaged. The cable crimp came out of the crimp pocket and was returned with the broken cable piece. Other cables at the wrist are not damaged. No other damge was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1636429 | 1596897 | 1519672 | 2955842201000048 | 1 | GEI | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2010-01-13 | 2010-02-09 | No | Yes | Health Professional | 42020705 | Device Returned To Manufacturer | Yes | 2010-01-13 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-01-12 | Other |  | It was reported that prior to using the da vinci s surgical system to perform surgical tasks for a nissen fundoplication procedure, the touchscreen monitor located on the patient cart displayed sdi searching. An isi technical support engineer (tse) attempted to troubleshoot the issue via telephone, however, the sdi searching message persisted and could not be cleared. The patient was under anesthesia when the tse advised the surgical staff that using the touchscreen monitor is optional and not required to perform the surgical procedure, however, the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported issue experienced by the customer was related to the system cable that connects the surgeon console and the patient cart. This cable passes video, audio, and data during system operation. The system was repaired by replacing the affected cable. The cable was returned to the original equipment manufacturer (oem) for evaluation. The oem observed a lot of exposed shield wire and that the coax and its associated shield are stressed. The cable also failed the oem's fiberoptic test, thus resulting in connectivity issues and generating the message displayed to the customer. It was determined that the cable was unrepairable and scrapped by the oem. As of february 4, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1635542 | 1595762 | 1518820 | 2955842201000051 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-12 | 2010-02-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2010-01-12 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-01-08 | Malfunction | Required Intervention | It was reported that approximately 1 hour post a da vinci si hysterectomy procedure, while cleaning the monopolar scissors (mcs) instrument, it was noted that one of the mcs instrument blades was sheared off. A flat x-ray of the patient's abdomen revealed that the scissor blade was located in the patient's pelvis. The patient was returned to the operating room and the instrument fragment was laparoscopically removed. On (b)(6), 2010, (b)(6) reported that there were no post operative complications experienced by the patient and the patient's hospitalization was not prolonged. | The monopolar curved scissors instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the instrument fragment was retrieved from the patient and there were no patient complications. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1635292 | 1595497 | 1518616 | 2955842201000041 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-08 | 2010-02-05 | Yes | Yes | Health Professional | 42017909 | No | Yes | 2010-01-08 | No | No | Reuse | 2010
IS2000 A51P8 | 2010-01-06 | Other | Other | It was reported that during a da vinci s segmentectomy procedure, the surgical staff experienced difficulty performing a guided tool change (gtc) and when the surgical staff installed a thoracic grasper instrument, the patient's pleura was nicked due to the instrument being advanced too far into the surgical field. With the assistance of a technical support engineer, the surgical staff performed the guided tool change correctly. The planned surgical procedure was completed and the affected area did not require repair or treatment of any kind. No patient complications were reported. | The investigation conducted by field service engineering (fse) found that there were no system errors and that the system was functioning to specification. It was concluded that the guided tool change (gtc) issue experienced by the customer was related to training as the site indicated to the fse that they were unaware of how to perform a guided tool change. On (b)(4) 2010, an isi clinical sales representative met with the surgeon to discuss the reported event and learned that the scrub technician assisting during the case was new and had not had a lot of experience operating the da vinci s system. On (b)(4) 2010 additional training was provided to the applicable surgical staff by isi regarding the gtc process. As of (b)(4) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1635293 | 1595498 | 1518617 | 2955842201000040 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-06 | 2010-02-05 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2010-01-06 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-01-07 | Other |  | It was reported that while starting the beginning of a da vinci surgical procedure the surgeon experienced an image with color bars on the left side view of the surgeon's console. With the assistance of an isi technical support engineer, the staff attempted to troubleshoot the issue by moving the camera cable, however, there was no change to the image. The site then moved the video output to the surgeon side console (ssc) and then to the camera control unit (ccu), however, the issue continued. The site then moved the video output to the synchronizer and the issue did not follow, indicating that issue was related to the video synchronizer. The surgeon had the option of continuing the procedure in 2d, however, made the decision to discontinue the case. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the system synchronizer. The synchronizer is designed as provide a 3d image after overlaying two images coming from the two separate camera control units with their own optical paths. The system was repaired by replacing the affected video synchronizer. As of (b)(6), 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1635294 | 1595499 | 1518618 | 2955842201000043 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-07 | 2010-02-05 | No | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-01-07 | Device Not Retu | No | Reuse | 2010
42011704 | 2010-01-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure a piece from the 5mm needle driver instrument broke off into the patient. The broken piece was retrieved and the procedure was successfully completed. No patient complications, harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1635296 | 1595502 | 1518620 | 2955842201000045 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-11 | 2010-02-05 | No | Yes | Health Professional | 42011704 | No | No | 2009-01-11 | No | No | Unkown | 2010
IS1200 A43P9 | 2010-01-15 | Other | Other | It was reported that when the surgeon sat down in the console to begin a da vinci prostatectomy procedure, system error code 20002 was observed on the screen. The patient was under anesthesia and the port incisions were made when the surgeon made the decision to abort the planned surgical procedure. On (b)(6) 2010 (b)(6) reported that after the patient was discharged, the patient reported feeling pain in the areas where the trocars had been placed and a burning sensation when urinating. Reportedly, the symptoms experienced by the patient have cleared. | The investigation conducted by field service engineering confirmed that the system error code 20002 experienced by the customer was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20002 appears when the da vinci safety systems determines that the joint angle, as measured by the primary sensor, is outside of the specified expected range. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The mtml was returned to isi for failure analysis investigation. Engineering evaluation found that the mtml was out of calibration, thus causing the system error code experienced by the customer. The mtml was recalibrated and as a precaution, the potentiometer and motor encoder were replaced. As of february 4, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1635298 | 1595504 | 1518622 | 2955842201000042 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-15 | 2010-02-05 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2010-01-15 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-01-07 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer was unable to focus the image on the high resolution stereo viewer (hrsv). The site tried multiple endoscopes, performed re-alignment, and power cycling the system, however, this did not resolve the issue. The site also reported that the patient cart was switching to battery when moved. The site agreed to continue with the procedure and call back after completion, however, they called back prior to completion and advised that they had swapped the camera head with no improvement. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv provides the video image for the surgeon console. The fse found that the power loss to the patient cart occurred when the power cord was stepped on while moving the cart, causing the power cord to become removed from the electrical outlet. The system was repaired by replacing the affected hrsv and reseating the power cord to the patient cart. The hrsv was returned to isi for evaluation. Engineering was able to replicate the reported vision issue and determined that the hrsv monitors had malfunctioned. As of january 28, 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1635299 | 1595505 | 1518623 | 2955842201000046 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-07 | 2010-02-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-01-07 | Device Not Retu | No | Reuse | 2010
42009307 | 2009-12-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, when the endowrist prograsp instrument was removed a silver object was observed inside of the patient. The silver object was retrieved and the planned surgical procedure was completed. Examination of the silver object by the surgical staff confirmed that it came from the endowrist prograsp instrument. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that instrument was returned with one distal idler pulley detached from the wrist and taped to the housing. The idler pin on the distal clevis does not appear to be swaged sufficiently to retain the pulley. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1635301 | 1595507 | 1518625 | 2955842201000044 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-12-31 | 2010-02-05 | No | Yes | Health Professional | 42009307 | Device Returned To Manufacturer | No | 2010-12-31 | Device Not Retu | No | Reuse | 2010
400077 | 2010-01-07 | Malfunction |  | It was reported that approximately one hour into a da vinci si gastric bypass procedure, while inserting the bowel grasper instrument through the cannula accessory, a piece of the 5mm cannula seal dropped into the patient's abdomen. The piece of the cannula seal was immediately retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Evaluation, conclusions: the 5mm cannula seal accessory has not been returned for evaluation, the root cause of the customer reported complaint cannot be determined. If the accessory is returned (post-evaluation) and/or if additional information is received, a follow up mdr will be submitted. | 5MM CANNULA SEAL | ENDOSCOPIC ACCESSORY | 1684334 | 1615624 | 1537047 | 2955842201000047 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-07 | 2010-02-05 | No | Yes | Health Professional | 400077 | No | Yes | 2010-01-07 | Device Not Retu | Yes | Unkown | 2010
42017909 | 2010-12-29 | Malfunction |  | It was reported while the surgeon performed a nerve sparing technique during a da vinci s radical retropubic prostatectomy procedure, the tip of the monopolar curved scissors instrument broke off and fell into the pt. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one scissor blade (left grip) is missing. The blade fractured at the cross section of the cable crimp. The fracture plane is straight. The broken blade presents damage related to material removal from the sharp edge near the base. The intact blade does not exhibit damage at the tip. The instrument passed electrical continuity testing. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1683726 | 1615360 | 1536786 | 2955842201000037 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-05 | 2010-02-03 | No | Yes | Health Professional | 42017909 | Yes | No | 2010-01-05 | Yes | No | Reuse | 2010
42017909 | 2009-12-30 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip of the monopolar curved scissors instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Results, conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one scissor blade (left grip) is missing. The blade fractured at the cross section of the cable crimp. The fracture plane is straight. The intact blade does not exhibit damage at tip. The instrument passed electrical continuity testing. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1654634 | 1615363 | 1536789 | 2955842201000039 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2010-01-05 | 2010-02-03 | No | Yes | Health Professional | 42017909 | Yes |  | 2010-01-05 | Yes | No | Reuse | 2010
40018009 | 2010-01-04 | Other | Other | It was reported that during a da vinvi si prostatectomy procedure, while using a monopolar curved scissors instrument with a tip cover accessory installed, a hole was observed in the tip cover accessory. Arcing from the mcs instrument with the tip cover accessory installed was not observed, however, tissue adjacent to the hole appeared "burned". On (b)(6) 2010, isi followed up with the laura woltz, the site's or material coordinator to verify if repair of the pt's burn injury was required, however, laura was unable to provide any additional information. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported failure mode. The tip cover was returned with a hole burned though the silicone. The silicone exhibits localized melting around the hole which is indicative of arcing. All tip cover damage is contained within one half of the silicone, between the two parting lines. The hole is approximately. 015" in diameter and is located. 145" above the silicone to sleeve interface. The silicone also exhibits various mechanical tears in the general vicinity of the hole. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1669546 | 1632753 | 1552896 | 2955842201000038 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-05 | 2010-02-03 | Yes | Yes | Health Professional | 40018009 | Yes | Yes | 2010-01-05 | Yes | Yes | Unkown | 2010
40018302 | 2010-01-05 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure a piece of the permanent cautery hook instrument jaw fell into the patient. The piece was retrieved and the planned surgical procedure was successfully completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument conductor cap had detached from the wrist and the boss feature of the yaw pulley that goes through the middle hole of the conductor cap is broken. The conductor cap does not appear to be damaged. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1630486 | 1590510 | 1514095 | 2955842201000027 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-05 | 2010-01-29 | No | Yes | Health Professional | 40018302 | Device Returned To Manufacturer | No | 2010-01-05 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-01-08 | Other | Other | It was reported that during a da vinci prostatectomy procedure the system's left master hand sensor separated from the grip and the instrument assigned to the left hand became unresponsive. The surgical staff was unable to reassemble the left master. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system issue experienced by the customer was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtml gimbal. The mtml gimbal was returned to isi for failure analysis investigation. Engineering observed that the flex cable had been snagged by someone/something, is torn at one end, and is hanging outside of the grip assembly. The damage thus resulted in the system issue experienced by the customer. As of (b)(6), 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1630487 | 1590511 | 1514096 | 2955842201000026 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2010-01-08 | 2010-01-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-01-08 | Device Not Retu | No | Reuse | 2010
42017909 | 2010-01-04 | Malfunction |  | It was reported that approximately 3 hours into a da vinci s hysterectomy procedure, while the surgeon was dissecting the patient's lymph nodes, the monopolar curved scissors (mcs) instrument with the tip cover accessory installed failed. Upon removal of the mcs instrument the surgical staff noticed the tip cover and instrument had a large burn mark on them. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. | Instrument and tip cover were returned and evaluated. Per the customer reported complaint engineering found the instrument proximal clevis and the distal end of the tube extension exhibit char marks. The proximal clevis is charred near the section of the exposed silicone seal. The seal has material removed typical of thermal damage. The tube extension has a. 175 inch by. 130 inch char mark directly below one of the keying features that mates with the proximal clevis, evidence that an arcing event likely occurred. Electrical continuity passed. No other damage found. Medwatch mfr report 2955842-2010-00030 was also sent to the fda regarding this event. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1630488 | 1590512 | 1514097 | 2955842201000029 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-04 | 2010-01-29 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-01-04 | Device Not Retu | No | Reuse | 2010
40018010 | 2010-01-04 | Malfunction |  | It was reported that approximately 3 hours into a da vinci s hysterectomy procedure, while the surgeon was dissecting the patient's lymph nodes, the monopolar curved scissors (mcs) instrument with the tip cover accessory installed failed. Upon removal of the mcs instrument the surgical staff noticed the tip cover and instrument had a large burn mark on them. The planned surgical procedure was completed with a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument were returned and evaluated. Per the customer reported complaint engineering found the tip cover has a large hole burned through the ultem sleeve and the silicone with the majority of the damage located on the ultem sleeve. The silicone exhibits localized melting around the hole, indicative of arcing. Engineering concluded that the arcing likely started on the silicone portion of the tip cover and spread down to the ultem sleeve. When the tip cover is installed on the suspected mcs instrument, the hole in the tip cover aligns with the char mark on the instrument's proximal clevis and tube extension. No other damage found. Medwatch mfr report 2955842-2010-00029 was also sent to the fda regarding this event. | TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 1630489 | 1590513 | 1514098 | 2955842201000030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-04 | 2010-01-29 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2010-01-04 | Device Not Retu | Yes | Initial | 2010
42017909 | 2009-12-23 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors instrument would not cauterize. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a reportable incident, however, during internal failure analysis evaluation of the instrument, engineering discovered evidence that an arcing event likely occurred. Although no patient harm was reported, we assessed the evaluation results and conservatively decided to report our findings as it cannot be determined whether or not the arcing event occurred during a surgical procedure. | The instrument was returned and evaluated. Per the customer reported complaint, engineering performed a functional test and found the instrument passed electrical continuity and cauterized at the tips as intended. Engineering also observed the instrument main tube to have a char mark at the distal end and the tube extension exhibits a charred section near the interface with the main tube. The main tube has material removed directly below the char mark. The instrument was disassembled to find the hypotubes contain charred material as well. The entire tube extension wall is circumferentially broken at the base of the axial keys, however, no pieces are missing. Engineering concluded that the breakage was likely due to excessive side loading of the instrument causing the breakage and contributing to the charred components. No other damage found. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1630490 | 1590514 | 1514099 | 2955842201000028 | 1 | NAY | Manufacturer | OtherUser facility | Health Professional | Initial | 2010-01-05 | 2010-01-29 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | Yes | 2010-01-05 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2009-12-30 | Injury | Required Intervention | It was reported, that while mobilizing the left internal mammary artery during a da vinci s cabg procedure, the micro bipolar forceps (mbf) instrument jumped forward and punctured the patient's artery. The damaged section of the artery was transected and the healthy portion of the artery was used to complete the bypass. It was also reported that throughout the case, pneumo was established and lost several times. The isi representative present during the procedure indicated that prior to the mbf instrument jump, the surgeon noted that when the master tool manipulator was moved, the instrument felt stuck and then moved. It is believed that the surgeon caught the mbf instrument in the patient's rib. The planned surgical procedure was completed with no further issues or patient harm. | (b) (4) - the system was evaluated by a field service engineer (fse). The fse found no error codes associated with the reported malfunction when reviewing the system error logs and was unable and to replicate the reported failure mode. Functional testing of the system revealed that the patient side manipulator (psm) on which the micro bipolar forceps instrument was installed functioned normally, however, an axis of the psm was found to be out of tolerance. Although the psm axis was found to be out of tolerance, it is highly unlikely that the failure mode in itself would have caused or contributed to the instrument jump experienced by the customer. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The affected psm was replaced to repair the system. A dvd of the recorded procedure was returned to isi for review. Engineering observed a sudden motion of the insertion axis of the psm causing the instrument to pierce the patient's artery. The distance traveled appeared to be larger than 1cm. Engineering also observed other instances where the instrument tip appeared to make sudden movements, however, they were of a smaller magnitude than the movement that caused damage to the patient's artery. Engineering concluded that the root cause of the instrument movement was due to motion being impeded by a large amount of static friction. The source of the friction is unknown. Some of the possible sources are sterile drape, a bent instrument, or a bent cannula. In this case, the surgeon would move the master tool manipulator (mtm) and the instrument would not move initially due to the friction, however, once the surgeon moved far enough for the instrument to overcome the friction, the instrument could make a quick jump to catch up with the mtm. The master tool manipulator refers to the master controller which provide the means for the surgeon to control the instrument and endoscope inside the patient. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure mode, however, the axis out of tolerance issue was confirmed and it was determined that the axis potentiometer and motor/encoder had malfunctioned. As of (b) (6) there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1630493 | 1590517 | 1514102 | 2955842201000031 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-12-30 | 2010-01-29 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2009-12-30 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-01-05 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site experienced a system error code 20008. With the assistance of an isi technical support engineer, the site undocked the endoscopic camera arm (ecm) and exercised the axis, however, the fault returned after pressing the override button. The site then powered down and exercised the ecm but the system error code occurred again after homing the system. The site attempted to troubleshoot a third time without success. The procedure had been underway for approximately an hour when the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering found that system error code 20008 experienced by the customer as well as system error code 21008 in the system log. The field service engineer concluded that the system error codes 20008 and 21008 were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. The 20008, and 21008 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a recoverable safe state. The ecm was returned to isi for failure analysis investigation. Engineering evaluation found that the potentiometer gear had malfunctioned, thus generating the system errors experienced by the customer. As of (b)(6) 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1630494 | 1590518 | 1514103 | 2955842201000032 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-05 | 2010-01-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | No | 2010-01-05 | Device Not Retu | No | Reuse | 2010
IS1200 A43P9 | 2010-01-04 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure, the customer experienced multiple system error codes 20009 and 21009 when initially inserting instruments into the patient after homing. The site attempted to override the system by exercising the patient side manipulator (psm) and by recycling power to the system, however, the error reoccurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found system error code 20009 and 21009 were associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 20009 and 21009 appear when the primary sensor in the specified manipulator measured an unexpected change in the angular position, resulting in system error code 20009 and 21009. Upon determining this condition, the safety systems put the da vinci in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the psm faulted immediately. The motor and encoder assembly were replaced. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1630495 | 1590519 | 1514104 | 2955842201000033 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2010-01-04 | 2010-01-29 | Yes | Yes | Health Professional | IS1200 A43P9 | Device Returned To Manufacturer | Yes | 2010-02-27 | Device Not Retu | No | Reuse | 2010
IS2000 A51P8 | 2010-01-05 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced system error code 23 and the led on one of the patient side manipulator (psm), arms illuminated red. The site attempted to repair the system by performing several troubleshooting techniques including restarting the system, however, the system was not able to complete start-up. The surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23 was associated with the remote arm controller (rac) and/or the embedded serilizer for setup joints (essj). The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The essj gathers position information and sends this information as serialized data to the rac. The system was repaired by replacing the affected rac and essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1630497 | 1590521 | 1514106 | 2955842201000034 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2010-01-05 | 2010-01-29 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2010-02-03 | Device Not Retu | No | Reuse | 2010
42000605 | 2009-11-10 | Malfunction |  | It was reported that during a da vinci si surgical procedure, fragments of a da vinci surgical endoscopic instrument were observed falling into the patient. The patient's affected area was irrigated and known pieces were removed. Upon removal of the instrument, the site observed the shaft of the instrument to appear as if it were scratched. No patient harm, adverse outcome or injury was reported. On december 24, 2009, the site returned two instruments for this event and on january 4, 2010, additional information was received stating that both instruments were used during the above mentioned procedure. See mdr 2955842-2009-00407. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube of the instrument to have sections with material removed and a rough surface finish. There is a scratch mark located. 925 above the proximal clevis that is not axially aligned with the tube. Above the scratch is a 1. 75 long section exhibiting wearing parallel to the tube axis. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1630498 | 1590522 | 1514107 | 2955842201000035 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2010-01-04 | 2010-01-29 | No | Yes | Health Professional | 42000605 | Device Returned To Manufacturer | No | 2010-01-04 | Device Not Retu | No | Unkown | 2010
420179 | 2009-11-05 | Other | Other | On (b)(6) 2010, intuitive surgical, inc. Received a copy of user facility medwatch (b)(4) from the fda. The event details are provided below: event desc: surgeon was operating with robotic scissors when it was observed that the scissors broke apart in the patient. The scissor arm was removed with what appeared to be half the scissor arm missing. Surgery was converted to laparotomy allowing for the robot to be removed and taken apart - missing scissor piece discovered in the robot. Rep. Notified, no known reason for arm breaking, the arm had 5 uses left. No injury noted to the patient at the time. Patient discharged and then returned to hospital with increased drainage one week later, abscess drained and treated for infection. Patient currently infection free and remains at home. Device usage problem: device failed (e. G. ) broke, couldn't get it to work or stopped working. | The monopolar curved scissors instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. The console surgeon and site's risk management department were contacted on three separate occasions to obtain additional information, however, no response has been received. The isi clinical sales representative (csr) for this site indicated that in december 2009 he was first contacted about the event. The csr was advised by the site that the procedure was challenging and that the site had considered converting to a laparotomy prior to the instrument breakage but made the final decision to convert once the instrument breakage occured. If the instrument is returned for evaluation or if additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1630499 | 1590523 | 1502450 | 2955842201000036 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2009-12-01 | 2010-01-29 | Yes | Yes | Health Professional | 420179 | No | No | 2009-12-01 | No | No | Unkown | 2010
40018010 | 2009-12-29 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while dissecting tissue with a monopolar curved scissors (mcs) instrument with a tip cover accessory installed, unintentional and superficial coagulation of the fat tissue near the patient's bowel was observed. Inspection of the mcs tip cover accessory by the surgical staff revealed that the tip cover had a hole. The planned surgical procedure was completed using a replacement tip cover accessory. The affected area did not require repair or treatment of any kind. No patient harm or adverse outcome was reported. | The mcs tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory and/or instrument is returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1629083 | 1589075 | 1512732 | 2955842201000025 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-29 | 2010-01-28 | No | Yes | Health Professional | 40018010 | No | Yes | 2009-12-29 | No | Yes | Unkown | 2010
42017909 | 2009-12-15 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors instrument broke and a fragment fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported by this account. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was returned detached from the instrument. The blade fracture is located at the cross section of the grip cable crimp. The fracture plane of blade is straight. The intact blade does not exhibit damage at the tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1627560 | 1586637 | 1510341 | 2955842201000023 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2009-12-31 | 2010-01-26 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-22 | Device Not Retu | No |  | 2010
42017909 | 2009-12-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors instrument broke and a fragment fell into the patient. The broken piece was retrieved and the planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported by this account. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was missing. The blade fracture is located at the cross section of the grip cable crimp. The fracture plane of blade is straight. The intact blade does not exhibit damage at the tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1627571 | 1586649 | 1510352 | 2955842201000024 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2009-12-31 | 2010-01-26 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-01-22 | Device Not Retu | No |  | 2010
42017909 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the blades of the monopolar curved scissors instrument would not open. No additional info was provided. No pt harm, adverse outcome or injury was reported. Although the customer reported malfunction does not in itself constitute a reportable incident, during internal evaluation of the instrument, engineering found evidence of arcing. No pt harm was reported, however, we assessed the evaluation results and conservatively decided to report our findings as it cannot be determined whether or not the arcing occurred during a surgical procedure. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as functional testing found that the instrument blades open and close properly and passed latex cut testing. The instrument moves intuitively with full range of motion in all directions. Engineering also observed charring on the tube extension and proximal clevis, indicative of arcing related damage. The distal end of tube extension has a. 105" x. 055" char mark at the interface with the proximal clevis. The proximal clevis also exhibits charring directly above the tube extension char mark and next to the exposed section of the proximal clevis silicone seal. A section of the proximal clevis seal has material removed as a result of thermal damage. The location of this damage is approx where the tip cover silicone meets the ultem sleeve. The instrument passed electrical continuity testing. No other damage was found. A mcs tip cover accessory was returned with the instrument, however, the tip cover accessory does not exhibit any signs of arcing. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1677870 | 1592781 | 1516169 | 2955842201000021 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-12-28 | 2010-01-26 | No | Yes | Health Professional | 42017909 | Yes | No | 2009-12-28 | Yes | No | Initial | 2010
IS2000 A51P8 | 2009-12-28 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced system error code #23008 and #23017. With the assistance of a technical support engineer, the site restarted the system and exercised the affected axis of the endoscopic camera arm, however, the site was unable to override the system error codes. The surgeon decided to convert the procedures to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system error codes #23008 and #23017 were associated with the endoscopic camera manipulator(ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. System error code #23017 occurs when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation. Engineering evaluation found that the potentiometer assembly and motor encoder were defective, thus generating the system error codes experienced by the customer. The potentiometer assembly and motor encoder were replaced to repair the ecm. As of (b)(6) 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1682717 | 1594122 | 1517448 | 2955842201000022 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-28 | 2010-01-26 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-12-28 | Yes | No | Reuse | 2010
42017206 | 2009-12-23 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon was unable to grasp using the maryland bipolar forceps instrument to control the patient's bleeding. The surgeon made the decision to convert the procedure using traditional open surgical techniques. | The instrument was returned and evaluated. Engineering investigation found that the grips of the instrument open and close, however, the grasping functionality is reduced in the instrument's current condition. One yaw pulley has a. 175" x. 060" piece broken off on the side opposite of the conductor wire. This breakage allows the grip to move within the yaw pulley and allows a larger range of motion in the yaw, and reduces the effectiveness of the grasping force. Engineering also observed that the conductor wire and main tube are damaged. The wire on the side opposite of the yaw pulley breakage has broken at the grip. The grip is crimped properly where the wire is installed. The distal end of the main tube has various scratches and abrasions with material removal and a rough surface finish. The scratches are not axially aligned with the tube and the abrasions are parallel to the tube axis. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENTS | ELECTROSURGICAL INSTRUMENT | 1683561 | 1594116 | 1517444 | 2955842201000014 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-12-23 | 2010-01-22 | Yes | Yes | Health Professional | 42017206 | Yes | Yes | 2009-12-23 | Yes | No | Reuse | 2010
| 2010-01-08 | Injury | Required Intervention | During the davinci instrument arm cleaning process, it was noted that one complete side of the monopolar scissor was sheared off. The csr supervisor reported the incident and the scissor to the clinical coordinator for surgical inventory for verification. Surgeon and pa were notified. Flat plate x-ray of abdomen was ordered and confirmed image was consistent with a scissor blade being seen in the pelvic. Pt returned to operating room for removal of the scissor fragment. |  | INTUITIVE SURGICAL | MONOPOLAR CURVED SCISSORS | 1628215 | 1584210 | 1507977 | MW5014415 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2010-01-19 | 2010-01-20 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2010
38003009 | 2010-01-04 | Malfunction |  | Patient brought to hospital for robot assisted laparoscopic prostatectomy. Underwent anesthesia, laparoscopic ports placed into patient. Robot docked to the patient, right arm of the robot would not function. Surgeon spoke with intuitive technician. As arm would not function, patient was awakened from anesthesia, discharged home and brought back at future date for surgery. Per surgeon, or staff, this has occurred before which necessitated the robot being "rebooted", this time the "reboot" did not correct the problem. |  | DAVINCI | ROBOT SURGICAL | 1625879 | 1584394 | 1508161 | 1584394 | 1 | NAY | User Facility |  |  | Initial | 2010-01-18 | 2010-01-18 | No | Yes | Physician | 38003009 | No | No Answer Provided |  |  |  |  | 2010
IS2000 A51P8 | 2009-12-18 | Other |  | It was reported that during a da vinci s prostatectomy procedure, the site experienced system error code #23008. The site pressed fault override and cycled power to the system, however, the error recurred during homing. The site cycled power to the system several more times, however, the error code recurred. The tse instructed the site to exercise an axis of the right master tool manipulator (mtm) several times and re-home, however, the system again faulted with system error code #23008. The tse instructed the site to power down and emergency power off the surgeon side console (ssc). After performing these tasks, the site observed the right mtm bumping into the right side of ssc during homing and system error code #23008 recurred. The tse instructed the site to power down and rotate the mtmr towards the center and restart the system, however, the mtmr again bumped into the ssc and system error code #23008 recurred. The pt was under anesthesia for approx 2 hours and the port placements were made when the surgeon made the decision to abort and reschedule the surgical procedure for a later date. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controller which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarms (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. The mtmr was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure mode, however, as a precaution, the potentiometer and motor were replaced. As of jan 15, 2010, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1628971 | 1584464 | 1508231 | 2955842201000010 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-12-18 | 2010-01-15 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-12-18 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2009-12-21 | Other |  | The isi rep on site reported that during a da vinci s mitral valve repair procedure, multiple occurrences of system error code #23008 were experienced. With the assistance of a technical support engineer, the site was instructed to exercise the axis of the right and left master tool manipulators (mtm) when error code #23008 recurred, however, when the error code recurred, the site was unable to exercise the axes of the mtm. The site power cycled the system, however, when the system powered up and homed, the system error code #23008 recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer (fse) was unable duplicate the error code #23008, however, as a precaution, the fse replaced the right master tool manipulator (mtm) gimbal. The mtm gimbal is a subsection of the complete mtm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "nonrecoverable safe state". The mtm gimbal was returned to isi for evaluation. Engineering evaluation found that the axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. The axis potentiometer was replaced. As of (b)(4) 2010 there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1628949 | 1584466 | 1508233 | 2955842201000019 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-12-21 | 2010-01-15 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-12-21 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2009-12-22 | Other |  | It was reported that while draping a pt side manipulator for a da vinci surgical procedure, the site experienced nonrecoverable fault error code #20013. The pt was under anesthesia when the surgical staff made the decision to abort the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the nonrecoverable system error code #20013 was associated with a pt side manipulator (psm) arm. The system was repaired by replacing the affected psm. The psm is an instrument arm located on the pt side cart and provides the sterile interface for the endowrist instruments. The system error code #20013 appears when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a "nonrecoverable safe state". The psm was returned to isi for evaluation. Engineering was unable to replicate the customer reported failure mode, however, an axis potentiometer and motor/encoder were replaced as a precaution since the error was confirmed in the system logs. As of (b)(6) 2010, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1629078 | 1584468 | 1508235 | 2955842201000020 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-22 | 2010-01-15 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-12-22 | Yes | No | Reuse | 2010
40018009 | 2009-12-18 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing was seen coming from the side of the mcs instrument. The mcs instrument was removed and the tip cover accessory was replaced. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the holes, indicating multiple arcing events occurred. The hole sizes range from under. 010 inches to. 030 inches with varying shapes (round and oblong) and are located from. 235 inches to. 560 inches above the silicone to sleeve interface. Mechanical tears are in the general vicinity of some of the holes. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 1628131 | 1584542 | 1508306 | 2955842201000016 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-21 | 2010-01-15 | No | Yes | Health Professional | 40018009 | Yes | Yes | 2009-12-21 | Yes | Yes | Initial | 2010
42017909 | 2009-12-18 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was using the monopolar curved scissor (mcs) instrument, the tip broke off and fell into the patient. The instrument fragment was immediately retrieved. The mcs was replaced with another instrument of the same type and the planned surgical procedure was completed with no patient harm, adverse outcome or injury reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1628106 | 1584544 | 1508308 | 2955842201000017 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-12-23 | 2010-01-15 | No | Yes | Health Professional | 42017909 | No | No | 2009-12-23 | Device Not Retu | No | Unkown | 2010
40011710 | 2009-12-17 | Malfunction |  | It was reported that during a da vinci gastric bypass procedure, the tip of the needle driver instrument broke off and fell into the patient. The broken piece was retrieved and the planned procedure was completed using a replacement instrument of the same type. No patient harm, injury or complications were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one needle driver grip returned detached from the grip assembly. The grip broke off it's base, where the grip hub meets the grip arm. The grip fracture is located near a 90 degree corner between the hub and arm. The outer surfaces of both grips exhibit scratch marks. No other damage was found. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1628117 | 1584545 | 1508309 | 2955842201000015 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-12-17 | 2010-01-15 | No | Yes | Health Professional | 40011710 | Yes | No | 2009-12-17 | Yes | No | Unkown | 2010
40017908 | 2009-12-17 | Malfunction |  | It was reported that during da vinci surgical procedure, the outside of the monopolar curved scissors (mcs) instrument was compromised, causing a burn to the patient. Debridement of the affected area was performed. | Results/conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of main tube has a 2" long section exhibiting charring, with a. 100" x. 065" hole burned through the tube indicative of arcing. The hole is located roughly. 500" above the tube extension and has a crack below it that is axially aligned with the tube. The crack is located halfway between (and 90 degrees from) two opposing slots that mate with the tube extension axial keys. The instrument passed electrical continuity testing. No other damage was found. Several attempts have been made to obtain additional information from the customer, however, the site reported that they are unable to provide any additional details related to this event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1628417 | 1584546 | 1508310 | 2955842201000018 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-17 | 2010-01-15 | Yes | Yes | Health Professional | 40017908 | Yes | Yes | 2009-12-17 | Yes | No | Initial | 2010
42017909 | 2009-10-12 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, a blade from the monopolar curved scissors instrument broke into the pt. The blade was retrieved and the planned surgical procedure was successfully completed. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the blade fractured at the cross section of the grip cable crimp with the fractured plane of the blade nearly straight. Half of the cable crimp is exposed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1674200 | 1588732 | 1512381 | 2955842201000011 | 1 | NAY | Manufacturer | ForeignUser facility | Other | InitialFollowup | 2009-12-15 | 2010-01-14 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-03-25 | Device Not Retu | No |  | 2010
42017909 | 2009-11-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, a blade from the monopolar curved scissors instrument broke off into the pt. The blade was retrieved adn the planned surgical procedure was completed using laparoscopic techniques. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- Additional investigation performed by an isi representative found that the original procedure was completed with a replacement instrument of the same type. It was originally reported that the procedure was converted to a traditional laparoscopic techniques for completion. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the blade to be fractured at the cross section of the grip cable crimp with the fractured plane of the blade nearly straight. Half of the cable crimp is exposed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1654624 | 1588734 | 1512383 | 2955842201000013 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupFollowup | 2009-12-15 | 2010-01-14 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2010-03-25 | Device Not Retu | No |  | 2010
420179 | 2009-12-15 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument was "burned". The site contacted an isi technical support engineer (tse) to troubleshoot and the tse concluded that the site may have improperly installed of the mcs tip cover accessory on the mcs instrument. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) found that the shaft of the mcs instrument used in conjunction with the mcs tip cover accessory has a melted hole. The fse tested the system and found that it functioned properly. The mcs tip cover accessory and mcs instrument were not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch report was sent to the fda for the mcs tip cover accessory: mfr report # 2955842-2010-00008. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1674254 | 1588735 | 1512384 | 2955842201000009 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-21 | 2010-01-14 | No | Yes | Health Professional | 420179 | No | No | 2009-12-21 | Device Not Retu | No | Unkown | 2010
42017908 | 2009-11-19 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, a blade from the monopolar curved scissors instrument broke off into the pt. The blade was retrieved and the planned surgical procedure was successfully completed. No pt harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the blade to be fractured at the cross section of the grip cable crimp with the fractured plane of the blade nearly straight. Half of the cable crimp is exposed. No other damage found. --/-- Additional Manufacturer Narrative: --/-- (b)(4). | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 1674209 | 1588737 | 1512386 | 2955842201000012 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupFollowup | 2011-06-06 | 2010-01-14 | No | Yes | Health Professional | 42017908 | Device Returned To Manufacturer | No | 2010-03-26 | Device Not Retu | No |  | 2010
400180 | 2009-12-15 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tip cover accessory installed on the monopolar curved scissors (mcs) instrument was burned. The site contacted an isi technical support engineer (tse) to troubleshoot and the tse concluded that the site may have improperly installed the tip cover accessory on the mcs instrument. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer (fse) found that the shaft of the mcs instrument used in conjunction with the tip cover accessory has a hole melted on the instrument shaft. The fse tested the system and found that it functioned properly. The mcs tip cover accessory and mcs instrument were not returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The follow medwatch report was sent to the fda for the mcs instrument: mfr # 2955842-2010-00009. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1670472 | 1588739 | 1512388 | 2955842201000008 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-15 | 2010-01-14 | No | Yes | Health Professional | 400180 | No | No | 2009-12-15 | Device Not Retu | Yes | Initial | 2010
50053 | 2009-11-06 | Malfunction |  | Dr. Operating with robotic scissors when it was observed that the scissors broke apart in patient. The scissors arm was removed with what appeared to be half the scissors arm missing. Surgery converted to laparotomy allowing for the robot to be removed and taken apart - missing scissors piece discovered in the robot. Rep. Notified, no known reason for arm breaking, the arm had 5 uses left. No injury noted to the patient at that time. Patient discharged and returned one week later to hospital with increased drainage, abscess drained and treated for infection. To date, the patient is infection free and remains at home. |  | MONOPOLAR CURVED SCISSORS | SCISSORS | 1620805 | 1578221 | 1502450 | 1578221 | 1 | NAY | User Facility |  |  | Initial | 2010-01-11 | 2010-01-11 | No | Yes | Invalid Data | 50053 | Yes | No Answer Provided |  |  |  |  | 2010
400180 | 2009-11-30 | Injury | Required Intervention | It was reported that during a da vinci s gynecological procedure, an arc was observed coming from the tip of the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. A small hole was observed in the pt's aorta and was repaired by suture. The planned surgical procedure was completed with no add'l pt harm, adverse outcome or injury. | The tip cover accessory and monopolar curved scissors instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument or accessory are returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1622944 | 1580190 | 1504336 | 2955842201000005 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-12-11 | 2010-01-08 | Yes | Yes | Health Professional | 400180 | No | No | 2009-12-11 | Device Not Retu | Yes | Unkown | 2010
IS2000 A60P2 | 2009-12-03 | Injury | Required Intervention | It was reported that the day following a successful da vinci si prostatectomy procedure, the pt developed compartment syndrome in their right arm. A surgical procedure was performed to relieve pressure to the pt's affected area. No add'l pt harm, adverse outcome or injury was reported. | An investigation conducted by an isi clinical sales rep found that the surgeon was unable to determine a root cause for the pt's injury, however, the surgical team evaluated the following factors as possible contributors. Pt side assistant was leaning on pt's right arm during part of the procedure. Surgical staff tucked pt's right arm too closely to the pt during procedure. Scaring tissue from pt's previous rotator cuff surgery on right arm. As of 2010, the hosp has continued to use the system with no reported recurrences of the reported issue. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1622964 | 1580191 | 1504337 | 2955842201000007 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-12-10 | 2010-01-08 | Yes | No | Health Professional | IS2000 A60P2 | No | No | 2009-12-10 | No | No | Reuse | 2010
400180 | 2009-12-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed arcing coming from the tip of the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover accessory was replaced to successfully complete the planned procedure with no pt harm, adverse outcome or injury reported. | The tip cover accessory and monopolar curved scissors instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument or accessory are returned for eval, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1661658 | 1582524 | 1506550 | 2955842201000006 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-12-11 | 2010-01-08 | No | Yes | Health Professional | 400180 | No | No | 2009-12-11 | Device Not Retu | Yes | Unkown | 2010
42018304 | 2009-11-30 | Malfunction |  | It was reported that during a da vinci s adrenalectomy procedure, two pieces of the permanent cautery hook instrument broke off and fell into the pt's abdomen. The broken pieces were removed using a grasper forceps instrument. No pt complications, harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the hook and ceramic sleeve to be detached from the yaw pulley with some pieces missing. The hook fractured at the cross section of one of the drilled holes in the straight shank and the ceramic sleeve is broken off as well. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1661482 | 1582525 | 1506551 | 2955842201000003 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2009-12-13 | 2010-01-08 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-12-13 | Yes | No | Reuse | 2010
42022702 | 2009-12-10 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the case was converted to traditional open surgical technique due ot the pt's anatomy. The entire procedure was reported to be lengthy and complex due to the pt's anatomy and required multiple teams, multiple trays, and multiple surgical techniques consisting or robotic assisted laparoscopic and traditional open techniques. As a precautionary measure, the surgical staff performed a ct scan and an x-ray which revealed what appeared to be a small, curly wire in the pt's left pelvic area. The surgical staff then observed the pk dissecting forceps instrument to be broken, suggesting that a piece of the instrument may have fallen into the pt. An exploratory laparoscopic procedure was performed, however, the piece observed under x-ray could not be located. No add'l info has been provided. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval, a follow-up mdr will be submitted. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1738280 | 1582526 | 1506552 | 2955842200900432 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-10 | 2010-01-08 | No | Yes | Health Professional | 42022702 | No | Yes | 2009-12-10 | Device Not Retu | No | Unkown | 2010
IS2000 A51P8 | 2009-12-09 | Other | Other | It was reported that when the console surgeon sat down in the console, to begin a da vinci s hysterectomy procedure, the left eye image through the high resolution stereo viewer was green and flickering. With the assistance of an isi technical support engineer, the site restarted the system several times, however, the issue persisted. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the vision issue experienced by the customer was related to the camera controller unit (ccu) and the camera cable. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the ccu and the camera cable. As of (b)(6), 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1627127 | 1582889 | 1506853 | 2955842201000002 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2009-12-09 | 2010-01-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-12-09 | Yes | No | Reuse | 2010
IS1200 A43 P9 | 2009-12-09 | Other | Other | It was reported that while starting a da vinci surgical procedure, the surgical staff had difficulty docking the system to the pt. The nurse was unable to unclutch the endoscopic camera manipulator (ecm) arm and restarted the system multiple times, however, they were unable to move the ecm. No attempt to contact isi technical support was made. The surgeon decided to convert to traditional open surgical techniques, due to the pt's age, to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering (fse) found that there were no system errors in the system log and that the ecm was functioning to specification. The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The site reported that the issue they experienced was associated with a non experienced surgical staff assisting during the procedure. The hospital has since ensured that a trained da vinci surgical staff will be present for all scheduled cases. As of (b)(6), 2010, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1627120 | 1582890 | 1506854 | 2955842201000004 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-09 | 2010-01-08 | Yes | No | Health Professional | IS1200 A43 P9 | Yes | Yes | 2009-12-09 | Yes | No | Reuse | 2010
IS1200 A43P9 | 2009-12-02 | Other | Other | On (b)(6) 2009, isi received a complaint related to the below report and on (b)(6), 2009, user facility medwatch report (b)(4) was received with additional info, this is a reportable event. It was reported that during a da vinci prostatectomy procedure, the site experienced a non-recoverable error code #20008. With the assistance of an isi technical support engineer (tse), the site exercised the right master tool manipulator (mtmr) on the surgeon side console, however, the issue persisted. The tse instructed the site to undock, emergency power off and restart the system, however, system error code #212 appeared. The site exercised the mtmr for a second time, but was unable to recover from the error code #212. The pt was under anesthesia and the port placements were made when the surgeon decided to convert to open surgical techniques. No pt harm, injury or complications were reported. It was reported that the pt spent one day in the icu due to blood loss and the length of the procedure. As of (b)(6) 2009, the pt was reported as doing well and has not returned to the hospital due to any post operative complications. | The investigation conducted by the field service engineer concluded that system error codes #20008 and #212 were associated with the right master tool manipulator. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #20008 appears when the da vinci safety system determines that the angular positon of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specific tolerance for agreement. Upon determining this condition, the safety systems put davinci in a "recoverable safe state. " system error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Since the motor's velocity is a component of the expected voltage calculation, this error can be caused by a faulty encoder. The mtmr was returned to isi for failure analysis investigation. Engineering eval found no issues with the mtmr; however, as a precaution, the potentiometer and motor were replaced. As of january 6, 2010 there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1629160 | 1586313 | 1489229 | 2955842201000001 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-12-11 | 2010-01-08 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-12-11 | Yes | No | Reuse | 2010
IS2000 A51P8 | 2009-12-03 | Other | Other | It was reported that during a da vinci s surgical procedure the customer experienced system error code #23008. The system was restarted, however, system error code #212 occurred during homing. By rotating an axis on the left master tool manipulator, restarting the system a couple of times, and hitting the emergency power off button, the surgical staff was able to get the system to work for an additional 2 hours. After two hours, system error code #23008 re-appeared. The site restarted the system, however, the error code continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes #23008 and #212 were associated with the left master tool manipulator. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state". System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. The mtml was returned to isi for failure analysis investigation. Engineering reviewed the system error logs and confirmed the system error codes experienced by the customer. Based on the system error logs, an axis potentiometer and motor encoder were replaced. As of december 23, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1744908 | 1576692 | 1501028 | 2955842200900429 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-12-03 | 2009-12-31 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2009-12-03 | Yes | No | Reuse | 2009
IS2000 A60P2 | 2009-12-01 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the left master tool manipulator (mtml) was unresponsive and a grinding noise was heard coming from the mtml. The surgical staff attempted to resolve the issue by swapping instruments, however, the mtml continued to remain unresponsive and the grinding noise continued to recur. The patient had been under anesthesia for approx 1 hour and the ports had been in place for approx 10 minutes when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the system issues experienced by the customer were associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The mtml was returned to isi for failure analysis investigation. Engineering evaluation discovered an axis cable with loose tension, thus causing the system issues experienced by the customer. As of december 23, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1743598 | 1576693 | 1501029 | 2955842200900428 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-01 | 2009-12-31 | Yes | Yes | Health Professional | IS2000 A60P2 | Yes | No | 2009-12-01 | Yes | No | Reuse | 2009
42018404 | 2009-12-02 | Malfunction |  | It was reported that during a da vinci s cardiac procedure, a plastic piece of the permanent cautery spatula instrument fell in the patient. No additional information was received despite attempts to contact the site. No patient harm, adverse outcome or injury was reported. | Conclusions - the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659432 | 1577665 | 1501950 | 2955842200900435 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-12-03 | 2009-12-31 | No | Yes | Health Professional | 42018404 | No | No | 2009-12-03 | Device Not Retu | No | Unkown | 2009
42018403 |  | Malfunction |  | It was reported that during a da vinci s procedure, the tip of the permanent cautery spatula instrument broke off, however, nothing fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the yaw pulley dislodged from the distal clevis and one broken yaw cable. Charred biodebris was noticed around the base of the distal clevis as well as on the yaw pulley, ceramic sleeve, and spatula. The ceramic sleeve is also cracked and missing parts. As previously mentioned, the initial reporter indicated no instrument pieces fell into the patient. Conclusions - engineering also observed charring and localized melting on the yaw pulley, indicating an arcing event occurred. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659423 | 1577666 | 1501951 | 2955842200900433 | 1 | NAY | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2009-12-07 | 2009-12-31 | No | Yes | Health Professional | 42018403 | Yes | No | 2009-12-07 | Yes | No | Reuse | 2009
42011005 | 2009-12-08 | Malfunction |  | It was reported that approximately one hour and 15 minutes into a da vinci si vaginal hysteroscopy (lavh) procedure, the precise bipolar forceps instrument tip was "fraying" and then broke. Several small pieces of the black shaft fell into the patient and were removed. The surgical staff stated that the instrument was difficult to remove from the cannula accessory and they had to cut the tip off to get the instrument out. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the cables at the distal end have been cut and the instrument wrist has been detached at the proximal clevis. Approximately half of the instrument's main tube interface wall is broken off and missing pieces. Engineering concluded that the proximal clevis likely dislodged from the main tube and a section of the distal end of the tube collapsed, requiring the instrument wrist to be cut off to remove the instrument from the cannula. No other damage was found. | PRECISE BIPOLAR FORCEPS | ENDOSCOPIC INSTRUMENT | 1738331 | 1577667 | 1501952 | 2955842200900430 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-08 | 2009-12-31 | No | Yes | Health Professional | 42011005 | Yes | Yes | 2009-12-08 | Yes | No | Reuse | 2009
42020705 | 2009-12-09 | Malfunction |  | It was reported that approximately 15 minutes into a da vinci si vaginal hysteroscopy (lavh) procedure, the tenaculum forceps instrument broke and pieces fell into the patient. The instrument fragments were retrieved. The surgeon tried to straighten the instrument wrist to remove it from the cannulae, however, it remained at a right angle. The surgical staff had to cut the tip of the instrument off in order to remove from the cannula accessory. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the cables at the distal end have been cut and the instrument wrist has been detached at the proximal clevis. Approximately half of the instrument's main tube interface wall is broken off and missing pieces. Engineering concluded that the proximal clevis likely dislodged from the main tube and a section of the distal end of the tube collapsed, requiring the instrument wrist to be cut off to remove the instrument from the cannula. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1754466 | 1577669 | 1501954 | 2955842200900431 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-12-09 | 2009-12-31 | No | Yes | Health Professional | 42020705 | Yes | Yes | 2009-12-09 | Yes | No | Initial | 2009
42018404 |  | Malfunction |  | It was reported that during a da vinci s procedure, the tip of the permanent cautery spatula instrument broke off, however, nothing fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument broken at the proximal clevis to main tube interface. The proximal clevis is dislodged from the main tube as a result of the damage. The proximal clevis and the main tube are missing pieces. Additionally, all cables are broken at the distal end of the hypotubes, as well as the conductor wire. The main tube was removed from the back end, exposing the hypotubes, flush tube and conductor wire tube. As previously mentioned, the initial reporter indicated no instrument pieces fell into the patient. Conclusions - engineering also observed charring and localized melting on one distal clevis ear, indicating an arcing event occurred. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659443 | 1577670 | 1501955 | 2955842200900434 | 1 | NAY | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2009-12-07 | 2009-12-31 | No | Yes | Health Professional | 42018404 | Yes | No | 2009-12-07 | Yes | No | Initial | 2009
42022702 |  | Malfunction |  | It was reported that during a da vinci s procedure, it was observed that the jaws of the pk dissecting forceps instrument were misaligned. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of the grips. There is an offset of approximately. 174" at the tips. Conclusions - engineering also observed charring at the conductor wire exit of the yaw pulley, indicating an arcing event occurred. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1659452 | 1577672 | 1501957 | 2955842200900427 | 1 | NAY | Manufacturer | Company RepresentativeOther | NOT APPLICABLE | Initial | 2009-12-02 | 2009-12-31 | No | Yes | Health Professional | 42022702 | Yes | No | 2009-12-02 | Yes | No | Reuse | 2009
40018302 | 2009-11-18 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, a piece of the permanent cautery hook instrument, near the hook, broke and fell into the pt. The broken piece was retrieved, and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval; therefore, the root cause of the customer reported failure mode cannot be determined. A f/u mdr will be submitted if the instrument is returned (post-evaluation) and/or if add'l info is received. Per the info provided by the initial reporter, the broken instrument component was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643857 | 1570641 | 1495284 | 2955842200900423 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-11-25 | 2009-12-24 | No | Yes | Health Professional | 40018302 | No | No | 2009-11-25 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-10-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the monopolar curved scissors (mcs) instrument was found to be splintering where the tip meets the shaft. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Additionally, the site confirmed that nothing fell into the patient during the surgical procedure. The customer reported malfunction does not itself constitute a mdr reportable incident, however, during internal evaluation of the instrument, engineering discovered the tip cover accessory to be installed on the mcs instrument. The tip cover exhibited evidence of an arcing event. Although no patient harm was reported, we assessed the evaluation results and conservatively decided to report our findings. | Conclusions - the mcs tip cover accessory was returned and evaluated. Engineering found that there is a hole burned through the silicone of the tip cover accessory. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The size of the hole is approximately. 014" in diameter and is located. 036" from the silicone to sleeve interface. In addition, the tip cover has multiple mechanical tears directly above the hole as well as at the opposite side of the hole, near the sleeve interface and around the midpoint. The endowrist instruments for use (ifu) specifically states: general precautions and warnings: do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. A follow up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1656563 | 1570642 | 1495285 | 2955842200900409 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-30 | 2009-12-24 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-11-30 | Yes | Yes | Initial | 2009
42011704 | 2009-11-30 | Malfunction |  | It was reported that approximately one hour into a da vinci si gastric bypass procedure, while the surgeon was suturing, with the 5mm needle driver instrument, the tip broke and fell into the patient. The broken piece was immediately retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported problem and found the instrument tip to be broken. The detached grip was returned with the instrument. The grip fractured at the transition between the grip arm and the grip hub. The cracking appears to have initiated where there is a 90 degree angle between the grip arm and hub. No other damage was found. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1648991 | 1570643 | 1495286 | 2955842200900425 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-12-03 | 2009-12-24 | No | Yes | Health Professional | 42011704 | Yes | Yes | 2009-12-03 | Yes | No | Initial | 2009
A43P9 IS1200 | 2009-11-26 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the site was unable to focus the image through the high resolution stereo viewer (hrsv). Trouble shooting with the assistance of an isi field service engineer (fse) revealed that the camera head was defective. The patient was under anesthesia when the surgical staff made the decision to abort and reschedule the planned procedure. There was no patient harm, adverse outcome or injury reported. | The investigation conducted by the isi field service engineer concluded that the focusing issue experienced by the customer was associated with the system camera head. With the assistance of an fse, the site replaced the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. The site performed, and successfully completed the procedure the same day. The camera head was evaluated by the original equipment manufacturer (oem). The oem found that the left cable loom and the stage are defective, thus preventing functionality of the camera head focus feature. As of (b)(4) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1649480 | 1570680 | 1495328 | 2955842200900426 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2009-11-26 | 2009-12-24 | Yes | Yes | Health Professional | A43P9 IS1200 | Yes | No | 2009-11-26 | Yes | No | Reuse | 2009
40018009 | 2009-11-05 | Other | Required Intervention | It was reported that during a da vinci hysterectomy procedure, the patient sustained a 1-2 mm superficial burn to the bowel due to a hole in the mcs tip cover accessory. As a precaution, the affected area was over-sewn with a 3-0 silk. The planned surgical procedure was completed. On (b)(6) 2009, (b)(6), the site's materials coordinator, reported that there were no patient complications and the patient was discharged from the hospital after 3 days. In addition, the patient has not returned to the hospital due to any post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the mcs tip cover accessory was returned with a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicative that arcing occurred. The hole is oblong, approximately. 020" x. 010", and is located. 280" above the silicone to sleeve interface. The tip cover also has a mechanical tear in the silicone, located 45 degrees radially from the hole. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments, or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1649537 | 1570681 | 1495329 | 2955842200900422 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-12-01 | 2009-12-24 | Yes | Yes | Health Professional | 40018009 | Yes | No | 2009-12-01 | Yes | Yes | Initial | 2009
400180 | 2009-10-02 | Other | Other | During a site visit the week of (b)(6), 2009, by an (b)(4) clinical sales rep (csr), the site reported that a tip cover arcing event had occurred. The csr performed a pin gage test on the site's cannulae and found one had a sharp burr on the inner tube. It was believed that the cannula caused a tear in the tip cover upon insertion of the monopolar curved scissor (mcs) instrument through the cannula and into the pt's body cavity. On (b)(6) 2009, isi rec'd the following user facility medwatch report # (b)(4) for this event that occurred on (b)(6) 2009. Event desc: during procedure, davinci monopolar curved scissors failed causing a burn to the small bowel. Instrument was also difficult to remove. Instrument was replaced without further incident or problem. Surgeon over sewed burn area as a precaution. | The tip cover accessory has not been returned for eval; however, the cannula accessory believed to have caused the tear in the tip cover was disposed off at the site. As of (b)(6), 2009, there have been no reported recurrences of the issues at this hosp. If add'l info is rec'd, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1879805 | 1785136 | 1692292 | 2955842200900424 | 1 | NAY | Manufacturer | Company RepresentativeOther | PHYSICIAN ASSISTANT | Initial | 2009-11-13 | 2009-12-24 | Yes | Yes | Health Professional | 400180 | No | No | 2009-11-13 | Device Not Retu | Yes | Unkown | 2009
370090625329 | 2009-11-05 | Malfunction |  | Instrument tip of the davinci was observed on the screen. A small piece, approximately 8 mm x 10mm, broke off from the tip. The piece was seen briefly but then lost in the abdominal cavity. An attempt was made to locate it but the piece was not found. The device was in use for approximately 2 hours before the event occurred. |  | ENDOWRIST | HOOK PERMANENT CAUTERY | 1610017 | 1566432 | 1491241 | 1566432 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-12-18 | 2009-12-18 | No | Yes | Physician | 370090625329 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
|  | Injury | Other | It was reported by the pt that she recently had a da vinci hysterectomy procedure and was under anesthesia for 8 hours instead of the expected 4 hours due to her being severely overweight. She stated that she woke up and was unable to move either of her arms, however, after time, she was able to get full use of her left arm. She stated her right arm is still in a lot of pain. | Due to the pt not providing the site name of where her procedure was performed, we are unable to confirm whether or not malfunctions of the da vinci surgical system were reported to have occurred during the procedure. Based on the info provided, it was determined that the da vinci surgical system likely did not malfunction in a way that would cause nor contribute to the pt injury, however, the pt's positioning and the extended procedure duration may have contributed to the event. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1614952 | 1570042 | 1494691 | 2955842200900410 | 1 | NAY | Manufacturer | Consumer | NOT APPLICABLE | Initial | 2009-11-19 | 2009-12-18 | Yes | No | Health Professional |  | No | No | 2009-11-19 | No | No | Reuse | 2009
1S2000 A51P8 | 2006-12-08 | Death | Death | It was reported several weeks after a da vinci s prostatectomy surgical procedure, the pt expired. No additional info was provided. | In 2009, an isi clinical sales rep contacted the proctor to obtain add'l info, however, the proctor indicated that at this time he is unable to provide any additional details related to the reported event. If additional info is received, a follow-up mdr will be submitted. Based on the limited info provided by the customer, it is indeterminable if the da vinci s surgical system contributed to the pt's demise. As of the following month, no adverse events have been experienced with the site's system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1614260 | 1570088 | 1494736 | 2955842200900411 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-11-18 | 2009-12-18 | Yes | No | Health Professional | 1S2000 A51P8 | No | Yes | 2009-11-18 | Device Not Retu | No | Reuse | 2009
42017909 | 2009-11-17 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure, the monopolar curved scissors instrument broke and a fragment fell into the pt. The broken piece was retrieved, and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. Per the info provided by the initial reporter, the broken piece was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1763445 | 1570588 | 1495233 | 2955842200900415 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-11-19 | 2009-12-18 | No | Yes | Health Professional | 42017909 | No | Yes | 2009-11-19 | Device Not Retu | No | Unkown | 2009
42017909 | 2009-11-18 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was cutting tissue, a blade from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved, and the planned surgical procedure was completed after a delay of 15 minutes and with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is fractured and the broken piece was not returned with the instrument. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is nearly straight. The intact blade does not exhibit any damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1763435 | 1570589 | 1495234 | 2955842200900419 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-19 | 2009-12-18 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2009-11-19 | Yes | No | Reuse | 2009
42020906 | 2009-11-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was suturing using the mega suturecut needle driver instrument, the surgical staff noticed a wire was broken on the instrument and pieces had fallen into the pt. The instrument pieces were removed from the pt and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one grip close to be broken near the proximal pulleys and the proximal clevis cable hole, causing the pulleys to spin freely. The proximal clevis hole exhibits wear on one edge. Other cables at the instrument wrist are not damaged. Per the info provided by the initial reporter, the known instrument fragments were successfully retrieved from the pt. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1643794 | 1570590 | 1495235 | 2955842200900413 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-20 | 2009-12-18 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | Yes | 2009-11-10 | Yes | No | Reuse | 2009
42020504 | 2009-11-23 | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the fenestrated bipolar forceps instrument broke off and fell into the pt. The broken piece was retrieved, and the planned surgical procedure was successfully completed. There was no pt harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint the yaw pulley is missing a. 159" x. 060" piece opposite to the conductor break. The missing piece allows the grip to move within the pulley and have a larger range of motion in yaw. One conductor wire is broken at the yaw pulley exit. The instrument does no exhibit any signs of arcing. The instrument failed electrical continuity testing. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1763515 | 1570591 | 1495236 | 2955842200900418 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-23 | 2009-12-18 | No | Yes | Health Professional | 42020504 | Yes | No | 2009-11-23 | Yes | No | Reuse | 2009
42017909 | 2009-11-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a blade from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is fractured and the broken piece was returned with the instrument. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is straight. The intact blade does not exhibit any damage at the tip but does present indentations around the midpoint and near the base. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1643804 | 1570644 | 1495287 | 2955842200900420 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-24 | 2009-12-18 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2009-11-24 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-11-19 | Other |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff experienced difficulty moving the endoscopic camera manipulator (ecm) and observed the right master tool manipulator (mtmr) drifting when it was released. The drift caused the associated patient side manipulator (psm) to move and resulted in a puncture to the patient's uterus from the instrument installed on the psm. The patient's uterus was scheduled to be removed as part of this procedure. The site undocked the system from the patient and restarted the system, however, the mtmr was reported as still drifting. The surgeon made the decision to convert to traditional laparoscopic surgical techniques to complete the planned procedure. | The investigation conducted by field service engineering (fse) consisted of reviewing system error logs and performing functional tests. The fse's review of the logs did not reveal any system errors and the system functioned as intended with no issues. Since this event, an isi clinical sales representative provided additional training to the console surgeon on december 10, 2009, and again on december 17, 2009. During training, it was determined that the event was caused by the surgeon removing his hands from the mtm's without removing his head from the surgeon console's head sensor. The system's psms are designed to follow the movements of the mtms during normal use, however, once the surgeon's head is removed from the head sensor, the psms are no longer in the following mode, resulting in no movement of the psms. As of december 16, 2009, the hospital has continued to use the system with no reported recurrences of the reported issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1649555 | 1570678 | 1495326 | 2955842200900412 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-11-19 | 2009-12-18 | Yes | No | Health Professional | IS2000 A51P8 | Yes | Yes | 2009-11-19 | Yes | No | Reuse | 2009
42020504 | 2009-11-20 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the "plastic piece" of the fenestrated bipolar forceps instrument broke off and fell into the pt. The piece was retrieved and the planned surgical procedure was successfully completed. There was no pt harm, adverse outcome or injury reported. | On (b)(6) 2009, (b)(4) the site's biomedical engineering mgr reported that during the surgical procedure, an endowrist instrument of an unk model was caught between the jaws of the fenestrated bipolar forceps instrument, causing the "plastic part" of the fenestrated bipolar forceps instrument to break off inside of the pt. While the initial report was that a piece broke off from the instrument, a post operative inspection of the fenestrated bipolar forceps instrument by the site revealed that it is possible that the instrument fragmented. The pt's abdomen was irrigated and there were no pieces of the instrument found. A copy of the site's voluntary report (b)(4) received by isi on (b)(6) 2009 indicates that the hospital believes that the surgeon retrieved all of the broken pieces. In addition, to the best of (b)(4)'s knowledge, the pt has not returned to the hospital due to any post operative complications related to the event. Conclusions: the fenestrated bipolar forceps instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. (b)(4). | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1754564 | 1576667 | 1478104 | 2955842200900417 | 1 | NAY | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2009-11-20 | 2009-12-18 | No | Yes | Health Professional | 42020504 | No | No | 2009-11-20 | Device Not Retu | No | Unkown | 2009
42022702 | 2009-11-17 | Malfunction |  | On 11/20/2009, isi was advised of the following event and on (b) (6)2009, user facility medwatch# (b) (4) was received with the following info: "event desc: the pt was in for laparoscopic myomectomy with robotic assistance and exploratory laparotomy, supracervical hysterectomy. The instruments for the laparotomy had not been counted prior to surgery because, the surgery was a laparoscopy converted to an open case. For this reason, a flat plate x-ray was taken of the pt. There was a 3-4 mm radiopaque structure in the left abdomen. It was not consistent with an instrument. All sponges and needles were then recounted twice, and the counts appeared to be correct. The pt was then placed on a clean table and another x-ray was taken and the substance persisted. Since it was not an instrument, sponge, or needle, it was decided not to pursue with re-exploration. The pt was then taken to the recovery room in stable condition. When the robotic instrument was going to set on another case a few days later, it was discovered that it was broken and the staff realized that this may have been the source of the retained foreign body. ". | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If add'l info is received, a f/u mdr will be submitted. | PK DISSECTING INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1754540 | 1576670 | 1484527 | 2955842200900416 | 1 | NAY | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2009-11-20 | 2009-12-18 | No | Yes | Health Professional | 42022702 | No | No | 2009-12-20 | Device Not Retu | No | Unkown | 2009
DA VINCI | 2009-11-29 | Malfunction |  | Equipment (robot) had minor visual distortions. Procedure was delayed because the circuit board was swapped out. Delay was approximately 15- 20 minutes of extra anesthesia time. No patient injury noted. |  | DA VINCI SURGICAL SYSTEM | SURGICAL SYSTEM ROBOTIC | 1607910 | 1564340 | 1489229 | 1564340 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-12-12 | 2009-12-12 | No | Yes | Physician | DA VINCI | Yes | No Answer Provided |  |  |  |  | 2009
42018304 | 2009-11-05 | Malfunction |  | It was reported that during a da vinci s thyroidectomy procedure, the site observed a foreign object in the surgical field. The piece was removed and the console surgeon continued with the procedure. The piece was determined to be a portion of a larger piece that was missing from the permanent cautery hook instrument. The site attempted to locate the other half of the missing piece using x-ray and sonogram testing. The attempt to locate the missing piece was unsuccessful and the surgeon made the decision to complete the schedule procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a piece of the conductor cap detached from the instrument. The cap is broken diagonally and roughly half of the entire cap is missing. Both sides of the yaw pulley have the extruded boss features that mate with slots on the inner surface of the distal clevis ears broken off, allowing yaw pulley to have extra lateral play within distal clevis. Per the info provided by the initial reporter, half of the broken conductor cap was retrieved, however, the site was unsuccessful in locating the other half. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1640854 | 1566764 | 1491571 | 2955842200900401 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-17 | 2009-12-11 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2009-11-17 | Yes | No | Reuse | 2009
IS1200 A4 3P9 | 2009-11-18 | Malfunction |  | It was reported that approx 30 minutes into a da vinci surgical procedure, the vision from one eye was unclear and the surgical staff was unable to correct the problem. The endoscope was removed, and a lens was observed missing from the endoscope tip. The missing lens could not be found and the site was unable to determine if the lens was lost in the pt or in the sterile field. No pt harm, adverse outcome or injury was reported. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment mfr (oem) for eval. The oem confirmed that a lens is missing and found the glue to be washed out at all windows and lens causing fluid invasion and several optical parts damaged by fluid. The oem concluded that the likely cause of the missing glue was due to unsuitable cleaning and operating room sterilization methods. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1639666 | 1566765 | 1491572 | 2955842200900408 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-18 | 2009-12-11 | No | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2009-11-18 | Yes | No | Reuse | 2009
42000605 | 2009-11-16 | Malfunction |  | It was reported that over two hours into and near the end of a da vinci si prostatectomy procedure, a piece of the large needle driver instrument shaft fell into the pt's abdomen. The fragment was removed and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of the main tube presents deep scratches and tube abrasions within an area extending approx 4. 85 inches from the interface with the proximal clevis. Material has been removal from the main tube as a result. The abrasions are parallel to the tube axis and have a rough surface finish. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2574864 | 2543844 | 2440576 | 2955842200900436 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-16 | 2009-12-11 | No | Yes | Health Professional | 42000605 | Yes | No | 2009-11-16 | Yes | No | Reuse | 2009
42018304 | 2009-11-05 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, a small piece from the tip of the permanent cautery hook instrument broke off and fell into the pt. The broken piece was seen briefly, however, an attempt to locate and remove the fragment was unsuccessful. No pt complications, harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one ear of the distal clevis to be broken off and missing, measuring approx. 240 inches by. 269 inches. The conductor wire cap is missing and both of the yaw pulley's boss features are broken and missing, measuring approx. 030 inches by. 055 inches and 040 inches by. 055 inches respectively. Electrical continuity passed. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1608629 | 1563091 | 1488059 | 2955842200900405 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-10 | 2009-12-10 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-11-10 | Yes | No | Reuse | 2009
| 2009-11-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure, fragments of a da vinci surgical endoscopic instrument were observed falling into the pt. The pt's affected area was irrigated and known pieces were removed. Upon removal of the instrument, the site observed the shaft of the instrument to appear as if it were scratched. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Despite attempts to contact the site, no additional info was received. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional info is received. Per the info provided by the initial reporter, the known instrument fragments were successfully retrieved from the pt. |  | ENDOSCOPIC INSTRUMENT | 1608617 | 1563092 | 1488060 | 2955842200900407 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-11-10 | 2009-12-10 | No | Yes | Health Professional |  | No | No | 2009-11-10 | Device Not Retu | No | Unkown | 2009
420179 | 2009-10-02 | Malfunction |  | Event Description: --/-- During procedure davinci monopolar curved scissors failed causing a burn to the small bowel. Instrument was also difficult to remove. Instrument was replaced without further incident or problem. Surgeon over sewed burn area as a precaution. --/-- Additional Event Description: --/-- During procedure davinci monopolar curved scissors failed causing a burn to the small bowel. Instrument was also difficult to remove. Instrument was replaced without further incident or problem. Surgeon over sewed burn area as a precaution. | Manufacturer Narrative: --/-- | HOT SHEARS |  | 1599889 | 1557344 | 1482671 | 1557344 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-12-04 | 2009-12-04 | No | Yes | Physician | 420179 | Yes | No Answer Provided |  |  |  |  | 2009
IS2000 A51P8 | 2009-11-10 | Other | Other | It was reported that during a da vinci s mitral valve repair surgical procedure, the surgeon experienced a pink and distorted image on the left side view of the surgeon's console during movement of the camera cable. The site swapped the camera cable ends in an attempt to resolve the issue, however, the issue moved to the right side image. Prior to contacting an isi technical support engineer for assistance the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable and the vision quality was verified. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of (b)(4), 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1633254 | 1558275 | 1483622 | 2955842200900406 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-11-10 | 2009-12-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-11-10 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-11-04 | Other | Other | It was reported that during start-up and prior to docking the da vinci s surgical system to the patient for a mitral valve repair procedure, the system started faulting and the led on the third arm turned red. An isi technical support engineer (tse) reviewed the error logs and found error codes #25515 & #25503. The tse had the surgical staff power cycle and emergency power-off the system, however, the system began faulting again during homing. The patient had been under anesthesia for 1 hour and 50 minutes when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes #25515 and #25503 found in the error logs were associated with the remote arm controller (rac) and/or the embedded serializer for setup joints (essj). The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The essj gathers position information and sends this information as serialized data to the rac. The system was repaired by replacing the affected rac and essj. The system alarm (system generated fault code) functioned as designed, and there was no injury to the patient. System error code #25515 indicates a communication problem between the rac and essj. System error code #25503 is a sympathetic error which occurs when a communication fault associated with the rac is detected. Upon determining this condition, the safety system puts da vinci s in a "non-recoverable safe state". As of december 3, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1633237 | 1558277 | 1483624 | 2955842200900393 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-04 | 2009-12-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-11-04 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-11-05 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the right master tool manipulator (mtmr) gimbal came apart due to a loose screw. The surgical procedure had been in progress for 30 minutes when the surgeon decided to abort, close the patient and reschedule the procedure for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluding that the issue experienced by the customer was associated with a loose set screw located on the right master tool manipulator (mtmr) gimbal. The mtmr gimbal is a subsection of the complete mtm arm, consisting of links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the right mtm gimbal. The mtmr gimbal was returned to isi for evaluation. Engineering determined that a misalignment in the assembly of the mtmr gimbal caused the set screw to become loose. As of (b)(4), 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1633226 | 1558280 | 1483627 | 2955842200900396 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-05 | 2009-12-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-11-05 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-11-06 | Other | Other | It was reported that during a da vinci s adrenalectomy procedure, the site experienced system message "cannot connect to the psc". The sites biomedical engineer reported that the cable that connects the surgeon side cart (ssc) to the patient side cart (psc) cable was damaged and that it had been pulled hard approximately two inches from the connector resulting in the cable breaking. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the issue experienced by the customer was associated with the red system cable. The red system cable connects the surgeon console and the patient cart and passes video, audio, and data during system operations. The system was repaired by replacing the affected cable. Failure analysis confirmed that the cable had been physically damaged and was scrapped. As of (b)(4), 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1633213 | 1558282 | 1483629 | 2955842200900400 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-06 | 2009-12-04 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-11-06 | Yes | No | Reuse | 2009
| 2009-11-17 | Malfunction |  | Event Description: --/-- The patient was in for laparoscopic myomectomy with robotic assistance and exploratory laparotomy, supracervical hysterectomy. The instruments for the laparotomy had not been counted prior to surgery because the surgery was a laparoscopy converted to an open case. For this reason, a flat plate x-ray was taken of the patient. There was a 3-4 mm radiopaque structure in the left abdomen. It was not consistent with an instrument. All sponges and needles were then recounted twice, and the counts appeared to be correct. The patient was then placed on a clean table and another x-ray was taken and the substance persisted. Since it was not an instrument, sponge, or needle, it was decided not to pursue with re-exploration. The patient was then taken to the recovery room in stable condition. When the robotic instrument was going to be used on another case a few days later it was discovered that it was broken and the staff realized that this may have been the source of the retained foreign body. --/-- Additional Event Description: --/-- The patient was in for laparoscopic myomectomy with robotic assistance and exploratory laparotomy, supracervical hysterectomy. The instruments for the laparotomy had not been counted prior to surgery because the surgery was a laparoscopy converted to an open case. For this reason, a flat plate x-ray was taken of the patient. There was a 3-4 mm radiopaque structure in the left abdomen. It was not consistent with an instrument. All sponges and needles were then recounted twice, and the counts appeared to be correct. The patient was then placed on a clean table and another x-ray was taken and the substance persisted. Since it was not an instrument, sponge, or needle, it was decided not to pursue with re-exploration. The patient was then taken to the recovery room in stable condition. When the robotic instrument was going to be used on another case a few days later it was discovered that it was broken and the staff realized that this may have been the source of the retained foreign body. | Manufacturer Narrative: --/-- | ENDOWRIST |  | 1601829 | 1559178 | 1484527 | 1559178 | 1 | NAY | User Facility |  |  | Initial | 2009-12-04 | 2009-12-04 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2009
R ARM 25832 | 2009-12-02 | Malfunction |  | During a davinci robotic assisted laparoscopic prostatectomy, the robot started alarming and an alert came on saying there was a recoverable error. Attempted to override the fault, unsuccessful, troubleshooting process/ undock done per rep instructions, also robot shut down and attempted to restart. The robot restarted, but the fault would not allow it to be homed, which made it impossible to continue the case robotically. Surgeon had to do an open prostatectomy. |  | DAVINCI | SYSTEM SURGICAL | 1609853 | 1566300 | 1491109 | 1566300 | 1 | NAY | User Facility |  |  | Initial | 2009-12-04 | 2009-12-04 | No | Yes | Physician | R ARM 25832 | Yes | No Answer Provided |  |  |  |  | 2009
40012608 | 2009-11-04 | Malfunction |  | It was reported that during a da vinci surgical procedure, the "wire" on the permanent cautery hook instrument broke. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the drive cables and conductor wire are not damaged at the wrist. The conductor wire cap is missing, thus causing the wire to stick out of the distal clevis. The instrument passed electrical continuity testing. Engineering also observed that the instrument presents arcing at the cable crimp. The yaw pulley has charring near the cable crimp and has a piece broken off. The hook insulation has material removal just above the crimp. No other damage was found. There are no indications from the site that the missing instrument components fell inside of a pt during a surgical procedure. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1762283 | 1566543 | 1491351 | 2955842200900404 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-10 | 2009-12-04 | No | Yes | Health Professional | 40012608 | Yes | No | 2009-11-10 | Yes | No | Reuse | 2009
42017908 | 2009-11-05 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, a piece of the blade from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) to be missing. The blade fractured at the cross section of the cable crimp, however, the fracture plane is straight. The intact blade does not exhibit damage at the tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1762258 | 1566545 | 1491353 | 2955842200900395 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-05 | 2009-12-04 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-11-05 | Yes | No | Reuse | 2009
40016906 | 2009-11-09 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the harmonic curved shears insert fell into the pt. The insert was retrieved and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, the insert was returned with the curved blade detached from the distal end. The blade is approx. 750" long with the fracture surface of the blade nearly straight. Additionally, engineering observed that the clamp arm was detached from the insert. One side of the clamp arm also attachment feature located near the distal end is slightly bent. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1762272 | 1566546 | 1491354 | 2955842200900403 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2009-11-09 | 2009-12-04 | No | Yes | Health Professional | 40016906 | Yes | Yes | 2009-11-09 | Yes | Yes | Unkown | 2009
| 2009-11-24 | Malfunction |  | Event Description: --/-- Metal 8mm long davinci trocar broke where shaft meets the top of trocar while attached to patient and robot. --/-- Additional Event Description: --/-- Metal 8mm long davinci trocar broke where shaft meets the top of trocar while attached to patient and robot. | Manufacturer Narrative: --/-- | DAVINCI |  | 1594585 | 1551799 | 1477654 | 1551799 | 1 | NAY | User Facility |  |  | Initial | 2009-12-01 | 2009-12-01 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2009
| 2009-11-20 | Malfunction |  | Event Description: --/-- Two arms of the davinci robot device collided in patient. One arm of the grasper device bent back causing a gear to break inside the patient. Surgeon is believed to have retrieved all pieces, but pieces are not radiopaque. --/-- Additional Event Description: --/-- Two arms of the davinci robot device collided in patient. One arm of the grasper device bent back causing a gear to break inside the patient. Surgeon is believed to have retrieved all pieces, but pieces are not radiopaque. | Manufacturer Narrative: --/-- | DAVINCI |  | 1595572 | 1552259 | 1478104 | 1552259 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-11-30 | 2009-11-30 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2009
IS2000 A51P8 | 2009-11-03 | Death | Death | It was reported that approximately six hours into a da vinci s mitral valve repair procedure, the surgeon decided to complete the procedure via a thoracotomy. It was reported that the pt expired shortly after the procedure. Reportedly, the pt's demise was unrelated to the da vinci s surgical system. | The hospital was contacted to obtain add'l info, however, despite several attempts, the site has decided to not provide any info related to this event. No malfunction of the da vinci s surgical system were reported to have occurred during the procedure, and the hospital has continued to used the system to perform surgery on patients. As of 2009, no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1598781 | 1553131 | 1478857 | 2955842200900338 | 1 | NAY | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2009-11-05 | 2009-11-25 | Yes | No | Health Professional | IS2000 A51P8 | No |  | 2009-11-05 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-10-29 | Other | Other | It was reported that approximately 1. 5 hrs into a da vinci s hysterectomy procedure the surgeon experienced double vision and left eye video loss. Prior to contacting an isi field service engineer to assist with troubleshooting the issue, the circulating nurse restarted the power to the vision side cart (vsc). Unaware that the circulating nurse had restarted power to the vsc, the surgeon decided to convert the procedure to traditional open surgical techniques because he thought that the entire vision system had failed. No pt harm was reported. | Investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera system to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The camera was also replaced as a precaution. The camera produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera head was returned to isi for failure analysis. Engineering eval found that the camera head stage is damaged. Engineering concluded that the camera head had likely been dropped or slammed as it exhibits marks on the housing. As of (b)(6), 2009, there have been no reported recurrences of the issue at this hosp. | AIDA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1757334 | 1554084 | 1479752 | 2955842200900386 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-29 | 2009-11-25 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-10-29 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-10-29 | Other | OtherRequired InterventionHospitalization | It was reported that while beginning a da vinci s abdominal penincal dissection procedure, the image was not aligned appropriately through the high resolution stereo viewer (hrsv). Isi's clinical sales rep (csr) present during the procedure advised an isi tech support engineer that the site will perform an emergency power off and call back for more troubleshooting assistance. Approximately 30 minutes later the csr called to report that during retraction of the pt's bowel, the pt's accessory vessel was cut or torn. The procedure was converted to traditional open surgical techniques to repair the affected vessel. The pt also experienced excessive bleeding, requiring the pt to undergo a blood transfusion. A monopolar curved scissors instrument, a fenestrated bipolar instrument and traditional laparoscopic instruments were used during the procedure, however, it is indeterminable as to if these instruments may have caused the damage to the pt's vessel. | Isi's clinical sales rep followed up with the surgeon to verify the pt's status. The surgeon reported that the damage to the pt's vessel extended the pt's hospitalization, however, the pt was discharged from the hosp after 5 days and has not returned to the hosp due to any post operative complications. The camera head was returned to isi for failure analysis. Engineering eval found that the camera head stage was loose. Engineering concluded that the stage was loose due to physical impact. As of (b)(6), 2009, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1757308 | 1554085 | 1479753 | 2955842200900385 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-29 | 2009-11-25 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2009-10-29 | Yes | No | Unkown | 2009
40018009 | 2009-10-12 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, after removal of the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed, an inspection of the mcs tip cover accessory by the surgical staff revealed that it had a hole. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the mcs tip cover to have various mechanical tears and holes burned through the silicone. The silicone exhibits localized melting around the holes. The hole sizes vary from. 008" to. 015" and are located from. 210" to. 365" above the silicone to sleeve interface. Engineering concluded that multiple holes in the silicone of the tip cover indicates that multiple arcing events occurred as two of the holes are in line with each other and are directly adjacent to each other. In addition, one of three holes exhibits a tear on one edge. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1638360 | 1559189 | 1484538 | 2955842200900391 | 1 | NAY | Manufacturer | User facilityOther | Nurse | Initial | 2009-11-02 | 2009-11-25 | No | Yes | Health Professional | 40018009 | Yes | Yes | 2009-11-02 | Yes | Yes | Unkown | 2009
42017908 | 2009-11-05 | Malfunction |  | It was reported that approximately 30 minutes into a da vinci s hysterectomy procedure, the tip of the monopolar curved scissors instrument broke off and fell into the patient's abdomen. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) was detached from the distal end of the blade. The blade fractured at the cross section of the cable crimp with the fracture plane nearly straight. The intact blade does not exhibit damage at tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1638253 | 1559190 | 1484539 | 2955842200900397 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-05 | 2009-11-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-11-05 | Yes | No | Reuse | 2009
42011704 | 2009-11-06 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, after removing fibroids and while the surgeon was suturing the uterus, half of the tip of the 5mm needle driver instrument broke and fell into the patient. The instrument was replaced with another of the same type to continue with the planned procedure, however, the replacement 5mm needle driver instrument tip broke and fell into the patient. Exploration and x-rays were performed to retrieve the instrument pieces, however, they could not be located. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch mfr report was sent to the fda regarding this event: 2955842-2009-00402. The 5mm needle driver instrument grip (referred to as the "tip" by the reporter) consists of (b)(4). This material has been tested for biocompatibility and was found to be inert and biocompatible for it's intended use. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1638313 | 1559191 | 1484540 | 2955842200900399 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-06 | 2009-11-25 | No | Yes | Health Professional | 42011704 | No | Yes | 2009-11-06 | Device Not Retu | No | Unkown | 2009
42018404 | 2009-10-16 | Malfunction |  | It was reported that during a da vinci s thoracoscopy procedure, the coating of the permanent cautery spatula instrument broke off and a piece fell into the patient. The broken piece was retrieved, and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the ceramic sleeve of the instrument to have a. 065" x. 065" piece broken off at the distal end. The sleeve also exhibits light scratches with no cracking. Other components at the wrist are not damaged. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1638262 | 1559192 | 1484541 | 2955842200900382 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-27 | 2009-11-25 | No | Yes | Health Professional | 42018404 | Yes | No | 2009-10-27 | Device Not Retu | No | Reuse | 2009
42011704 | 2009-11-06 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, after removing fibroids and while the surgeon was suturing the uterus, half of the tip of the 5mm needle driver instrument broke and fell into the patient. The instrument was replaced with another of the same type to continue with the planned procedure, however, the replacement 5mm needle driver instrument tip broke and fell into the patient. Exploration and x-rays were performed to retrieve the instrument pieces, however, they could not be located. The planned surgical procedure was successfully completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch mfr report was sent to the fda regarding this event: 2955842-2009-00399. The 5mm needle driver instrument grip (referred to as the "tip" by the reporter) consists of 17-4 ph stainless steel. This material has been tested for biocompatibility and was found to be inert and biocompatible for it's intended use. | 5MM NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1638326 | 1559193 | 1484540 | 2955842200900402 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-11-06 | 2009-11-25 | No | Yes | Health Professional | 42011704 | No | Yes | 2009-11-06 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-10-29 | Malfunction |  | It was reported that during a da vinci s cystectomy procedure, while dropping the patient's bladder, the "side wall" of the permanent cautery hook instrument broke apart and fell into the patient. The broken pieces were retrieved and an inspection of the patient's pelvis area was performed to ensure that there were no remaining pieces left inside of the patient. The planned surgical procedure was successfully completed with a replacement instrument of the same type. No patient complications, harm, adverse outcome or injury was reported. The site's robotic coordinator reported that the patient was discharged on (b)(6) 2009 and to the best of her knowledge, the patient has not returned to the hospital due to any post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one ear of the distal clevis and the conductor cap are detached. The conductor wire is exposed on yaw pulley as a result of the damage. The broken ear piece is approximately. 240" x. 285. " the yaw pulley boss feature that mates with slot feature on inner surface of ear piece is broken off. The proximal clevis exhibits a hairline crack halfway between it's ears that is located on the same side as the distal clevis breakage. Engineering has concluded that the damages to the instrument are most likely due to overloading of the tip while the wrist is pitched. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1638279 | 1559194 | 1473164 | 2955842200900384 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-10-29 | 2009-11-25 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2009-10-29 | Yes | No | Reuse | 2009
42017909 | 2009-10-27 | Malfunction |  | It was reported that approximately 1 hour into a da vinci prostatectomy procedure, the jaw of the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) to be missing. The blade fractured at the cross section of the cable crimp with the fracture plane nearly straight. The intact blade does not exhibit damage at tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1638176 | 1559195 | 1484544 | 2955842200900383 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-10-28 | 2009-11-25 | No | Yes | Health Professional | 42017909 | Yes | Yes | 2009-10-28 | Yes | No | Reuse | 2009
40018009 | 2009-10-12 | Malfunction | Other | It was reported that during a da vinci s prostatectomy procedure, after removal of the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed, an inspection of the mcs tip cover accessory by the surgical staff revealed that it had a hole. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that there is a hole burned through the silicone of the tip cover accessory. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The size of the hole is approximately. 015" in diameter and is located. 185" above the silicone to sleeve interface. In addition, the tip cover has a mechanical tear directly below the hole and a. 140" long tear at the tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1638385 | 1559196 | 1484545 | 2955842200900390 | 1 | GEI | Manufacturer | User facilityOther | Nurse | Initial | 2009-11-02 | 2009-11-25 | No | Yes | Health Professional | 40018009 | Yes | Yes | 2009-11-02 | Yes | Yes | Unkown | 2009
IS1200 A43P9 | 2009-10-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff placed the endoscope in the patient camera port and noticed that the image was blurry. The endoscope was removed and a lens was observed missing from the endoscope tip. The missing lens could not be found and the planned surgical procedure was completed. An mri was later performed, and the missing lens was not located. No patient harm, adverse outcome or injury was reported. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment manufacturer (oem) for evaluation. The oem found that the scope had been damaged from an impact of a stroke / hit at the proximal and distal end. As a result, the objective is damaged and the distal window is lost. Per the information provided by the initial reporter, the blurry image was noticed upon endoscope insertion, suggesting that the lens could have already been missing before the endoscope was placed in the patient. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1628101 | 1548614 | 1474643 | 2955842200900368 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2009-10-25 | 2009-11-24 | No | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-10-25 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-07-16 | Death | Death | It was reported that during a da vinci s renal procedure, the console surgeon inadvertently perforated one of the patient's arteries resulting in blood loss beyond the point of revival. Reportedly, the patient's demise was unrelated to the da vinci s surgical system. | The hospital was contacted to obtain additional information and advised that there were no malfunctions of the da vinci s surgical system which occurred during this procedure. The hospital has continued to use the system to perform surgery on patients. Additionally, the surgeon involved in this event is considered an experienced user who has performed many successful procedures prior to this event. Despite several attempts, the site has decided to not provide any additional information related to this event. As of 2009, no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1592080 | 1546455 | 1472529 | 2955842200900381 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-08-12 | 2009-11-20 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2009-08-12 | Device Not Retu | No | Unkown | 2009
400180 | 2009-10-22 | Other | Other | It was reported that approximately 3 hours into a da vinci s hysterectomy procedure, while the surgeon was performing a right and left lymphadenectomy, a "flash" was observed from the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. When the surgical staff removed the mcs instrument, a small hole was observed in the tip cover. The surgeon then discovered that there was a surface burn on the patient's bowel. Although the burn had not penetrated the bowel, the surgeon decided to suture the burned area. The planned surgical procedure was completed and no additional patient's harm, adverse outcome or injury was reported. | The mcs instrument and tip cover accessory have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Members of isi's clinical sales team followed up with the surgeon and robotics coordinator to discuss the following recommendations/suggestions for prevention of mcs tip cover damage or arcing: cannula inspection prior to use. Careful installation of tip cover, do not exceed maximum cautery settings. Avoid instrument collisions, straight wrist prior to removal. Limit use of cautery when not in contact with tissue, cases with extended cautery use. Surgical tasks around critical vessels and structures. A follow up mdr will be submitted if the instrument or accessory are returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1625779 | 1546721 | 1472791 | 2955842200900372 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-22 | 2009-11-20 | Yes | Yes | Health Professional | 400180 | No | No | 2009-10-22 | Device Not Retu | Yes | Unkown | 2009
IS2000 A51P8 | 2009-10-21 | Other | Other | It was reported that during a da vinci s surgical procedure, power issues were experienced with the pt cart. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineer (fse) concluded that the power issues experienced by the customer was associated with the red system cable. The red system cable connects the surgeon console and the pt cart and passes video, audio, and data during sys operations. The sys was repaired by replacing the affected cable. As of (b)(6), 2009, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1624216 | 1546722 | 1472792 | 2955842200900369 | 1 | NAY | Manufacturer | ForeignHealth Professional | Other Health Care Professional | Initial | 2009-10-21 | 2009-11-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2009-10-21 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-10-30 | Other | Other | It was reported that during a da vinci s surgical procedure, the surgical staff was unable to focus the camera from the foot pedal or from the focus controller. With the assistance of an isi technical support engineer, the staff reseated the cable to the focus controller and to the camera, reset power to the vision cart, and replaced the camera and camera cable, however, the issue remained. The staff then removed the back cover to the vision cart, reseated the 12v power and cable on the rear of the focus controller and then restarted the system, however, the issue could not be resolved. The patient was under anesthesia when the surgeon made the decision to abort the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the system focus controller. The focus controller is designed to adjust the endoscope lens focus to sharpen the surgical image. The system was repaired by replacing the affected focus controller. The focus controller was returned to isi for evaluation. Engineering discovered that a printed circuit assembly board within the focus controller had malfunctioned, thus causing the vision issue experienced by the customer. As of (b)(6) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1625789 | 1546723 | 1472793 | 2955842200900387 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-30 | 2009-11-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-10-30 | Yes | No | Reuse | 2009
420093 | 2009-10-23 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the distal end of the prograsp forceps instrument broke, resulting in a piece falling into the pt. The piece was successfully retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1622865 | 1547090 | 1473159 | 2955842200900376 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-10-23 | 2009-11-20 | No | Yes | Health Professional | 420093 | No | Yes | 2009-10-23 | Device Not Retu | No | Unkown | 2009
40018009 | 2009-10-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the physician was performing a colpotomy a "flash" was observed from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. When the surgical staff removed the mcs instrument, a hole was observed in the tip cover. The surgeon proceeded with the procedure using a replacement tip cover accessory; however, a "flash" from the replacement mcs tip cover accessory was observed. Examination of the replacement tip cover revealed a hole. The surgical procedure was completed with a third replacement tip cover accessory. No pt harm, adverse outcome or injury was reported. | The mcs instrument and tip cover accessory have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Members of isi's clinical sales team followed up with the surgeon and robotics coordinator to discuss the following recommendations/suggestions for prevention of mcs tip cover damage or arcing: cannula inspection prior to use. Careful installation of tip cover. Do not exceed maximum cautery settings. Avoid instrument collisions. Straight wrist prior to removal. Limit use of cautery when not in contact with tissue. Cases with extended cautery use. Surgical tasks around critical vessels and structures. Med watch mfr report #2955842-2009-00389 and was also sent to the fda regarding this event. A follow up mdr will be submitted if the instrument or accessory are returned (post engineering evaluation) or if additional info is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1624508 | 1547091 | 1473161 | 2955842200900380 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-26 | 2009-11-20 | No | Yes | Health Professional | 40018009 | No | No | 2009-10-26 | Device Not Retu | Yes | Initial | 2009
40018009 | 2009-10-26 | Malfunction |  | It was reported that the monopolar curved scissor (mcs) tip cover accessory was used as a replacement to complete a da vinci s hysterectomy procedure, in which the original mcs tip cover accessory presented arcing during the procedure. It was reported that a "flash" was observed coming from the tip cover accessory installed on the mcs instrument. No pt harm, adverse outcome or injury was reported. | The mcs instrument and tip cover accessory have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Members of isi's clinical sales team followed up with the surgeon and robotics coordinator to discuss the following recommendations/suggestions for prevention of mcs tip cover damage or arcing: cannula inspection prior to use. Careful installation of tip cover. Do not exceed maximum cautery settings. Avoid instrument collisions. Straight wrist prior to removal. Limit use of cautery when not in contact with tissue. Cases with extended cautery use. Surgical tasks around critical vessels and structures. Med watch mfr report # 2955842-2009-00380 was also sent to the fda regarding this event. A follow up mdr will be submitted if the instrument or accessory are returned (post engineering evaluation) or if additional info is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1624499 | 1547092 | 1473161 | 2955842200900389 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-26 | 2009-11-20 | No | Yes | Health Professional | 40018009 | No | No | 2009-10-26 | Device Not Retu | Yes | Initial | 2009
42018304 | 2009-10-26 | Malfunction |  | It was reported that during the beginning of a da vinci s prostatectomy procedure, while dropping the pt's bladder, the "side portion" of the permanent cautery hook instrument broke off and fell into the pt. Laparoscopic, robotic and x-ray searches were performed, however, the broken piece was not found. The procedure was completed with an instrument of the same type. No pt complications, harm, adverse outcome or injury was reported. The site's rn robotic coordinator reported that the pt was discharged on (b) (6) 2009 and to the best of her knowledge, the pt has not returned to the hospital due to any post operative complications. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one ear of the distal clevis and the conductor cap are missing. The missing ear piece measures approx. 235" x. 298". The yaw pulley is dislodged from the distal clevis. Engineering concluded that dislodgement of the yaw pulley caused the ear of the distal clevis to break off. Electrical continuity passed. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1624493 | 1547095 | 1473164 | 2955842200900379 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-10-26 | 2009-11-20 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2009-10-26 | Yes | No | Reuse | 2009
40016906 | 2009-10-26 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the harmonic curved shears insert fell into the pt. The piece was retrieved and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was detached from the insert. One side of the clamp arm attachment hinge feature is severely bent. The other hinge feature is also bent slightly outwards. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1624243 | 1547097 | 1473166 | 2955842200900392 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-11-02 | 2009-11-20 | No | Yes | Health Professional | 40016906 | Yes | No | 2009-11-02 | Yes | Yes | Initial | 2009
40017908 | 2009-10-21 | Malfunction |  | It was reported that during a da vinci surgical procedure, the tip of the monopolar curved scissors instrument broke off and fell into the surgical field. The broken piece was retrieved and the planned surgical procedure was successfully completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left blade was returned detached from the instrument. The blade fractured at the cross section of the grip cable crimp, thus exposing half of the cable crimp. No other damage was found. Per the info provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1624237 | 1547099 | 1473168 | 2955842200900378 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2009-10-21 | 2009-11-20 | No | Yes | Health Professional | 40017908 | Yes | Yes | 2009-10-21 | Yes | No | Reuse | 2009
42018304 | 2009-08-04 | Malfunction |  | On october 23, 2009, user facility medwatch (b) (4) was received from the fda with the following event description: "volun 19-aug-2009: during a robotic davinci case the instrument arm grasper broke inside pt, but staying intact, after completion of case, a cutdown was performed at the incision site allowing removal of instrument, op report; instrument through 3rd davinci arm snapped at level of port and given angulation; was difficult to remove through poa. At the end of the case, arm was manipulated and instrument was pulled along with arm after extending incision by 2-3 cm, pelvic cavity was irrigated, fragments from instrument sheath were removed. ". | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. The user facility is unk at this time and additional investigation can not be performed. If additional info becomes available or the instrument is returned for eval, a follow-up mdr will be submitted. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1599669 | 1552430 | 1371113 | 2955842200900377 | 1 | NAY | Manufacturer | Health ProfessionalOther | Physician | Initial | 2009-10-23 | 2009-11-20 | No | Yes | Health Professional | 42018304 | No | Yes | 2009-10-23 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-10-19 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | Conclusions - the instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A f/u mdr will be submitted if the instrument is returned (post-eval) and/or if add'l info is received. Per the info provided by the initial reporter, the broken instrument component was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1624234 | 1547803 | 1473870 | 2955842200900366 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-10-19 | 2009-11-18 | No | Yes | Health Professional | 42017908 | No | Yes | 2009-10-19 | Device Not Retu | No | Unkown | 2009
IS2000 A51P8 | 2009-10-15 | Other | Other | It was reported that at the beginning of a da vinci s prostatectomy procedure, the lamp module would not start. The site attempted to use another lamp module; however, the issue persisted. The pt was under anesthesia when the surgeon made the decision to abort the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering concluded that the vision issue experienced by the customer was associated with the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. As of (b)(4), 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1591709 | 1546571 | 1472643 | 2955842200900365 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-15 | 2009-11-13 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-10-15 | Yes | No | Reuse | 2009
42017909 | 2009-09-30 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a blade from the monopolar curved scissors instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed after a delay of 21 minutes and with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is broken and the broken piece was returned with the instrument. The broken blade had fractured at the cross section of the cable crimp and the fracture plane is not straight. The intact blade does not exhibit any damage at the tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1755891 | 1549913 | 1475899 | 2955842200900373 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-22 | 2009-11-13 | No | Yes | Health Professional | 42017909 | Yes | No | 2009-10-22 | Yes | No | Reuse | 2009
42017206 | 2009-10-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a clear amber pulley disc broke from the maryland bipolar forceps instrument and fell into the patient's abdomen. The piece was retrieved from the patient with a laparoscopic grasper. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1755881 | 1549914 | 1475900 | 2955842200900371 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-10-21 | 2009-11-13 | No | Yes | Health Professional | 42017206 | No | Yes | 2009-10-21 | Device Not Retu |  | Invalid Data | 2009
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a cable on the maryland bipolar instrument broke. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of a broken cable. Performance testing was conducted and the instrument passed recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly and electrical continuity passed. Engineering observed one yaw pulley exhibited charring and localized melting near the grip base. The damage suggests an arcing event has occurred. | MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1755872 | 1549917 | 1475903 | 2955842200900374 | 1 | NAY | Manufacturer | OtherUser facility | NOT APPLICABLE | Initial | 2009-10-23 | 2009-11-13 | No | Yes | Health Professional | 42017206 | Yes | No | 2009-10-23 | Yes | No | Reuse | 2009
40018009 | 2009-10-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from a hole in the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The event occurred while the physician was working on the left side of the patient and while skeletonizing the area. The tip cover accessory was replaced and the planned surgical procedure was completed with no patient harm, adverse outcome or injury. | The mcs tip cover accessory and mcs instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1755857 | 1549918 | 1475904 | 2955842200900375 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-16 | 2009-11-13 | No | Yes | Health Professional | 40018009 | No | No | 2009-10-16 | Device Not Retu | Yes | Unkown | 2009
IS2000 A51P7 | 2009-10-21 | Other | Other | It was reported that during a da vinci s surgical procedure, blurry and foggy vision were experienced on the left and right images of the surgeon console. The camera and endoscope were cleaned and then the endoscope was replaced, however, the issue continued to recur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issues experienced by the customer were associated with their endoscopes. Three of the site's four endoscopes were unable to simultaneously focus both eyes. The endoscope projects light from the illuminator onto the surgical site via fiber optic cables. The video image of the surgical site captured by the endoscope is sent back through the left and right channels to the camera head. The fse recommended that the site replace the affected endoscopes prior to performing their next case. Additionally, the fse replaced the camera head and camera cable as a precaution. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. As of (b)(4), 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1885540 | 1784965 | 1692149 | 2955842200900370 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-21 | 2009-11-13 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-10-21 | Yes | No | Reuse | 2009
42018304 | 2009-10-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tip of the permanent cautery hook instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken instrument component was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1622589 | 1541871 | 1468076 | 2955842200900367 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-10-13 | 2009-11-12 | No | Yes | Health Professional | 42018304 | No | No | 2009-10-13 | Device Not Retu | No | Unkown | 2009
IS2000 | 2009-11-05 | Injury | DisabilityOther | Pt was sedated and abdominal incisions were made for placement of trocars for a robotic prostatectomy. At that time, the right hand controller on the control unit fell off. The surgical procedure was immediately stopped and the pt taken from the or to recovery. Upon investigation, a set screw that holds the hand unit to the controller arm became loose. While the surgeon was manipulating the hand unit to begin the surgical procedure, the handle separated from the support shaft on the controller arm. The mfr's clinical rep was present in the or at the time, and immediately called the company's hotline for assistance. The arm could not be repaired. A repair tech was called and came to the or. Upon examination, he determined the set screw was loose and that was the cause of the controller unit separation. There was no explanation regarding how the set screw became loose. The surgical unit had only been used three times - add'l use was made during staff training-. All studies had been performed within the last 45 days with no errors. On startup, there were no mechanical or computer systems error codes that displayed. |  | DAVINCI S SURGICAL SYSTEM | NONE | 1584635 | 1531942 | 1458244 | MW5013513 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2009-11-10 | 2009-11-10 | Yes | Yes | Health Professional | IS2000 | Device Returned To Manufacturer | No |  |  |  |  | 2009
IS2000 A51P8 | 2009-10-07 | Injury | Required Intervention | It was reported that toward the end of a da vinci s sacralcopopexy procedure, the console surgeon experienced user message "lighten grip". The console surgeon attempted to troubleshoot the system error by releasing the grips on the master tool manipulators (mtm); however, the system error persisted. The console surgeon then placed their head back into the high resolution stereo viewer (hsrv) located on the surgeon's console and accidently clutched and changed control of the patient side manipulator (psm) from psm 2 to psm 3. The surgeon was unclear why the instrument on the psm 2 was not moving when she noticed that she had accidently clutched and gained control of the instrument in psm 3. At this point, the surgeon had already moved psm 3 with instrument down causing the instrument to make contact and injuring the patient's sacral. The case was converted to a traditional open procedure and the patient's injury was repaired and the procedure was completed. As of 2009, the patient was reported as doing well with no further issues. | The investigation conducted by field service engineering (fse) consisted of reviewing system error logs and performing functional test. The fse's review of system error logs did not reveal any system errors and the system functioned as intended with no issues. The user message "lighten grip" appears when the system attempts to drive the master to the desired alignment position but the master cannot reach this positon. Lightening the hold of the mtm will allow the user to continue. The master alignment is designed to ensure the master grips are appropriately positioned relative to the tips of the instruments in the high definition stereo viewer. Master alignment happens in the following conditions: whenever a new instrument is installed, whenever a new scope angle is selected, when the patient cart operator moves the instrument or camera arms, instrument arm swapping. Since this event, the isi clinical sales representative has provided additional training to the console surgeon regarding appropriate troubleshooting of this user message. As of 2009, the hospital has continued to use the system with no reported recurrences of the reported issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1575901 | 1531172 | 1457489 | 2955842200900363 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-07 | 2009-11-06 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2009-10-07 | Device Not Retu | No | Reuse | 2009
IS2000 A51P8 | 2009-10-07 | Other | Other | It was reported that halfway through a da vinci s prostatectomy procedure, the customer experienced multiple occurrences of system error code #212. The surgical staff attempted to override the error and then restarted the system, however, the error code continued to recur. An isi technical support engineer was contacted and troubleshooting techniques were performed, however, the issue could not be resolved. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #212 was associated with the right master tool manipulator. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a "non-recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering eval found that the harness conductor had malfunctioned, thus generating the system error experienced by the customer. As of (b)(4) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1769282 | 1531420 | 1457735 | 2955842200900362 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-10-07 | 2009-11-06 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2009-10-07 | Yes |  | Invalid Data | 2009
ISI200 A43P9 | 2009-10-08 | Other | Other | It was reported that during a da vinci lower anterior resection procedure, the customer experienced a non-recoverable system error code #1. The surgeon advised that the case may be converted due to the pt's previous surgeries, however, the nature of the surgeries and the condition of the pt at the time of surgery is unk to isi. The system had been in use for approx one hour when the surgeon decided to convert to traditional open surgical technique to finish the planned surgery. No pt harm, adverse outcome, or injury was reported. | The investigation conducted by isi field service engineering found a defective printed circuit assembly, fan power supply (pca, fps). The pca fps converts the 48 volt dc supplied by the primary power supply into +12v, -12v, +5v, -5v and +3v which are used by other boards in the system for various purposes. The +12v supply drives the fans. The system was repaired by replacing the affected pca, fps. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #1 appears when the da vinci safety system determines that a power supply voltage was out of range. Upon determining this condition, the safety system puts da vinci s in a "non-recoverable safe state". The pca, fps was returned to isi for failure analysis investigation. Engineering confirmed the customer reported complaint as they were able to duplicate the system error code #1. As a result the pca, fps was scrapped. As of (b)(4) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1771609 | 1531422 | 1457737 | 2955842200900364 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-10-08 | 2009-11-06 | Yes | Yes | Health Professional | ISI200 A43P9 | Yes | No | 2009-10-08 | Yes | No | Reuse | 2009
42018304 |  | Malfunction |  | It was reported that after cleaning and sterilization a frayed cable was observed on the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of a frayed cable at the instrument wrist as all pitch and yaw cables are intact and undamaged. The instrument moves as intended in pitch and yaw when the input discs are manually rotated. Engineering also observed that the conductor wire is damaged, which resulted in arcing to nearby components at the instrument wrist. The yaw pulley, distal clevis and conductor cap exhibit charring and localized melting indicating an arcing event occurred. The section of conductor wire adjacent to the charred areas has damaged insulation, exposing bare wire. Electrical continuity is intermittent, passing when the instrument wrist is straight and failing when the wrist is in yaw to either side. No other damage was found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1586917 | 1534748 | 1461033 | 2955842200900358 | 1 | GEI | Manufacturer | Health ProfessionalOther | Nurse | Initial | 2009-10-12 | 2009-11-06 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2009-10-12 | Yes | No | Reuse | 2009
| 2009-08-17 | Malfunction |  | Event Description: --/-- When removing the robotic arms, the prograsp instrument was not coming out. The surgical technician looked with the camera to see what was impeding removal of the instrument. The end of the instrument (at the wrist) was cracked and caught on trocar. The instrument was then manipulated and advanced slightly to facilitate removal. The pieces of the end of the prograsp were described as being "shaved off. " there was no known patient harm. The prograsp is a reprocessed instrument and number of uses is documented. This prograsp was on use 9 out of 10. --/-- Additional Event Description: --/-- When removing the robotic arms, the prograsp instrument was not coming out. The surgical technician looked with the camera to see what was impeding removal of the instrument. The end of the instrument (at the wrist) was cracked and caught on trocar. The instrument was then manipulated and advanced slightly to facilitate removal. The pieces of the end of the prograsp were described as being "shaved off. " there was no known patient harm. The prograsp is a reprocessed instrument and number of uses is documented. This prograsp was on use 9 out of 10. | Manufacturer Narrative: --/-- | ENDOWRIST PROGRASP |  | 1564781 | 1522487 | 1449016 | 1522487 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-10-30 | 2009-10-30 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2009
IS2000 A51P8 | 2009-10-07 | Other | Other | It was reported that a few mins into a da vinci s prostatectomy procedure, the surgeon side console (ssc) switched to battery power and then turned off after approx 2-3 mins. The isi rep present for the procedure reset the system breaker, tried a different power outlet and secured the power cable, however, the ssc would not power on. The surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that the customer reported complaint was related to the primary power supply (pps). The pps is a +48v power supply with battery backup and is mounted outside of the system control electronics chassis. The system was repaired by replacing the pps. As of 10/29/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1569321 | 1523695 | 1450161 | 2955842200900361 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-10-07 | 2009-10-30 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-10-07 | Yes | No | Reuse | 2009
40017908 | 2009-09-23 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the instrument was in the patient but not in use, the monopolar curved scissors instrument snapped near the base of the jaw and one piece fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade (left grip) detached from and returned with the instrument. The blade fractured at the cross section of the cable crimp with the fracture plane nearly straight. The intact blade exhibits chipping on the blade edge, roughly halfway between the midpoint and tip. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1618233 | 1523879 | 1450344 | 2955842200900360 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2009-10-06 | 2009-10-30 | No | Yes | Health Professional | 40017908 | Yes | No | 2009-10-06 | Yes | Yes | Reuse | 2009
42017907 | 2009-10-06 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the monopolar curved scissors (mcs) instrument with the mcs tip cover installed was removed for cleaning. It was noted that the gray hard plastic section of the mcs tip cover accessory was open and charred. Removal of the mcs tip cover accessory revealed a corresponding charred and eroded section on the mcs instrument. The surgeon stated that he noticed that the coag did not work normally when made aware by the surgical staff of the condition of the mcs instrument and tip cover accessory. The planned surgical procedure was completed using a replacement mcs instrument and tip cover accessory. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of tube extension has a hole approximately. 265" x. 143" burned through it. Localized melting and material removal around the hole is indicative of arcing. Disassembly of the instrument by engineering revealed that the tube extension, proximal clevis, and cables are charred in the same location as the hole. The tube extension does not exhibit cracking damage. The instrument also passed electrical continuity testing. No other damage was found. The site returned the mcs tip cover accessory used in the conjunction with the mcs instrument. Engineering observed that the ultem sleeve of the tip cover has an l shaped hole extending approximately. 367" from the silicone to sleeve interface. The sleeve exhibits arcing damage that aligns with the arcing damage located on the mcs instrument tube extension. There is no damage to the silicon portion of the tip cover. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1619639 | 1523880 | 1450345 | 2955842200900359 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-10-06 | 2009-10-30 | No | Yes | Health Professional | 42017907 | Yes | Yes | 2009-10-06 | Yes | No | Reuse | 2009
IS3000 A60P2 | 2009-09-18 | Injury | Required Intervention | It was reported that 4 days post op, a da vinci si hysterectomy procedure, a 5mm perforation of the pt's bowel was observed during a laparoscopic colostomy procedure. In 2009, the surgeon provided the following additional info: pre-operatively, the pt experienced complications of severe bowel impaction (hard stool) throughout the entire large intestine due to diverticulosis. The pt's bowel was infused with an enema and contrast all the way to the cecum to aid in the movement of the pt's bowel. In addition, the pt had a colostomy tube that presented leakage thus causing the pt to develop an infection. It is unk how damage to the pt's bowel occurred; however, the perforation may have been introduced during the enema and contrast procedure. The pt remains hospitalized as of the same day. In addition, the infection is under control and the pt is reported to be doing well. | Per the info provided by the customer, it is indeterminable as to how injury to the pt's bowel occurred. As of 10/22/2009, there have been no reported recurrence of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1564830 | 1520149 | 1446892 | 2955842200900347 | 1 | NAY | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2009-09-28 | 2009-10-23 | Yes | No | Health Professional | IS3000 A60P2 | No | No | 2009-09-28 | Device Not Retu | No | Reuse | 2009
40018009 | 2009-09-25 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was performing the pelvic lymph node dissection, arcing was observed from the monopolar curved scissors instrument with tip cover accessory installed. No patient harm, adverse outcome or injury was reported, and the planned surgical procedure was completed. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617312 | 1520306 | 1447036 | 2955842200900352 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018009 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-01-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. It was reported that during this procedure, the electrosurgical unit (esu) was set to fulgarate mode at 50 watts, which is above use recommendations. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Recommended power settings for a given mode are provided to the customer in the electrosurgical unit (esu) settings and footswitch cables instructions for use. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617288 | 1520307 | 1447037 | 2955842200900353 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-02-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. It was reported that during this procedure, the electrosurgical unit (esu) was set to fulgarate mode at 50 watts, which is above use recommendations. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Recommended power settings for a given mode are provided to the customer in the electrosurgical unit (esu) settings and footswitch cables instructions for use. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617277 | 1520308 | 1447038 | 2955842200900356 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Unkown | 2009
400180 | 2009-09-24 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with mcs tip cover accessory installed. The instrument and tip cover accessory were removed and replaced. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the tip cover accessory is returned for evaluation, a follow-up mdr will be submitted. The monopolar curved scissors (mcs) instrument used with the affected tip cover was returned. Engineering's evaluation of the instrument could not locate signs of arcing or charring. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617332 | 1520309 | 1447039 | 2955842200900349 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-09-28 | 2009-10-23 | No | Yes | Health Professional | 400180 | Device Returned To Manufacturer | Yes | 2009-09-28 | Device Not Retu | Yes | Unkown | 2009
40018304 | 2009-09-29 | Malfunction |  | It was reported that during a da vinci s low anterior resection procedure, the permanent cautery hook instrument broke apart and pieces fell into the patient. The broken pieces were retrieved and the surgical staff confirmed that the field was visually clear. The planned surgical procedure was completed, and no patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1617316 | 1520310 | 1447040 | 2955842200900350 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-29 | 2009-10-23 | No | Yes | Health Professional | 40018304 | No | Yes | 2009-09-29 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-09-30 | Malfunction |  | It was reported that 24 minutes into a da vinci si hysterectomy procedure, while the surgeon was dissecting, arcing to the patient's uterus was observed coming from the monopolar curved scissors instrument with tip cover accessory installed. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and one hole burned through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is roughly. 025 inches in diameter and is located. 315 inches above the silicone to sleeve interface with a tear on one side of the hole. Engineering observed various mechanical tears in other locations on the tip cover, including at the tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1764926 | 1520312 | 1447042 | 2955842200900351 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-09-30 | Yes | Yes | Initial | 2009
40018005 | 2009-03-19 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. It was reported that during this procedure, the electrosurgical unit (esu) was set to fulgarate mode at 50 watts, which is above use recommendations. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Recommended power settings for a given mode are provided to the customer in the electrosurgical unit (esu) settings and footswitch cables instructions for use. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617308 | 1520314 | 1447044 | 2955842200900357 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-02-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. It was reported that during this procedure, the electrosurgical unit (esu) was set to fulgurate mode at 50 watts, which is above use recommendations. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Recommended power settings for a given mode are provided to the customer in the electrosurgical unit (esu) settings and footswitch cables instructions for use. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617302 | 1520316 | 1447046 | 2955842200900355 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-02-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no patient harm, adverse outcome or injury. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617295 | 1520318 | 1447048 | 2955842200900354 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-30 | 2009-10-23 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-09-30 | Device Not Retu | Yes | Initial | 2009
IS2000 A51P8 | 2009-09-28 | Other |  | It was reported that during a da vinci s prostatectomy procedure the customer experienced non-recoverable system error code #45049. A loud "click" or "snap" was heard and the pt side manipulator arm stopped working. An isi technical support engineer attempted to troubleshoot the issue via telephone and requested that the site perform a power cycle and restart the system, however, the system faulted with system error code #281. Further troubleshooting actions were performed, however, the system continued to fault. The pt was under anesthesia approx 2. 25 hours and the port incisions were made when the surgeon made the decision abort and reschedule the planned surgical procedure. Add'l info provided by an isi rep on oct 22, 2009, reported that the rescheduled da vinci s surgical procedure was performed and successfully completed. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering (fse) found that the site also experienced system error code #23. Engineering concluded that error code #23 experienced by the customer was associated with a pt side manipulator (psm) embedded serializer set up joint (essj). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The essj is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The system was repaired by replacing the affected psm's essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23 is reported by software to denote communication faults in the low voltage differential signal carrying info about the psm. System error codes #281 and #45049 were generated as a result of communication lost between the system video processor (svp) board that controls the video displays and the master supervisory controller (msc) when the essj failed. In the event of these communication issues, the system records the fault and transitions to a safe state. The essj was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the essj did not communicate correctly with the psm or system. As of october 22, 2009, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1772996 | 1520506 | 1447213 | 2955842200900348 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-09-28 | 2009-10-23 | No | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-09-29 | Yes | No | Reuse | 2009
40018402 | 2009-09-22 | Malfunction |  | It was reported that during a da vinci gynecology procedure the permanent cautery hook instrument broke and pieces fell into the patient. One broken piece was retrieved, however, another broken piece could not be found. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1777301 | 1517178 | 1444036 | 2955842200900342 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-22 | 2009-10-21 | No | Yes | Health Professional | 40018402 |  | Yes | 2009-09-22 | Device Not Retu | No | Unkown | 2009
| 2009-10-13 | Malfunction |  | During robotic assisted supracervical hysterectomy hinged grasper on intuitive harmonic curved insert worked initially, then malfunctioned making grasping action impossible. Staff reports this is the second such event with this product. Product was safely removed, with no injury to patient. However, when product was being cleaned after the procedure, the grasper became completely separated from the remainder of the device. Piece is similar in shape to a ball-point pen clip and roughly one-fourth the size. In addition to intra-procedural malfunction, potential exists for retained foreign body. |  | HARMONIC CURVED SHEARS INSERT | HARMONIC CURVED INSERT | 1559310 | 1515365 | 1442305 | MW5013133 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2009-10-20 | 2009-10-20 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2009
IS2000 |  | Injury | Other | On september 18, 2009, a reuters article titled "medical societies push standards for robotic surgery" was published. The article includes quotes from dr. Kevin zorn and mentions his recount of a case where a patient died post a da vinci prostatectomy procedure: "he recounted one case of a surgeon who was using the system for the fourth time. After eight hours of surgery, the proctor -- an experienced surgeon who supervises the operation -- told the surgeon that progress was too slow. He recommended the surgeon switch to conventional surgery, where an incision is made from the navel to the pubic bone to access the prostate. After the proctor left the operating room, the surgeon continued using the robot. The patient later died from complications. " on september 29, 2009, upon following up, it was discovered that another dr was the proctor. | On september 29, 2009, dr stated that the prostatectomy procedure took an extended amount of time and resulted in conversion to traditional open surgical techniques to complete, however, the patient did not die. On october 12, 2009, dr confirmed the procedure took place at the hospital and that the patient suffered nerve injury due to the length of the procedure. Based on the information provided, it was determined that the da vinci s surgical system, instruments or accessories did not malfunction in a way that would cause nor contribute to the patient injury. The case was the surgeon's fourth procedure using the da vinci s surgical system, and it was determined that the patient's injury was likely due to the prolonged length of the procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1558872 | 1515425 | 1442363 | 2955842200900346 | 1 | NAY | Manufacturer | Health ProfessionalOther | Physician | Initial | 2009-09-18 | 2009-10-16 | Yes | No | Health Professional | IS2000 | No | Yes | 2009-09-18 | Device Not Retu | No | Reuse | 2009
40017908 | 2009-09-24 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, one of the blades from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval, a f/u mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1618483 | 1516429 | 1443328 | 2955842200900345 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-09-25 | 2009-10-16 | No | Yes | Health Professional | 40017908 | No | No | 2009-09-25 | Device Not Retu | No | Unkown | 2009
42017909 | 2009-09-16 | Malfunction |  | It was reported that while cutting tissue during a da vinci s hysterectomy procedure, the monopolar curved scissors instrument tip broke and fell into the pt. The broken piece was retrieved, and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the left grip blade is missing and the blade had fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit any damage at the tip, however, material pitting can be observed on the side opposite to the sharp edge. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1618510 | 1516431 | 1443330 | 2955842200900344 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-09-23 | 2009-10-16 | No | Yes | Health Professional | 42017909 | Yes | No | 2009-09-23 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-09-21 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced system error code #3500. The site pressed the fault override button and disabled one of the pt side manipulator (psm) arms, however, the fault recurred. The site contacted isi and a technical support engineer (tse) attempted to instruct the site on restarting the system to gain access to the arm, when the site made the decision to discontinue troubleshooting. Approximately 4 hours into the procedure, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering did not find system error code #3500, but did find system error code #23005 and #26005 present in the system error logs. It was suspected that system error code #23005 and #26005 were due to a user induced arm collision during use. The field service engineer tested the system and found it to be working to manufacturers specifications with no defects or malfunction. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23005 appears when the da vinci s safety system determines that a condition exist where the system was generating an unexpectedly high amount of output power. This error is intended to detect extremely rare failures in the internal memory or the processor that controls the robot motors. System error code #26005 appears when a manipulator has been disabled by the user. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". As of october 15, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1766525 | 1516438 | 1443337 | 2955842200900340 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-21 | 2009-10-16 | Yes | No | Health Professional | IS2000 A51P8 | Yes | Yes | 2009-09-21 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-09-21 | Other | Other | It was reported that at the start of a da vinci s internal mammary artery (ima) harvest procedure, the right image on the high resolution stereo viewer (hrsv) was black with no visible icons, text, or pictures. A technical support engineer (tse) recommended that the site place the system in 2-d; however, the vision issue remained. The site then recycled power to the system multiple times as well as raised and lowered the hrsv; however, the issue persisted. The pt was under anesthesia for approximately 1 hour and port incisions had been created when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv provides the video image for the surgeon side console (ssc). The system was repaired by replacing the affected hrsv. The hrsv was returned to isi for evaluation. Engineering was able to replicate the report, vision issue determined that the hrsv monitor had malfunctioned. As of october 15, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1766540 | 1516439 | 1443338 | 2955842200900341 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2009-09-21 | 2009-10-16 | Yes | No | Health Professional | IS2000 A51P8 | Yes | No | 2009-09-21 | Yes | No | Reuse | 2009
40016906 | 2009-09-22 | Malfunction |  | It was reported that while working on a vessel during a da vinci colectomy procedure, the harmonic curved shears insert tip broke and fell into the patient. The insert had been in use for 3 minutes when the incident occurred. The broken piece was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, the insert was returned with a piece of the curved blade broken off. The broken piece is approximately. 425" x. 08" in size. Engineering observed that the blade fractured near its transition point from straight to curved. A gouge mark was also observed on one side of the blade, located slightly above the fracture plan. The intact section of the blade (still attached to the insert) exhibits scratches near the fracture. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1672380 | 1589615 | 1513223 | 2955842200900343 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-09-22 | 2009-10-16 | No | Yes | Health Professional | 40016906 | Yes | No | 2009-09-22 | Yes | Yes | Initial | 2009
40018005 | 2009-09-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, arcing was observed on the monitor. Upon closer examination, 4 "micro" holes were observed on the monopolar curved scissors (mcs) tip over accessory while installed on a mcs instrument. The mcs tip cover accessory was removed and replaced. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. The silicone exhibits localized melting around the holes, indicating that arcing events occurred. The holes are all under. 030" with varying shapes (round and oblong), and are located from. 145" to. 460" above the silicone to the sleeve interface. The customer also returned the mcs instrument that was used during the procedure. Examination of the instrument revealed a char mark at the proximal clevis ear. When the tip cover accessory was installed on the instrument, it was observed that the location of one of the hole matches with the location of the char mark on the instrument. A video of the procedure was also returned for analyses. The video showed several collisions between the tip of the mcs instrument with tip cover installed and another instrument and accessories used in the procedure, prior to the arcing event being observed. It is inconclusive, however, these observed collisions may have created damage to the tip cover which allowed arcing to occur. The endowrist instruments instructions for use specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1617514 | 1516368 | 1443267 | 2955842200900339 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-09-17 | 2009-10-14 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-09-17 | Yes | Yes | Initial | 2009
IS1200 A43P9 | 2009-09-02 | Other | Other | It was reported that during a da vinci procedure, the customer experienced a nonrecoverable system error code #21013. The site recycled power to the system; however, the system error code persisted. The site then exercised axis 3 of the left master tool manipulator (mtml), but was not successful in clearing the system error code. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering found system error code #21013 in the system error logs as well as system error code #20013 and determined that both error codes were associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed, and there was no injury to the pt. System error codes #20013 and #21013 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a "recoverable safe state". The mtml was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, an axis potentiometer and motor encoder assembly were replaced. As of 10/01/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1614959 | 1512251 | 1439417 | 2955842200900326 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2009-09-09 | 2009-10-09 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2009-09-09 | Yes | No | Invalid Data | 2009
IS1200 A43P9 | 2009-09-09 | Other | Other | It was reported that during a da vinci surgical procedure, the customer experienced multiple occurrences of system error codes #21013 and #21008. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering (fse) concluded that system error codes #21013 and #21008 experienced by the customer were associated with a pt side manipulator (psm) arm. The psm is an instrument arm located on the pt side cart and provides the sterile interface for the endowrist instruments. The fse also found error code #21008, associated with the right master tool manipulator (mtmr) gimbal, in the system error logs. The mtm gimbal is a subsection of the complete mtm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected psm and mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error codes #21013 and #21008 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts da vinci in a "nonrecoverable safe state. " the psm and mtmr gimbal were returned to isi for eval. Engineering determined that an encoder had malfunctioned on the psm and axis potentiometer had malfunctioned on the mtmr gimbal, thus generating the system error codes experienced by the customer. As of 10/01/2009, there have been no reported recurrences of the issue at the hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1615011 | 1512252 | 1439418 | 2955842200900324 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2009-09-09 | 2009-10-09 | No | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-09-10 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, no picture was displayed on the left view image of the surgeon's console and the issue continued to recur after the system was restarted. Prior to contacting an isi field service engineer (fse) to assist with troubleshooting the issue, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | Investigation conducted by the fse concluded that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The camera was also replaced as a precaution. The camera produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of 10/01/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1614899 | 1512253 | 1439419 | 2955842200900328 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2009-09-10 | 2009-10-09 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-09-10 | Yes | No | Reuse | 2009
IS2000 A51P8 | 2009-09-15 | Injury | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, the surgeon activated that monopolar curved scissors (mcs) instrument while it was out of their field of view and accidently lacerated the patient's right external iliac artery. The planned surgical procedure was converted to an open procedure and the damage to the patient's artery was repaired using a primary closure technique. The patient was reported to be doing well and has not experienced any post operative complications. | The investigation conducted by engineering to review the site's system error logs for the reported event concluded that the system did not experience any faults while in use. In addition, an isi clinical sales representative discussed the event with the surgeon and indicated that the laceration of the patient's right iliac artery was likely due to movement of the mcs instrument outside of the surgeon's view during the procedure. The surgeon has undergone additional da vinci s system training in 2009. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1556279 | 1512169 | 1439336 | 2955842200900337 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-15 | 2009-10-06 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2009-09-15 | Device Not Retu | No | Reuse | 2009
IS3000 A60P2 | 2009-09-11 | Other | Other | It was reported that during a da vinci si surgical procedure, the customer experienced a system error and the led of one of the system arms turned red. The site power cycled the system a couple of times and then called isi for assistance. An isi field service engineer (fse) found a loose connection on a pca board related to the affected system arm and the issue appeared to have been resolved after the connector was reseated. The site then called back indicating that they experienced system error code #281 and the error code continued to recur after they restarted the system. An isi technical support engineer had the site power down the system, emergency power off and reset the breaker of the patient side cart, and had the surgical staff reseat a cable connection. The system was restarted, however, system error code #25595 occurred upon restart. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering (fse) concluded that the system error codes experienced by the customer were associated with a setup joint (suj) located on the pt side cart (psc) and the isi core controller (icc). The psc contains four setup joints and system commands are communicated between the various pca's of the four setup joints. The icc cfg is the central control board of the da vinci si, containing the master system controller and the control and transform processor. The system was repaired by replacing the affected suj and icc. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error codes #281 occurs when a processor did not complete a step during system startup within the allotted time and #25595 appears when communication errors are detected between two circuit boards. Upon determining these conditions, the safety system puts da vinci in a "nonrecoverable safe state". As of 10/08/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1615102 | 1512249 | 1439415 | 2955842200900330 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-09-11 | 2009-10-06 | Yes | Yes | Health Professional | IS3000 A60P2 | Yes | No | 2009-09-11 | Yes | No | Reuse | 2009
IS2000 A51P8 |  | Death | Death | It was reported that eleven days post op a da vinci s mitral valve repair procedure, the pt expired. Reportedly, the pt's demise was unrelated to the da vinci s surgical procedure. On september 24, 2009, an isi representative provided the following add'l info: the pt experienced complications related to acute lung injury and pulmonary embolism. The surgeon made the decision to convert the planned surgical procedure to a thoracotomy to complete the planned surgical procedure expeditiously to allow for taking the pt off of bypass. | On september 30, 2009, the hospital risk management coordinator confirmed that the pt's demise was unrelated to the da vinci s system. No other details were provided. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1552209 | 1503746 | 1431152 | 2955842200900323 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-09-09 | 2009-10-02 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2009-09-09 | No | No | Reuse | 2009
42003305 | 2009-09-09 | Malfunction |  | It was reported that during a da vinci s tubal reanastomosis procedure, the tip of the black diamond micro forceps instrument broke off and fell into the patient. The planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is missing. The grip had fractured at the junction with the grip hub and as a result, the cable crimp is partially exposed. The fracture surfaces do not exhibit any obvious damage when viewed under 25x magnification. The intact grip does not exhibit cracking in the same location as the broken grip. No other damage was found. | BLACK DIAMOND MICRO FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1555031 | 1508435 | 1435670 | 2955842200900327 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42003305 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
42019006 | 2009-09-04 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure a "shard" from the cobra grasper instrument shaft may have fell inside of the patient. The "shard" was retrieved and the surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. It was noted during the procedure that the cobra grasper instrument "hit" the trocar. A maryland bipolar instrument and an endowrist prograsp instrument were also used during the procedure. It is unknown from which instrument the "shard" fell from. An inspection of the instrument shafts by the site revealed that the surface of the instruments were damaged. The sterile processing department and the robotics coordinator found that when they rubbed the instrument surface with their hands, "slivers" from the instrument shafts stuck in their hands. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube has one side that contains various scratches with material removal. The scratches are narrow and not quite aligned with the tube axis. The surface finish of the scratches is rougher than the undamaged sections of the tube. No other damage was found. The following med watch mfr reports were sent to the fda regarding this event: 2955842-2009-00332, -00333, -00335. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1781069 | 1508437 | 1435672 | 2955842200900334 | 1 | GEI | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42019006 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
42009307 | 2009-09-04 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure a "shard" from the endowrist prograsp instrument shaft fell into the patient. The "shard" was retrieved and the surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. It was noted during the procedure that the endowrist prograsp instrument "hit" the trocar. A cobra grasper instrument and a maryland bipolar instrument were also used during the procedure. It is unknown from which instrument the "shard" fell from. An inspection of the instrument shafts by the site revealed that the surface of the instruments were damaged. The sterile processing department and the robotics coordinator found that when they rubbed the instrument surface with their hands, "slivers" from the instrument shafts stuck in their hands. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube presents deep scratches within an area extending approximately 1" from the interface with the proximal clevis where material was removed from the main tube. The damaged sections have a rough surface finish. No other damage was found. The following med watch mfr reports were sent to the fda regarding this event: 2955842-2009-00332, -00334, -00335. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1780253 | 1508439 | 1435674 | 2955842200900333 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
42017206 | 2009-09-04 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure a "shard" from the endowrist prograsp instrument shaft fell into the patient. The "shard" was retrieved and the surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. It was noted during the procedure that the maryland bipolar forceps instrument "hit" the trocar. During the procedure a cobra grasper instrument and an endowrist prograsp instrument were used, and it is unknown from which instrument the "shard" fell from. An inspection of the instrument shafts by the site revealed that the surface of the instruments were damaged. The sterile processing department and the robotics coordinator found that when they rubbed the instrument surface with their hands, "slivers" from the instrument shafts stuck in their hands. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube has a section with material removal. The damaged area includes a couple of scratch marks not aligned with the tube axis and has an. 850" long,. 060" wide scratch parallel to the tube axis. No other damage was found. The following medwatch mfr reports were sent to the fda regarding this event: 2955842-2009-00333, -00334, -00335. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1780273 | 1508441 | 1435676 | 2955842200900332 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
42018404 | 2009-09-09 | Malfunction |  | It was reported that during a da vinci s atrial septal and patent foramen procedure, a "shard" from the permanent cauter spatula instrument fell inside the patient. The "shard" was retrieved and the surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. It was noted that when the instrument was manipulated during the procedure the instrument "hit" the trocar. An inspection of the instrument shaft by the site revealed that the surface of the instrument was damaged. The sterile processing department and the robotics coordinator found that when they rubbed the instrument surface with their hands, "slivers" from the instrument shafts stuck in their hands. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube has a section with material removed on one side. The damaged area includes a couple of wide scratches parallel to the tube axis. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1555012 | 1508443 | 1435678 | 2955842200900331 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2009-09-09 | Yes | No | Reuse | 2009
40016906 | 2009-09-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the moveable jaw of the harmonic curved shears insert became loose and fell off inside of the patient during removal. The jaw was retrieved with the use of fluoroscopy and a laparoscopic grasper. | The harmonic shears insert has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1555043 | 1508446 | 1435681 | 2955842200900329 | 1 | LFL | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-11 | 2009-10-02 | No | Yes | Health Professional | 40016906 | No | Yes | 2009-09-11 | Device Not Retu | Yes | Initial | 2009
42019407 | 2009-09-04 | Malfunction |  | It was reported that during a da vinci myomectomy procedure, the cable on the mega needle drive instrument broke. Reportedly, no part of the instrument fell into the pt. A post op inspection of the mega needle driver instrument by the site revealed that the instrument shaft surface has damage. The sterile processing dept and the robotics coordinator found that after rubbing the instrument with their hands, "silvers" from the instrument shaft stuck in their hands. | Results & conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube presents a deep scratch within an area extending approx. 75" from the interface with the proximal clevis where material was removed from the main tube. The damaged section has a rough surface. No other damage was found. The following medwatch mfr reports were sent to the fda regarding this event: 2955842-2009-00332, 2955842-2009-00333, 2955842-2009-00334. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1555093 | 1508447 | 1435682 | 2955842200900335 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-09-09 | 2009-10-02 | No | Yes | Health Professional | 42019407 | Yes | Yes | 2009-09-09 | Yes | No | Initial | 2009
42017908 | 2009-08-21 | Malfunction |  | It was reported that while cutting the endopelvic fascia during a da vinci s prostatectomy procedure, one of the blades from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | Conclusions - the instrument and detached left grip blade were returned and evaluated. Per the customer reported complaint, engineering observed that the blade fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit any damage at the tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1555075 | 1508449 | 1435684 | 2955842200900325 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2009-09-02 | 2009-10-02 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-09-02 | Yes | No | Reuse | 2009
| 2009-10-01 | Malfunction |  | Piece of forceps melted during procedure. Manufacturer contacted. |  | INTUITIVE TENACULUM FORCEPS | FORCEPS | 1551925 | 1501180 | 1428609 | MW5012919 | 1 | GCJ | Voluntary |  | RISK MANAGER | Initial | 2009-10-01 | 2009-10-01 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2009
420179 | 2009-09-29 | Malfunction |  | No injury to pt. Piece of forceps melted during procedure. |  | MONOPOLAR CURVED SCISSORS | FORCEPS | 1551933 | 1501181 | 1428610 | MW5012920 | 1 | HTD | Voluntary |  | RISK MANAGER | Initial | 2009-10-01 | 2009-10-01 | No | Yes | Health Professional | 420179 | Yes | Yes |  |  |  |  | 2009
| 2009-05-13 | Malfunction |  | No injury to pt. Piece of forceps began to melt during the procedure. |  | MONOPOLAR CURVED SCISSORS | FORCEPS | 1551935 | 1501182 | 1428611 | MW5012921 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2009-10-01 | 2009-10-01 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2009
|  | Malfunction | Other | While attempting to remove instrument i was unable to withdraw it completely through the port cannula. The port cannula instrument were visualized using the camera and it was noted that the scissor end had separated from the shaft of instrument. It appeared to have gotten caught on the cannula (port) and broken at that point. |  | INTUITIVE | METZ SCISSOR INSTRUMENT | 2633702 | 2602390 | 2499077 | 2602390 | 1 | JOK | User Facility | Unknown | Other | Initial | 2009-08-24 | 2009-09-30 | No | Yes | Health Professional |  | No | Yes |  | No Answer Provi |  | Invalid Data | 2009
42017908 | 2009-09-02 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a piece of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was not retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have the tip of one scissor blade broken off. The size of missing piece is roughly. 090" x. 060". The fracture plane of broken blade is not straight and the same blade has a chip in the blade edge, below the midpoint. The intact blade does not exhibit any obvious cracking. The proximal clevis exhibits a large char mark. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1602170 | 1505572 | 1432917 | 2955842200900321 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-09-04 | 2009-09-29 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-09-02 | Yes | No | Reuse | 2009
| 2009-07-23 | Other | Hospitalization | It was reported post a successful da vinci s prostatectomy procedure, the pt was found to have a small bowel injury. The isi distributor rep reviewed the case video and was unable to identify when this injury may have occurred. The site's instruments were inspected for cracked insulation or other types of damage that may have caused the pt injury, however, nothing was observed. The nature and cause of the injury was reported as being unclear by the surgeon. The pt was required to stay in the hospital for an additional day for observation, however, no additional pt harm, adverse outcome or injury was reported. | Conclusions - the site has been asked to provide isi with the procedure video for further analysis and evaluation, however, no info has been received as of september 25, 2009. The root cause of the customer reported event cannot be determined. A follow-up mdr will be submitted if the procedure video is returned (post-evaluation) and/or if additional info is received. | 8MM CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1774646 | 1490172 | 1418139 | 2955842200900311 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2009-08-27 | 2009-09-25 | Yes | No | Health Professional |  | No | No | 2009-08-27 | Device Not Retu | No | Unkown | 2009
IS1200 A43P9 | 2009-08-24 | Malfunction | Other | It was reported that during a da vinci surgical procedure the customer experienced system error code #20013. The pt was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. Additional info provided by the distribution rep on sept 24, 2005 reported that the pt underwent surgery with the da vinci surgical system. The surgical procedure was completed and the pt is reported to be doing well. | The investigation conducted by field service engineering found system error code #20013 in the system error logs as well as system error code #21013 and determined that both error codes were associated with a pt side manipulator (psm). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error codes #20013 and #21013 appear when the da vinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system, puts da vinci in a "nonrecoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of sept 24, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1587255 | 1490174 | 1418141 | 2955842200900317 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2009-08-31 | 2009-09-25 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-08-31 | Yes | No | Reuse | 2009
42017908 | 2009-08-27 | Malfunction |  | It was reported that during a da vinci s surgical procedure, one of the blades from the monopolar curved scissors instrument broke off and fell into the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument and detached left grip blade were returned and evaluated. Per the customer reported complaint, engineering observed that the blade fractured at the cross section of the cable crimp. The fracture plane is straight and the intact blade does not exhibit any damage at the tip, however, material corrosion was observed on both blades. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1609250 | 1505506 | 1432852 | 2955842200900313 | 1 | GEI | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-08-28 | 2009-09-25 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-08-28 | Yes | No | Reuse | 2009
40018008 | 2009-07-24 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, and while cutting through tissue, arcing was observed coming from the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was replaced to successfully complete the planned surgical procedure with no additional pt harm, adverse outcome or injury. The tip cover was discarded. | The tip cover accessory was discarded by the site and will not be returning to intuitive surgical. The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1609264 | 1505507 | 1432853 | 2955842200900319 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-01 | 2009-09-25 | No | Yes | Health Professional | 40018008 | No | Yes | 2009-09-01 | Device Not Retu | Yes | Unkown | 2009
42018304 | 2009-09-01 | Malfunction |  | It was reported that halfway through a da vinci s general surgery procedure, a plastic piece from the permanent cautery hook instrument broke off and fell inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the ceramic sleeve (which the customer referred to as the "plastic piece") of the instrument is almost entirely missing. A. 080" x. 065" piece of the sleeve is still attached to the yaw pulley. Per the info provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1609281 | 1505508 | 1432854 | 2955842200900318 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-09-01 | 2009-09-25 | No | Yes | Health Professional | 42018304 | Yes | Yes | 2009-09-01 | Yes |  | Reuse | 2009
40017908 | 2009-09-01 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, one of the monopolar curved scissors instrument blades broke off and fell into the pt. The broken piece was not retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one scissor blade (left grip) is missing. The blade fractured at the cross section of the cable crimp. The fractured plane is nearly straight and the intact blade does not exhibit damage. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1609301 | 1505510 | 1432855 | 2955842200900320 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-09-02 | 2009-09-25 | No | Yes | Health Professional | 40017908 | Yes | No | 2009-09-01 | Yes | No | Reuse | 2009
40009308 | 2009-08-04 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure the "bolt that holds the grasper" on the endowrist prograsp instrument fell off inside of the patient. The "bolt" was not retrieved. The planned surgical procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip axis pin is missing form the instrument, resulting in loose grips that will not respond intuitively when the instrument is driven on a system. One grip exhibits scratch marks on the inner mating surface of the grip, near the pin hole. No other damage was found. On (b) (6) 2009, additional information regarding the event was provided by the site's interim director of perioperative services, (b) (6). (b) (6) reported that the patient is doing well and did not experience any post operative complications. In addition, to the best of (b) (6)'s knowledge, the patient has not returned to the hospital due to complications related to the event. | ENDOWRIST PROGRASP INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1609262 | 1505052 | 1432401 | 2955842200900303 | 1 | NAY | Manufacturer | User facility | Other Health Care Professional | Initial | 2009-08-20 | 2009-09-24 | No | Yes | Health Professional | 40009308 | Yes | Yes | 2009-08-20 | Yes | No | Reuse | 2009
42018304 | 2009-07-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the permanent cautery hook instrument's energy current conducted away from the instrument tip. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on the yaw pulley and one distal clevis ear indicating an arcing event occurred. The yaw pulley charring is located on the side opposite the conductor wire and directly below the ceramic sleeve. The distal clevis ear that contains charring is also missing a. 180 inch by. 060 inch piece. The conductor wire is not damaged and electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1609307 | 1505053 | 1432402 | 2955842200900223 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-08-14 | 2009-09-24 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer |  | 2009-08-14 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-07-31 | Death | Death | It was reported that 2 days post a da vinci sacrocolpopexy with hysterectomy procedure, the pt developed a peritonitis infection and expired. | The following additional information was provided by the site: the procedure went well; however, the pt had significant omentum adhesions that were difficult to remove. The following morning, a post op examination of the pt found that the pt was doing well; however, later that evening, it was noted that the pt's blood pressure dropped. The 2nd day post op, the pt's condition became worse and the pt expired later that evening. There was no injury to the pt's bowel or bladder, and there was no vaginal separation observed. The pt's development of peritonitis was not related to the robotic procedure. No malfunction of the da vinci surgical system was alleged by the hospital, and the information provided by the site suggests the pt's preexisting medical condition (s) contributed to this event. As of september 17, 2009 no adverse events or system malfunctions have been experienced with site's da vinci surgical system. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1541308 | 1482258 | 1410498 | 2955842200900310 | 1 | NAY | Manufacturer | User facilityCompany Representative | NOT APPLICABLE | Initial | 2009-08-25 | 2009-09-18 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2009-08-25 | Device Not Retu | No | Reuse | 2009
IS2000 A51P8 | 2009-08-28 | Other |  | It was reported that during a da vinci s surgical procedure, the customer experienced a non recoverable system error code #23. With the assistance of an isi technical support engineer (tse) the site powered down the system and cleaned the surgeon side cart (ssc) and patient side cart (psc) fiber cable, however, when the system was restarted system error code #297 appeared. The tse had the site emergency power off the surgeon side cart (ssc) and reset the breaker multiple times, however, the system error code #297 continued to occur. Approximately 1 hour into the procedure, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluding that the system error codes #23 and #297 experienced by the customer were associated with a patient side manipulator (psm) embedded serializer set up joint (essj). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The essj is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The system was repaired by replacing the affected psm's essj. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23 is reported by software to denote communication faults in the low voltage differential signal carrying information about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The essj was returned to isi for failure analysis investigation. Engineering confirmed the customer reported a problem and found the essj did not communicate correctly with the psm or system. The essj inductor was found to be cracked and therefore was replaced to repair the essj. As of september 17, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1773165 | 1487674 | 1415732 | 2955842200900315 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-08-28 | 2009-09-18 | Yes | No | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer |  | 2009-08-28 | Yes |  | Invalid Data | 2009
IS2000 S51P7 | 2009-08-21 | Other |  | It was reported that during a da vinci s hysterectomy procedure, the site experienced unrecoverable system error code #23017 and one of the patient side manipulator (psm) arms led was lit red. With the assistance of an isi technical support engineer (tse), the site power cycled the system, however upon restarting, system error code #25588 occurred. The tse had the site emergency power off the system, cycle the breakers on the surgeon side cart (ssc) and patient side cart (psc), however upon powering up, the system error code #25588 recurred. The tse had the site undrape the affected psm and stow the arm, emergency power off the system, and cycle the ssc and psc breakers, however, the system error codes continued to occur. Approximately 30 minutes into the procedure, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found system error code #23017 and system error code #25588 to be associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state. " system error code #25588 is a sympathetic error and occurs during the self test upon system power up when a loop response test fails. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported a problem and found the psm faulted immediately. The psm axis 2 motor encoder was replaced and axis 2 and axis 4 potentiometers were replaced. As of september 17, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1773198 | 1487676 | 1415734 | 2955842200900304 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-08-21 | 2009-09-18 | Yes | No | Health Professional | IS2000 S51P7 | Device Returned To Manufacturer |  | 2009-08-21 | Yes |  | Invalid Data | 2009
IS3000 A60 P2 | 2009-08-25 | Other |  | It was reported that during a da vinci si surgical procedure, the customer experienced intermittent bipolar and monopolar cautery energy functionality. The isi representative present for the procedure had the site exchange instruments, cautery cables and the generator (esu); however, they were unable to get the cautery function to consistently work. With the assistance of an isi technical support engineer (tse), the site power cycled the system, re-seated all esu related cable connections, and installed an external bipolar foot switch, however, the intermittent energy issue continued to occur. The system had been in use for approximately two hours when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) verified proper cautery functionality with the site's instruments, cautery cables and esus used during the procedure. The isi representative present during the procedure stated that the esu setting was below the isi recommended level. It was also observed that the site had the vision side cart (vsc), esu and two other devices plugged into one power strip, which was plugged into the wall with a single power cord. The plug was not totally secured in the wall outlet, therefore, may not have been well grounded. The fse advised the site to plug esu and vsc directly into wall outlets. The fse concluded that cautery issue experienced by the site was most likely caused by improper settings and system set up. As of september 17, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1773229 | 1487677 | 1415735 | 2955842200900309 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-08-25 | 2009-09-18 | Yes | No | Health Professional | IS3000 A60 P2 |  |  | 2009-08-25 | Yes |  | Invalid Data | 2009
IS2000 A51P8 | 2009-08-28 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure, the site experienced system error #23008 when moving the camera. The site decided to continue with the procedure; however, with the assistance of an isi clinical sales representative (csr), a call was placed to an isi technical support engineer for assistance to recycle power to the system. At that time, the site advised that they had pushed back the camera arm and a member of the surgical staff was physically holding the endoscope and camera head assembly. The site was unable to continue with the procedure. The patient was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineer confirmed the occurrence of system error code 23008. System error code #23017 was also found in the system error logs. Engineering concluded that system error codes #23008 and #23017 were associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. System error code #23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal simulation of the motor. Upon determining these conditions, the safety systems put da vinci s in a "nonrecoverable safe state. " the ecm was returned to isi for evaluation. Engineering confirmed the customer reported problem and found the ecm faulted immediately. The axis 4 assembly motor and potentiometers were replaced. As of september 17, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1773173 | 1487678 | 1415736 | 2955842200900314 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-08-28 | 2009-09-18 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-08-28 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-08-27 | Other | Other | It was reported that during a da vinci surgical procedure, the customer detected a mechanical problem with one of the system arms. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the issue experienced by the customer was associated with a patient manipulator (psm). The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for evaluation. Engineering discovered that a piece of the cannula (lock/bearing mechanism) had broken off and was rattling around the elbow area of the psm. As of september 17, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1773293 | 1487679 | 1415737 | 2955842200900312 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2009-08-27 | 2009-09-18 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2009-08-27 | Yes | No | Reuse | 2009
42017908 | 2009-08-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, one of the blades from the monopolar curved scissors (mcs) instrument broke off and fell into the pt's abdomen. The broken blade was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the left blade was returned detached from the mcs instrument. Engineering observed that the blade had fractured at the cross section of the grip cable crimp, exposing half of the cable crimp. The fracture plane of the blade is nearly straight. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1609421 | 1503650 | 1431056 | 2955842200900308 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2009-08-24 | 2009-09-18 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-08-24 | Yes | No | Reuse | 2009
40018005 | 2009-06-04 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure a tear was observed on the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The tip cover was removed and replaced and the planned surgical procedure was completed. No pt harm, adverse outcome or injury. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the tip over. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 024" x. 025" hole in the silicone, located. 026" above the silicone to sleeve interface. The silicone around the hole exhibits localized melting, indicating an arcing event occurred. The mcs instrument used in conjunction with the tip cover accessory was also returned and evaluated. Engineering observed a char mark on the instrument at the proximal clevis ear. With the tip cover installed, the hole position aligns closely to the char mark. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1601432 | 1503651 | 1431057 | 2955842200900307 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-08-24 | 2009-09-18 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-08-24 | Yes | Yes | Initial | 2009
IS2000 A51P7 | 2009-08-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the electrocautery cable connecting the monopolar curved scissors (mcs) instrument and electrosurgical unit (esu) "caught on fire". The "fire", the size of a candle flame, was located near the connector of the mcs instrument. The fire was contained and the electrocautery cable was removed and replaced. Prior to the fire, the surgical staff experienced inconsistent cautery energy functionality. The system was used for an additional 15-20 minutes when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure due to the patient's size. No patient harm was reported. | The affected cautery cable is not manufactured nor distributed by isi. The cable was purchased by the hospital's surgical services department from an external vendor not associated with isi. Numerous attempts have been made to obtain additional information from the hospital regarding the cable, however, no information has been received to date. The instrument used in conjunction with the cable was returned and evaluated. Engineering found instrument to be within specification and fully functional. No thermal damage was observed on the instrument. The investigation conducted by the field service engineer found no functional/performance issues with the system, with the system passing all verification tests. It was determined that the da vinci s surgical system, instruments or accessories did not malfunction in a way that would cause nor contribute to the reported event. A follow-up mdr will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1603437 | 1503652 | 1431058 | 2955842200900301 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2009-08-20 | 2009-09-18 | No | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-08-20 | Yes | No | Reuse | 2009
42018304 | 2009-08-11 | Malfunction |  | It was reported that during a da vinci s surgical procedure, part of the permanent cautery hook instrument broke off and fell inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional info is received. Per the info provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1609449 | 1503653 | 1431059 | 2955842200900316 | 1 | GEI | Manufacturer | User facility | Nurse | Initial | 2009-08-24 | 2009-09-18 | No | Yes | Health Professional | 42018304 | No | Yes | 2009-08-24 | Device Not Retu | No | Unkown | 2009
40016906 | 2009-08-14 | Malfunction |  | It was reported that while resecting during a da vinci s surgical procedure, a piece of the harmonic curved shears insert fell into the pt. The piece was retrieved and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusions - the insert has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the insert is returned (post-eval) and/or if add'l info is rec'd. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the info provided by the initial reporter, the broken piece was successfully retrieved from the pt. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL ACC | 1609343 | 1503984 | 1431390 | 2955842200900300 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-08-18 | 2009-09-17 | No | Yes | Health Professional | 40016906 | No | No | 2009-08-18 | Device Not Retu | Yes | Unkown | 2009
| 2009-05-22 | Malfunction |  | During a robotically assisted myomectomy the staff was using a robotic instrument called a permanent cautery spatula. Near the end of the case the physician was removing the uterine fibroids and staff noticed a foreign object in the patient. When the object was taken out and explored it looked like it may have broken off the permanent cautery spatula during the procedure. |  | DAVINCI S | ROBOTIC DEVICE SPATULA CAUTERY | 1539625 | 1481261 | 1409510 | 1481261 | 1 | NAY | User Facility |  | Nurse | Initial | 2009-09-16 | 2009-09-16 | No | Yes | Invalid Data |  | Yes | Yes |  |  |  |  | 2009
42017908 | 2009-08-06 | Malfunction |  | It was reported that while dissecting during a da vinci s gynecology procedure, the tip of the monopolar curved scissors instrument broke and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1596785 | 1501661 | 1429088 | 2955842200900288 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-08-11 | 2009-09-16 | No | Yes | Health Professional | 42017908 | No | Yes | 2009-08-11 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-06-30 | Malfunction |  | It was reported that during a da vinci s procedure, the electrosurgical unit (esu) was incorrectly set to fulgrate mode instead of desiccate mode when using the monopolar curved scissors instrument (mcs) with a tip cover accessory installed. The tip cover accessory was removed and replaced to complete the planed procedure. No pt harm, adverse outcome or injury was reported. The customer reported malfunction does not itself, constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal evaluation of the accessory. | The tip cover accessory was returned and evaluated. Engineering observed a hole burnt through the silicone portion of the tip cover. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is -. 035" in diameter and located. 315" above the silicone to sleeve interface. The tip cover also has a mechanical tear at the tip, on the opposite side of the hole. Recommended power settings for a given mode are provided to the customer in the electrosurgical unit (esu) settings and footswitch cables instructions for use. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1597465 | 1502421 | 1429839 | 2955842200900295 | 1 | NAY | Manufacturer | OtherHealth Professional | Other Health Care Professional | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-29 | Yes | Yes | Initial | 2009
40018005 | 2009-01-05 | Malfunction |  | During a da vinci s surgical procedure, arcing was observed to have come from the monopolar curved scissors instrument with tip cover accessory attached. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | Conclusions - the tip cover accessory was returned and evaluated. Engineering found evidence of arcing. The tip cover was returned with a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is oblong, roughly. 030" x. 020, located. 200" above the silicone to sleeve interface. In addition, a mechanical tear was found bisecting one side of the hole. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1596883 | 1502423 | 1429841 | 2955842200900297 | 1 | GEI | Manufacturer | User facilityOther | Other Health Care Professional | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-29 | Yes | Yes | Initial | 2009
40018005 | 2009-01-05 | Malfunction |  | During a da vinci s surgical procedure, arcing was observed to have come from the monopolar curved scissors instrument with tip cover accessory attached. No pt harm, adverse outcome or injury was reported, and the planned surgical procedure was completed. | The tip cover accessory was returned and evaluated. Engineering found evidence of arcing. The tip cover was returned with a large hole burned through the silicone and ultem sleeve. The silicone exhibits localized melting around the hole and has a. 450" long section of charring on the sleeve, indicating an arcing event occurred. The hole is oblong, roughly. 105" x. 075", located at the silicone to sleeve interface. The site returned an mcs instrument with a char mark and material removal on the proximal clevis silicone seal. With the tip cover installed on the instrument, the hole position is approximate to the char mark. The tip cover also exhibits various mechanical tears in the silicone, including adjacent to one edge of the hole and at the tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1597475 | 1502424 | 1429842 | 2955842200900298 | 1 | NAY | Manufacturer | User facilityOther | Other Health Care Professional | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-29 | Yes | Yes | Initial | 2009
40018005 | 2009-01-09 | Malfunction |  | During a da vinci s surgical procedure, arcing was observed to have come from the monopolar curved scissors instrument with tip cover accessory attached. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found evidence of arcing. The tip cover was returned with multiple holes burned through the silicone. There is a. 315" long tear in the silicone at the distal tip with multiple small, irregularly shaped holes branched off of the main tear. Silicone exhibits localized melting around the holes, indicating an arcing event occurred. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1596876 | 1502425 | 1429843 | 2955842200900299 | 1 | GEI | Manufacturer | User facilityOther | Other Health Care Professional | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-29 | Yes | Yes | Initial | 2009
40018005 | 2009-12-26 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering observed a hole burnt through the silicone portion of the tip cover. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is approx. 510 inches above the silicone to sleeve interface. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1596872 | 1502427 | 1429845 | 2955842200900305 | 1 | NAY | Manufacturer | Other | Other Health Care Professional | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-29 | Yes | Yes | Initial | 2009
40018005 | 2009-02-02 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a collision was observed to have occurred between the maryland bipolar instrument and the monopolar curved scissors instrument with tip cover accessory installed. During this procedure two tears in the tip cover accessory were also observed. One tear was visible throughout the procedure; however, the surgeon continued to use the tip cover accessory. Two hours and 49 minutes into the procedure, a second tear was observed in the tip cover accessory. The damaged tip cover was used to complete the procedure without harm to the pt. The esu setting as 60 watts, dessicate. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found evidence of arcing. The tip cover was returned with what appears to be two adjacent holes burned through the silicone. The silicone exhibits localized melting around the holes, indicating an arcing event occurred. The holes are roughly. 035" and. 015" in diameter, located. 025" and. 040" above the silicone to sleeve interface. The tip cover does not exhibit mechanical tears in the silicone. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The monopolar curved scissors instrument instructions for use (ifu) specifically states: general precautions and warnings: inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1597460 | 1502429 | 1429847 | 2955842200900306 | 1 | NAY | Manufacturer | OtherHealth Professional | NOT APPLICABLE | Initial | 2009-07-29 | 2009-09-16 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-07-29 | Yes | Yes | Invalid Data | 2009
40018009 | 2009-08-20 | Injury | Other | It was reported that during a da vinci s procedure, arcing was observed coming from a hole in the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The arcing resulted in damage to the patient's iliac vein and required repair. The planned surgical procedure was completed with no additional patient harm, adverse outcome or injury. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone, located. 160" above the silicone to sleeve interface. The hole is approx 020" in diameter and the silicone around the hole exhibits localized melting, indicating an arcing event occurred. Engineering also observed a small tear adjacent to an edge of the hole, located on the inside surface of the silicone. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1537980 | 1477730 | 1405993 | 2955842200900302 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2009-08-20 | 2009-09-14 | Yes | Yes | Health Professional | 40018009 | Yes | Yes | 2009-08-20 | Yes | Yes | Initial | 2009
VERSION 8 | 2009-09-01 | Malfunction |  | Robotic scissors were being used in conjunction with the intuitive surgical robot. One tip of the scissors broke off. All fluids from the suction cannisters were drained in an attempt to locate the missing piece with no luck. Performed an x-ray and could not locate the piece by that method either. |  | ENDOWRIST | ROBOTIC DEVICE SCISSORS CURVED | 1536366 | 1476613 | 1404935 | 1476613 | 1 | NAY | User Facility |  |  | Initial | 2009-09-04 | 2009-09-11 | No | Yes | Physician | VERSION 8 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420205 | 2009-08-05 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, visualization had not yet been obtained when the surgeon inadvertently pressed the clutch button, causing the fenestrated bipolar forceps instrument to advance and pierce the patient's large intestine. The damage was repaired and the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No additional patient harm was reported. | Based on the information provided by the isi representative in attendance during the surgical procedure, it was determined that the da vinci si surgical system, instruments or accessories did not malfunction in a way that would cause nor contribute to the patient injury. The case was the surgeon's first procedure using the da vinci si surgical system, and it was determined that the patient's injury was directly caused through user error by the surgeon. As of 2009, no adverse events or system malfunctions have been experienced with the site's da vinci si surgical system. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1526554 | 1461928 | 1391354 | 2955842200900285 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-08-05 | 2009-09-01 | Yes | No | Health Professional | 420205 | No | No | 2009-08-05 | Device Not Retu | No | Unkown | 2009
42020906 | 2009-08-03 | Malfunction |  | It was reported that during a da vinci s av fistula repair procedure, one of the wires of the mega suturecut needle instrument snapped and "wires" fell inside of the patient. The "wires" were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the instrument wires were successfully retrieved from the patient. | MEGA SUTURECUT NEEDLE INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1596036 | 1497180 | 1424986 | 2955842200900283 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-08-04 | 2009-09-01 | No | Yes | Health Professional | 42020906 | No | Yes | 2009-08-04 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-08-03 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, part of the permanent cautery hook instrument wrist broke off and fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved form the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1625467 | 1497183 | 1424989 | 2955842200900290 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-08-07 | 2009-09-01 | No | Yes | Health Professional | 42018304 | No | No | 2009-08-07 | Device Not Retu | No | Unkown | 2009
420179 | 2009-08-27 | Malfunction |  | During the procedure, the monopolar scissor wire broke while the instrument was in the patient. The instrument was removed, and the surgical field was inspected by the physician to ensure there was no retained foreign object. There was no patient injury and the instrument is being returned to the manufacturer. The device had been used 6 times; the instrument is designed for 10 uses. Manufacturer response (as per reporter) for scissor, monopolar, curved da vinci swaiting on arrival of instrument to evaluate. |  | DA VINCI S | ROBOTIC INSTRUMENT SCISSOR MON | 1528167 | 1464228 | 1393159 | 1464228 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Physician | 420179 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420205 | 2009-08-20 | Malfunction |  | The wire on the fenestrated bipolar forcep broke. There was no reported injury to the patient. The instrument is being returned to the manufacturer for evaluation. Manufacturer response (as per reporter) for fenestrated bipolar forceps, da vinci s:will evaluate upon receipt of instrument. |  | DA VINCI S | FORCEPS BIPOLAR FENESTRATED | 1528169 | 1464229 | 1393160 | 1464229 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Invalid Data | 420205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420209 | 2009-08-27 | Malfunction |  | During the procedure, the wire at the end of the needle driver broke. This was the fifth use of this re-usable instrument, which is supposed to last for ten uses. There was no patient harm and the instrument is being returned to the manufacturer. |  | ENDOWRIST | ROBOTIC INSTRUMENT NEEDLE DRIV | 1529109 | 1465263 | 1394054 | 1465263 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Physician | 420209 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420209 | 2009-08-21 | Malfunction |  | The wire on the mega suture cut needle driver broke intra-operatively. There was no patient harm reported, and the device is being returned to the manufacturer for evaluation. This was the first time the instrument had been used. The instrument is designed to last for 10 uses. Manufacturer response (as per reporter) for mega suture cut needle driver, da vinciwill evaluate when product received. |  | ENDOWRIST 8MM | ROBOTIC INSTRUMENT NEEDLE DRIV | 1529110 | 1465264 | 1394055 | 1465264 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Physician | 420209 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420205 | 2009-08-19 | Malfunction |  | The wire on the da vinci fenestrated bipolar forceps broke. There was no reported injury to the patient. The instrument had been used 6 times; it is designed for 10 uses. The instrument is being returned to the manufacturer for evaluation. Manufacturer response (as per reporter) for fenestrated bipolar forceps, da vinci:will evaluate upon receipt of instrument. |  | DA VINCI | FORCEPS BIPOLAR FENESTRATED | 1529113 | 1465265 | 1394056 | 1465265 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Physician | 420205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420205 | 2009-08-21 | Malfunction |  | The wire on the fenestrated bipolar forceps broke intra-operatively. No harm to the patient was reported. The instrument is being returned to the manufacturer for evaluation. The instrument had been used 6 times; it is designed for 10 uses. Manufacturer response (as per reporter) for forceps, fenestrated, bipolar, da vinciwill evaluate when the instrument is received. |  | DA VINCI | ROBOTIC INSTRUMENT FORCEPS FEN | 1529384 | 1465655 | 1394409 | 1465655 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-08-28 | 2009-08-28 | No | Yes | Physician | 420205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
IS2000 A51P8 | 2009-08-03 | Other | Other | It was reported that during a da vinci s surgical procedure, the insertion axis of the camera arm would not move. The surgical staff undocked the camera arm, undraped it, and re-homed, however, the insertion axis still would not move when pressing the clutch switch on top of the camera arm. The pt was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No pt harm adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the issue experienced by the customer was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the pt side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The ecm was returned to isi for eval. Engineering discovered that the lower part of the insertion axis carriage rails, thus preventing the carriage from moving downward. As of 08/27/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1546710 | 1466297 | 1395046 | 2955842200900280 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-08-03 | 2009-08-28 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2009-08-03 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-08-05 | Other | Other | It was reported that at the beginning of a da vinci s prostatectomy procedure, the site experienced system error code #23008. The site recycled power to the system and back drove the left master tool manipulator (mtml), however, the system fault recurred. The pt was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 was associated with the left master tool manipulator (mtml) gimble. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtm gimble. The system alarm (system generate fault code) functioned as designed and there was no injury to the pt. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "nonrecoverable safe state". The mtml gimble was returned to isi for failure analysis investigation. Engineering eval found the potentiometer produced a noisy signal, likely generating the system error code experienced by the customer. As an add'l precaution, the motor and encoder were also replaced. As of 08/27/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1546746 | 1466298 | 1395047 | 2955842200900286 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-08-05 | 2009-08-28 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-08-05 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-08-03 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced non-recoverable system error code #23008. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | During the investigation conducted by the field service engineer (fse), a review of the system error logs was performed and the occurrence of system error code #23008 was confirmed. The fse determined that system error code #23008 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers, which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeons' left hand (mtml) and one to his right (mtmt). The system was repaired by replacing the affected mtml. The system alarm (alarm generated fault code) functioned as designed and there was no injury to the pt. System error code #23008, appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "nonrecoverable safe state". The mtml was returned to isi for eval, however, engineering unable to replicate the customer reported failure mode. Based on the system error logs, the gimbal hall effect sensor was replaced as a precaution. As 08/27/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1546767 | 1466299 | 1395048 | 2955842200900271 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2009-08-03 | 2009-08-28 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-08-03 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-08-05 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, a cable was torn on the right master tool manipulator (mtmr). A system error #23008 also occurred. The site restarted the system, however, the system was not able to home, due to the damaged mtmr. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers, which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed, and there was no injury to the pt. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci s in a "nonrecoverable safe state". The mtmr was returned to isi for eval and engineering confirmed that the customer reported failure mode was caused by a broken cable. As of 08/27/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1546874 | 1466300 | 1395049 | 2955842200900287 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-08-05 | 2009-08-28 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-08-05 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-08-07 | Other | Other | It was reported during a da vinci prostatectomy procedure, the site experienced the jaws of instruments not responding in a pt side manipulator (psm). The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure without calling intuitive surgical for technical support. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the system problem was related to a setup issue. An in service on proper system set up has been performed, and the field service engineer monitored for two weeks without further problem. As of 08/28/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1546802 | 1466301 | 1395050 | 2955842200900291 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-08-07 | 2009-08-28 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2009-08-07 | Yes | No | Reuse | 2009
IS3000 A51P8 | 2009-08-06 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the endoscopic camera was removed for cleaning. Upon re-inserting the camera into the pt, the surgeon noticed abnormal bleeding in the right iliac area. The camera was immediately removed for cleaning a second time however, upon re-inserting the camera into the pt, the bleeding appeared to have worsened. The surgeon decided to convert to traditional open surgical techniques to stop the pt's bleeding and complete the planned procedure. The surgeon found an accessory vein in the right iliac area to have been torn loose. The pt required a blood transfusion and prolonged hospitalization. As of august 18, 2009, the pt was recovering as expected and was in stable condition. | No malfunction of the da vinci si system was alleged by the hospital. The surgeon stated he "could not definitively determine how the pt's injury occurred". The hospital has continued to use the da vinci si system to perform surgery on pts. As of august 27, 2009, no adverse events or system malfunctions have been experienced with the site's da vinci si system, and as of august 27, 2009, no similar instances of this event have been reported to isi. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1798136 | 1466302 | 1395051 | 2955842200900289 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-08-06 | 2009-08-28 | Yes | No | Health Professional | IS3000 A51P8 | No | No | 2009-08-06 | Device Not Retu | No | Reuse | 2009
IS3000 A51P8 | 2009-07-16 | Other | Other | On july 16, 2009, it was reported that during a da vinci si prostatectomy procedure, the site experienced a loss of video in the left eye of the surgeon's console and system error #45311. With the assistance of a technical support engineer, the site swapped the camera head; however, the issue continued and the displayed was now in 2-d. The site powered down the system, cycle all breaker and applied emergency power off to the system and let the system sit for 30 seconds before powering the system back on. This system fault continued; however; the site decided to continue laparoscopically. On august 5, 2009, user facility medwatch report (b)(4) was received with the following information. "event desc: the patient was undergoing a robotic prostatectomy. During the procedure, the 3d camera component of the surgery robot failed resulting in the conversion of the surgery to an open procedure. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no. What was the original intended procedure? robotic. Prostatectomy device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working. ". | The investigation conducted by the field service engineer concluded that the vision issue and system error codes #45311 were associated with the system's personality module video acquisition (pmva). During the investigation, the engineer also replaced the generic cart controller (gcc), embedded serializer for setup-joint (essj) and surgeon head sensors as a preventative action for unrelated issues. The pmva is a printed circuit assembly (pca) which receives the stereo video images from the camera controller and outputs the audio and video for the touchscreen monitor. The gcc is the pca which is the central processor in the patient side cart. The essj is the pca inside the non-motorized setup joint arm that monitors the encoder for each of four joints and their associated backup potentiometers. The surgeon head sensors are located in the stereo viewer in the surgeon console and determines if the system is in use or not. The system was repaired by replacing the affected pmva, gcc, essj, surgeon head sensors. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code #45311 appear when the da vinci si safety system detects a loss of one or both surgeon's video inputs. Upon determining this condition, the safety system puts da vinci si in a "non-recoverable safe state". The system was repaired by replacing the affected svp pca. As of august 27, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1537656 | 1476144 | 1360323 | 2955842200900284 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2009-08-05 | 2009-08-28 | Yes | Yes | Health Professional | IS3000 A51P8 | Yes | Yes | 2009-08-05 | Yes | No | Reuse | 2009
42020504 | 2009-07-24 | Malfunction |  | It was reported that one hour into a da vinci s cystectomy procedure, a piece of wire was found in the patient. The fragment was retrieved and the planned surgical procedure was completed using the same instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed no damage to any cables, conductor wires, or other wrist components when examined under magnification. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1593029 | 1494777 | 1422605 | 2955842200900274 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-29 | 2009-08-28 | No | Yes | Health Professional | 42020504 | Yes | Yes | 2009-07-29 | Yes | No | Reuse | 2009
40018005 | 2009-07-02 | Malfunction |  | It was reported that the monopolar curved scissor (mcs) tip cover accessory was used as a replacement to complete a da vinci s hysterectomy procedure, in which the original mcs tip cover accessory presented arcing during the procedure. (refer to med watch mfr report # 2955842-2009-00245). It was reported that this replacement developed a hole during the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering found that the tip cover has a hole burned through the silicone. The silicone exhibits localized melting around the hole, indicative of arcing. The hole is oblong in shape,. 030" x. 020" and is located. 090" above the silicone to sleeve interface. The tip cover also has one small mechanical tear in the general vicinity of the hole. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1593064 | 1494779 | 1414351 | 2955842200900279 | 1 | NAY | Manufacturer | Other | Physician | Initial | 2009-08-03 | 2009-08-28 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-08-03 | Yes | Yes | Initial | 2009
42018404 | 2009-08-03 | Malfunction |  | It was reported that during a da vinci s assisted lobectomy procedure, arcing and a frayed cable were observed with the permanent cautery spatula instrument. The robotic portion of the procedure was successfully completed with no report of patient harm, adverse outcome or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring and localized melting on the distal clevis hub and in multiple locations on the proximal clevis. The damage indicates arcing events occurred. Engineering also observed that the conductor wire has damaged insulation, which exposes the bare wire near the wire hole in the proximal clevis. One of the grip/yaw cables is frayed near the distal idler pulley, and the cable strands stick out slightly from the yaw pulley, however, the instrument is still able to move in yaw direction. The proximal clevis has a hairline crack in the axial direction, halfway between the ears. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1593044 | 1494781 | 1422609 | 2955842200900281 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-06-17 | 2009-08-28 | No | Yes | Health Professional | 42018404 | Yes | No | 2009-06-17 | Yes | No | Reuse | 2009
42020504 | 2009-05-28 | Malfunction | Other | On july 31, 2009, medwatch uf/importer report (b) (4) was received: "event desc: the da vinci s fenestrated grasper broke while being used in a case. The md continued to use the instrument during the cuff closure per the md request. The broken instrument was retrieved from the surgical field after the cuff closure. The md and the da vinci rep were in the room during the incident. The instrument had been used 9 times. Device usage problem: device failed (e. G broke, couldn't get it to work or stopped working). | On aug 4, 2009, (b) (4) confirmed that pieces of the broken instrument fell into the pt but all were retrieved. Per (b) (4), the instrument was returned to intuitive surgical, however, the tracking # provided pertains to a different instrument type and lot#. Upon further investigation with the isi rep noted as being present for the procedure on (b) (6) 2009, he was not made aware that the instrument broke or that pieces fell into the pt. The isi rep was between 2 cases that day and may not have been present when this specific event occurred. The fenestrated bipolar forceps instrument has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the instrument fragments were retrieved from the pt and the procedure was successfully completed. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1550129 | 1496105 | 1356339 | 2955842200900282 | 1 | GEI | Manufacturer | Other | Other | Initial | 2009-07-31 | 2009-08-28 | No | Yes | Health Professional | 42020504 | No | No | 2009-07-31 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-07-29 | Other | Other | It was reported that the system was powered on and the surgical staff was ready to begin the da vinci s prostatectomy procedure when the system started alarming with system error codes #45310 & #45312 and a vision loss message was displayed. The surgeon console viewer and touchscreen monitor showed vertical color bars, with black and gray color bars across the center. With assistance of an isi technical support engineer (tse) the system was powered down and troubleshooting attempts were performed, however, the same issues recurred when the system was restarted. The pt was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the vision loss and system error codes #45310 and #45312 were associated with the system video processor printed circuit assembly (svp pca) board. The svp pca is in the video processor for the system which creates the display for the high resolution stereo viewer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. System error codes #45310 and #45312 appear when the da vinci s safety system detects a loss of one or both surgeon's video inputs. Upon determining this condition, the safety system put da vinci s in a "non-recoverable safe state". The system was repaired by replacing the affected svp pca. The svp pca was returned to isi for eval, however, engineering was unable to replicate the customer reported failure mode. As of august 20, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1539128 | 1455708 | 1385299 | 2955842200900272 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-07-29 | 2009-08-21 | Yes | No | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-07-29 | Yes | No | Reuse | 2009
IS2000 A5 1P8 | 2009-07-30 | Other | Other | It was reported during a da vinci s prostatectomy procedure, the site experienced system error code #212. With the assistance of an isi technical support engineer the site restarted the system, however the system error code #212 reoccurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #212 was associated with the right master tool manipulator (mtmr) and a multiple servo driver (msd). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr) boards which provides drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the effected mtml and msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety system put the da vinci in a "recoverable safe state". The mtml and msd were returned to isi for failure analysis investigation. The error was traced to the axis 4 encoder assemblies on the mtml. The corresponding motor/encoder assembly was replaced as well as the wiring harness. The msd was also evaluated and found to be within specification with no issues noted. As of august 20, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINICI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1539275 | 1455709 | 1385300 | 2955842200900276 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-30 | 2009-08-21 | Yes | Yes | Health Professional | IS2000 A5 1P8 | Device Returned To Manufacturer |  | 2009-07-30 | Yes |  | Invalid Data | 2009
| 2009-03-25 | Malfunction |  | On july 27, 2009, isi received medwatch report (b) (4) from the fda. The event details are provided below: "pt undergoing robotic prostatectomy, when monopolar shears sparked. Removed from pt and large crushing hole noted in shaft of instrument. Later thought to have been processing error. ". | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1532816 | 1464468 | 1357077 | 2955842200900267 | 1 | GEI | Manufacturer | Other | Other | Initial | 2009-07-27 | 2009-08-21 | No | Yes | Health Professional |  | No | No | 2009-07-27 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-07-28 | Malfunction |  | It was reported that while cutting tissue during a da vinci s prostatectomy procedure, the monopolar curved scissors instrument came apart and the tip fell into the patient. The piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument and detached left blade were returned and evaluated. Per the customer reported complaint, engineering observed that the left blade fractured at the cross section of the grip cable crimp. Half of the cable crimp is exposed and the fracture plane of the blade is nearly straight. No other damage was found. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1590340 | 1494504 | 1422332 | 2955842200900269 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-07-28 | 2009-08-21 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-07-28 | Yes | No | Reuse | 2009
40016906 | 2009-07-27 | Malfunction |  | It was reported that during a da vinci anterior posterior resection procedure, while the surgeon was dissection the colon, the "instrument jaws came apart" and a piece fell into the patient. The surgeon removed all pieces of the instrument from the patient and completed the dissection with a replacement instrument of a different type. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering confirmed that the clamp arm is detached from the insert. The distal end of the shaft has one side of the feature that mates with the clamp arm pins, bent backwards. This damage suggests that overloading at the clamp arm occurred, which likely caused the clamp arm to hyperextend and detach. No other damage was found. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1590343 | 1494505 | 1422333 | 2955842200900270 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-07-27 | 2009-08-21 | No | Yes | Health Professional | 40016906 | Device Returned To Manufacturer | Yes | 2009-07-28 | Yes |  | Invalid Data | 2009
42018304 | 2009-07-29 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the surgeon discovered that a piece of the black casing around the shaft of the permanent cautery hook instrument had broken off and fallen into the patient. The piece was immediately removed from the patient and the planned surgical procedure was completed. The site had difficulty removing the instrument from the cannula due to the breakage, therefore, the cannula and instrument were removed from the patient. No patient harm, adverse outcome or injury was reported. | The permanent cautery hook instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the missing pieces were retrieved from the patient and the procedure was successfully completed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1591189 | 1494507 | 1422337 | 2955842200900273 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-29 | 2009-08-21 | No | Yes | Health Professional | 42018304 |  |  | 2009-07-29 | Device Not Retu |  | Invalid Data | 2009
42018304 | 2009-07-31 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon attempted to move the prograsp forceps instrument, however, the instrument would not respond. It was then noticed that two of the patient side manipulator (psm) arms were externally crossed. The site had difficulty removing the instrument and found that the instrument shaft had bent and that fragments were sticking out from the shaft. The planned surgical procedure was completed at which time the surgeon extended a port incision to remove the instrument and irrigate the patient anatomy. No patient harm, adverse outcome or injury was reported. | The permanent cautery hook instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Any lateral pressure on the instrument during removal may damage, break or disconnect the working tip or bend the shaft. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1588290 | 1494509 | 1422337 | 2955842200900278 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-31 | 2009-08-21 | No | Yes | Health Professional | 42018304 |  |  | 2009-07-29 | Device Not Retu |  | Invalid Data | 2009
42018304 | 2009-06-30 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the permanent cautery hook instrument tip was badly damaged. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed charring on the yaw pulley, distal clevis, and the conductor wire, indicating that an arching event occurred. Most of the damage on the components is located on the side opposite to the conductor wire. The conductor's insulated jacket is melted down, exposing the wires at the instrument wrist. Electrical continuity passed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1588309 | 1494511 | 1422339 | 2955842200900268 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-28 | 2009-08-21 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer |  | 2009-07-28 | Yes |  | Invalid Data | 2009
40018302 | 2009-06-19 | Malfunction |  | It was reported that during a da vinci surgical procedure, the tips of the permanent cautery hook instrument were bent and had wires sticking out. No instrument pieces fell into the patient and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, during internal evaluation of the instrument, engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not observe the tip or hook to be bent. Engineering did observe that the wrist assembly did present charring around the cable crimp pocket area and at the proximal clevis which is indicative of arcing. The tip is dislodged from the distal clevis and is missing the conductor cap. As a result, both grip cables are derailed. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1591170 | 1494513 | 1422341 | 2955842200900277 | 1 | NAY | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2009-07-31 | 2009-08-21 | No | Yes | Health Professional | 40018302 | Yes | No | 2009-07-31 | Yes | No | Reuse | 2009
40018005 | 2009-07-27 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed to have come from the monopolar curved scissors instrument (mcs) with a tip cover accessory installed. The tip cover accessory was replaced to continue with the procedure. After use, the second tip cover was discovered to have a hole and was replaced with a third tip cover to complete the procedure. No additional information provided. The tip covers accessory were discarded by the site. No patient harm, adverse or injury was reported. | The instrument and tip cover accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1588320 | 1494515 | 1422343 | 2955842200900264 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-27 | 2009-08-21 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-07-27 | Device Not Retu | Yes | Initial | 2009
IS2000 A5 1P7 |  | Injury | Disability | In 2008, it was reported that after a successful da vinci s mitra valve repair procedure, the pt developed complications related to a heparin induced thrombocytopenia. As a result, the pt's leg was amputated. No additional info has been provided by this site despite attempts. | Additional investigation conducted with the site's risk management department included a review of records for a healthcare. As of 2009, there have been no reports of malfunction of the da vinci s surgical system related to a case involving a leg ischemia requiring leg amputation. If additional info becomes available, a follow-up mdr will be provided. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1515924 | 1447116 | 1376783 | 2955842200900221 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial |  | 2009-08-17 | Yes | No | Health Professional | IS2000 A5 1P7 | No | Yes |  | No Answer Provi | No | Reuse | 2009
42017908 | 2009-07-02 | Malfunction |  | On 7/27/2009, isi received user facility medwatch (b)(4) reporting the following. "event desc: the monopolar scissor tip broke while it was being used in the operative field. The surgeon was aware of the break and was able to recover the tip. There was no harm to the pt. These instruments are suppose to last for 10 uses; this particular one had been used four times prior. Mfr response (as per reporter) for monopolar scissors, da vinci s. The site had already reported the event to isi on 07/22/2009. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade broken off and taped to the housing. The blade fractured at the cross section of the grip cable crimp and half of the cable crimp is exposed. The fracture plane of the blade is nearly straight and the blades do not exhibit damage at tips. Five uses remaining. No other damage found. Per the case notes, the broken piece was retrieved from the pt and the procedure was successfully completed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1532857 | 1464473 | 1356351 | 2955842200900265 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-22 | 2009-08-14 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-07-22 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-08-10 | Other |  | It was reported that during a da vinci s surgical procedure, while docking to the pt, the site experienced non recoverable system error code #22003. With the assistance of an isi technical support engineer the site back drove the pt side manipulator (psm) robotic joint, and power cycled the system; however, the system fault recurred. The pt was under anesthesia and port incisions had been created when the surgeon made the decision to abort the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code #22003 experienced by the customer was associated with a pt side manipulator (psm) set-up joint (suj). The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The suj supports the instrument and camera arms and are also used to set up the initial positions of the arms. The system was repaired by replacing the suj potentiometer. The system error code #22003 functioned as designed and there was no injury to the pt. The da vinci s safety system determined the position of one or more joints on the specified suj, as measured by that joint's primary sensor, disagreed with the measurement from secondary sensor by more than the specified tolerance. Upon determining these conditions, the safety system put da vinci s in a recoverable safe state. As of (b)(4) 2009, there have been no reported recurrence of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1792707 | 1473419 | 1401804 | 2955842200900275 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-08-10 | 2009-08-13 | No | Yes | Health Professional | IS2000 A51P7 |  |  | 2009-08-10 | Yes |  | Invalid Data | 2009
IS1200 A43P9 | 2009-08-12 | Other | Other | It was reported that during a da vinci procedure the customer experienced recoverable system error code# 20008. The error code continued to recur after fault override was attempted. The surgeon decided to convert to traditional open surgical technique to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #20008 was associated with the right master tool manipulator (mtmr) gimbal. The mtm gimbal is a subsection of the complete mtm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtmr gimbal. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #20008 appears when the da vinci safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a "recoverable safe state". The mtmr gimbal was returned to isi for eval. Engineering discovered that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of sept 3, 2009, there have been no reported recurrence of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1792688 | 1473421 | 1401806 | 2955842200900294 | 1 | NAY | Manufacturer | ForeignDistributor | NOT APPLICABLE | Initial | 2009-08-12 | 2009-08-13 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2009-08-12 | Yes | No | Reuse | 2009
42017908 | 2009-07-27 | Malfunction |  | It was reported that during a da vinci s procedure, the monopolar curved scissors instrument was dull. No instrument pieces fell into the pt and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute a fda reportable event, however, engineering discovered evidence that an arcing event occurred during internal eval of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering tested the instrument on an in-house system and observed that the scissors do not cleanly cut through a. 006" piece of latex. The latex gets snagged at the scissor tips. The scissor blades are not damaged, however, the blade edges are slightly rounded at the tips, which negatively affects cut performance. Engineering also observed cracking and thermal damage on the instrument tube extension. The tube extension is cracked in a couple of places at the proximal end. As a result of the cracks, the tube extension flexes more than usual when the wrist is loaded. The tube extension also has a. 090" x. 170" area with localized melting, located directly below the metal ring. One of the cracks in the tube extension intersects the melted area. The localized melting indicates that an arcing event occurred. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1603446 | 1505065 | 1432414 | 2955842200900296 | 1 | NAY | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2009-08-13 | 2009-08-13 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-08-13 | Yes | No | Reuse | 2009
420172 | 2009-08-12 | Malfunction | Other | It was reported that while completing a da vinci s hysterectomy procedure, while the surgical staff was removing the instruments and cannulae from the pt, a burn was found at one of the cannula port incision sites. The burn was considered superficial and the port incision was sutured close. The pt did not require an add'l procedure or treatment for the burn. No add'l pt harm, adverse outcome or injury was reported. | The instrument used during the surgical procedure has not been returned for eval, however, the isi rep on site during the procedure inspected the instrument and was unable to find any defects. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the pt side manipulator arms to ground, through the cannula accessory to the pt, if the pt is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if add'l info is rec'd. As of (b) (6) 2009, there have been no reported recurrences of the issue at this hosp. | MARYLAND BIOPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1603434 | 1505066 | 1432415 | 2955842200900293 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-08-12 | 2009-08-13 | Yes | Yes | Service Personnel | 420172 |  |  | 2009-08-12 | Device Not Retu |  | Invalid Data | 2009
40017908 | 2009-08-11 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, the jaws of the monopolar curved scissors instrument broke off and fell into the pt. The piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument and detached right blade were returned and evaluated. Per the customer reported complaint, engineering observed that the right blade fractured at the cross section of the grip cable crimp and the cable crimp has dislodged from the grip hub. The left blade has a piece of the tip broken off. Corrosion was exhibited on both grips and on the left blade shearing surface. No other damage was found. Per the info provide by the initial reporter, the broken piece was successfully retrieved from the pt. Per the endowrist instruments instructions for use (ifu) specifically states: general instruments cleaning instruction: do not expose instruments to hydrogen peroxide, bleach, or alkaline based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1603425 | 1505068 | 1432417 | 2955842200900292 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other Health Care Professional | Initial | 2009-08-12 | 2009-08-13 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer |  | 2009-08-12 | Yes |  | Invalid Data | 2009
8 MM | 2009-08-04 | Injury | Required Intervention | During a robotic davinci case the instrument arm grasper broke inside pt, but staying intact. After completion of case a cutdown was performed at the incision site allowing removal of instrument. Op report: instrument through 3rd davinci arm snapped at level of port and given angulation; was difficult to remove through poa. At the end of the case, arm was manipulated and instrument was pulled along with arm after extending incision by 2-3cm. Pelvic cavity was irrigated, fragments from instrument sheath were removed. |  | DA VINCI SURGICAL SYSTEM | PRO GRASP FORCEPS | 1511533 | 1441342 | 1371113 | MW5012442 | 1 | NAY | Voluntary |  | Nurse | Initial | 2009-08-11 | 2009-08-12 | No | Yes | Service Personnel | 8 MM |  | Yes |  |  |  |  | 2009
IS2000 A51 P7 | 2009-07-13 | Injury | Required Intervention | It was reported that three days after a successful da vinci s hysterectomy procedure, the pt was re-admitted to the hosp with a temperature and peritonitis symptoms. Two days after being admitted, a diagnostic laparoscopy procedure was performed, resulting in the discovery of eight sites on the pt's sigmoid colon needing over sewing. Six of the eight sites were identified as thermal burns. No additional pt harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer could not identify a system malfunction related to the pt's injury. The site has been asked to return instruments used during the procedure for evaluation, however, the instruments have not been received as of august 10, 2009. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1513468 | 1443458 | 1373229 | 2955842200900263 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-17 | 2009-08-10 | Yes | No | Health Professional | IS2000 A51 P7 | No | Yes | 2009-07-17 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-07-13 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced unrecoverable system error code #281. With the assistance of an isi technical support engineer (tse), the site restarted the system and encountered system error code #903 during start up. The tse instructed the site power down the system, perform an emergency power off, and reset the circuit breaker multiple times; however, system error code #903 continued. Add'l troubleshooting techniques were performed; however, the issue was not resolved. The pt was under anesthesia and the port placements were made when the surgeon made the decision to abort the planned surgical procedure. No pt harm was reported. | The investigation conducted by field service engineering (fse) concluded that the system error codes #281 and #903 experienced by the customer were associated with the remote arm controller (rac). The rac consists of five printed circuit assembly (pca) boards, which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #281 indicates no response was received from one of the rac pca boards during power-up. System error code #903 was a signal error indicating that the system error log was overflowing with error messages since no communication was being received from the rac pca. The rac was returned to isi for failure analysis. Engineering eval discovered the cable connections were loose, thus generating the error codes experienced by the customer. As of 08/07/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1534785 | 1448917 | 1378580 | 2955842200900258 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-07-13 | 2009-08-10 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-07-13 | Yes | No | Reuse | 2009
42017908 | 2009-07-14 | Malfunction |  | It was reported that during a da vinci s hernia repair procedure the scissor tip of the monopolar curved scissors instrument broke off and fell into the pt. The scissor tip was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type and without significant delay. No pt harm, adverse outcome or injury was reported. | The monopolar curved scissors instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with one scissor blade broken off. The broken scissor was taped to the instrument housing. The blade is fractured at cross section of the grip cable crimp and the half of the cable crimp is exposed. The fracture plane of the blade is straight and the intact blade is chipped in three different locations of the blade edge. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1773081 | 1486658 | 1414726 | 2955842200900261 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-16 | 2009-08-10 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-07-16 | Yes | No | Reuse | 2009
40017908 | 2009-06-29 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the right blade of the monopolar curved scissors instrument became detached and fell into the pt. The piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade broken off. The detached blade was also returned. The blade is fractured at the cross section of the cable crimp. The fractured plane of the blade is straight and the remaining blade tip does not exhibit damage. Electrical continuity passed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1773070 | 1486661 | 1414729 | 2955842200900260 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-07-16 | 2009-08-10 | No | Yes | Health Professional | 40017908 | Yes | Yes | 2009-07-16 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-07-14 | Other | Other | It was reported that during a da vinci s gynecological procedure, the operating room experienced power loss. The image viewed through the surgeon side console (ssc) was black. An isi technical support engineer attempted to provide the site with trouble shooting techniques to resolve the issue; however, the surgeon decided to convert to traditional open surgical techniques to finish the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering confirmed that the vision issue experienced by the customer was due to the hospital's loss of power and not related to a system malfunction. In addition, the customer was able to resolve the vision issue after power cycling the system. As of (b)(4) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2577734 | 2548257 | 2444970 | 2955842200900259 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-07-14 | 2009-08-10 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-07-14 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-07-02 | Death | Death | It was reported that during da vinci s radical nephrectomy procedure performed in 2009, the surgeon accidently ligated and cut the pt's superior mesenteric artery. The procedure was converted to a traditional hand assisted open procedure. The pt expired postoperatively in the same day. | Conclusions: no malfunction of the da vinci s surgical system was alleged by the hospital and the info provided by the site states that the surgeon accidently caused this event. As of july 31, 2009, no adverse events or system malfunctions have been experienced with the site's da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1498830 | 1428554 | 1360521 | 2955842200900249 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-07-06 | 2009-07-31 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2009-07-06 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-07-07 | Other | Other | It was reported that during a da vinci s gynecological procedure, the site experienced system error #212 during system homing. The site attempted to override the fault without success. With the assistance of an isi technical support engineer (tse), the site powered down the system, reset the breaker and emergency powered off the system before exercising the right master tool manipulator (mtmr) through its range of motion. The system was then powered on and the fault reoccurred prior to homing. The troubleshooting steps were repeated without change. The procedure had been underway for approximately 1. 5 hours when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #212 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarms (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety system puts da vinci in a "recoverable safe state". The mtmr was returned to isi for evaluation. Engineering was able to confirm the occurrence of system error code #212, thus confirming the reported failure mode experienced by the customer. The mtmr axis 3 potentiometer, the motor/encoder assembly, and the motor harness were replaced. As of july 30, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1523713 | 1428735 | 1360556 | 2955842200900250 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-07 | 2009-07-31 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-07-07 | Yes | No | Reuse | 2009
40017908 | 2009-06-22 | Malfunction |  | It was reported that during da vinci s surgical procedure, the tip of the monopolar curved scissors instrument came off and fell into the patient. The instrument fragment was retrieved and the procedure continued with a replacement instrument of the same type. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade broken off and missing. The blade is fractured at the cross section of the grip cable crimp exposing half of the cable crimp. The fractured plane of the blade is straight and the blade tips do not exhibit damage. Engineering also noted, the presence of debris around the damaged area. The fragments are white / clear and appear similar to the tip cover accessory material. No other damage was found. Per the case notes, the instrument fragment was retrieved from the patient and the procedure was successfully completed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1564317 | 1485017 | 1413106 | 2955842200900257 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2009-07-10 | 2009-07-31 | No | Yes | Health Professional | 40017908 | Yes | No | 2009-07-10 | Yes | No | Reuse | 2009
40017908 | 2009-06-22 | Malfunction | Other | It was reported that during a da vinci surgical procedure, a piece of the monopolar curved scissors instrument broke off and fell inside of the patient. The piece was retrieved and the planned surgical procedure was completed using a replacement instrument. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade broken off and missing. The blade is fractured at the cross section of the grip cable crimp exposing half of the cable crimp. The fractured plane of the blade is straight and the remaining blade tip does not exhibit damage. Engineering also observed the presence of debris around the damaged area. The fragments are white / clear and similar to the silicone material used in the tip cover accessory. Electrical continuity passed. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1564324 | 1485018 | 1413106 | 2955842200900253 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2009-07-10 | 2009-07-31 | No | Yes | Health Professional | 40017908 | Yes | No | 2009-07-10 | Yes | No | Reuse | 2009
42018304 | 2009-06-18 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the permanent cautery hook instrument worked intermittently. No instrument pieces fell into the patient and no patient harm, adverse outcome or injury was reported. The customer reported malfunction does not itself constitute an fda reportable event, however, during internal evaluation of the instrument engineering discovered evidence that an arcing event occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering testing found that electrical continuity failed. The instrument was disassembled and engineering discovered the conductor wire broken at the interface with the cap coupling tube. The section before the breakage appears to be smashed. Engineering also observed char marks on the instrument distal clevis, on the opposite side of the conductor wire. The char marks indicate an arcing event occurred prior to the conductor wire breaking. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1774578 | 1485019 | 1422341 | 2955842200900252 | 1 | NAY | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2009-07-09 | 2009-07-31 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-07-09 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-07-10 | Malfunction |  | It was reported that during set up for a da vinci s surgical procedure, the surgeon experienced difficulty focusing the image through both eyes of the high resolution stereo viewer simultaneously. The site informed an isi technical support engineer (tse) that they would swap to their backup camera head and call back. The site called back to report that swapping to their back up camera head did not resolve the focusing issue. The tse verified the camera control unit settings and had the site re-perform the white balance, however, the issue persisted. Further troubleshooting actions were performed, however, the focus issue was not resolved. The patient was under anesthesia before the surgical staff made the decision to abort the procedure. No port incisions had been created. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the reported vision issue experienced by the customer was related to the camera heads. The system was repaired by replacing the affected camera heads. The camera heads were returned to isi for evaluation. Engineering evaluation revealed that the first camera head was dropped; thereby causing the focus adjustment to be out of tolerance. Engineering evaluation of the back up camera head revealed that the stage was damaged. As of 07/30/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1774606 | 1485021 | 1413110 | 2955842200900254 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-07-10 | 2009-07-31 | No | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-07-10 | Yes | No | Reuse | 2009
42017908 | 2009-07-02 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, "part of the scissor" of the monopolar curved scissors (mcs) instrument broke off and fell into the patient. The fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Results & conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the blades are not entirely broken off at the cable crimp cross section, however, both blade edges exhibit a small amount of material removal near the midpoint. The damaged sections on each blade acts as a hardstop, thereby preventing the scissors from fully closing. The instrument also failed a latex cut test. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1564189 | 1485022 | 1413111 | 2955842200900251 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-07-08 | 2009-07-31 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-07-08 | Yes | No | Reuse | 2009
40018005 | 2009-07-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure while dissection of the vas and seminal vesicles was being performed, the nurse practitioner assisting in the procedure observed arcing through the mcs tip cover accessory. The surgeon succeeded with the procedure using a replacement tip cover; however, arcing was observed. Please see mdr 2955842-2009-00256 for the replacement tip cover accessory used. The planned surgical procedure was completed with a replacement mcs tip cover accessory and without significant delay. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint engineering observed the tip cover has two holes burned though the silicone, adjacent to one parting line. The silicone exhibits localized melting around the holes and is indicative that multiple arcing events occurred. One hole is oblong in shape and is. 030" x. 015", and is located. 135" above the silicon to sleeve interface. The second hole is located directly above the first hole and is. 015" in diameter and. 170" above the silicone to sleeve interface. The following additional information was provided by the site: the monopolar curved scissors (mcs) instrument used with the affected tip cover can not be identified as there were multiple mcs instruments used during the procedure, however, inspection of the mcs instrument and mcs tip cover accessory was performed prior to use, and the tip cover was installed on the mcs instrument with an installed tool. The application of electrolube was performed after installation of the tip cover accessory. As of (b)(6) 2009, there have been no reported recurrences of the issue at this hospital. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1774619 | 1485023 | 1413112 | 2955842200900255 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-10 | 2009-07-31 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-10 | Yes | No | Initial | 2009
40018005 | 2009-07-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while dissection of the vas and seminal vesicles was being performed, the nurse practitioner assisting in the procedure observed arcing through the mcs tip cover accessory. The mcs tip cover accessory was used as a replacement for the original mcs tip cover accessory that presented arcing. Please see mdr #2955842-2009-00255 concerning the original tip cover accessory used. The planned surgical procedure was completed with a replacement mcs tip cover accessory and without significant delay. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint engineering observed the tip cover has two holes burned through the silicone, adjacent to one parting line. The silicone exhibits localized melting around the holes and is indicative that multiple arcing events occurred. One hole is oblong in shape and is. 040" x. 025", and is located. 375" above the silicone to sleeve interface. The second hole. 015" in diameter and. 285" above the silicone to sleeve interface. In addition, there is. 035" long axial tear located at the distal tip of the silicone. No other damage was found. The following additional information was provided by the site: the monopolar curved scissors (mcs) instrument used with the affected tip cover cannot be identified as there were multiple mcs instruments used during the procedure, however, inspection of the mcs instrument and mcs tip cover accessory was performed prior to use, and the tip cover was installed on the mcs instrument with an installation tool; the application of electrotube was performed after installation of the tip cover accessory. As of (b)(6) 2009, there have been no reported recurrences of the issue at this hospital. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1774566 | 1485024 | 1413112 | 2955842200900256 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-07-10 | 2009-07-31 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-10 | Yes | No | Initial | 2009
MODEL S | 2009-07-16 | Malfunction |  | The patient was undergoing a robotic prostatectomy. During the procedure, the 3d camera component of the surgery robot failed resulting in the conversion of the surgery to an open procedure. This was a new da vinci robot. The vendor determined that the vision acquisition module needed to be replaced. They also stated that head sensor beam on the surgeon console "may be weak. ". |  | DA VINCI S | ROBOTIC SURGICAL CAMERA | 1497012 | 1427732 | 1360323 | 1427732 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-29 | 2009-07-29 | No | Yes | Physician | MODEL S | Yes | No Answer Provided |  |  |  |  | 2009
| 2009-06-18 | Malfunction |  | Upon insertion of robotic camera endoscope, the color in the left eye of the surgeon console was noted to be abnormal, -right eye normal-. Company contacted and after about 20 min of troubleshooting it was determined that the left monitor in surgeon console was bad and would need to be replaced. The surgery was converted from laparoscopic to an open procedure since the robot could not be used. |  | DA VINCI S SURGICAL SYSTEM | ROBOTIC CAMERA | 1499122 | 1429331 | 1361022 | MW5012122 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2009-07-29 | 2009-07-29 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2009
42017908 | 2009-06-29 | Malfunction | Other | It was reported that approx 2 hours into da vinci s hysterectomy procedure, the monopolar curved scissors (mcs) instrument "sleeve" made contact with the pt's bowel serosa, causing a small thermal char injury. The injury to the pt's bowel serosa was repaired with a single vicryl suture. The mcs instrument was used as a replacement for the original monopolar scissors (mcs) instrument that broke. (refer to mdr# 2955842-2009-00246). The planned surgical procedure was completed with a replacement mcs instrument and without significant delay. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | MONOPOLAR SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1565753 | 1486271 | 1414348 | 2955842200900247 | 1 | NAY | Manufacturer | User facilityCompany Representative | Other | Initial | 2009-06-30 | 2009-07-29 | Yes | Yes | Health Professional | 42017908 | No | Yes | 2009-06-30 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-07-01 | Malfunction |  | It was reported that while performing the splenic flexure takedown during a da vinci s low anterior resection procedure, a piece of plastic from the permanent cautery hook instrument broke off and fell inside of the pt. The broken piece was not retrieved. The instrument had been in use for approx 1. 5 hours when the break occurred. The instrument was removed and replaced with another instrument and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. The materials used in the permanent cautery hook instrument have been tested for biocompatibility. All materials were found to be appropriately biocompatible for their intended use. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1774575 | 1486272 | 1414349 | 2955842200900244 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-07-01 | 2009-07-29 | No | Yes | Health Professional | 42018304 | No | No | 2009-07-01 | Device Not Retu | No | Unkown | 2009
40018301 | 2009-06-01 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the leading wires of the permanent cautery hook instrument broke. The planned surgical procedure was completed with a replacement permanent cautery hook instrument. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported failure mode engineering found that the conductor wire is broken near the proximal pulley and the wire insulation exhibits damage as it exits the conductor cap. In addition, the instrument passed electrical continuity testing. Engineering also observed that the proximal clevis and the distal clevis hub exhibit charring which is indicative that arcing occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1774543 | 1486273 | 1414350 | 2955842200900242 | 1 | NAY | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2009-07-01 | 2009-07-29 | No | Yes | Health Professional | 40018301 | Yes | No | 2009-07-01 | Yes | No | Reuse | 2009
40018005 | 2009-07-02 | Malfunction |  | It was reported that midway through a da vinci s hysterectomy procedure, arcing was observed from the "crease" of the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The surgeon proceeded with the procedure using a replacement tip cover accessory. After some time of use, the replacement tip cover developed a hole; however, arcing through the tip cover was not observed. In addition, the surgeon reported that the isi recommended electrical surgical unit settings were used and collision of the mcs instrument with another instrument was not observed. The procedure was completed with a replacement tip cover accessory. No pt harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1565759 | 1486274 | 1414351 | 2955842200900245 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-02 | 2009-07-29 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-07-02 | Device Not Retu | Yes | Unkown | 2009
42017908 | 2009-06-29 | Malfunction |  | It was reported that approx 2 hours into a da vinci s hysterectomy procedure while performing the colpotomy step in the procedure, the surgeon noticed a char mark on the pt's posterior uterus. At this time a break in the monopolar curved scissors (mcs) instrument was observed. The surgeon proceeded with the procedure using a replacement mcs instrument; however, sparks coming from the proximal edge of the "instrument sleeve" made contact with the pt's bowel serosa causing a small thermal char injury. The injury to the pt's bowel serosa was repaired with a single vicryl suture. Please see mdr 2955842-2009-00247 for the replacement mcs instrument used. The planned surgical procedure was completed with a replacement mcs instrument and without significant delay. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1774585 | 1486275 | 1414348 | 2955842200900246 | 1 | NAY | Manufacturer | User facilityCompany Representative | Other | Initial | 2009-06-30 | 2009-07-29 | Yes | Yes | Health Professional | 42017908 | No | Yes | 2009-06-30 | Device Not Retu | No | Unkown | 2009
42017206 | 2009-06-29 | Malfunction |  | It was reported that approximately 30 minutes into a da vinci s prostatectomy surgical procedure a small metal "disc" was observed inside of the pt. The site initially thought that the "disc" was a stapler; however, removal and examination of the "disc" revealed that is was a piece from the maryland bipolar forceps instrument. The site examined the instrument and confirmed that the instrument has a missing piece. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, one of the distal idler pulleys is missing. One side of the idler pin has wear marks that line up with the close grip cable. The instrument also passed electrical continuity testing. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the pt. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1784100 | 1468467 | 1397129 | 2955842200900239 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-06-29 | 2009-07-27 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2009-06-29 | Yes | No | Reuse | 2009
42020504 |  | Malfunction |  | It was reported that prior to use, the fenestrated bipolar forceps instrument was observe to have cables exposed at the tip. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the mechanical drive cables to be intact. No cable strands are frayed. The grips open and close and the wrist pitches appropriately. Engineering also observed one wire to have damaged insulation at the yaw pulley, exposing bare wire. The yaw pulley exhibits localized melting on either side of the wire damage which is indicative or arcing. Electrical continuity passed. No other damage found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1784096 | 1468470 | 1397131 | 2955842200900243 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-07-01 | 2009-07-27 | No | Yes | Health Professional | 42020504 | Yes | No | 2009-07-01 | Yes | No | Reuse | 2009
40018005 | 2009-05-05 | Malfunction | Other | It was reported that during a da vinci prostatectomy procedure the monopolar curved scissors (mcs) instrument did not react to a closure movement. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. The mcs instrument returned to intuitive surgical with a tip cover accessory installed. The tip cover accessory exhibited evidence of arcing, therefore the complaint was upgraded to reportable. | The mcs tip cover accessory was returned and evaluated. Engineering observed a 0. 65 inch by. 050 inch oval shaped hole at the silicone to sleeve interface with a. 400 inch long charred section on the ultem sleeve just below the hole. There is a. 012 inch diameter hole along one parting line, located. 390 inches above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. The tip cover also has various mechanical tears adjacent to and directly above the larger hold, and a. 165 inch long tear at the distal tip. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The mcs instrument returned with the tip cover exhibits a char mark at the proximal clevis. With the tip cover installed the hole is positioned very close to the char mark on the instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1784085 | 1468473 | 1397134 | 2955842200900241 | 1 | NAY | Manufacturer | Other | Other | Initial | 2009-07-01 | 2009-07-27 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-07-01 | Yes | Yes | Initial | 2009
42009307 | 2009-06-29 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed "cable shards" in the pt. The fragments were retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the cables at the wrist under magnification and found no fraying or damage. However, unidentified debris was found on different sections of the instrument cables, pulleys and cable holes at the proximal clevis. The debris is not consistent with the customer reported" cable shards" as it is somewhat clear in color and round in shape. Multiple attempts have been made to obtain additional information from the site including the procedure video, however, at the time of this report no additional information has been received. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. The instrument is fully functional with no trouble found. Per the case notes, the fragments were retrieved from the pt and the procedure was successfully completed. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1784090 | 1468476 | 1397137 | 2955842200900238 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-06-29 | 2009-07-27 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2009-06-29 | Yes | No | Initial | 2009
IS2000 A51P7 | 2009-06-26 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced a non-recoverable system error code #23020. With the assistance of an isi technical support engineer, the site was instructed to disable the affected pt's side manipulator #3, however, the fault recurred. Further troubleshooting actions were performed, however, the issue persisted. The pt had been under anesthesia, and the port incisions had been made when the surgeon decided to abort the planned procedure. No pt harm, adverse outcome, or injury was reported. | The investigation conducted by isi field service engineering concluded that the system error code #23020 experienced by the customer was associated with the clutch button of the pt side manipulator (psm). The pt side manipulator is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. System error code #23020 appears when the da vinci safety system determines one of the switches in an specific manipulator is showing inconsistent signals on two switch leads. The system alarm (system generated fault code) functioned as designed, and there was no injury to the pt. Upon determining this condition, the safety system put da vinci in a "recoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering was able to replicate the customer reported failure mode and concluded that the flat flex cable caused the error. As of 07/23/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1515155 | 1422956 | 1356067 | 2955842200900237 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-26 | 2009-07-24 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-06-26 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-06-26 | Other | Other | It was reported that during a da vinci s gynecological procedure, the monopolar cautery energy function did not work. Prior to contacting an isi technical support engineer (tse), the site reported that they exchanged instruments, cautery cables, and the generator; however, they were unable to get the cautery function to work. The tse instructed the site to install an external footswitch and check the grounding pad. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure, as the pt had been under anesthesia approx 1. 5 hours. No pt harm was reported. | The site called an isi technical support engineer (tse), to report that after the case was completed, their biomedical dept. Was unable to duplicate their reported failure mode as functional testing using test equipment and components from the case revealed that the monopolar foot pedal functioned properly. The tse explained to the site that the da vinci foot peddle only takes the place of the foot peddle of the electrical surgical unit (esu), and concluded that the system may have been improperly set up. The investigation conducted by the field service engineer (fse), found the customer was able to resolve the cautery function issue. The fse concluded that cautery issue experienced by the site was most likely caused by an improperly positioned grounding pad. As of 07/22/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1515170 | 1422957 | 1356068 | 2955842200900236 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-06-26 | 2009-07-24 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-25 | No Answer Provi | No | Reuse | 2009
420207 | 2009-06-29 | Malfunction | Other | It was reported that approx 20 minutes into a da vinci hysterectomy procedure while dissecting the uterine ligaments and while using the tenaculum forceps instrument to manipulate the uterus, the instrument bent at the shaft and shavings from the main tube fell into the pt. The instrument fragments were retrieved and the planned surgical procedure was completed with approx a 20 minutes delay. No pt harm, adverse outcome or injury was reported. | The tenaculum forceps instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the instrument fragments were retrieved from the pt and the procedure was successfully completed. | TENACULUM FORCEPS INSTRUMENTS | ENDOSCOPIC INSTRUMENT | 1792565 | 1468453 | 1397116 | 2955842200900240 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-30 | 2009-07-24 | No | Yes | Health Professional | 420207 | No | No | 2009-06-30 | Device Not Retu | No | Unkown | 2009
| 2009-03-06 | Malfunction |  | Shortly after insertion, the physician reported that the wire controlling movement of the tip was broken. Manufacturer response (as per reporter) for tenaculum forceps, da vinci smost likely broke under tensile loading. |  | DA VINCI S FORCEPS | INSTRUMENT ROBOTIC SURGICAL FO | 1492274 | 1421943 | 1355071 | 1421943 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-22 | 2009-07-22 | No | Yes | Invalid Data |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
| 2009-03-10 | Malfunction |  | The grip close cable broke at the distal idler during use. All pieces were retrieved from surgical field. The instrument had been used 7 times; they are designed to be used 10 times. Manufacturer response (as per reporter) for needle driver, mega suturecutmost likely broke under tensile loading. |  | MEGA SUTURECUT | INSTRUMENT ROBOTIC SURGICAL NE | 1492276 | 1421944 | 1355072 | 1421944 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-22 | 2009-07-22 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
| 2009-05-28 | Malfunction |  | The da vinci s fenestrated grasper broke while being used in a case. The md continued to use the instrument during the cuff closure per the md request. The broken instrument was retrieved from the surgical field after the cuff closure. The md and the da vinci rep were in the room during the incident. The instrument had been used 9 times. |  | DA VINCI S GRASPER | INSTRUMENT ROBOTIC SURGICAL GR | 1493745 | 1423236 | 1356339 | 1423236 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-22 | 2009-07-22 | No | Yes | Invalid Data |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
| 2009-02-28 | Malfunction |  | The sheath on the monopolar scissors fractured. The device was returned to the manufacturer. |  | DA VINCI | INSTRUMENT ROBOTIC SURGICAL MO | 1495267 | 1425300 | 1358371 | 1425300 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-22 | 2009-07-22 | No | Yes | Invalid Data |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
420179 |  | Malfunction |  | During a site visit performed on (b) (6) 2009, an isi representative observed video of a da vinci s hysterectomy procedure with a monopolar curved scissors instrument which arced through the distal end of the main tube and 10 mm above the proximal clevis. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. During review of procedure video, the mcs was observed to have arced. Per phone history notes, no patient harm was reported. | MONOPOLAR CURVER | ELECTROSURGICAL INSTRUMENT | 1546784 | 1465905 | 1394658 | 2955842200900229 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-22 | 2009-07-22 | No | Yes | Health Professional | 420179 | No | No | 2009-06-22 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-06-15 | Malfunction |  | It was reported that while performing the colpotomy during a da vinci s total hysterectomy procedure the surgeon noticed the yellow distal portion of the permanent cautery hook instrument break and fall inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No additional pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1798088 | 1466019 | 1394772 | 2955842200900227 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-22 | 2009-07-22 | No | Yes | Health Professional | 42018304 | No | No | 2009-06-22 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-06-08 | Malfunction |  | It was reported that while dissecting the lymph node during a da vinci s total hysterectomy procedure, the surgeon noticed the yellow distal portion of the permanent cautery hook instrument break and fall inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No additional pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1798060 | 1466021 | 1394774 | 2955842200900228 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-22 | 2009-07-22 | No | Yes | Health Professional | 42018304 | No | No | 2009-06-22 | Device Not Retu | No | Unkown | 2009
420183 |  | Malfunction |  | During a site visit performed on (b)(6), 2009, and isi representative observed video of a da vinci s myomectomy procedure with a permanent cautery hook instrument breaking and fragments falling into the pt. Prior to the instrument breakage and approximately 2:14 minutes into the video, the instrument was observed to have a crack in the proximal clevis. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The permanent cautery hook instrument has not been returned for evaluation, however, while performing a site visit, an isi representative performing a site visit found the site to use off label cleaning methods that are associated with damage to instruments. Upon these finding, the isi representative made several recommendations for proper cleaning and sterilization methods as well as the implementation of a pre-surgery checklist designed to look for damage instruments prior to use. Central processing also received an in-service on how to inspect endowrist instruments during processing to find damaged instruments. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1802980 | 1466023 | 1394776 | 2955842200900231 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-22 | 2009-07-22 | No | Yes | Health Professional | 420183 | No | No | 2009-06-22 | Device Not Retu | No | Unkown | 2009
420183 |  | Malfunction |  | During a site visit performed on (b)(6) 2009, an isi representative observed video of a da vinci s myomectomy procedure with a permanent cautery hook instrument breaking and fragments falling into the pt. Prior to the instrument breakage and approximately 5;30 minutes into the video, the instrument was observed to have one cable off of the pitch pulley. This is an indication of structural damage as the cables are tensioned during manufacturing and cannot jump off a pulley unless damage is present. No add'l info was observed or provided by this site. No pt harm, adverse outcome or injury was reported. | The permanent cautery hook instrument has not been returned for evaluation, however, while performing a site visit, an isi representative investigation found the site to use off label cleaning methods that are associated with damage to instruments. Upon these findings, the isi representative, made several recommendations for proper cleaning and sterilization methods as well as the implementation of a pre-surgery checklist designed to look for damaged instruments prior to use. Central processing also received an in-service on how to inspect the end of endowrist instruments during processing to find damaged instrument during the cleaning and sterilization process. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1803007 | 1466024 | 1394777 | 2955842200900230 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-22 | 2009-07-22 | No | Yes | Health Professional | 420183 | No | No | 2009-06-22 | Device Not Retu | No | Unkown | 2009
HARMONIC CURVED | 2009-05-28 | Malfunction |  | During a robotic supracervical hysterectomy, tip of harmonic shears broke off. |  | HARMONIC CURVED SHEARS INSERT | NONE | 1493248 | 1420999 | 1354130 | MW5012083 | 1 | GEI | Voluntary |  | Other | Initial | 2009-07-21 | 2009-07-21 | No | Yes | Health Professional | HARMONIC CURVED | Yes | No |  |  |  |  | 2009
| 2009-03-25 | Malfunction |  | Patient undergoing robotic prostatectomy when monopolar shears sparked. Removed from patient and large crushing hole noted in shaft of instrument. Later thought to have been processing error. |  | DAVINCI SURGICAL SYSTEM | INSTRUMENT ROBOTIC SURGICAL MO | 1494337 | 1423986 | 1357077 | 1423986 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-06-23 | 2009-07-21 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2009
| 2009-07-02 | Malfunction |  | The monopolar scissor tip broke while it was being used in the operative field. The surgeon was aware of the break and was able to recover the tip. There was no harm to the pt. These instruments are supposed to last for 10 uses; this particular one had been used four times prior. Manufacturer response (as per reporter) for monopolar scissors, da vinci s. |  | HOT SHEARS | INSTRUMENT ROBOTIC SURGICAL SC | 1493751 | 1423248 | 1356351 | 1423248 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-20 | 2009-07-20 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2009
IS2000 A51P7 | 2009-06-23 | Other |  | It was reported that during a da vinci s surgical procedure the customer experienced multiple occurrences of system error code #23008. With the assistance of a field support engineer, the site power cycled the system and swapped the pt side manipulators, however, the issue persisted. The surgeon decided to convert to traditional open surgical techniques to complete the procedure. No pt harm was reported. | The investigation conducted by field service engineering found that system error code #23008 experienced by the site was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "nonrecoverable safe state". As of july 16, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1489547 | 1416440 | 1350724 | 2955842200900234 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2009-06-23 | 2009-07-17 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-06-23 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-06-23 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the site observed liens in the right eye image of the high stereo resolution viewer. With the assistance of an isi technical support engineer the surgical staff reseated the vision cable on both ends and moved the camera cable, however, the vision issue was not resolved. The pt was under anesthesia for approximately 45 minutes when the surgical staff made the decision to convert to open surgical techniques to complete the planned procedure. No pt harm or adverse outcome was reported. | The investigation conducted by the field service engineer (fse) discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the defective camera cable. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of july 16, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1489557 | 1416441 | 1350725 | 2955842200900233 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2009-06-23 | 2009-07-17 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-06-23 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2008-06-23 | Other | Other | It was reported that while setting up for a da vinci gynecological procedure, the site experienced a system error code #20013. With the assistance of an isi technical support engineer (tse) the site restarted the system; however, while draping the system, the system error code #20013 recurred. The issue was resolved after two attempts of moving the endoscopic camera manipulator (ecm) through it's full range of motion and restarting the system. Approximately 2. 5 hours into the procedure, the site called back to report that they were experiencing a nonrecoverable system error code #20008. The tse instructed the site to restart the system, however, the fault recurred during homing. Further troubleshooting actions were performed, however, the issue persisted. The surgeon decided to covert the procedure to traditional open surgical techniques to complete the planned surgery. No pt harm was reported. | The investigation conducted by field service engineering found that system error codes #20013 and #20008 were experienced by the customer as well as system error code #21008 and #21013. The field service engineer concluded that the system error codes #20008, #20013, #21008 and #21013 experienced by the customer were associated with the endoscopic camera manipulator. The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault codes) functioned as designed and there was no injury to the pt. The #20008, #20013, #21008 and #21013 system error codes appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state". The ecm was returned to isi for failure analysis investigation. Engineering evaluation found that the potentiometer gear had broken, thus generating the system errors experienced by the customer. As of july 15, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1489570 | 1416442 | 1350726 | 2955842200900235 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-23 | 2009-07-17 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-06-23 | Yes | No | Reuse | 2009
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the monopolar curved scissors instrument was not recognized by the patient side manipulator (psm) arm. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as instrument recognition testing was successful and revealed that the instrument has 2 uses remaining. Engineering also observed that the instrument presents arcing through the tube extension and the tube extension is broken. The black section of the tube extension below the metal ring exhibits charring and melting. The orange warning band of the tube extension has a small section that exhibits melting where it meets with the black section of the tube extension. In addition, the tube extension is broken circumferentially at the junction between the orange warning band and the black section. The main tube and tube extension are no longer coaxial due to breakage. It was concluded that breakage of the tube extension most likely created a path for arcing to occur. The instrument also passed electrical continuity testing. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1781694 | 1481848 | 1410096 | 2955842200900232 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-06-23 | 2009-07-17 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-06-23 | Yes | No | Reuse | 2009
42020702 | 2009-07-24 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the tenaculum forceps instrument broke and fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument clevis is missing the entire lower section that mates with the main tube interface wall. The distal end of the main tube is broken and missing nearly the entire interface wall. The instrument was returned with the distal and cables cut, and the wrist detached at the proximal clevis, suggesting that the wrist was most likely cut off in order to remove the instrument from the cannula. Per the information provided by the initial reporter, the broken piece was successfully retrieved from the patient. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1570017 | 1490421 | 1418388 | 2955842200900266 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-07-24 | 2009-07-14 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer |  | 2009-07-24 | Yes |  | Invalid Data | 2009
40018005 | 2009-07-09 | Malfunction |  | It was reported that during a da vinci s nephrectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The tip cover accessory was removed and replaced, and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument and tip cover accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted, if the instrument and/or tip cover accessory are returned (post-evaluation) and/or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1569999 | 1490423 | 1418390 | 2955842200900262 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-07-17 | 2009-07-14 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-07-17 | Yes | Yes | Initial | 2009
IS2000 A51P7 | 2009-06-08 | Death | Death | It was reported that post a da vinci s lung biopsy surgical procedure, the pt expired. The site stated that the planned da vinci s procedure was successfully completed, however, the pt was in overall poor health. | In 2009, the site's operating room manager stated that she believes the da vinci s system in no way contributed to the demise of the patient. The hospital has continued to use the da vinci s surgical system to perform surgery on pts. As of the following month, no adverse events have been experienced with the site's system. The same day, the site's risk mgr. Stated she is unable to release the pt's pre-existing medical condition information nor is she able to provide the operating report related to this event. Multiple attempts to obtain additional information have been made, however, no additional information has been provided by the site related to this event. If additional information is received, a follow-up mdr will be submitted. Based on the information provided by the site, it was determined that the da vinci s surgical system did not malfunction in a way that would have led or contributed to the pt's death. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1484840 | 1413107 | 1347944 | 2955842200900217 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-11 | 2009-07-10 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2009-06-11 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-06-12 | Death | Death | It was reported that in 2009, after a successful da vinci s mitral valve repair procedure, the patient experienced difficulties coming off of cardiopulmonary bypass. The pt's heart was said to not be contracting. Two days later, the pt was transferred to another hospital for "ventricular system. " the pt was removed from life support after it was deemed that it would not be possible to perform "ventricular system. " the pt expired in the same day. | No malfunction of the da vinci s system was alleged by the hospital and the information provided indicates that this event is related to the inherent risks associated with surgery and the pt's medical history. The hospital has continued to use the da vinci s system to perform surgery on pts. As of july 10, 2009, no adverse events or system malfunctions have been experienced with the site's da vinci s system and as of july 10, 2009, no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1484832 | 1413108 | 1347945 | 2955842200900226 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-06-14 | 2009-07-10 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2009-06-14 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-06-18 | Other | Other | It was reported that at the start a da vinci s surgical procedure, the left image on the high resolution stereo viewer (hrsv) had a green tint. An technical support engineer (tse) recommended that the site swap the left and right camera cables to determine if the issue could be related to the camera cables, however, swapping the camera cables did not resolve the issue. With the assistance of the tse, further troubleshooting actions were performed, however, the issue persisted. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv provides the video image for the surgeon side console (ssc). During the investigation, the fse also experienced intermittent static in the left image along with the green tint. The system was repaired by replacing the affected hrsv. The hrsv was returned to isi for evaluation. Engineering was able to replicate the fse report of intermittent static and determined that the left hrsv monitor had malfunctioned. The green tint was not replicated. As of july 8, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1502046 | 1413837 | 1361022 | 2955842200900224 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-18 | 2009-07-10 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-06-18 | Yes | No | Reuse | 2009
40016906 | 2009-05-28 | Malfunction | Other | On june 16, 2009, medwatch uf/importer report (b) (4) was received: "event desc" pt undergoing laparoscopic assisted lower anterior resection with davinci robotics and sigmoidoscopy. It was noticed that an instrument of the intuitive davinci robot, a harmonic curved shears, had half of its jaw missing. Prior to noticing the missing piece of the harmonic curved shears, the instrument had been functioning properly and intact. Surgeon unable to locate the piece during the procedure. The pt was ordered to have an x-ray. Upon review of the x-ray, a metallic radiopaque foreign body density was noted in the right upper quadrant of the pt's abdomen. Decision made to not remove the harmonic scalpels piece. Full disclosure made to pt following procedure. Harmonic curved shears was visually inspected by clinical engineering's director on 6/5/2009: the fixed, or immobile, part of the shears has snapped off at the metal post which held the shear piece now missing. The shear piece that pivots is still in place. ". | Conclusions - the harmonic curved shears insert accessory has not been returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". The harmonic curved shears insert accessory consists of stainless steel, teflon and titanium materials. These materials have been tested for bio compatibility in accordance with iso (b) (4) standards for there intended use. If additional info is received, a follow-up mdr will be submitted. | HARMONIC CURVED SHEAR INSERT ACCESSORY | ELECTROSURGICAL INSTRUMENT | 1488834 | 1416362 | 1339160 | 2955842200900222 | 1 | NAY | Manufacturer | Other | NOT APPLICABLE | Initial | 2009-06-16 | 2009-07-10 | Yes | Yes | Health Professional | 40016906 | No | No | 2009-06-16 | Device Not Retu | Yes | Unkown | 2009
42020503 | 2009-06-16 | Malfunction |  | On (b) (6) 2009 the site reported the below event to isi and on july 9, 2009 user facility medwatch #(b) (4) was received. It was reported that while dissecting a fibroid during a da vinci s myomectomy procedure the fenestrated bipolar forceps instrument broke toward the tip of the shaft resulting in splinters falling inside of the pt. The pieces were successfully retrieved and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to be broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as result. The clevis is intact, but the main tube interface wall collapsed and is missing pieces. Per the case notes, the missing pieces were retrieved from the pt without further incident. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1488825 | 1416366 | 1342101 | 2955842200900220 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-16 | 2009-07-10 | No | Yes | Health Professional | 42020503 | Yes | Yes | 2009-06-16 | Yes | No | Unkown | 2009
| 2009-06-23 | Malfunction |  | Robotic small clip applier failed. There are 91 out of 100 lives left on the instrument. |  | SMALL CLIP APPLIER | INSTRUMENT ROBOTIC SURGICAL CL | 1492384 | 1422061 | 1355185 | 1422061 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-07-10 | 2009-07-10 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2009
40018402 | 2009-06-16 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, while dissecting a fibroid, part of the permanent cautery spatula instrument's distal clevis broke off and fell into the pt. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No pt harm, adverse outcome or injury was reported. | The permanent cautery spatula instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the missing pieces were retrieved from the pt and the procedure was successfully completed. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1795269 | 1473446 | 1401831 | 2955842200900219 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-06-16 | 2009-07-10 | No | Yes | Health Professional | 40018402 | No | No | 2009-06-16 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-06-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while the surgeon was transacting the pt's bladder neck, one of the monopolar curved scissors (mcs) instrument blades broke in half and fell into the pt's bladder. The instrument fragment was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. The procedure was lengthened by approx 30 minutes while the fragment was retrieved. No pt harm, adverse outcome or injury was reported. | The monopolar curved scissors instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Per the case notes, the missing piece was retrieved from the pt and the procedure was successfully completed. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1795287 | 1473447 | 1401832 | 2955842200900218 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-16 | 2009-07-10 | No | Yes | Health Professional | 42017908 | No | No | 2009-06-16 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-06-19 | Malfunction |  | It was reported that 1. 5 hours into a da vinci s surgical procedure, while dissecting with cautery, the surgical staff noticed that approx 1-3 cm of the "yellow plastic" between the shaft and tip was missing. An x-ray was performed, however, the missing piece could not be located. The planned surgical procedure was completed with a replacement instrument of the same type. No additional pt harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. Materials used in the intuitive surgical permanent cautery hook instruments have been tested for biocompatibility. All materials were found to be appropriately biocompatible for their intended use. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1795261 | 1473448 | 1401833 | 2955842200900225 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-19 | 2009-07-10 | No | Yes | Health Professional | 42018304 | No | No | 2009-06-19 | Device Not Retu | No | Unkown | 2009
| 2009-07-01 | Malfunction |  | A small piece of plastic that holds the tip of the da vinci robot cautery hook broke off inside the pt during a robotic-assisted laparascopic extensive lysis of abdominal and pelvic adhesions. An x-ray was taken and there was no evidence of plastic was seen. The surgeon searched for the piece and nothing was found. |  |  | ROBOT CAUTERY HOOK | 1482640 | 1412344 | 1347215 | 1412344 | 1 | GEI | User Facility |  |  | Initial | 2009-07-08 | 2009-07-08 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2009
420278 | 2009-05-15 | Injury | OtherRequired Intervention | It was reported that during a da vinci s sigmoidectomy procedure, while the surgeon was utilizing the graptor instrument to retract the patient's bowel, the bowel lost vasculature and was diluted. The damaged bowel was repaired and the planned surgical procedure was completed. No additional patient harm or adverse outcome was reported. | The instrument has not been returned for evaluation, therefore, it cannot be determined if the instrument malfunctioned in a way that would have caused or contributed to the reported patient injury. Multiple attempts to obtain additional information has been made, however, no add'l info has been provided by the site or the surgeon related to this event. If additional information is received, a follow-up mdr will be submitted. | GRAPTOR GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1482411 | 1410656 | 1345684 | 2955842200900216 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-06-10 | 2009-07-02 | Yes | No | Health Professional | 420278 | No | Yes | 2009-06-10 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-06-01 | Other |  | It was reported that while completing a da vinci s left side pyeloplasty-cystoretrograde procedure, the 8. 5mm camera port with the camera installed became dislodged from the pt while attached to the camera arm. The surgical staff immediately turned off the light source, reintroduced the camera port and turned the light source back on. The surgical procedure was completed; however, while closing the port sites, the surgeon observed an 5mm "superficial" burn located above the pt's umbilicos. Bacitracin ointment was applied to treat the burn injury. | On 06/30/2009, the surgeon reported that the pt is in good condition and the pt's hospitalization was not prolonged. The following additional info was provided: the cannula and port placements were sutured in place. There was no loss of insufflation and no pt movement was observed; however, there was movement of the camera. The surgeon reported that the camera pulled out with the port and burned the pt's abdomen; however, the events happened so fast, the surgeon is unsure if contact was made. Based on the info provided, it was determined that movement of the camera likely contributed to the event and that the da vinci s surgical system did not malfunction in a way that would cause or contribute to the pt's injury. As of 06/30/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1501946 | 1410761 | 1345787 | 2955842200900203 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-03 | 2009-07-02 | No | Yes | Health Professional | IS2000 A51P7 | No | Yes | 2009-06-03 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-06-10 | Other | Other | It was reported that during a da vinci s prostatectomy procedure the customer experienced recurring system error code #23013. With the assistance of an isi technical support engineer (tse), the site exercised the right master tool manipulator (mtmr) and was able to override the fault and continue with the case, however, system error #23013 later returned. The surgical procedure had been in progress for 60 minutes when the surgeon decided to abort, close the pt and reschedule the procedure for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code #23013 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #23013 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The mtmr was returned to isi for evaluation. Engineering determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of 07/01/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1501958 | 1410762 | 1345788 | 2955842200900215 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-10 | 2009-07-02 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-06-10 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-06-09 | Other | Other | It was reported that during a da vinci surgical procedure, the surgical staff was unable to focus the camera image from the foot pedal nor from the focus controller. The isi rep onsite attempted to troubleshoot the issue by reseating the camera cable and then replacing the camera cable, however, these actions did not resolve the issue. The pt was under anesthesia for 45 minutes and the port incisions had been made when the surgical staff made the decision to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the focus issue experienced by the customer was associated with the system camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1501926 | 1410763 | 1345789 | 2955842200900214 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-06-09 | 2009-07-02 | Yes | Yes | Service Personnel | IS1200 A43P9 | Yes | No | 2009-06-09 | Yes | No | Reuse | 2009
42018304 | 2009-05-28 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, wires were sticking out of the tip of the permanent cautery hook instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed fraying of one pitch cable at the proximal idlers. The tip can still move, however, movement may not be precise. Engineering also found the conductor wire to present exposed wires at the proximal idlers (opposite side of cable fraying). Charring and burned plastic can be seen at the distal clevis' base, which is in contact with the damaged section of the wire and next to the pitch cable tract. Electrical continuity passed. No other damage found. The charring and burnt plastic indicates an arcing event occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1545359 | 1462268 | 1391688 | 2955842200900213 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-06-09 | 2009-07-02 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-06-09 | Yes | No | Reuse | 2009
42017206 | 2009-05-26 | Malfunction |  | It was reported that during a da vinci s esophagectomy procedure the surgeon applied bipolar cut energy to the maryland bipolar instrument and the tip sparked. The surgical staff noticed that the instrument jaws were burned and replaced the instrument with another instrument of the same type. The surgeon applied bipolar cut energy to the replacement maryland bipolar instrument and the tip sparked. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported, however, the surgical staff reported that the pt appeared to bleed more than previous pt's whom underwent similar procedures. The following medwatch report was sent to the fda for the second maryland bipolar instrument: mfr report # 2955842-2009-00211. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the insulating material at the tip was burned. Both yaw pulleys exhibit charring and localized melting at the grip base. Electrical continuity passed. No other damage found. On 05/27/2009, intuitive surgical's manager of clinical development engineering emailed the surgeon, (b)(6), and highly recommended he discontinue use of the bipolar cut function with intuitive surgical's instruments, as this function has never been tested for use and can cause damage to the instruments. The cut function has the potential to allow for a higher energy current to go through the instrument tips. On 05/30/2009, (b)(6) responded to confirm that he will abstain from using bipolar cut mode in future surgical procedures. He stated that he believes the spark occurred because the instrument jaws had fatty tissue and blood attached to them as well as the instrument was angled upward. | MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT ACC | 1557492 | 1476306 | 1404631 | 2955842200900212 | 1 | NAY | Manufacturer | DistributorForeign | Other | Initial | 2009-06-08 | 2009-07-02 | No | Yes | Health Professional | 42017206 | Yes | No | 2009-06-08 | Yes | Yes | Reuse | 2009
47017206 | 2009-05-26 | Malfunction |  | It was reported that during a da vinci s esophagectomy procedure the surgeon applied bipolar cut energy to the maryland bipolar instrument and the tip sparked. The surgical staff noticed that the instrument jaws were burned and replaced the instrument with another instrument of the same type. The surgeon applied bipolar cut energy to the replacement maryland bipolar instrument and the tip sparked. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported, however, the surgical staff reported that the pt appeared to bleed more than previous pt's whom underwent similar procedures. The following medwatch report was sent to the fda for the second maryland bipolar instrument: mfr report # 2955842-2009-00212. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the insulating material at the tip was burned. Both yaw pulleys exhibit charring and localized melting at the grip base. Electrical continuity passed. No other damage found. On 05/27/2009, intuitive surgical's manager of clinical development engineering emailed the surgeon, (b) (6), and highly recommended he discontinue use of the bipolar cut function with intuitive surgical's instruments, as this function has never been tested for use and can cause damage to the instruments. The cut function has the potential to allow for a higher energy current to go through the instrument tips. On 05/30/2009, dr (b) (6) responded to confirm that he will abstain from using bipolar cut mode in future surgical procedures. He stated that he believes the spark occurred because the instrument jaws had fatty tissue and blood attached to them as well as the instrument was angled upward. | MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT ACC | 1557487 | 1476308 | 1404631 | 2955842200900211 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2009-06-08 | 2009-07-02 | No | Yes | Health Professional | 47017206 | Yes | No | 2009-06-08 | Yes | Yes | Initial | 2009
| 2009-06-19 | Malfunction |  | A robotic thymectomy was underway. The doctor was working at the console. A piece of the robot arm broke off. There was no injury to the pt, but a surgical delay of 45 minutes while the manufacturer's rep went to another hospital to get the needed park. Surgery resumed after part replaced with no other complications. |  | DAVINCI S | NONE | 1478767 | 1408599 | 1343763 | MW5011687 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2009-06-29 | 2009-06-29 | No | Yes | Health Professional |  | No | No |  |  |  |  | 2009
IS2000 A51P7 | 2009-06-02 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon experienced stiffness in the patient side manipulator (psm) arm grip closure and resistance in the master tool manipulator (mtm) grip movement. The surgeon was unable to control instrument closure. The surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with the left master tool manipulator (mtml) gimble. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtm gimble. The mtml gimble was returned to isi for failure analysis investigation. Engineering evaluation of the mtml gimble found a missing spring on axis-8, thus confirming the reported complaint experienced by the customer. As of june 25, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1506205 | 1408773 | 1343935 | 2955842200900201 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-02 | 2009-06-26 | Yes | No | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-06-02 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-06-05 | Other | Other | It was reported that during a da vinci s single vessel small thoracotomy cardiac procedure, the surgeon's image had lines and flickered when the camera was moved. The isi representative onsite for the procedure found the camera head cable connection was loose and tightened the connection, however, the image issue continued. The site obtained their secondary vision cart to continue the procedure. The site received system error code # 23017 on a patient side manipulator (psm) arm and could not over ride the fault. The affected psm was disabled and the system was rebooted, however, the system would not pass the power up sequence and presented system error code # 25588. The system was rebooted multiple times, however, the system error codes continued to occur. Approximately 2. 5 hours into the procedure, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) found system error code #23017 and system error code #25588 were associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The fse also replaced the camera cable due to the vision issue experienced by the customer earlier in the procedure. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected psm and the camera cable. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23017 appears when the da vinci s safety system determines that a motor did not respond as expected and the measured motion did not match the internal stimulation of the motor. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". System error code #25588 is a sympathetic error and occurs during the self test upon system power up when a loop response test fails. The psm and camera cable were returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found the psm faulted immediately. The logic board assembly as well as multiple axis' motor / encoder, potentiometers, and main harness assembly were replaced. The camera cable was tested and found to be working to specification without physical damage or flaws. As of june 25, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1522386 | 1408774 | 1343936 | 2955842200900209 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-05 | 2009-06-26 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-06-05 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-06-01 | Other |  | It was reported that during a da vinci partial nephrectomy procedure, the image through the right and left eye high resolution stereo viewer (hrsv) (high resolution stereo viewer) was dark. The site installed a new light guide, a new illuminator bulb and adjusted the brightness; however, the dark image issue persisted. With the assistance of an isi technical support engineer (tse) the site re-white balanced the cameral control unit (ccu); however, the image through the hrsv remained dark. The tse recommended that the site install a second endoscope and illuminator, and call isi back if the issue remained. The site called back to report that the procedure was converted to traditional open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer found that the customer reported failure mode was attributed to defective light guide cables. Functional testing of both light guide cables revealed that they emitted between 10 to 50 lumens. The cable ends were observed to be discolored and one end is missing the lens. Functional testing of the illuminator bulb found that it is within specification as it emitted 2006 lumens. The light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The system was repaired by replacing the affected light guide cable. As of june 25, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1500385 | 1408775 | 1343937 | 2955842200900200 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2009-06-01 | 2009-06-26 | No | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2009-06-01 | Yes | No | Unkown | 2009
IS1200 A43 P9 | 2009-06-05 | Other | Other | It was reported that during a da vinci thoracoscopy procedure the bifurcated light guide cable melted. The cable was replaced and the replacement bifurcated light guide cable began to melt. Approximately 2 hours into the procedure, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found the two bifurcated light guide cables were brown at the ends from over heating. The lumen light from the illuminator was tested and confirmed within specification. The bifurcated light guide cable provides a path to focus light from the illuminator into the light ports of the endoscope. The system was repaired by replacing the two affected bifurcated light guide cables and the illuminator was replaced as a precaution. The bifurcated light guide cables were returned to isi for additional evaluation. Engineering observed both cables having evidence of thermal damage due to the proximal end connectors exhibiting brown discoloration. The glass tip of one cable is missing and the fiber strands are darkened 0. 285 inches below the face end and one of the fittings at the bifurcated end is missing. The other cable's glass tip shows evidence of melting and there is a gap between the outer metal connector and the metal tube that encases the fiber strands that is partially filled with charred material. The light output of both cables is slightly dimmer than usual. The illuminator was also returned to isi for evaluation, however, at the time of the report failure analysis has not been completed. As of june 25, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1500393 | 1408776 | 1343938 | 2955842200900210 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2009-06-05 | 2009-06-26 | Yes | Yes | Health Professional | IS1200 A43 P9 | Yes | No | 2009-06-05 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-06-04 | Other | Other | It was reported that during the beginning of a da vinci s surgical procedure the site experienced system error #23008 during system homing. With the assistance of an isi technical support engineer (tse), the site attempted to exercise axis 7 on the left master tool manipulator (mtml) without success. The site then powered down the system, exercised the axis, and restarted the system, however, system error #212 appeared. The system was then powered down, the breaker was reset, and the surgeon side console (ssc) had the emergency power off (epo) button applied before restarting the system. At this time system error #23008 re-appeared during homing. The patient had been under anesthesia for approximately 30 minutes with ports placed when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 and #212 were associated with the left master tool manipulator (mtml) gimble. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The gimbal is separated from the rest of the mtm by removing the platform link from the forearm link. The system was repaired by replacing the affected mtml gimble. The system alarms (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinic s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety system puts da vinci in a "recoverable safe state". The mtml gimble was returned to isi for evaluation. Engineering was not able to reproduce the customer reported problem, however, as a precaution, axis 7 potentiometer and the motor/encoder assembly were replaced. As of june 25, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1500396 | 1408777 | 1343939 | 2955842200900205 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-06-04 | 2009-06-26 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-06-04 | Yes | No | Reuse | 2009
42018404 | 2009-05-29 | Malfunction |  | It was reported that in the middle of a da vinci s hysterectomy procedure, the spatula broke off of the permanent cautery spatula instrument and fell inside of the patient. The broken piece was retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional information is received. Per the information provided by the initial reporter, the broken instrument piece was successfully retrieved from the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1804790 | 1459221 | 1388708 | 2955842200900198 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-05-29 | 2009-06-26 | No | Yes | Health Professional | 42018404 | No | No | 2009-05-29 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-05-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was removed and the instrument and tip cover were replaced. Shortly after the replacement tip cover was in use, arcing was observed from the second tip cover. The mcs instrument was removed and the tip cover was replaced. The procedure was completed with a replacement tip cover accessory. No pt harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch report was sent to the fda for the second tip cover accessory: mfr report # 2955842-2009-00208. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1804794 | 1459244 | 1388731 | 2955842200900207 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-05 | 2009-06-26 | No | Yes | Health Professional | 40018005 | No | No | 2009-06-05 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-05-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was removed and the instrument and tip cover were replaced. Shortly after the replacement tip cover was in use, arcing was observed from the second tip cover. The mcs instrument was removed and the tip cover was replaced. The procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The following medwatch report was sent to the fda for the second tip cover accessory. Mfr report # 2955842-2009-00207. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1804713 | 1459245 | 1388731 | 2955842200900208 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-06-05 | 2009-06-26 | No | Yes | Health Professional | 40018005 | No | No | 2009-06-05 | Device Not Retu | Yes | Unkown | 2009
400180 | 2009-06-03 | Malfunction |  | It was reported that during a da vinci s gynecological procedure, the surgeon observed an arc from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The procedure was completed with a replacement tip cover accessory. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1804717 | 1459246 | 1388733 | 2955842200900206 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-06-04 | 2009-06-26 | No | Yes | Health Professional | 400180 | No | Yes | 2009-06-04 | Device Not Retu | Yes | Unkown | 2009
42018904 | 2009-05-18 | Malfunction |  | It was reported that during a da vinci s cardiac surgical procedure, the double fenestrated grasper instrument tip was blocked and not working. The surgical staff found the instrument "plastic cover" had fallen into the patient. The instrument piece was found and removed from the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation a follow-up mdr will be submitted. Per the information provided, the instrument fragment was retrieved from the patient and the procedure was successfully completed without incident. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1804716 | 1459249 | 1388736 | 2955842200900202 | 1 | GEI | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2009-06-03 | 2009-06-26 | No | Yes | Health Professional | 42018904 | No | No | 2009-06-03 | Device Not Retu | No | Unkown | 2009
42018304 | 2009-06-01 | Malfunction |  | It was reported that 1. 5 hours into a da vinci s left side oophorectomy surgical procedure, while using a blunt dissection technique to dissect the patient's left ovary from the pelvic side wall, two pieces of the permanent cautery hook instrument broke off and fell inside the patient. Both pieces were retrieved and the planned surgical procedure was completed with a replacement instrument of the same type. The surgical procedure was lengthened approximately 45 minutes; however, there was no patient harm, adverse outcome or injury reported. | The instrument along with the two broken pieces from the instrument were returned and evaluated. Per the customer reported complaint, engineering observed that the instrument has one distal clevis ear approximately. 284" x 248" broken off. The tip is dislodged as a result of the damage, which caused the conductor wire cap to loosen. There was no damage to the conductor cap observed. The instrument also passed electrical continuity testing. Engineering concluded that breakage of the distal clevis ear was most likely caused by overloading at the tip. Per the information provided, the instrument pieces were retrieved from the patient and the procedure was successfully completed without incident. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 1804621 | 1459251 | 1388737 | 2955842200900199 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-06-01 | 2009-06-26 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-06-01 | Yes | No | Reuse | 2009
42009307 | 2009-06-04 | Malfunction |  | It was reported that during da vinci s hysterectomy procedure, while manipulating the uterus, frayed cables were observed on the endowrist prograsp instrument and cable shard was noticed on the patient's tissue. The shard was removed and the planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported failure mode cannot be confirmed as there was no physical damage found. Examination of the instrument under 20x magnification found that there was no missing material from the main tube, the components (main tube, proximal/distal clevis) did not exhibit any damage and the cable is not frayed. Engineering was unable to link the alleged metal pieces to this instrument. Per the information provided by the initial reporter, the instrument cable shard were successfully retrieved from the patient. An intra operative photographic image of the instrument reveals damage, however, isi contacted the site and verified with the rn robotics coordinator that the instrument in the photographic image was the same instrument. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1804625 | 1459252 | 1388738 | 2955842200900204 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-06-04 | 2009-06-26 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2009-06-04 | Device Not Retu | No | Reuse | 2009
420179 | 2009-05-05 | Injury | OtherRequired Intervention | In 2009, intuitive surgical, inc. , received a user facility medwatch for from reading hospital and medical center for patient. The event details are provided below: "left ureter transected during robotic tah/cystoscopy/stent placement, monopolar curved scissors removed from service. Cautery unit checked and cleared by biomed. ". | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | MONOPOLOR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1474235 | 1406297 | 1341565 | 2955842200900197 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-27 | 2009-06-19 | Yes | Yes | Health Professional | 420179 | No | Yes | 2009-05-27 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-05-27 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced reoccurring system error code #23013. With the assistance of an isi technical support engineer (tse), the site exercised axis 3 on the right master tool manipulator (mtmr) and the system was successfully able to enter following mode, however, approximately 30 seconds later, the system error code #23013 reoccurred. The surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #23013 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one of his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #23013 system error code appears when the da vinci s safety system determines the rate of change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, an axis potentiometer and motor encoder assembly were replaced. As of june 19, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1502090 | 1406316 | 1341583 | 2955842200900176 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-05-27 | 2009-06-19 | Yes | No | Health Professional | IS2000 A51P7 | Yes | No | 2009-05-27 | Yes | No | Reuse | 2009
420205 VER 03 | 2009-06-16 | Malfunction |  | Fenestrated bipolar wrist instrument broke inside patient. Had to remove instrument and trocar together to remove the instrument. Small fragments of graphite were left in patient. |  | FENESTRATED BIPOLAR | WRIST INSTRUMENT | 1473067 | 1406848 | 1342101 | 1406848 | 1 | GEI | User Facility |  |  | Initial | 2009-06-18 | 2009-06-18 | No | Yes | Physician | 420205 VER 03 | No | No Answer Provided |  |  |  |  | 2009
420178 VER 04 | 2009-06-16 | Malfunction |  | Curved scissor wrist instrument was stuck in the right arm of the robot & could not get it to slide out of the sterile adapter. |  | CURVED SCISSOR | ROBOTIC SCISSOR WRIST INSTRUME | 1474121 | 1407363 | 1342574 | 1407363 | 1 | LRW | User Facility |  |  | Initial | 2009-06-18 | 2009-06-18 | No | Yes | Physician | 420178 VER 04 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2009
|  | Other |  | Patient's leg was amputated due to postoperative leg ischemia following a mitral valve operation using davinci system. |  | DAVINCI | NONE | 1470724 | 1403765 | 1339374 | MW5011531 | 1 | NAY | Voluntary |  |  | Initial | 2009-06-17 | 2009-06-17 | No | No | Service Personnel | [7 empty] | 2009
|  | Other |  | Patient's leg was amputated due to postoperative leg ischemia following a mitral valve operation using davinci system. |  | DAVINCI | NONE | 1470715 | 1403767 | 1339376 | MW5011532 | 1 | NAY | Voluntary |  |  | Initial | 2009-06-17 | 2009-06-17 | No | No | Service Personnel | [7 empty] | 2009
|  | Other |  | Patient's leg was amputated due to postoperative leg ischemia following a mitral valve operation using davinci system. |  | DAVINCI | NONE | 1470718 | 1403768 | 1339377 | MW5011533 | 1 | NAY | Voluntary |  |  | Initial | 2009-06-17 | 2009-06-17 | No | No | Service Personnel | [7 empty] | 2009
|  | Other |  | Patient had nerve paralysis and foot drop following a mitral valve operation using davinci system. |  | DAVINCI | NONE | 1470720 | 1403770 | 1339379 | MW5011534 | 1 | NAY | Voluntary |  |  | Initial | 2009-06-17 | 2009-06-17 | No | No | Service Personnel | [7 empty] | 2009
IS2000 A51P7 |  | Injury | Other | It was reported that post a da vinci hysterectomy procedure, the pt developed right shoulder paralysis. Reportedly, prior to starting the planned surgical procedure, the site experienced difficulties placing the pt in the trendelenburg position and the pt experienced labored breathing. | Several attempts have been made to obtain additional info concerning the reported event to determine the root cause. Based on the info provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the pt injury, however, it was determined that the pt's positioning and the pt's weight likely contributed to the event. If additional info is received, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1470073 | 1404063 | 1339663 | 2955842200900196 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-05-22 | 2009-06-17 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes |  | Device Not Retu | No | Reuse | 2009
400003 | 2009-06-16 | Malfunction |  | Clip applier hard to load clip and needed to be replaced. The instrument had 90 uses (out of 100). Instrument was removed from service. |  |  | ROBOTIC CLIP APPLIER | 1471740 | 1406038 | 1341328 | 1406038 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-06-17 | 2009-06-17 | No | Yes | Physician | 400003 | Yes | No Answer Provided |  |  |  |  | 2009
IS2000 A51P7 | 2009-05-26 | Other | Other | It was reported that system error code #23008 occurred while the system was being prepped for a da vinci s myomectomy procedure. The surgical staff attempted to resolve the error by replacing the cable connecting the surgeon console and patient cart, however, system error code #23008 returned prior to docking the system to the patient. The patient was under anesthesia when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder), and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a "recoverable safe state. " the ecm was returned to isi for failure analysis investigation. Engineering determined that an axis motor encoder had malfunctioned, thus generating the system error code experienced by the customer. As of june 11, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1502639 | 1404024 | 1339624 | 2955842200900154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-26 | 2009-06-12 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-05-26 | Yes | No | Reuse | 2009
42017908 | 2009-05-21 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, a piece of the monopolar curved scissors instrument broke off and fell inside of the pt. The piece was retrieved and the planned surgical procedure was completed using a replacement instrument. No pt harm, adverse outcome of injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument to have one scissor blade broken off and taped to the housing. The blade is fractured at the cross section of the grip cable crimp exposing the cable crimp. The fractured plane of the blade is straight and the blade tips do not exhibit damage. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1550912 | 1470697 | 1399264 | 2955842200900195 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-21 | 2009-06-12 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-05-21 | Yes | No | Reuse | 2009
40016906 | 2009-05-12 | Malfunction |  | It was reported that during a da vinci hysterectomy surgical procedure, the tip of the harmonic curved shears insert fell into the pt. The tip was retrieved. The planned surgical procedure was successfully completed. There was no pt harm, adverse outcome or injury reported. | The accessory was returned and evaluated. Per the customer reported complaint, engineering found that the clamp arm is missing from the distal end of the insert. One side of the hinge feature that accepts the clamp arm pins is bent. Engineering concluded that bending of the pins is consistent with overloading of the clamp arm. No other damage was found. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. | HARMONIC CURVED SHEARS INSERT | ELECTROSURGICAL ACCESSORY | 1550917 | 1470698 | 1399265 | 2955842200900193 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2009-05-20 | 2009-06-12 | No | Yes | Health Professional | 40016906 | Yes | No | 2009-05-20 | Yes | Yes | Unkown | 2009
40018005 | 2009-05-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed coming from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed, and no pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. All of the damage is contained within one-half of the silicone, between the two parting lines. The holes are located. 150" to. 410" above the silicone to sleeve interface and their sizes range from under. 010" to. 025" with varying shapes (round and oblong). Multiple holes indicate multiple arcing events occurred. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1551007 | 1470700 | 1399267 | 2955842200900190 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-18 | 2009-06-12 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-05-18 | Yes | Yes | Initial | 2009
400180 | 2009-05-12 | Malfunction |  | It was reported that near the completion of a da vinci s hysterectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The procedure had been underway for approx 1 to 1. 5 hours. The tip cover accessory was replaced to successfully complete the procedure. No pt harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MSC TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1551002 | 1470701 | 1399268 | 2955842200900189 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-18 | 2009-06-12 | No | Yes | Health Professional | 400180 | No | Yes | 2009-05-18 | Device Not Retu | Yes | Unkown | 2009
42017908 | 2009-05-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the monopolar curved scissors (msc) instrument sparked. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was not returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post-evaluation) and/or if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1550922 | 1470702 | 1399269 | 2955842200900194 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-05-20 | 2009-06-12 | No | Yes | Health Professional | 42017908 | No | No | 2009-05-20 | Device Not Retu | No | Unkown | 2009
42009307 | 2009-06-01 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while the pt's ureter was being cut, the endowrist prograsp instrument and another instrument rubbed together and a 1" piece of fiberglass from the shaft of the endowrist prograsp instrument fell into the pt's abdomen. The fiberglass fragment was immediately removed and the same endowrist prograsp instrument was used to complete the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a 1" long x. 040" wide section with material removed on the distal end of the instrument main tube. The damaged area is parallel to the tube axis and has a rough surface finish. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the info provided by the initial reporter, the instrument fragment was successfully retrieved from the pt. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1550998 | 1470703 | 1399270 | 2955842200900169 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-06-01 | 2009-06-12 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2009-06-01 | Yes | No | Reuse | 2009
400180 | 2009-05-11 | Malfunction |  | It was reported that during a da vinci s gynecological procedure, interoperative collisions were observed with the monopolar curved scissors (mcs) instrument with tip cover accessory installed and another instrument. Arcing was then observed from the tip of the mcs and the tip cover accessory was removed and replaced to complete the procedure. No pt harm, adverse outcome or injury was reported. | The tip cover accessory has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the additional info is received, a follow-up mdr will be submitted. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1550992 | 1470704 | 1399271 | 2955842200900191 | 1 | NAY | Manufacturer | Health ProfessionalOther | Other | Initial | 2009-05-19 | 2009-06-12 | Yes | Yes | Health Professional | 400180 | No | Yes | 2009-05-19 | Device Not Retu | Yes | Unkown | 2009
42018304 |  | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the plastic tip of the permanent cautery hook instrument fell inside the pt and was retrieved. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint and found the instrument to be broken at the distal clevis. One of the ears has broken off at its base. A piece of clevis ear (approx. 280" x. 215" in size) and the conductor cap are missing. The yaw pulley clevis is loose because it is no longer constrained by the clevis. The extruded boss feature on the central yaw pulley hub that mates with the slot on the inner surface of the clevis ear is broken off. The piece of yaw pulley that holds the cable crimp in place is also broken, allowing the crimp to partially come out of the yaw pulley. The fracture location of the clevis ear is consistent with overloading the tip with the wrist pitched. No other damage found. Per the case notes, the missing pieces were retrieved from the pt and the procedure was successfully completed. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1550990 | 1470706 | 1399273 | 2955842200900192 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-05-19 | 2009-06-12 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-05-19 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-15 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the customer experienced system error code #252. With the assistance of an isi field service engineer (fse), the surgical staff restarted the system, however, system error code #252 recurred 10 minutes after restart. The isi fse, then instructed the surgical staff to emergency power off the system, however, the system did not automatically perform a self-test after it was restarted. The monitor screen was black and system error code #45040 was displayed when the system suddenly shut down without prompting. The surgical staff restarted the system two more times with the same result. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer, concluded that system error code #252 was associated with the system control processor printed circuit assembly (scp pca) board and system error code #45050 was associated with the system video processor printed circuit assembly (svp pca) board. The scp pca provides real-time servo processing and control, system supervisor and fault monitoring, and external system communication control. The s5vp pca is the video processor for the system, which creates the display for the high-resolution stereo viewer. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code #252 appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the safety system puts da vinci s in a "non-recoverable safe state. " system error code #45050 appears when the svp pca interface is not available. The system was repaired by replacing the affected scp pca and svp pca. | DA VINICI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1466795 | 1399732 | 1335752 | 2955842200900187 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-04-15 | 2009-06-05 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-15 | Yes |  | Reuse | 2009
PS2000 | 2009-05-28 | Malfunction |  | Patient undergoing laparoscopic assisted lower anterior resection with davinci robotics and sigmoidoscopy. It was noticed that an instrument of the intuitive davinci robot, a harmonic curved shears, had half of its jaw missing. Prior to noticing the missing piece of the harmonic curved shears, the instrument had been functioning properly and intact. Surgeon unable to locate the piece during the procedure. The patient was ordered to have an x-ray. Upon review of the x-ray, a metallic radiopaque foreign body density was noted in the right upper quadrant of the patient's abdomen. Decision made to not remove the harmonic scalpel's piece. Full disclosure made to patient following procedure. |  | DAVINCI SURGICAL ROBOT | ROBOTIC SURGICAL DEVICE | 1467769 | 1403415 | 1339160 | 1403415 | 1 | NAY | User Facility |  |  | Initial | 2009-06-03 | 2009-06-05 | No | Yes | Physician | PS2000 | Yes | No Answer Provided |  |  |  |  | 2009
PS2000 | 2009-05-28 | Malfunction |  | Patient undergoing laparoscopic assisted lower anterior resection with davinci robotics and sigmoidoscopy. It was noticed that an instrument of the intuitive davinci robot, a harmonic curved shears, had half of its jaw missing. Prior to noticing the missing piece of the harmonic curved shears, the instrument had been functioning properly and intact. Surgeon unable to locate the piece during the procedure. The patient was ordered to have an x-ray. Upon review of the x-ray, a metallic radiopaque foreign body density was noted in the right upper quadrant of the patient's abdomen. Decision made to not remove the harmonic scalpel's piece. Full disclosure made to patient following procedure. |  | DAVINCI SURGICAL ROBOT | ROBOTIC SURGICAL DEVICE | 1467769 | 1403415 | 1339160 | 1403415 | 1 | NAY | User Facility |  |  | Initial | 2009-06-03 | 2009-06-05 | No | Yes | Physician | PS2000 | Yes | No Answer Provided |  |  |  |  | 2009
40018005 | 2009-05-11 | Other |  | On may 12, 2009, isi received the complaint from the site and on may 26, 2009, isi received medwatch uf / importer report #(b)(4) from the site regarding this event: "on (b)(6) 2009 a (b)(6) female underwent a robotic hysterectomy procedure. During the procedure, a hot shear instrument tip arced to the bowel. The physician examined the patient and did not feel there was any patient injury at that time. The case procedure and was completed without further event. The robotic instrument was changed out and found that the one time use plastic protective tip had a hole. It is unknown if the patient will have any undesired effects from the incident. ". | The tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On may 15, 2009, additional information regarding the event was provided by the robotic service leader. She stated that the procedure was delayed by approx 45 minutes while a general surgeon was called to check the patient's bowel. It was determined that the burn was superficial and no repair was needed. It was also noted that the surgical team observed intraoperative instrument collisions prior to the arcing event between the mcs instrument with the tip cover installed, and a pk dissector forceps instrument. If the additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1472566 | 1403537 | 1399271 | 2955842200900091 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-05-12 | 2009-06-05 | Yes | Yes | Health Professional | 40018005 | No | No | 2009-05-12 | Device Not Retu | Yes | Unkown | 2009
40018005 | 2009-05-13 | Malfunction |  | It was reported that during a da vinci s lower anterior ressection procedure, the isi clinical sales representative (csr) who was in attendance, observed arcing coming from the monopolar curved scissors (mcs) tip cover accessory which was installed on the mcs instrument. This occurred during the starting phase of the mesorectal dissection. The csr advised the surgeon to stop using the instrument and replace the tip cover to continue the procedure. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a very small. 010" x. 005" oval shaped hole in the silicone of the tip cover. The hole is located. 250" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. There is a tear that bisects either side of the hole. The tip cover also has various mechanical tears and gouge marks, roughly 90 degrees from the hole. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1529592 | 1439119 | 1368971 | 2955842200900185 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2009-05-13 | 2009-06-05 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-05-13 | Yes | Yes | Unkown | 2009
42000605 | 2009-05-04 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while suturing the urethral to bladder anastomosis, the surgical staff observed a small piece of fiber glass fragment clinging to the shaft of the large needle driver instrument. The fiber glass fragment was removed immediately with a grasper and the procedure resumed without delay. Upon removal of the large needle driver instrument, the scrub and circulator observed a small white sheared area on the shaft, close to the working end of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a 1. 2" long section with light material removed from the distal end of the main tube. The glass fibers are exposed and the surface finish in the affected area is rough. The damaged area is roughly parallel to tube axis. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1529606 | 1439120 | 1368972 | 2955842200900186 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-05-11 | 2009-06-05 | No | Yes | Health Professional | 42000605 | Yes | Yes | 2009-05-11 | Yes | No | Reuse | 2009
42017908 | 2009-05-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the scrub technician noticed the shaft of the monopolar curved scissors instrument was melted. The melting was observed during an instrument exchange, after the instrument had been in use for 4 hours to cut fibroids off the uterus. The planned surgical procedure was completed with a replacement instrument of the same type and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 421" x. 322" hole burnt through the distal end of the main tube. The hole center is located approximately. 537" from the interface with the tube extension. Localized material removed around the hole indicates an arcing event occurred. The instrument was disassembled to find the tube extension and hypotubes charred in the same location as the hole. Cracking in the distal end of the tube was also observed, starting at the slot features and running along the side of the hole. It was concluded that the cracking in the tube most likely created a path for arcing to occur. The source of the cracking cannot be determined. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1532641 | 1439122 | 1368974 | 2955842200900188 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-05-13 | 2009-06-05 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-05-13 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-05-05 | Other | Other | It was reported that approx half way through a da vinci prostatectomy procedure, the customer experienced system error codes #7 and #153. With phone support provided by an isi field service engineer (fse), the site restarted the system. Emergency power off (epo) was applied multiple times and the system cables were reseated, however, the attempt to resolve the issue was unsuccessful. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the isi field service engineer concluded that system error codes #7 and #153 were associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. The ria pca performs numerous tasks related to the function, control and hardware fault monitoring for a specific arm. There is a ria pca assigned to each system arm. The system was repaired by replacing the affected ria pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #7 is reported by hardware to denote a power supply fault. System error code #153 is reported by hardware to denote a high side switch fault. This indicates that there was a short on the board that triggered this. In the event of these power supply issues, the system records the fault and transitions to a safe state. The ria pca was returned to isi for failure analysis investigation. Engineering was not able to reproduce the customer reported problem, which leads engineering to conclude that the system error codes experienced by this site were likely due to a poorly seated cable. As of may 29, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1465302 | 1400136 | 1336119 | 2955842200900180 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-05-05 | 2009-06-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-05-05 | Yes | No | Reuse | 2009
| 2009-05-21 | Malfunction |  | Small robotic clip applier unable to load clips correctly. Instrument has 69 uses left, more than half of its life. |  | ENDOWRIST SMALL CLIP APPLIER | ACCESSORY ROBOTIC CLIP APPLIER | 1464859 | 1401056 | 1336940 | 1401056 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-05-29 | 2009-05-29 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2009
400180 | 2009-05-05 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The tip cover accessory was replaced to successfully complete the procedure. No pt harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1548169 | 1467717 | 1396408 | 2955842200900183 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-06 | 2009-05-29 | No | Yes | Health Professional | 400180 | No | Yes | 2009-05-06 | Device Not Retu | Yes | Unkown | 2009
400180 | 2009-05-09 | Malfunction |  | It was reported that approx 10 mins into a da vinci s hysterectomy procedure, arcing was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The wrist of the mcs instrument was bent when the spark was observed. The mcs instrument was immediately removed and the tip cover accessory was replaced. The procedure was completed and no pt harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1548163 | 1467718 | 1399268 | 2955842200900182 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-05-09 | 2009-05-29 | No | Yes | Health Professional | 400180 | No | Yes | 2009-05-05 | Device Not Retu | Yes | Unkown | 2009
40017908 | 2009-05-06 | Malfunction |  | It was reported that approx one hour into a da vinci s prostatectomy procedure, the surgeon noticed arcing from the monopolar curved scissors (mcs) instrument. The mcs instrument was removed and replaced with another instrument of the same type. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The monopolar curved scissors (mcs) instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1546956 | 1467719 | 1396410 | 2955842200900184 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-05-06 | 2009-05-29 | No | Yes | Health Professional | 40017908 | No | No | 2009-05-06 | Device Not Retu | No | Unkown | 2009
42017206 | 2009-05-04 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a piece of "wire" from the maryland bipolar forceps instrument fell into the patient and was retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Engineering confirmed the customer reported complaint and found a segment of the pitch up cable to be missing, appearing to have slipped out of the clevis after breakage. Both pitch cables are broken at the distal clevis hub. The pitch down cable has approximately a. 400 inch long cable segment that includes the cable crimp still attached to the distal clevis. It is undetermined as to how both cables broke. No other damage found. Per the case notes, the missing piece of cable was retrieved from the patient and the procedure was successfully completed without incident. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1548151 | 1467720 | 1396411 | 2955842200900181 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-05-05 | 2009-05-29 | Yes | Yes | Health Professional | 42017206 | Yes | No | 2009-05-05 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-27 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, the site experienced a message of right eye video loss. The left eye image in the surgeon side console (ssc) was normal, however, the left eye had color bars. With the assistance of an isi technical support engineer (tse) the site verified the camera control unit (ccu) settings and verified camera cable connections. The site reported that the camera head cable was fully seated but the connector was loose and could be wiggled. The tse had the site power off both ccu's and switched the left and right camera cables, however, upon power up the left eye continued to show only color bars. The tse recommended changing the camera cable. The pt was under anesthesia for approximately 45 minutes when the surgical staff made the decision to reschedule the planned procedure. No pt harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera and processed independently. The system was repaired by replacing the affected camera head. As of may 21, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1493068 | 1420453 | 1353656 | 2955842200900174 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-27 | 2009-05-22 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-27 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-28 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, no text, icons or images were displayed on the left monitor of the surgeon's console. The pt had been under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule it for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected hrsv. The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The hrsv was returned to isi for eval. Engineering discovered that the left side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. As of may 21, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1537789 | 1420454 | 1353657 | 2955842200900178 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-28 | 2009-05-22 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-28 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-04-24 | Other | Other | It was reported that prior to starting a da vinci prostatectomy procedure, double vision or fusion issues were experienced on the surgeon side console (ssc). There were no reported issues during scope alignment. The site replaced the endoscope with no change, and it was also reported that the icons and overlays in the console were not aligned properly. With the assistance of an isi technical support engineer (tse) the site verified all camera head connections and cables on the back of the ssc with no issues found. The site then power cycled the camera control units (ccu) and the vision side cart (vsc) with no change. The site then, removed the scope from the camera head and performed the alignment on an object in the room with no issues. The surgeon placed the scope back on the system and the alignment fusion issue persisted at the ssc. The site was asked to undock and power cycle the system, however, the site made the decision to discontinue troubleshooting. The pt was under anesthesia when the surgical staff made the decision to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer was associated with the camera head. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera heads. As of may 21, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1537836 | 1420455 | 1353658 | 2955842200900171 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-04-24 | 2009-05-22 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-04-24 | Yes | No | Reuse | 2009
IS2000 A5 1P7 | 2009-04-23 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a system error code #23. The site attempted to restart the system multiple times, however, the system would not complete power on self test. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system error code #23 experienced by the customer was associated with a pt side manipulator (psm), an embedded serializer set up joint (essj), and a remote arm controller (rac). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The embedded serializer for set-up joint is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected psm, essj and rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23 is reported by software to denote communication faults in the low voltage differential signal carrying info about the psm. In the event of these communication issues, the system records the fault and transitions to a safe state. The components were returned to isi for failure analysis investigation. The psm registered info error code #30 messages, which is caused by the same communication issues as error code #23 at a lower threshold. This indicates the psm was the likely cause of the error code #23. It was repaired by replacing the espm board which transmits the low voltage differential signal. No trouble was found on the rac. The essj is awaiting rma. As of may 21, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1493150 | 1420457 | 1353660 | 2955842200900170 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-23 | 2009-05-22 | Yes | Yes | Health Professional | IS2000 A5 1P7 | Yes | Yes | 2009-04-23 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-28 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, a black image was experienced on the right side of the surgeon side console (ssc). The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned surgery. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the reported issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The hrsv provides the video image for the ssc. The hrsv was returned to isi for eval. Engineering eval revealed that the hrsv has no power, thus causing right eye vision loss of the hrsv. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1493159 | 1420458 | 1353661 | 2955842200900177 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2009-04-28 | 2009-05-22 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-28 | Yes | No | Reuse | 2009
IS1200 A4 3P9 | 2009-04-27 | Other | Other | It was reported that prior to starting a da vinci surgical procedure, the site experienced a system error code #212 when starting up the system. With the assistance of an isi technical support engineer and a field service engineer, the site reseated the cables at the pt side cart (psc). The site then performed an emergency power off (epo) and exercised all axis on the affected pt side manipulator (psm) arm while the power was off. The system was successfully homed and the site was advised that they could continue with the scheduled procedure. The system error 212 reoccurred as the site was preparing to dock the system to the pt. The pt was under anesthesia and port incisions had been replaced when the surgical staff made the decision to abort the planned surgical procedure and reschedule to a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that system error code #212 was associated with a pt side manipulator (psm) arm and multiple servo driver (msd). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instrument. The system contains two multiple servo driver (msd) pca boards which provides drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected psm and msd. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a "recoverable safe state. " the components were returned to isi for failure analysis investigation. Engineering evaluation could not duplicate the customer reported system error, however, based on the system error code log, the axis potentiometer and motor encoder assemblies were replaced on the psm as a precautionary measure. On the msd board, the offset voltage was found to be out of spec. The board was recalibrated to correct the offset voltage. As of may 21, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1493110 | 1420459 | 1353662 | 2955842200900173 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-27 | 2009-05-22 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2009-04-27 | Yes | No | Reuse | 2009
420006 | 2009-04-29 | Malfunction |  | It was reported that approximately 2 hours into a da vinci s hysterectomy procedure, the large needle driver instrument and the graptor grasping retractor instrument rubbed together causing "scrappings" from the shaft of the instrument's to fall into the patient. The surgeon irrigated and removed all of the instrument fragments and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instruments have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragments were retrieved from the patient and the procedure was successfully completed without incident. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1534402 | 1446227 | 1375896 | 2955842200900179 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-04-29 | 2009-05-22 | No | Yes | Health Professional | 420006 | No | No | 2009-04-29 | Device Not Retu | No | Unkown | 2009
42015703 | 2009-04-16 | Malfunction |  | It was reported that during a da vinci s cardiac revascularization procedure, the blade accessory installed on the snap-fit scalpel instrument detached and fell onto the pt's myocardium. The blade accessory was removed and the planned surgical procedure was completed with another instrument. No pt harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was returned and evaluated. Eval could not be performed on the snap-fit blade accessory used in conjunction with the instrument when the event occurred as the site informed isi that the accessory cannot be located. Per the customer reported complaint, engineering installed in-house snap-fit blade accessories on the instrument and conducted performance tests. The blades remained properly installed after the instrument tip and blades were intentionally collided against other instruments. No physical damage was found and the instrument moved intuitively with full range of motion in all directions. The root cause of the customer reported failure mode cannot be determined. Per the info provided by the initial rptr, the blade accessory was successfully retrieved from the pt. | SNAPFIT SCALPEL INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1534566 | 1446230 | 1375899 | 2955842200900175 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-04-27 | 2009-05-22 | No | Yes | Health Professional | 42015703 | Yes | No | 2008-04-27 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-16 | Death | Death | It was reported that during a da vinci s prostatectomy procedure, the patient was very sick and the case was difficult. The patient was also large and the prostate was located deep inside the patient's pelvis. It was the surgeon's first prostatectomy procedure using the da vinci s surgical system and the procedure was going slowly, with the surgeon fighting bowel the entire case. No system malfunction occurred and all surgical tasks had been performed correctly, however, the surgeon made the decision to convert to traditional open surgical techniques to complete the planned procedure due to the length of time it was taking to complete the procedure. Two days post-op, the patient expired. | Based on the information provided by the surgeon, it was determined that the da vinci s surgical system worked as intended, no system malfunction occurred and the system did not cause or contribute to the patient's demise. A copy of the patient's autopsy report has been requested from the hospital, however, as of 2009, no information has been received regarding the cause of the patient's death. A follow-up mdr will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1453376 | 1387516 | 1323923 | 2955842200900172 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-04-18 | 2009-05-18 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2009-04-18 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-04-22 | Other |  | It was reported that prior to starting a da vinci s bilateral ima mobilization procedure, the site experienced right eye video loss and double vision on the pt side cart (psc) touchscreen monitor. The site was instructed by an isi technical support engineer to change the right camera control unit (ccu) settings; however, the right eye video remained black. The site was instructed to switch the ccu connections from left to right and swap to a backup camera head; however, the issue moved to the left eye. The site swapped back to the original camera head using a back up camera head cable and corrected the right eye video loss issue. While modifying the psc settings the site experienced a double image and digital noise. The pt was under anesthesia with no port placements for approx. One hour when the surgeon made the decision to convert the procedure to traditional open surgical techniques. The planned procedure was completed and no pt harm was reported. | The investigation conducted by field service engineering (fse) found the site corrected the right eye video issue with replacement of the camera head cable; however, the fse found that the touchscreen monitor video transmission menu was set incorrectly. The psc touchscreen monitor allows the pt cart operator to view the operative site. Adjustment of the menu settings resolved the touchscreen monitor issue. As of 05/13/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1465521 | 1400186 | 1336169 | 2955842200900168 | 1 | NAY | Manufacturer | Company Representative |  | Initial | 2009-04-22 | 2009-05-15 | No | Yes | Health Professional | IS2000 A51P7 | No | No | 2009-04-22 | Yes | No | Reuse | 2009
40018005 | 2009-03-20 | Malfunction |  | It was reported that during a da vinci s total hysterectomy procedure, the monopolar curved scissors (mcs) instrument was broken at the wrist; however, no pieces fell into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension dislodged from the main tube and the hypotubes are bent. The entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. A. 200" x. 110" piece of the tube extension is missing. Per the information provided by the initial reporter, no instrument fragments fell into the patient during the surgical procedure. Engineering concluded the instrument damage may be due to excessive side loading. The instrument was returned with a mcs tip cover accessory installed. Engineering observed a. 015' x. 030" oblong hole and a. 015" diameter hole in the silicone. The holes are located. 040'" and. 090" above the silicone to sleeve interface. The silicone exhibits localized melting around the holes, indicating arcing events occurred. The tip cover also has various mechanical tears in the silicone. It cannot be determined whether or not the instrument breakage led to the arcing events. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1801269 | 1461456 | 1390887 | 2955842200900167 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-04-08 | 2009-05-15 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-04-08 | Yes | Yes | Initial | 2009
400180 |  | Malfunction |  | It was reported that in (b)(6) , during a da vinci s hysterectomy procedure, a spark was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was bent at a 90 degree angle when the spark was observed. The electrical surgical unit (esu) cautery setting was slightly higher than the isi recommended esu setting. The procedure was completed and no additional patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1801437 | 1461457 | 1390888 | 2955842200900166 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2009-04-21 | 2009-05-15 | No | Yes | Health Professional | 400180 | No | Yes | 2009-04-21 | Device Not Retu | Yes | Unkown | 2009
42017908 | 2009-04-16 | Malfunction |  | It was reported that during a da vinci s total abdominal hysterectomy procedure, the monopolar curved scissors instrument began to malfunction. The electrocautery cable was replaced and shortly after, smoke was observed coming from the connection site between the instrument and electrocautery cable. The instrument and cable were removed and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a 0. 210" x 0. 100" oblong hole burnt through the distal end of the main tube. The hole center is located approximately 0. 650" above the tube extension. Localized material removed around the hole indicates an arcing event occurred. The instrument was disassembled to find the tube extension and hypotubes charred in the same location as the hole. A crack in the distal end of the tube was observed, with the crack starting at one of the slot features and running to the bottom of the hole. It was concluded that the crack in the tube most likely created a path for arcing to occur. The source of the crack cannot be determined. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1801432 | 1461458 | 1390889 | 2955842200900161 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-04-17 | 2009-05-15 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-04-17 | Yes | No | Reuse | 2009
400180 |  | Malfunction |  | It was reported that in (b)(6) 2009, during a da vinci s hysterectomy procedure, a spark was observed coming from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was bent at a 90 degree angle when the spark was observed. The electrical surgical unit (esu) cautery setting was slightly higher than the isi recommended esu setting. The procedure was completed and no additional patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1801419 | 1461459 | 1390890 | 2955842200900165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-20 | 2009-05-15 | No | Yes | Health Professional | 400180 | No | Yes | 2009-04-20 | Device Not Retu | Yes | Unkown | 2009
42017908 | 2009-04-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, "sparking" was observed coming from the monopolar curved scissors (mcs) instrument at the hand piece. No additional info was provided. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no evidence of "sparking" at the hand piece. The instrument was inspected for signs of arcing at the back end; however, no evidence was found. The housing was removed to find that the conductor wire was not damaged at the back end. The electrical continuity passed. Instrument was returned with a tip cover installed and the tip cover does not exhibit any signs of arcing or mechanical tearing. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1543945 | 1461460 | 1390891 | 2955842200900164 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-20 | 2009-05-15 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-04-20 | Yes | No | Unkown | 2009
IS2000 A51P7 | 2009-04-13 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, no image was displayed on the patient cart (pc) touchscreen and the image displayed on the vision cart (vc) monitor was blurry. The endoscope was replaced, however, the vision on the vc monitor remained blurry. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. No patient harm was reported. | During the investigation conducted by the field service engineer (fse), the fse power cycled the system and the vision on the touchscreen was restored. The fse was unable to replicate the customer reported touchscreen failure mode, however, the touchscreen was replaced as a precaution. Per the customer reported complaint of "blurry" vision on the vc monitor, the fse also evaluated the site's vision system. During the evaluation, the fse discovered debris and water at the end of an endoscope lens. The site was able to clear the debris and water from the endoscope. As of may 7, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1503559 | 1416735 | 1351019 | 2955842200900155 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-04-13 | 2009-05-13 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-13 | Yes | No | Reuse | 2009
| 2009-01-30 | Injury | Required InterventionHospitalization | Patient underwent a robotic assited laparoscopic hysterectomy for uterine fibroids. Eight days post-operatively, she spiked a fever and there was an increase in her white blood count. An abdominal ct scan was suspicious for a small bowel perforation. She underwent an exploratory laparotomy where a foreign body was found. It was believed to be a piece of the sheath from the maryland bipolar forceps used during the hysterectomy. The patient had a perforated jejunum which was resected. It is unclear whether the retained foreign body contributed to the perforation; however, we have been experiencing a number of problems - particularly breakage - with various instruments in the davinci s series. ====================== manufacturer response for forceps, maryland bipolar forceps======================although this product was not available to be returned, several other products which broke were returned. The manufacturer evaluated all the returned instruments and could not discern a pattern. |  | MARYLAND BIPOLAR FORCEPS | ELECTROSURGICAL CUTTING COAGU | 1441916 | 1381029 | 1317540 | 1381029 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-05-11 | 2009-05-11 | Yes | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2009
IS2000 A51P7 | 2009-04-10 | Injury | Other | It was reported that post a da vinci s sacrocolpopexy procedure, the pt developed deep vin thrombosis (dvt). No additional pt harm or adverse outcome was reported. | The surgeon reported that the procedure duration was 2. 5 hours from start to finish and that he was not sure if the steep positioning of the pt during the procedure was related to the dvt she experienced. Based on the info provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the pt injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1440558 | 1378509 | 1315076 | 2955842200900157 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-04-14 | 2009-05-08 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2009-04-13 | Device Not Retu | No | Reuse | 2009
IS1200 A43P9 | 2000-04-17 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the customer experienced non-recoverable system error code #212. With the assistance of an isi technical support engineer (tse), the surgical staff exercised the right master tool manipulator (mtm) and restarted the system several times, however, system error code #212 continued to recur. The pt was under anesthesia and the port incisions had been made when the surgical staff decided to abort the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #212 was associated with the right master tool manipulator and/or the multiple servo driver (msd) pca board. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system contains six multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected mtmr and msd pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a "non-recoverable safe state". The mtmr and msd pca were returned to isi for failure analysis investigation. Engineering eval found no issues with the msd pca, however, during eval of the mtmr, engineering found that the harness wiring had malfunctioned, thus generating the system error experienced by the customer. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1537824 | 1378578 | 1315144 | 2955842200900162 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-17 | 2009-05-08 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-04-17 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-16 | Other |  | It was reported that prior to starting a da vinci s prostatectomy procedure, the message "left eye video loss" was intermittently displayed on the video monitor. The site was instructed by an isi technical support engineer (tse) to wiggle the camera cable; however, the failure mode persisted. The tse then instructed the site to power down the sys, swap the cable connections on the camera and camera control unit (ccu), reseat the vision cable on both ends, and power up the system; however, the message recurred upon power up. The site recovered; however, a few minutes later the issue returned. The site was instructed to swap to a back up camera head and cables; however, this did not resolve the issue. The site powered down and swapped both ends of the vision cable and the system powered up with no issues. The patient was under anesthesia for approximately an additional thirty minutes while the vision issue was being resolved. The planned surgical procedure was completed and no pt harm was reported. | The investigation conducted by an isi field service engineer found that the console vision cable was defective. The console vision cable connects the vision cart to the surgeon side console (ssc) and passes the video signals between the ssc and the vision cart. The system was repaired by replacing the console vision cable. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1492191 | 1378579 | 1315145 | 2955842200900159 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-04-16 | 2009-05-08 | No | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-04-16 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-04-17 | Other |  | It was reported that prior to starting a da vinci surgical procedure, the site was unable to power on the system. The site confirmed the surgeon side console ac power indicator was on, confirming the system was receiving power. The site changed out the power cable, however, the system would not power on. The pt was under anesthesia when the surgeon decided to abort the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the customer reported complaint was related to the primary power supply (pps). The pps is a +230v power supply with battery backup and contains logic and protection circuits to provide monitoring of under voltage, over voltage, over current, ac line status, over temperature, power supply fan faults, and detection of battery operation. The system was repaired by replacing the pps. As of may 7, 2009, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1492206 | 1378580 | 1315146 | 2955842200900163 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2009-04-17 | 2009-05-08 | No | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-04-17 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-04-14 | Other |  | It was reported that half way through da vinci s hysterectomy procedure, the site experienced an error and disabled the affected pt side manipulator (psm) arm. The site was unable to provide isi technical support engineer (tse) with the specific error code. The site also reported they had difficulty moving the second and third psm arms. With the assistance of the tse, the site performed a system over ride and regained control of psm2 and psm3. The site continued with the case and stated they would call back to review the system error logs. Approximately 1. 5 hours later, the site called back to report that due to the system related problems and pt complications, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering found system error code #2307 in the system error logs, and was associated with a pt side manipulator (pms) arm. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The sys was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The system error code#23017 appears when the da vinci s safety sys determined that the motor voltage tracking did not respond as expected and that the measured motion did not match the internal simulation. Upon determining these conditions, the safety system put da vinci s in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering was unable to duplicate system error code #23017, however as a precaution, the psm logic board, axis 4 motor, and axis 1, 2, and 3 motor/encoder assemblies were replaced. As of may 7, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1492115 | 1378581 | 1315147 | 2955842200900158 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-14 | 2009-05-08 | No | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-04-14 | Yes | No | Reuse | 2009
420183 | 2009-04-16 | Malfunction |  | It was reported that 2 hours into a da vinci s myomectomy procedure, the first joint cable on the permanent cautery hook instrument broke off and fell inside of the pt. The surgeon was able to retrieve the broken cable pieces and successfully complete the procedure with a replacement instrument of the same type; however, the procedure was lengthen in duration. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval, a follow-up mdr will be submitted. Per the info provided, the instrument fragment was retrieved from the pt, and the procedure was successfully completed without incident. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1524196 | 1430240 | 1361737 | 2955842200900160 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-16 | 2009-05-08 | No | Yes | Health Professional | 420183 | No | No | 2009-04-16 | Device Not Retu | No | Unkown | 2009
400180 | 2009-01-30 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, arcing was observed from the tip of the monopolar curved scissors (mcs) instrument with tip cover accessory installed. The arcing was reported as coming in contact with the external iliac artery, however, no injury occurred. The procedure was completed with a replacement mcs instrument tip cover accessory and without significant delay. The pt has been seen during follow up visits with no post-operative complications reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1524188 | 1430242 | 1361739 | 2955842200900156 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2009-04-13 | 2009-05-08 | Yes | Yes | Health Professional | 400180 | No | Yes | 2009-04-13 | Device Not Retu | Yes | Unkown | 2009
IS2000 A51P7 | 2009-04-02 | Death | Death | It was reported that during a da vinci s nephrectomy procedure, while the surgeon was performing the hilar dissection technique, the patient began to hemorrhage. The surgeon decided to convert the procedure to traditional open surgical technique. Despite efforts from the surgical staff, the patient hemorrhaging was not able to be controlled and the patient expired. Based on the information provided by the site, it was determined that the da vinci s surgical system did not malfunction and did not contributed to the patient's death. | The attending isi representative stated that the system in no way contributed to the demise of the patient. The hospital has continued to use the da vinci s surgical system to perform surgery on patients. As of 2009, no adverse events have been experienced with the site's system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1435835 | 1373630 | 1310311 | 2955842200900150 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-04-06 | 2009-05-01 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2009-04-06 | Device Not Retu | No | Reuse | 2009
IS2000 A51P7 | 2009-04-08 | Other | Other | It was reported that during a da vinci s mitral valve repair procedure, the message "patient cart not detected" appeared. With the assistance of an isi clinical sales representative (csr) and technical support engineer (tse), the site attempted to emergency power off the system multiple times, however, the message continued to appear. A fiber optic cleaning kit was used to clean the fiber connections on the red cable connecting the surgeon side console (ssc) to the patient side cart (psc) multiple times, however, the issue continued. The site was asked to swap the ends of the red cable, however, the cable was not fully accessible. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) concluded that the "patient cart not detected" message experienced by the customer was associated with the red system cable. The red system cable connects the surgeon console and the patient cart and passes video, audio, and data during system operations. The fse discovered that the cable fibers were dirty, thereby, preventing the psc and ssc to connect properly. The system was repaired by replacing the affected cable. As of april 30, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1491780 | 1373670 | 1310351 | 2955842200900151 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-08 | 2009-05-01 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-04-08 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2008-04-03 | Other | Other | It was reported that two hours into a da vinci s sacrocolpopexy w/hysterectomy procedure, the system's surgeon side console (ssc) circuit breaker tripped. An isi technical support engineer (tse) attempted to troubleshoot the issue and had the customer reboot the system, however, the ssc did not power back on. The planned surgical procedure was completed using the site's spare system and with an approximate 30 minute delay. No patient harm was reported. | The investigation conducted by field service engineering concluded that the system issue experienced by the customer was associated with the primary power supply (pps). The pps is a +115v power supply with a battery backup and contains logic and protection circuits to provide monitoring of under voltage, over voltage, over current, ac line status, over temperature, power supply fan faults, and detection of battery operation. The system was repaired by replacing the affected pps. As of april 30, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1491789 | 1373674 | 1310355 | 2955842200900148 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-04-03 | 2009-05-01 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-03 | Yes | No | Reuse | 2009
42018404 | 2009-04-09 | Malfunction |  | It was reported that during a da vinci s cardiac surgical procedure, a piece of the permanent cautery spatula instrument broke off and fell into the patient. The instrument fragment was retrieved and the planned procedure was completed. No patient harm, adverse outcome or injury were reported. | The permanent cautery spatula instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Per the information provided, the instrument fragment was retrieved from the patient and the procedure was successfully completed without incident. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1521511 | 1427413 | 1360026 | 2955842200900153 | 1 | GEI | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-04-10 | 2009-05-01 | No | Yes | Health Professional | 42018404 | No | Yes | 2009-04-10 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-03-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure with bilateral node dissection, a tear was noticed on the monopolar curved scissors (mcs) tip cover accessory that was installed on the mcs instrument. The tip cover accessory was removed and replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. Additionally, the initial reporter indicated that the electrosurgical unit (esu) used during the procedure was set to desiccate mode and below the recommended maximum power setting for that mode, however, the exact power setting used is unknown. | The replacement tip cover accessory was returned and evaluated. Engineering observed a. 030" x. 020" hole in the silicone, located. 215" above the silicone to sleeve interface. Approximately 30 degrees radial from this hole is a. 040" x. 020" oval shaped hole located. 195" above the silicone to sleeve interface. The silicone around both holes exhibits localized melting, indicating arcing events occurred. Engineering also observed various mechanical tears and scratches on the opposite side of the two holes. The mcs instrument used during the surgical procedure was also returned and evaluated. Engineering discovered a possible char mark at the proximal clevis ear, which closely matches up to the location of one of the holes. | A8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1522534 | 1427414 | 1360027 | 2955842200900149 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-04-06 | 2009-05-01 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-04-06 | Yes | Yes | Initial | 2009
IS12000 A51P7 | 2009-04-09 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the focus controller was not operating properly resulting in the site being unable to focus the camera. With the assistance of an isi clinical service manager (csm), the site was instructed to replace the camera head; however, the issue persisted. An isi field service engineer (fse) was contacted and trouble shooting found that the issue was related to with the camera focus controller. The surgical staff decided to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by field service engineering confirmed that the camera head focus controller did not work. The focus controller adjusts the endoscope lens focus to sharpen the surgical image. The system was repaired by replacing the camera focus controller. As of 04/29/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2642157 | 2605079 | 2501752 | 2955842200900152 | 1 | NAY | Manufacturer | User facilityCompany Representative | BIOMEDICAL ENGINEER | Initial | 2009-04-09 | 2009-05-01 | Yes | Yes | Health Professional | IS12000 A51P7 | Yes | Yes | 2009-04-09 | Yes | No | Reuse | 2009
| 2009-02-11 | Other | Required Intervention | It was reported that when the surgeon applied energy to the monopolar instrument during a da vinci s hysterectomy surgical procedure, the bipolar instrument also became energized. As a result, arcing occurred and the pt received two burns on their iliac artery. The burns did not completely perforate the artery, however, 75% of the artery wall was burnt. A vascular surgeon was brought in to repair the damage. The planned surgical procedure was completed and no add'l pt harm or adverse outcome was reported. It is also reported that at the time of the event, the bipolar instrument was incorrectly plugged into one of the monopolar jacks or the electrosurgical unit (esu), with the monopolar instrument also plugged into the other monopolar jack. | On april 21, 2009, (b) (6) (initial reporter) informed isi that no post-operative complications have been experienced by the pt. Ms. (b) (6) also indicated that the occurrence of arcing was a result of user error as the surgical staff incorrectly plugged the bipolar instrument into one of the monopolar jacks of the esu, with the monopolar instrument plugged into the other monopolar jack. Consequently, the bipolar instrument was also activated when the surgeon applied energy to the monopolar instrument. Ms. (b) (6) provided isi the eval results of the valleylab force fx esu used during the time of the event. Per valleylab's eval, the esu functioned normally and was within specifications. No issues were found with the esu. Additionally, ms. (b) (6) indicated that the bipolar instrument and monopolar instrument being used when the event occurred would not be returning to isi for eval as they were not set aside and put back into circulation. On april 22, 2009, covidien valleylab representative (b) (4) informed isi that he had conducted site in-service training for the valleylab force fx esu the morning of the event, prior to surgical procedure being performed. Mr. (b) (4) also indicated that the esu jacks are clearly labelled "bipolar", "monopolar 1", and "monopolar 2". The info provided indicates that user error caused the event. As of april 24,2009, the hospital has continued to use the da vinci s surgical system to perform surgery on patients. |  | ELECTROSURGICAL INSTRUMENT | 1533769 | 1371300 | 1308011 | 2955842200900147 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-04-03 | 2009-04-28 | Yes | Yes | Health Professional |  | No | Yes | 2009-04-03 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-04-01 | Malfunction | Other | It was reported that at the beginning of a da vinci s right partial nephrectomy procedure, it was observed that the fiber coating of the monopolar curved scissors (mcs) instrument was smoking and melting into the pt's colon. Reportedly, the event lengthened the surgical procedure. In 2009, intuitive surgical, inc. , received a copy of the user facility's voluntary report: "event desc: pt was undergoing a robotic right partial nephrectomy. During the procedure, the surgeon noticed smoke coming out of the laparoscopic port of the monopolar scissors. The scissors were removed to reveal the protective coating was hot to touch with areas of coating missing. Pieces of the fiber coating burned areas within the pt's peritoneal cavity. " the suspect medical device referenced in the report concerned an electrosurgical generator (esu). | The instrument has not been returned for evaluation. The root cause of the customer reported a failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional info is received. In 2009, isi provided the customer with a letter confirming once the mcs instrument is returned a formal response letter will be provided detailing our failure analysis findings. As of three days later, there have been no reported recurrences of the issue at this hospital. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1439233 | 1373214 | 1301759 | 2955842200900142 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2009-04-01 | 2009-04-24 | Yes | Yes | Health Professional | 42017908 | No | No | 2009-04-01 | Device Not Retu | No | Unkown | 2009
IS2000 A51 P7 | 2009-04-02 | Other | Other | It was reported that approximately 2. 5 hours into a da vinci s prostatectomy procedure, the left eye of the surgeon side console viewer went black. The surgical staff made the decision to convert to traditional open surgical techniques to complete the planned procedure. | Conclusion - the investigation conducted by the isi field service engineer discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of (b)(6) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1505603 | 1416632 | 1350916 | 2955842200900145 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Nurse | Initial | 2009-04-02 | 2009-04-24 | Yes | Yes | Health Professional | IS2000 A51 P7 | Yes | Yes | 2009-03-06 | Yes | No | Reuse | 2009
400180 | 2009-03-31 | Other | Other | It was reported that during a da vinci s partial nephrectomy procedure, the pt experienced increased bleeding. The surgeon increased the electrical surgical unit (esu) cautery setting from desiccate 30w cut/30 w coagulation to desiccate 70w cut/120w coagulation. Arcing was observed from the tip cover accessory installed on the monopolar curved scissors (mcs) instrument, causing a superficial burn to the pt's vena cava. No repair for the burn tissue was required. The procedure was completed with a replacement mcs instrument and tip cover accessory. No add'l pt harm, adverse outcome or injury was reported. | Conclusion - the tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If add'l info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1504637 | 1416633 | 1350917 | 2955842200900144 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-04-01 | 2009-04-24 | Yes | Yes | Health Professional | 400180 | No | Yes | 2009-04-01 | Device Not Retu | Yes | Unkown | 2009
42017908 | 2009-03-31 | Other | Other | It was reported that approximately 15 minutes into a da vinci s hysterectomy procedure, a burning smell was noticed. The surgical staff changed the monopolar cord and electrosurgical unit (esu), however, the burning smell persisted. At this time, the surgeon noticed debris had been removed from the mcs instrument shaft. The mcs instrument was removed and a hole burned through the shaft was discovered. The surgeon irrigated and suctioned the pt's affected area to remove debris. The planned procedure was completed with a replacement mcs instrument and no add'l pt harm, adverse outcome or injury was reported. | Conclusion - the monopolar curved scissors (mcs) instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If add'l info is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1504645 | 1416634 | 1350918 | 2955842200900141 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-04-01 | 2009-04-24 | Yes | Yes | Health Professional | 42017908 | No | Yes | 2009-04-01 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-04-01 | Other | Other | It was reported that approximately 20 minutes into a da vinci s gynecological surgical procedure while docking the system, the vision side cart (vsc) touch screen image went black with the message "sdi searching" displayed in the lower right corner. The external monitor connected to the console also went black. With assistance of a technical support engineer (tse) the site moved the touch screen through it's range of motion, checked the video cables located at the rear of the digitizer and the console, connected the composite cable to the digitizer, connected another monitor to the composite cable, and tried the vsc and console connection; however, the image did not return. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | Results and conclusions - the investigation conducted by an isi field service engineer (fse) confirmed the customer reported failure as the assistant monitor did not display any video and the message "composite searching" was displayed. The fse switched the monitor setting to sdi video; however, there was no image display. Review system error logs for the day did not reveal any unusual errors. Examination of the connections and settings on the vsc light found that the green light for the digitizer was not present. Removal of the side connections revealed that the site's electrosurgical unit (esu) was pressed against 3 plugs on the isolation transformer and were unseated. One of the affected plugs was the digitizer power cord. The fse concluded that esu pressed against the 3 plugs causing the loss of power. The video was restored to the assistant and external monitors when the esu and plugs were reseated. As of (b)(6) 2009 there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1505593 | 1416635 | 1350919 | 2955842200900143 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-04-01 | 2009-04-24 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-04-01 | Yes | No | Reuse | 2009
40015402 | 2008-12-03 | Malfunction | Other | It was reported that during a da vinci surgical procedure, the permanent cautery spatula instrument could not be removed through the cannula accessory. The cannula and instrument were removed from the pt. Outside of the pt cavity, the instrument broke off when removing the trocar. The procedure continued with a replacement instrument. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned for eval. Per the customer, reported complaint engineering found the instrument broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result. The clevis is intact, however, half of the main tube interface wall has collapsed. Engineering also observed charring on the yaw pulley, at the base of the spatula insulation, suggesting that an arcing event occurred. No other damage was found. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1529848 | 1437826 | 1367701 | 2955842200900136 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-03-26 | 2009-04-24 | Yes | Yes | Health Professional | 40015402 | Yes | Yes | 2009-03-26 | Yes | No | Reuse | 2009
42017908 | 2009-04-01 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors (mcs) instrument would not cauterize. Per the surgeon's request, the cautery settings were increased to 80 watts on the electrosurgical unit (esu) and shortly after, arcing, smoking, a burning scent and melting of the instrument shaft were observed. The surgical staff also noticed a hole burnt through the instrument shaft, located one inch from where the mcs tip cover accessory was installed. The instrument and esu were removed and replaced and the planned surgical procedure was completed without any further issues. No pt harm, adverse outcome or injury was reported. Additionally, the site reported that the instrument issue was caused by a broken esu. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a. 150" x. 165" oblong hole burnt through the distal end of the main tube. The hole center is located roughly. 600" above the tube extension. Localized material removed, around the hole indicates an arcing event occurred. The instrument was disassembled to find the tube extension and hypotube charred in the same location as the hole. A crack in the distal end of the tube was also observed, with the crack starting halfway between the two slot features and running along the side of the hole. It was concluded that the crack in the tube most likely created a path for arcing to occur. The source of the crack cannot be determined. No other damage was found. The esu being used when the instrument malfunction occurred is not manufactured by isi. On apr 21, 2009, the site indicated that the esu had been sent to the vendor for eval, however, the results of the eval had not yet been received. A f/u mdr will be submitted if additional info is provided by the site. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1529858 | 1437827 | 1367702 | 2955842200900146 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-04-01 | 2009-04-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-04-01 | Yes | No | Reuse | 2009
40018005 | 2009-03-24 | Other |  | It was reported that prior to starting a da vinci s surgical procedure with the monopolar curved scissor (mcs) instrument inserted through the cannula and with the tip of the mcs instrument in the patient's body cavity, isi's clinical sales representative observed that the tip of the instrument did not look appropriate. Removal of the mcs instrument revealed that the silicone distal end of the tip cover accessory was missing. The hospital's central processing center verified that prior to sterilizing the mcs instrument, the original tip cover was intentionally cut off at the silicone to ultem sleeve interface. Reportedly, the remaining piece of the tip cover accessory, along with the instrument, went through the hospital's sterilization process. The mcs instrument with the partial tip cover piece was immediately removed. The surgical procedure was completed with a replacement mcs instrument and tip cover accessory. | Isi's clinical sales representative instructed the hospital on proper removal of the tip cover accessory from the mcs instrument and how to identify that the tip cover accessory has been properly installed on the mcs instrument. The tip cover accessory has not been returned for evaluation; however, isi's investigation revealed that the root cause of the event is related to the improper removal of the tip cover accessory by the hospital's central processing center. The mcs tip cover is a one time use accessory that has not been validated for multi-use after sterilization. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1493118 | 1378358 | 1314927 | 2955842200900138 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-03-24 | 2009-04-17 | No | Yes | Health Professional | 40018005 | No | No | 2009-03-24 | Device Not Retu | Yes | Reuse | 2009
42017908 | 2009-03-27 | Other |  | It was reported that approximately 2 hours into a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, the patient developed a thermal burn on the abdominal aorta. The mcs instrument was removed and replaced and the planned surgical procedure was completed without significant delay. No additional patient harm or adverse outcome was reported. The hospital sterile processing department checked the integrity of the mcs instrument insulation using an olympus gun and stated that the integrity was compromised on the instrument in the mid shaft location. | The mcs instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1493108 | 1378359 | 1314928 | 2955842200900137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-03-27 | 2009-04-17 | Yes | Yes | Health Professional | 42017908 | No | Yes | 2009-03-27 | Device Not Retu | No | Unkown | 2009
ISI200 A51P7 | 2009-03-26 | Other |  | It was reported that prior to starting a da vinci s liver resection procedure, the site experienced an led warning light that was generated on the left camera control unit (ccu) and the image in the left eye of the surgeon side cart (ssc) displayed flashes of green and red. With the assistance of a technical support engineer (tse), the site powered off the left ccu and reseated the camera cable; however, the failure mode persisted after the ccu was re-powered. The tse instructed the site to power down the ccu and swap the camera cables; however, the issue transitioned to the right ccu. The tse instructed the site to power down and switch the cables back to their original positions and the issue moved back to the left eye. The issue remained after the tse instructed the site to shut down the left ccu and reseat the camera cable on the camera head side. The patient was under anesthesia approximately 30 minutes when the surgeon decided to abort and complete the procedure using another da vinci surgical system. No patient harm, adverse outcome or injury was reported. | Per the case notes, the investigation conducted by the isi tse discovered that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by providing the site with a new camera cable. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of april 16, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1493100 | 1378360 | 1314929 | 2955842200900140 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-03-26 | 2009-04-17 | No | Yes | Health Professional | ISI200 A51P7 | Yes | No | 2009-03-26 | Yes | No | Reuse | 2009
40018005 | 2009-03-24 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, arcing was observed when the monopolar curved scissors (mcs) instrument was bent against the pt's uterus with a large fibroid. A hole was noticed on the mcs tip cover accessory that was installed on the mcs instrument and the tip cover accessory was removed and replaced. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. The pt was discharged, the next day. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a hole burnt through the silicone. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole is approx. 020" in diameter and located. 190" above the silicone to sleeve interface. A small tear bisecting one edge of the hole was also observed, including various wear marks on the inner surface at the tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1532724 | 1443710 | 1373481 | 2955842200900135 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-03-24 | 2009-04-17 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-03-24 | Yes | Yes | Initial | 2009
42017908 |  | Malfunction |  | It was reported that the insulation chip of the monopolar curved scissors instrument, works intermittently. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube (which the customer referred to as the "insulation") has a. 065" x. 085" oblong hole burnt through it at the distal end, thus allowing energy to escape and causing the cautery at the tip to work intermittently. The hole center is located roughly. 550" above the tube extension and localized material removed around the hole indicates an arcing event occurred. The instrument was disassembled to find the tube extension and hypotubes charred in the same location as the hole. A crack in the distal end of the tube was observed starting at one of the slot features and running along the side of the hole. It was concluded that the crack in the tube most likely created a path for arcing to occur. The source of the crack cannot be determined. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1532734 | 1443711 | 1373482 | 2955842200900134 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-03-24 | 2009-04-17 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-03-24 | Yes | No | Reuse | 2009
42020504 | 2009-03-24 | Malfunction |  | It was reported that during da vinci total hysterectomy procedure, a piece of wire at the tip of the fenestrated biploar forceps instrument fell off into the patient. The wire was removed without incident. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint of a piece of wire at the tip of the instrument to have fallen off. Both pitch cables are broken at the distal clevis hub and one segment of the pitch down cable that is threaded through the distal clevis is missing. The pitch up segment (roughly. 350' long) is still threaded through the distal clevis, however, it is undetermined as to how both cables broke. Other cables at wrist are not damaged. No other damage found. Per the case notes, the missing piece of cable was retrieved from the patient and the procedure was successfully completed without incident. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1533726 | 1443713 | 1373484 | 2955842200900139 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2009-03-26 | 2009-04-17 | No | Yes | Health Professional | 42020504 | Yes | No | 2009-03-26 | Yes | No | Reuse | 2009
40018005 | 2009-02-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the console surgeon noticed a hole on the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. When the hole was observed, the electrosurgical unit (esu) was in fulgurate mode and the recommended power settings were not exceeded. The mcs tip cover accessory was removed and replaced and the esu was changed to desiccate mode to complete the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | Conclusions: the replacement tip cover accessory was returned and evaluated. Engineering observed a. 025" x. 040" oval shaped hole in the silicone. The hole is located along one silicone parting line, approximately. 225" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The inner surface of the silicone appears to have some light scratches/wear marks directly above the hole and 180 degrees from the hole. The mcs instrument used in conjunction with the tip cover accessory was also returned, however, no char marks were found on the instrument. The initial tip cover accessory used during the surgical procedure was also returned and based on the results of the evaluation performed on this tip cover accessory, mfr report #2955842-2009-00130 was submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1507665 | 1420013 | 1353285 | 2955842200900131 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-03-17 | 2009-04-13 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-03-17 | Yes | Yes | Initial | 2009
40018005 | 2009-02-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the console surgeon noticed a hole on the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. When the hole was observed, the electrosurgical unit (esu) was in fulgurate mode and the recommended power settings were not exceeded. The mcs tip cover accessory was removed and replaced, and the esu was changed to desiccate mode to complete the planned surgical procedure. No pt harm, adverse outcome or injury was reported. | Conclusions: the tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 020" diameter hole in the silicone. The hole is located halfway between the silicone parting lines, approx. 150" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. An adjacent mechanical tear was found on one side of the hole and two additional tears were also found approximately 180 degrees from the hole. The mcs instrument used in conjunction with the tip cover accessory was also returned, however, no char marks were found on the instrument. The replacement tip cover accessory was also returned. Based on the results of the evaluation performed on the replacement tip cover accessory, mfr report #2955842-2009-00131 was submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1507671 | 1420014 | 1353285 | 2955842200900130 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-03-17 | 2009-04-13 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-03-17 | Yes | Yes | Initial | 2009
40018005 | 2009-03-13 | Other | Other | It was reported that 2 hours during a da vinci s partial nephrectomy procedure, a hole developed in the monopolar curved scissor (mcs) instrument tip cover accessory. When the surgeon applied energy, electricity was seen escaping from the tip cover laterally causing a minor first degree burn to the pt's tissue. The isi representative in attendance recommended that the mcs instrument be removed and for the customer to replace the tip cover. The procedure was completed with the replacement tip cover. No additional pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed three holes burned through the silicone. The silicone exhibits localized melting around the holes, indicating multiple arcing events occurred. Two holes are approximately 180 degrees apart with the third hole halfway between the other two. The hole diameters range from. 012 inches to. 018 inches and are located. 355 inches to. 445 inches above the silicone to sleeve interface. The tip cover also has a section between two of the holes with various mechanical tears and scratches in the silicone. The site also returned the mcs instrument used in conjunction with the mcs tip cover accessory during the surgical procedure. Engineering observed the tip cover accessory installed on the instrument, the holes line up with the instrument's distal clevis and idler pulley edges. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1476604 | 1359576 | 1296788 | 2955842200900127 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-03-13 | 2009-04-10 | Yes | Yes | Health Professional | 40018005 | Yes | Yes | 2009-03-13 | Yes | Yes | Initial | 2009
400180 | 2009-03-06 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the patient suffered a minor first degree burn on the right uretheral orifice and when the monopolar curved scissor (mcs) instrument was removed, a hole was seen in the tip cover accessory. The planned surgical procedure was completed and no additional patient harm, adverse outcome or injury was reported. | The tip cover accessory was discarded by the customer and will not be returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On april 4, 2009, additional information regarding the event was provided by the surgeon. He stated that the pt burn was not likely caused by the hole in the tip cover, but rather due to the patient's anatomy, due to the pt's large prostate extending toward his urethra. The surgeon did confirm that cracks were found on the tip cover after the procedure. If the additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1476611 | 1359578 | 1296790 | 2955842200900128 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-03-17 | 2009-04-10 | Yes | Yes | Health Professional | 400180 | No | Yes | 2009-03-17 | Device Not Retu | Yes | Unkown | 2009
420189 | 2009-02-20 | Injury | Other | In 2009, intuitive surgical, inc. , received a user facility medwatch form from hospital for a pt. The event details are provided below: "during robotic assisted pneumonectomy while surgeon was using the fenestrated grasper, bleeding was observed coming from the left atrial appendage. Attempts to ligate or clip, it was unsuccessful and the procedure was converted to an open procedure. Surgeon did a lateral thoracotomy to control bleeding using pledgeted 5-0 prolene suture. The original procedure was completed and the pt was transferred to icu post operatively. The pt recovered without further incident. The surgeon believes that the hinge on the grasper "caught" the atrium and caused the bleeding. ". | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Attempts have been made to contact the initial reporter for additional info and for the return of the instrument without success. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1426024 | 1363068 | 1300207 | 2955842200900132 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2009-03-17 | 2009-04-10 | Yes | Yes | Health Professional | 420189 | No | Yes | 2009-03-17 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-03-10 | Malfunction |  | It was reported, that a hole in the insulation of the monopolar curved scissors (mcs) instrument was observed during a da vinci s surgical procedure. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the distal end of the main tube has a. 260" x. 011" burned section with a hole. The burn mark's center is located approximately. 600" above the tube extension. Localized material removal around the hole indicates an arcing event occurred. Disassembly of the instrument revealed that the tube extension and hypotubes were charred in the same location as the hole. A crack in the distal end of the tube was observed halfway between two of the slot features and leads to the lower edge side of the burned area. Engineering concluded that the crack in the tube most likely created a path for arcing. The source of the crack is indeterminable. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1504460 | 1418851 | 1352146 | 2955842200900133 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-03-19 | 2009-04-10 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-03-19 | Yes | No | Reuse | 2009
400180 |  | Malfunction | Other | It was reported that during a da vinci s prostatectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing was witnessed by the surgical staff. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was discarded by the customer and will not be returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1505749 | 1418853 | 1352148 | 2955842200900129 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-03-17 | 2009-04-10 | Yes | Yes | Health Professional | 400180 | No | Yes | 2009-03-17 | Device Not Retu | Yes | Unkown | 2009
REF 8MM 400184 | 2009-03-26 | Malfunction | Other | During laparoscopic davinci assisted hysterectomy, the cautery spatula small "wing" was noted to be missing after use. The md was unable to locate piece upon visualization. It was determined not to be radiopaque so subsequent x-rays were not completed. Md feels may not have been present prior to insertion but not noted prior to use. |  | INTUITIVE SURGICAL | PERMANENT CAUTERY SPATULA | 1421777 | 1358359 | 1295626 | MW5010643 | 1 | GEI | Voluntary |  | Nurse | Initial | 2009-04-07 | 2009-04-07 | No | Yes | Health Professional | REF 8MM 400184 |  | Yes |  |  |  |  | 2009
IS2000 A51P7 | 2009-02-24 | Injury | Required InterventionHospitalization | It was reported that post a da vinci s prostatectomy procedure, the patient developed bilateral leg nerve palsy. No additional patient harm or adverse outcome was reported. | In 2009, additional information regarding the event was provided by the console surgeon who reported that the procedure duration was 8 hours from start to finish, with the patient docked at the da vinci s patient side cart for 5 hours and 50 minutes. The patient was in a steep trendelenburg position and secured to the patient table with velcro straps and shoulder restraints with an "egg crate" underneath him. The surgeon speculated that the patient's legs were too abducted and lowered too far. The patient remained in the hospital for an additional 7 days for rehabilitation. As of the same day, the patient had 25% of his full capabilities. It was noted that the patient recovered well from the original da vinci s prostatectomy procedure. Based on the information provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the patient injury, however, it was determined that the patient's positioning and the extended procedure duration likely caused the event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1420735 | 1356892 | 1294223 | 2955842200900118 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-03-04 | 2009-04-03 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2009-03-04 | Device Not Retu | No | Reuse | 2009
42017206 | 2009-01-30 | Other | Required Intervention | It was reported that a da vinci s hysterectomy surgical procedure was performed on (b)(6) 2009, and eight days later, the patient spiked a fever and had an increase in white blood cells. An abdominal ct scan was performed for suspicion of bowel perforation. The patient was then taken to the or for an exploratory laparoscopy, during which a foreign body was found and removed. The site indicated that the foreign body was a component of the maryland bipolar forceps instrument which was used during the hysterectomy procedure performed with the da vinci s surgical system. The instrument was discarded by the site, however, the foreign body was sent to isi for evaluation. No additional patient harm, adverse outcome or injury was reported. | Conclusions - a black fragment was returned and evaluated. Per the customer reported complaint, engineering observed that the fragment is a partial cylindrical shell with an outer diameter of approx. 327" and a wall thickness of approx. 024". There are three fracture planes and one of the fracture lines that runs along the top circumference of the piece is jagged. The instrument component most similar to this fragment in terms of color, size and shape is the plastic proximal clevis. The fragment is similar to the part of the clevis that is located below the ears and mates with the main tube interface wall. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1503680 | 1416586 | 1350870 | 2955842200900122 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-03-10 | 2009-04-03 | Yes | Yes | Health Professional | 42017206 | Yes | Yes | 2009-03-10 | No | No | Unkown | 2009
IS2000 A51P9 | 2009-03-13 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, the site was unable to calibrate the endoscope as the alignment target cross hair was not viewable. With the assistance of an isi technical support engineer, the site reseated the endoscope adapter and tried using both a 30 degree and 0 degree endoscope. The site then reseated the camera cables and restarted the system. The setting on the camera control unit (ccu) were found to be normal and the white balance was completed. The site again attempted to look at the alignment target but the cross hairs were very faint on the screen. The color bars displayed properly, however, the site tried to view the alignment target with the camera head only and could not get a clear picture. The patient was under anesthesia when the surgical staff made the decision to abort the planned procedure. No patient harm, adverse outcome, or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer were associated with the camera head and camera cable. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera heads and camera cable. As of april 2, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1503751 | 1416587 | 1350871 | 2955842200900126 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2009-03-13 | 2009-04-03 | Yes | Yes | Health Professional | IS2000 A51P9 | Yes | No | 2009-03-13 | Yes | No | Reuse | 2009
40018005 | 2009-03-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a hole was observed on the monopolar curved scissors (mcs) tip cover accessory, which was installed on the mcs instrument. The surgeon decided to test the tip cover and put the instrument against the uterus that was being removed. Energy was applied to the instrument and arcing was observed from the instrument tip to the uterus. The mcs tip cover accessory was removed and replaced and the planned surgical procedure was completed. The site also indicated that their valley lab fx electrosurgical unit (esu) setting was at 40w in coag fulgurate mode. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various mechanical tears and holes burnt through the silicone. Two adjacent. 010" holes are located. 365" and. 380" above the silicone to sleeve interface. The holes exhibit localized melting, indicating that arcing events occurred. The tears include a. 150" long tear at the tip and a couple of smaller tears on the opposite side of the holes. Per the reported esu settings used, the high settings may have contributed to the arcing event. The electrosurgical unit (esu) settings and footswitch cables instructions for use specifically states: warning: do not exceed the 3kv peak limit. Do not use monopolar instruments with "cut" settings. Doing so may result in electrical arcs and alternate site burns. The corresponding maximum generator power settings to stay below the 3kv limit are as follows: esu model: valley lab fx. Mode: coag fulgurate. Max power setting: 38 watts. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1522223 | 1427569 | 1360182 | 2955842200900124 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-03-12 | 2009-04-03 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-03-12 | Yes | Yes | Initial | 2009
42018304 |  | Malfunction |  | It was reported that prior to starting da vinci s surgical procedure, the customer became concerned about "energy transfer" with the permanent cautery hook instrument. The instrument was returned for evaluation. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the yaw pulley, distal clevis and proximal clevis exhibit char marks on the side opposite of the conductor wire. The charring indicates an arcing event occurred. There was no damage to the conductor wire. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1522172 | 1427572 | 1360185 | 2955842200900121 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2009-03-09 | 2009-04-03 | No | Yes | Health Professional | 42018304 | Yes | No | 2009-03-09 | Yes | No | Reuse | 2009
42018404 | 2009-03-12 | Malfunction |  | It was reported that during a da vinci s aortobifemoral bypass procedure, the "yellow plastic part" of the permanent cautery spatula instrument broke off and fell into the pt. The broken piece was immediately retrieved. The planned surgical procedure was successfully completed and not significantly lengthened. No pt harm, adverse outcome or injury was reported. | The permanent cautery spatula instrument was returned and evaluated. Per the customer reported complaint, engineering found that the proximal clevis is missing a. 220" x. 165" piece at the interface to the main tube. The missing piece was observed to be located in between the two clevis ears. Further examination found that the clevis is not dislodged from the main tube and the main tube interface wall is intact. No other damages were found. | PERMANENT CAUTERY SPATULA INSTRUMETN | ELECTROSURGICAL INSTRUMENT | 1521697 | 1427573 | 1360186 | 2955842200900125 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-03-12 | 2009-04-03 | No | Yes | Health Professional | 42018404 | Yes | Yes | 2009-03-12 | Yes | No | Reuse | 2009
4200179 | 2009-03-11 | Injury | Required Intervention | On (b) (4) 2009, an isi representative reported this event and on april 1, 2009, medwatch uf/importer report #(b) (4) was rec'd from the user facility which stated the following: "pt had robotic laparoscopic hysterectomy with bilateral salpingo-oophorectomy and uterus morcellation. As reported by the surgeon: the pt was undergoing a robotic hysterectomy. The pt had large uterine fibroids. The hysterectomy was going well; however, the pt was found to have some cautery injury to the rectosigmoid junction. Colorectal surgical consultation was obtained to evaluate the area of cauterized segment. On careful inspection, using the robotic system, it was found, there were a few areas of cautery with some deeper areas of white discoloration around the cautery sites. It was deemed unsafe to leave this area alone. It was recommended to proceed with segmental resection of this area located within the rectosigmoid junction and to proceed with a primary anastomosis. A segmental distal sigmoid colon resection with end-to-end coloproctostomy was performed. Post op dx: rectosigmoid cautery injury during robotic hysterectomy. Pt informed, discharged in good condition. Monopolar curved scissors: 4th use of allowed 10 times. Expiration date entered in not exact. Tip cover accessory appears to have several holes in it. Also, used during procedure in another robotic arm that contains a bipolar forcep, pk dissector. | The instrument and accessory used during the surgical procedure have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If add'l info becomes available, a follow-up mdr will be submitted. As of (b) (6) 2009, there have been no reported recurrences of this issue at this hosp. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1781716 | 1512826 | 1289207 | 2955842200900123 | 1 | NAY | Manufacturer | User facilityCompany Representative | Other | Initial | 2009-03-11 | 2009-04-03 | Yes | Yes | Health Professional | 4200179 | No | No | 2009-03-11 | Device Not Retu | No | Unkown | 2009
| 2009-04-01 | Malfunction |  | Patient was undergoing a robotic right partial nephrectomy. During the procedure the surgeon noticed smoke coming out of the laproscopic port of the monopolar scissors. The scissors were removed to reveal the protective coating was hot to touch with areas of coating missing. Pieces of the fiber coating burned areas within patient's peritoneal cavity. ====================== manufacturer response for electrosurgical generator, force fx======================manufactured representative requested the device, we have refused until manufacturer follows our procedures for obtaining the device. ====================== manufacturer response for grounding pad, rem polyhesive ii======================manufacturer has requested device, once manufacturer complies with our requirements we will release the device. |  | FORCE FX | ELECTROSURGICAL GENERATOR | 1425906 | 1364708 | 1301759 | 1364708 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-04-02 | 2009-04-02 | No | Yes | Invalid Data |  |  | No Answer Provided |  |  |  |  | 2009
40018005 | 2009-02-19 | Injury | Other | On march 5, 2009, intuitive surgical, inc. , received a user facility medwatch form from a medical center. The event details are provided below: while using scissors to dissect, it was noted somewhere on shaft of instrument that it was allowing electricity (monopolar) to escape somewhere on the instrument out of vision. Instrument had burned the aorta, no repair was needed (close call). Scissors taken out of service immediately. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 010" diameter hole in the silicone. The hole is located. 160" above the silicone to sleeve interface and the silicone exhibits localized melting around the hole, indicating that an arcing event occurred. The tip cover also has various, partial mechanical tears above the hole on the inner surface of the silicone. There are also other scratches and indentations on the inner surface of the silicone, roughly 180 degrees from the hole. The site also returned the mcs instrument used in conjunction with the mcs tip cover accessory during the surgical procedure. Engineering observed a char mark on the instrument at the proximal clevis ear. With the tip cover accessory installed on the instrument, the hole position closely aligns to the char mark. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1416896 | 1353071 | 1290535 | 2955842200900117 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2009-03-05 | 2009-03-30 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2009-03-05 | Yes | Yes | Initial | 2009
IS2000 A51 P7 | 2009-03-06 | Other | Other | It was reported that during the beginning of a da vinci s hysterectomy procedure, the surgeon experienced intermittent video flickering on the touchscreen and assistance monitors, while in the process of moving the camera for insertion into the endoscopic camera manipulator (ecm). The surgical staff made the decision to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by the isi technical support engineer (tse) discovered that the vision issue experienced by the customer was associated with a faulty camera head and camera cable. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the systems vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera head and camera cable. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " the camera head and cable were also evaluated by the original equipment mfr (oem). The oem discovered the camera stage to be damaged due to physical impact and the cable to have intermittent connection issues causing color changing and flicker. As of march 27, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1520915 | 1424221 | 1357308 | 2955842200900119 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Nurse | Initial | 2009-03-06 | 2009-03-27 | Yes | Yes | Health Professional | IS2000 A51 P7 | Yes | Yes | 2009-03-06 | Yes | No | Reuse | 2009
40018005 | 2009-03-05 | Malfunction |  | It was reported that towards the end of a da vinci s hysterectomy procedure, the surgeon saw a spark coming from the tip of the monopolar curved scissors (msc) instrument with a tip cover accessory installed. When the surgical staff removed the mcs instrument, a hole was observed in the mcs tip cover accessory. The surgeon proceeded with the procedure using a replacement tip cover accessory; however, another spark coming from the replacement mcs tip cover accessory was observed. Examination of the replacement tip cover revealed that the tip cover accessory had a hole. Reportedly, the mcs instrument and tip cover accessory were inspected prior to use. The electrosurgical generator unit (esu) , coag and cut settings were 35 watts. Mfr report #2955842-2009-00116 was submitted for the second occurrence of the "sparking" observed by the customer. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory and/or instrument is returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1516061 | 1424367 | 1357454 | 2955842200900104 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-03-05 | 2009-03-27 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-03-05 | Device Not Retu | Yes | Initial | 2009
40016906 | 2009-03-03 | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the harmonic curved shears insert fell into the patient. The broken piece was immediately retrieved. The planned surgical procedure was successfully completed and not significantly lengthened. No patient harm, adverse outcome or injury was reported. | The harmonic curved shears insert has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | HARMONIC CURVED SHEARS INSERT | ELECTROSURGICAL INSTRUMENT ACC | 1515252 | 1424369 | 1357456 | 2955842200900103 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Unknown | Initial | 2009-03-03 | 2009-03-27 | No | Yes | Health Professional | 40016906 | No |  | 2009-03-03 | Device Not Retu | No | Unkown | 2009
42017908 | 2009-02-16 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the shaft of the monopolar curve scissors instrument was burnt. The surgical staff did not observe the instrument touching nor come into contact with other instruments and no parts were observed falling into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube has a. 230" x. 160" oblong hole burnt through it. The hole center is located roughly. 550" above the tube extension. Localized material removed around the hole indicates that an arcing event occurred. The instrument was disassembled to find the tube extension and hypotubes charred in the same location as the hole. A crack in the distal end of the tube was observed, with the crack starting at one of the slot features and running along the side of the hole. It was concluded that the crack in the tube most likely created a path for arcing to occur. The source of the crack cannot be determined. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1515150 | 1424370 | 1357457 | 2955842200900120 | 1 | NAY | Manufacturer | ForeignUser facility | Unknown | Initial | 2009-03-06 | 2009-03-27 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-03-06 | Yes | No | Reuse | 2009
40018005 | 2009-03-05 | Malfunction |  | It was reported that toward the end of a da vinci s hysterectomy procedure the surgeon saw a spark coming from the tip of the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. When the surgical staff removed the mcs instrument, a hole was observed in the mcs tip cover accessory. The surgeon proceeded with the procedure using a replacement tip cover accessory; however, another spark coming from the replacement mcs tip cover accessory was observed. Examination of the replacement tip cover revealed that the tip cover accessory had a hole. Reportedly, the mcs instrument and tip cover accessories were inspected prior to use. The electrosurgical generator unit (esu), coag and cut settings were 35 watts. Mfr report #2955842-2009-00104 was submitted for the first occurrence of the "sparking" observed by the customer. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The mcs tip cover accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory and/or instrument is returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2517797 | 2492702 | 1357454 | 2955842200900116 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-03-05 | 2009-03-27 | No | Yes | Health Professional | 40018005 | No | Yes | 2009-03-05 | Device Not Retu | Yes | Initial | 2009
IS2000 A51P7 | 2008-12-16 | Death | Death | In 2009, intuitive surgical was forwarded voluntary report by the fda for a pt death at hospital. The report details are provided below: cdrh maude event report 2009. Report outcome attributed to event: death. Date received: 2009. Event date: 2009. Event description the following month: healthy female for ovarian cystectomy using the davinci robot. Pt suffered fatal cardiac arrest in the middle of the procedure. History of htn, but other wise healthy. No autopsy done. Device information: brand: davinci robot. Device operator: health professional. Device available for evaluation: y. Expiration date: is reporter a health professional?: y. | Additional information regarding the event was provided to isi by the attending surgeon's office manager. Rep indicated that no pt death occurred in 2009. The pt death they have on record actually occurred in 2008, and the surgical procedure using the da vinci s surgical system was performed five days prior. Rep mentioned that forty minutes into the surgical procedure the pt's blood pressure dropped and the pt went into cardiac arrest. At that point, the procedure was terminated. The surgical staff was able to regain vital signs, however, the pt had to be put on life support. The pt remained in the icu on life support until her husband decided to take her off. The pt's official date of death is five days after, and the pt's family opted not to have an autopsy performed. Rep also indicated that the attending surgeon believes the system in no way contributed to the demise of the pt. The hospital has continued to use the da vinci s surgical system to perform surgery on patients. As of 2008, no adverse events have been experienced with the site's system. As of 2008, no similar instances of this event have been reported to isi. Based on the information provided by the site and surgeon, it was determined that the da vinci s surgical system did not malfunction in a way that would have led or contributed to the pt's death. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1417476 | 1354232 | 1247574 | 2955842200900110 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2009-02-23 | 2009-03-25 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2009-02-23 | Device Not Retu | No | Reuse | 2009
8MM REF 400179 | 2009-03-11 | Injury | Required Intervention | Pt had robotic laparoscopic hysterectomy with bilateral salpingo-oophorectomy, and uterus morcellation. As reported by the surgeon: the patient was undergoing a robotic hysterectomy. The patient had large uterine fibroids. The hysterectomy was going well; however, the patient was found to have some cautery injury to the rectosigmoid junction. Colorectal surgical consultation was obtained to evaluate the area of cauterized segment. On careful inspection, using the robotic system, it was found there were a few areas of cautery with some deeper areas of white discoloration around the cautery sites. It was deemed unsafe to leave this area alone. It was recommended to proceed with segmental resection of this area located within the rectosigmoid junction and to proceed with a primary anastomosis. A segmental distal sigmoid colon resection with end-to-end coloproctostomy was performed. Post op dx: rectosigmoid cautery injury during robotic hysterectomy. Pt informed, discharged in good condition. Monopolar curved scissors: 4th use of allowed 10 times. Expiration date entered is not exact. Tip cover accessory appears to have several holes in it. Also used during procedure is another robotic arm that contains a bipolar forcep, pk dissector. |  | INTUITIVE SURGICAL MONOPOLAR CURVED SCISSORS | MONOPOLAR CURVED SCISSORS | 1415052 | 1351701 | 1289207 | 1351701 | 1 | GEI | User Facility | Unknown | RISK MANAGER | Initial | 2009-03-19 | 2009-03-24 | Yes | No | Health Professional | 8MM REF 400179 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2009
IS2000 A51P7 | 2008-02-25 | Other | Other | It was reported that during a da vinci s surgical procedure, the site experienced a loss of vision on the right video of the high resolution stereo viewer (hrsv). The 2d signal was on and the led indicator was lit. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No additional information was provided. No patient harm was reported. | An investigation conducted by the field service engineer found the customer reported problem to be associated with the digitizer. The digitizer converts the analog video images recorded by the camera into a digital format, so that it can be transmitted over the fiber-optic cable and combined with other images. The system was repaired by replacing the affected digitizer. As of march 20, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1471613 | 1353511 | 1290972 | 2955842200900114 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2008-02-25 | 2009-03-20 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-25 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-02-24 | Other | Other | It was reported that during a da vinci s sacrocolpopexy surgical procedure, the customer experienced system error code #1, which they were not able to override. The surgeon decided to convert to traditional open surgical techniques to finish the planned surgery prior to contacting isi for troubleshooting assistance. No patient harm was reported. | The investigation conducted by isi field service engineering discovered a broken axis cable on a patient side manipulator (psm). The psm is an instrument arm located on the patient side chart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #1 appears when the da vinci s safety system determines that a power supply voltage was out of range. Upon determining this condition, the safety system puts da vinci s in a "non-recoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering observed that the psm axis-7 cables were wrapped around pulleys, thus causing an inconsistent power supply voltage level that led to the system error experienced by the customer. As of 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1471608 | 1353513 | 1290974 | 2955842200900112 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-02-24 | 2009-03-20 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-02-24 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-02-23 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the patient side cart (psc) switched to battery power use. The customer plugged the psc into several different power outlets, however, the psc continued to use power from just the battery. With the assistance of an isi technical support engineer (tse) the customer attempted to power cycle the psc. When the psc shut off it would not power back on. The tse had the customer emergency power off (epo) on the psc, however, the system would not turn back on. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found that the power issue experienced by the customer was associated with the psc primary power supply (pps). The pps is a 48v power supply with battery backup and is mounted inside of the system control electronics chassis. The system was repaired by replacing the affected pps. As of march 19, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1473946 | 1353517 | 1290978 | 2955842200900109 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-02-23 | 2009-03-20 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-23 | Yes | No | Reuse | 2009
42018404 | 2009-02-26 | Malfunction |  | It was reported that during a da vinci s aortic coarctation repair surgical procedure, the tip of the permanent cautery spatula instrument broke off and fell into the surgical field. The broken piece was immediately retrieved. The planned surgical procedure was successfully completed. There was no patient harm, adverse outcome or injury reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1480375 | 1364288 | 1301351 | 2955842200900115 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-27 | 2009-03-20 | No | Yes | Health Professional | 42018404 | No | Yes | 2009-02-27 | Device Not Retu | No | Unkown | 2009
400179 | 2008-04-11 | Injury | Required InterventionHospitalization | On february 24, 2009, medwatch uf/importer report was received: " event desc: the pt was reevaluated 4 days after hysterectomy for abdominal pain and distension. She was found to have a perforated sigmoid colon. She was brought emergently to the operating room where she underwent the repair of perforation and end-colostomy. The pt was admitted to the intensive care unit on ventilator and blood pressure support medications. She had an extensive stay in the intensive care unit where she suffered from bilateral pleural effusions and pneumonia. She was eventually able to be transferred to the general medical floor. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no". | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. On march 20, 2009, additional information regarding the event was provided by the site's risk management coordinator. The surgeon noted the cause of the pt's perforated sigmoid colon was a thermal burn, believed to have been caused by the monopolar curved scissors instrument. Coordinator stated that no test results confirming this information were included in the report. No additional information was provided. No additional pt harm or adverse outcome was reported. If additional information is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1504089 | 1436242 | 1253608 | 2955842200900113 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-02-24 | 2009-03-20 | Yes | No | Health Professional | 400179 | No | No | 2009-02-24 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-02-24 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, the surgeon side console (ssc) turned off while the site was prepping the pt. The led indicator on the primary power source was low and under voltage. With the assistance of an isi technical support engineer (tse), the site attempted to emergency power off (epo) the system multiple times, however, the system would not turn on. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that the customer reported complaint was related to the primary power supply (pps). The pps is a +48v power supply with battery backup and is mounted outside of the system control electronics chassis. The system was repaired by replacing the pps. As of 03/20/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2642224 | 2605059 | 2501746 | 2955842200900111 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2009-02-24 | 2009-03-20 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-24 | Yes | No | Reuse | 2009
IS1200 A4 3P9 | 2009-02-12 | Injury | Other | It was reported that during a da vinci surgical procedure, the endoscope tip became very hot and burnt the patient's bowel. Surgical intervention was not required, and the planned procedure was completed without any further issues. | Investigation was conducted by isi field service engineering, who performed extended testing on the system illuminator and illuminator lamp module. The illuminator and lamp module (located within the illuminator) provides light, which is then delivered to the endoscope via a fiber optic light guide cable and projected onto the surgical site. After approx 1 hour of testing, no change to the lumens output was observed and the endoscope tip did not become abnormally hot. The illuminator was replaced as a precaution. Further investigation revealed that the site recently changed their single light guide to a bifurcated light guide, which was used during the procedure. Since the bifurcated light guide emits a higher light output than the single light guide, engineering concluded that the site's unfamiliarity with the bifurcated light guide may have contributed to the high heat experienced at the endoscope tip as the proper working distance and setting adjustments may not have been made to accommodate for the change. As of 2009, there have been no reported recurrences of the issue a this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1409713 | 1345068 | 1282690 | 2955842200900106 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-02-12 | 2009-03-13 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | Yes | 2009-02-12 | Yes |  | Reuse | 2009
IS2000 A51P7 | 2009-02-18 | Other | Other | It was reported that approximately 75 mins into a da vinci s hysterectomy procedure, the customer experienced unrecoverable system error code #23002. The site tried to clear the system error code by selecting fault override multiple times, as well as, restarting the system; however, the system error code #23002 recurred during homing. With the assistance of an isi technical support engineer (tse) the site exercised the master tool manipulator (mtm) grip, overrode the fault and was able to complete homing the system. The instruments were inserted and the surgeon took control of the mtms; however, the system error code #23002 immediately returned. The tse had the site emergency power off the system, yet the error persisted upon start up. After approximately 45 mins of troubleshooting, the surgeon decided to convert to traditional laparoscopic surgical techniques to finish the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23002 was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml), and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23002 appears when the da vinci s safety system determines that the position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder ) and the secondary sensor (potentiometer), were beyond mechanical limits. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state. " the mtmr was returned to isi for failure analysis investigation. Engineering evaluation found an axis potentiometer to encoder assembly had malfunctioned. This malfunction was found to have generated the system errors experienced by the customer. As of the (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1471950 | 1353721 | 1291160 | 2955842200900076 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-02-18 | 2009-03-13 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-18 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-02-17 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the camera cannula mount accessory attached to the endoscopic camera manipulator (ecm) arm broke in half as the surgeon was docking the ecm to the pt. The pt had been under anesthesia for approx 30 mins and the port incisions had been made when the surgeon decided to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer confirmed that the camera cannula mount accessory had broken. The ecm is the camera arm located on the pt side cart which provides the sterile interface for the 3d endoscope. The camera cannula mount accessory is attached to the ecm and designed to hold the cannula securely in place, outside of the pt cavity. The system was repaired by replacing the affected ecm cannula mount accessory. The camera cannula mount was returned to isi for failure analysis investigation. Engineering evaluation observed that the camera cannula mount is broken in two at the base and one of the fracture planes occurs where the base has a 90 degree angle, directly behind the back surface of the top metal plate. As of 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1471932 | 1353722 | 1291161 | 2955842200900100 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-02-17 | 2009-03-13 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-02-17 | Yes | No | Reuse | 2009
40017908 | 2009-02-17 | Malfunction | Other | It was reported that approximately 20 minutes into a da vinci hysterectomy procedure, while cutting the round ligament, the surgeon reported that the monopolar curved scissors instrument was not functioning. The instrument was removed for inspection and it was found that the scissor tip had "fragmented". The procedure continued with a replacement instrument of the same type. The patient underwent x-rays to identify the position of the instrument fragment. An additional incision was made to re-enter the patient's abdomen and the instrument fragment was retrieved resulting in approximately one hour of additional surgical time. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1474748 | 1357747 | 1295002 | 2955842200900099 | 1 | GEI | Manufacturer | Company RepresentativeHealth Professiona | Nurse | Initial | 2009-02-17 | 2009-03-13 | Yes | Yes | Health Professional | 40017908 | No | Yes | 2009-02-17 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-02-23 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer experienced system error codes #23002 and #25551. System error code #23002 occurred while docking the system to the patient and the surgical staff was unable to clear the system error code. An isi technical support engineer (tse) was contacted to assist with resolving the issue and the isi tse instructed the site to shut down and power off the system. The system was powered up and homed with no issues; however, system error code #23002 recurred when the surgical staff attempted to re-dock the system. The site was able to clear the error by pressing "fault override"; however, system error code #25551 occurred while the surgical staff was attempting to dock an arm instrument. The instrument arm was disabled and the remaining instrument arms were docked. After the instrument arms were docked, system error code #23002 recurred. The patient had been under anesthesia for approximately 2 hours and the patient port incisions had been made when the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by the field service engineering concluded that system error codes #23002 and #25551 experienced by the customer were associated with the left master tool manipulator (mtml) and the remote arm controller (rac). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's right hand (mtmr) and one to his left (mtml). The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The system was repaired by replacing the mtml and the rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #23002 system error code appeared when the da vinci s safety system determined that one or more joints had moved outside the expected range of motion as measured by the primary sensor. System error code #25551 occurred when one or more racs are unable to initialize the fault recovery logic. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state". The mtml and rac were returned to isi for failure analysis investigation. Engineering's evaluation concluded that mtml grip axis was out of calibration, thus generating the system error codes experienced by the customer. No issues were found with the rac. As of 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1474817 | 1357748 | 1295003 | 2955842200900108 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-02-23 | 2009-03-13 | No | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-02-23 | Yes | No | Reuse | 2009
42018304 | 2009-02-06 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, a piece of the permanent cautery hook instrument fell into the patient. The broken piece was immediately retrieved and the procedure continued with a replacement permanent cautery hook instrument. The planned surgical procedure was successfully completed and not significantly lengthened. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The customer reported that failure mode cannot be determined. If the instrument is returned for evaluation, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1473763 | 1357749 | 1295004 | 2955842200900107 | 1 | NAY | Manufacturer | User facility |  | Initial | 2009-02-13 | 2009-03-13 | No | Yes | Health Professional | 42018304 | No | No | 2009-02-13 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-02-17 | Malfunction |  | It was reported that 1. 5 hours into a da vinci s prostatectomy procedure and while dissecting the neck of the patient's bladder, arcing was observed from the tip of the monopolar curved scissor (mcs) instrument with tip cover accessory attached. Reportedly, during the procedure the electrical generator unit (esu) mode setting used was desiccate with the cut setting at 1, the coag setting was 30 watts, and the bipolar setting was 50 watts. The planned surgical procedure was completed with a replacement tip cover accessory and without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover accessory has a. 035" x. 035" oval shaped hole in the silicone. The hole is located. 200" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. The tip cover also has various mechanical tears that are concentrated in an area that is approximately 90 degrees from the location of the hole. The customer also returned a mcs instrument believed to be used in the procedure. Examination of the instrument revealed a char mark located at the proximal clevis ear. When the tip cover accessory was installed on the instrument, it was observed that the location of the hole closely matches with the location of the char mark on the instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1474767 | 1357750 | 1295005 | 2955842200900102 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-18 | 2009-03-13 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-02-18 | Yes | No | Initial | 2009
40018302 | 2009-02-02 | Malfunction |  | It was reported that during a da vinci myomectomy procedure, while the surgeon was removing a fibroid, a "clear amber spool" fell out of the distal end of the permanent cautery hook instrument. It was additionally reported "the cables of the lateral side appear to be looser than the other side". The broken piece was immediately retrieved and the procedure continued with a replacement permanent cautery hook instrument. The planned surgical procedure was successfully completed and not significantly lengthened. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If the instrument is returned for eval, a follow-up mdr will be submitted. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1476071 | 1359944 | 1297152 | 2955842200900096 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-06 | 2009-03-06 | No | Yes | Health Professional | 40018302 | No | Yes | 2009-02-06 | Device Not Retu | No | Unkown | 2009
40018005 | 2009-02-06 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, arcing occurred through the monopolar curved scissor (mcs) tip cover accessory, which was installed on a mcs instrument. The planned surgical procedure was completed with a replacement tip cover accessory and without significant delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the tip cover accessory has three holes burned through the silicone. The holes exhibit localized melting, indicative of arcing. One of the holes is. 025" in diameter and is located. 185" above the silicone to sleeve interface. Adjacent to this hole are a couple of mechanical tears and gouge marks. There are also two smaller oblong holes in the silicone located approximately. 210" and. 250" above the silicone to sleeve interface. The customer also returned a mcs instrument believed to have been used in the procedure. Examination of the instrument revealed that there is a char mark located below the proximal pin. When the tip cover accessory was installed on the instrument, it was observed that the location of hole closely matches with the location of the char mark on the instrument. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1476063 | 1359946 | 1297154 | 2955842200900090 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-02-07 | 2009-03-06 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-02-07 | Yes | Yes | Initial | 2009
42017206 | 2009-02-09 | Malfunction |  | It was reported that after the da vinci s surgical procedure was completed, shavings were noticed on the shaft of the maryland bipolar forceps instrument. It was also reported that nothing was observed falling into the pt during the procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal part of the main tube has a 5" long section with material removed all around the tube's outer surface. The damaged area is parallel to the tube axis and has a rough surface finish. No other damage was found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1476051 | 1359948 | 1297156 | 2955842200900098 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-02-20 | 2009-03-06 | No | Yes | Health Professional | 42017206 | Yes | No | 2009-02-20 | Yes | No | Reuse | 2009
42020703 | 2009-02-11 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the "plastic end part" of the tenaculum forceps instrument, broke and fell into the pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. Multiple attempts have been made to obtain confirmation from the customer that all pieces were retrieved from the pt, however, at this time the customer has not responded. If additional info is received, follow-up mdr will be submitted. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument is broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result. Both the proximal clevis and main tube are missing pieces, with the proximal clevis missing a piece approximately. 189 inches by. 277 inches. No other damage found. Engineering included that based on the breakage, the damage was caused by overloading at the tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1476079 | 1359950 | 1297158 | 2955842200900105 | 1 | GCJ | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-11 | 2009-03-06 | No | Yes | Health Professional | 42020703 | Yes | Yes | 2009-02-11 | Device Not Retu | No | Reuse | 2009
| 2009-02-19 | Injury | Required Intervention | While using scissors to dissect, it was noted somewhere on shaft of instrument that it was allowing electricity (monopolar) to escape somewhere on the instrument out of vision. Instrument had burned the aorta, no repair was needed (close call). Scissors taken out of service immediately. |  | DA VINCIS | MONOPOLAR SCISSORS | 1400725 | 1334331 | 1272102 | MW5010163 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2009-02-23 | 2009-03-04 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2009
42004905 |  | Malfunction |  | It was reported that gauges were observed on the bottom shaft of the cadiere forceps instrument. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed deep scratches with material removed on the instrument main tube, located in an area extending approx. 570" from the intersection with the proximal clevis. The location and appearance of the scratches suggests that they may have been caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1476333 | 1360056 | 1297263 | 2955842200900084 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-01-28 | 2009-03-02 | No | Yes | Health Professional | 42004905 | Yes | No | 2009-01-28 | Yes | No | Reuse | 2009
42019006 | 2009-01-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the coating of the cobra grasper instrument was rubbing off and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a 2. 5" long section on one side of the distal end of the instrument main tube with material removed. The damaged section appears to consist of tube abrasions and deep scratches, with the tube abrasions parallel to the main tube. Based on the location and appearance, the tube abrasion damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. There are also various scratch marks that are not aligned with the tube axis in the same location as the abrasions. The wear pattern of the deep scratches suggests that the damage was most likely caused by instrument collisions and/or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1460443 | 1338221 | 1275916 | 2955842200900097 | 1 | GCJ | Manufacturer | User facility | Unknown | Initial | 2009-01-29 | 2009-02-27 | No | Yes | Health Professional | 42019006 | Yes | No | 2009-01-29 | Yes | No | Reuse | 2009
40018302 | 2009-02-05 | Malfunction |  | It was reported that during a da vinci prostatectomy surgical procedure, a piece of the permanent cautery hook instrument broke off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument cautery hook is broken off at the base of the yaw pulley and fractured through one of the machined holes in the shank. The hook and ceramic sleeve were not returned with the instrument. As indicated in the complaint description, the broken piece was successfully retrieved from the patient. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1460061 | 1338222 | 1275917 | 2955842200900095 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-02-05 | 2009-02-27 | No | Yes | Health Professional | 40018302 | Yes | Yes | 2009-02-05 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-01-30 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced system error code #297 and the led on one of the system pt side manipulators (psm) became red or yellow. With the assistance of an isi technical support engineer (tse), the customer moved the psm behind the system column, however, the fault recurred and the psm and set-up joint both became unresponsive to movement and without lights illuminated. The site then attempted to power down the system and stow the affected psm to continue with the procedure with remaining psms. The system was homed with the affected psm disabled and the procedure continued. After continuing, the fault recurred. The site then powered down the system and restarted the system multiple times. The fault continued to occur during the system self testing. The site decided to covert to traditional open surgical techniques to finish. | The investigation conducted by field service engineering (fse) confirmed that the site experienced system error code #297. While reviewing the system error logs, the fse also found that the system error code #23 had occurred. Engineering concluded that the system error codes experienced by the customer were associated with a configured embedded serializer setup joint (cfg, essj) and the psm. The embedded serializer for setup-joint is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four joints and their associated backup potentiometers. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected cfg, essj, and psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23 indicates a communication failure in the low voltage differential signal chain. In the process of transitioning to a safe state due to this error, a sympathetic communication error (#297) between the pt side cart and the surgeon side cart was generated. The cfg, essj, and psm were returned to isi for failure analysis investigations, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, the cfg, essj were scrapped. As of (b)(4) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1466923 | 1342476 | 1280135 | 2955842200900087 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-01-30 | 2009-02-27 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-01-30 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-02-05 | Other | Other | It was reported that prior to starting a da vinci prostatectomy procedure, the customer experienced an unrecoverable system error code #2 and #21008. The pt was under anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule for the following day. | The investigation conducted by field service engineering determined that the system error code #2 and 21008 were associated with a specific pt side manipulator (psm) arm. The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The error code #21008 appeared when the da vinci safety system determined the angular position of one or more robotic joints on the specified manipulator, as measure by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1466904 | 1342477 | 1280136 | 2955842200900094 | 1 | NAY | Manufacturer | ForeignCompany Representative | NOT APPLICABLE | Initial | 2009-02-05 | 2009-02-27 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-02-05 | Yes | No | Reuse | 2009
IS2000 A51 P6 | 2009-01-28 | Other | Other | It was reported that during a da vinci s surgical procedure, the surgeon experienced a blurry image on the right side viewer of the surgeon's console while using the 30 degree endoscope. With the assistance of an isi clinical sales representative (csr), and an isi technical support engineer (tse), the site attempted to troubleshoot the issue. Prior to determining the root cause, the surgical staff made the decision to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by the isi csr and tse discovered that the vision issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable and the vision quality was verified. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1466932 | 1342478 | 1280137 | 2955842200900086 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2009-01-28 | 2009-02-27 | Yes | Yes | Health Professional | IS2000 A51 P6 | Yes | No | 2009-01-28 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-02-03 | Other | Other | It was reported that during a da vinci s cardiac surgical procedure, the surgical staff decided to reposition a system arm for better port placement. While the system arm was undocked from the pt, the instrument that was installed on the system arm was still connected to the generator and inadvertently energized. The system faulted and error code #1 was displayed. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #1 was generated when the instrument installed on the system arm was energized during the time it was undocked from the pt. Extensive system testing was performed with no issues found, and system performance was verified. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #1 appears when the da vinci s safety system determines that a power supply voltage was out of range. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state". As of (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1466943 | 1342479 | 1280138 | 2955842200900092 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-02-03 | 2009-02-27 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-02-03 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-01-26 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgeon experienced a dark image and a black halo around both eyes in the stereo viewer of the surgeon side console (ssc). With the assistance of the isi clinical sales representative (csr) onsite, and an isi technical support engineer (tse), the site checked their camera control unit (ccu), tried multiple scopes, reseated the camera head sterile adaptor, and re-performed the black and white balance. This slightly improved the image quality. The csr changed out the camera head and the image in the left eye of the stereo viewer became purple, however, this was corrected by replacing the camera head cable. The ccu settings were then set to "low" to increase the brightness. The image was acceptable by the surgical staff and they decided to continue with the procedure. The site later reported that the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by isi field service engineering concluded that the dark image and black halo issue experienced by the customer were due to a non approved refurbished illuminator lamp module located within the illuminator. The illuminator lamp module provides light, which is then delivered to the endoscope via a fiber optic, bifurcated light-guide cable and projected onto the surgical site. The system was repaired by replacing the refurbished illuminator lamp module. The fse instructed the site to replace the illuminator bulb with isi approved bulbs as the refurbished bulb outputs 150 lumens as opposed to the 1400 lumens output from an isi approved bulb. As of (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1459442 | 1329106 | 1266925 | 2955842200900080 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-01-26 | 2009-02-26 | No | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-01-26 | Yes | No | Reuse | 2009
42017206 | 2009-01-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, "splinters" were observed on the shaft of the maryland bipolar forceps instrument. Reportedly, no fragments fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the instrument shaft has "splinters". One side of the distal end of the main tube has a couple of sections with light material removed. The damaged areas are 2. 5" and 1. 5" long, parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1459156 | 1334714 | 1272465 | 2955842200900093 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2009-01-27 | 2009-02-26 | No | Yes | Health Professional | 42017206 | Yes | No | 2009-01-27 | Yes | No | Reuse | 2009
400179 | 2009-02-17 | Other | Other | During laparoscopic hysterectomy, the infundibulopelvic ligaments were cauterized and then excised with the laparoscopic scissors and then later excision of the uterine vessels. At this point the surgeon noted that scissors would not open. They were removed and it was noted that 1 blade of the scissors was missing. An xray was performed and the piece noted to be in the pelvis. The piece was retrieved by reentry suprapubically. |  | MONOPOLAR CURVED SCISSORS | SCISSORS LAPAROSCOPIC | 1400958 | 1337196 | 1274891 | 1337196 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2009-02-24 | 2009-02-26 | Yes | Yes | Physician | 400179 | Yes | No Answer Provided |  |  |  |  | 2009
| 2009-02-17 | Malfunction |  | Da vinci small clip applier tip broke off inside of this patient. It was unable to be seen in the patient. Hours later, dr. Found the tip in tissue which he had to pull to remove. |  | DAVINCI 8MM SMALL CLIP APPLIER | NONE | 1397785 | 1327334 | 1265257 | MW5010101 | 1 | GDO | Voluntary |  | Other | Initial | 2009-02-25 | 2009-02-25 | No | Yes | Health Professional |  | Yes | No |  |  |  |  | 2009
IS1200 A43P9 | 2009-01-26 | Other | Other | It was reported that prior to starting a da vinci surgical procedure, the site experienced no video on their surgeon side console (ssc). With the assistance of the isi clinical sales representative (csr) onsite, the camera cable was changed, however, this resulted in a no power indication of the right camera control unit (ccu). The csr called an isi technical support engineer who advised the site to power cycle the ccu, however, no change occurred. The site then checked the power connection to the direct current (dc) adapter and power strip, as well as, swapped out the right camera cable going to ccu; however, no change was observed. The site then replaced the camera cable that had previously been working, which resulted in neither of the ccu's to powering on. The color bars were checked on both ccu units and neither could output. The patient was under anesthesia when the surgical staff made the decision to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer were associated with the right and left camera control units (ccu), camera heads and camera cables. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the right ccu, camera heads and camera cables. As of (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1460636 | 1329124 | 1266943 | 2955842200900083 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-01-26 | 2009-02-20 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-01-26 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-01-26 | Other | Other | It was reported that prior to starting a da vinci s hysterectomy procedure, the site experienced system error code #23020. The site contacted the isi field service engineer, however, they were unable to clear the system error code. The patient was under anesthesia for approximately 30 minutes when the site decided to convert to traditional laparoscopic techniques to finish the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that system error code # 23020 experienced by the customer was associated with a patient side manipulator (psm). The patient side manipulator is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. System error code #23020 appears when the da vinci safety system determines one of the switches in a specific manipulator is showing inconsistent signals on two switch leads. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put da vinci in a "recoverable safe state. " the psm was returned to isi for failure analysis investigation. Engineering was unable o replicated the customer reported problem, however, the link 3 switch and board were replaced as a precaution. As of february 19, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1460583 | 1329125 | 1266944 | 2955842200900081 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-01-26 | 2009-02-20 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | Yes | 2009-01-26 | Yes | No | Reuse | 2009
40017208 | 2008-12-20 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure a small piece of the conductor wire of the maryland bipolar forceps instrument broke off and fell into the surgical field. The broken piece was immediately retrieved. The planned surgical procedure was successfully completed. There was no patient harm, adverse outcome or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one of the conductor wires to be broken. A. 625" long segment of the conductor wire is missing from the yaw pulley exit to the proximal clevis. Disassembly of the instrument by engineering revealed that the wire had broken near the proximal clevis wire hole. Engineering concluded that it is indeterminable as to when the breaks occurred, however, necking observed at the break point likely occurred due to stretching or the wire being smashed on a sharp edge. Engineering also concluded that detachment of the wire at the yaw pulley exit likely occurred as a result of not being securely attached to the grip. No other damage found. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1459109 | 1334816 | 1272567 | 2955842200900085 | 1 | GEI | Manufacturer | Company RepresentativeDistributor | Other | Initial | 2009-01-28 | 2009-02-20 | No | Yes | Health Professional | 40017208 | Yes | No | 2009-01-28 | Yes | No | Reuse | 2009
40018005 | 2009-01-26 | Malfunction |  | It was reported that the monopolar curved scissor (mcs) tip cover accessory was used as a replacement to complete a da vinci s prostatectomy procedure, in which the original mcs tip cover accessory presented arcing during the procedure. It was not reported that the accessory was being returned due to failure, however, evaluation of the returned accessory by engineering revealed evidence of arcing. | The tip cover accessory was returned and evaluated. Engineering found that the tip cover has a small hole approximately. 010" in diameter burned through the silicone. The hole is located. 215" above the silicone to the sleeve interface. The silicone exhibits localized melting around the hold, indicative of arcing. The tip cover also has various mechanical tears approximately 120 degrees from the hole. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength caused by damage to the silicone. The site also returned an mcs instrument and another tip cover accessory with evidence of arcing. Refer to med watch mfr report #s 2955842-2009-00082 and 29552842-2009-00088. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1459113 | 1334817 | 1272568 | 2955842200900089 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-01-26 | 2009-02-20 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-02-03 | Yes | Yes | Initial | 2009
42017908 | 2009-01-26 | Malfunction |  | It was reported while taking down the patient's bladder during a da vinci prostatectomy s procedure, the monopolar curved scissor (mcs) instrument with tip cover installed collided with another instrument of an unknown model. The surgeon noticed that the monopolar curved scissor (mcs) instrument with tip cover was not energizing properly when cautery energy was applied. Reportedly, arcing took place during dissection of the seminal vesicle. When the surgical staff removed the instrument a hole was observed in the tip cover. The planned surgical procedure was completed with a replacement tip cover and without significant delay. No patient harm, adverse outcome or injury was reported. It was not reported that the instrument was being returned due to failure, however, evaluation of the returned instrument by engineering revealed material missing from the proximal clevis seal. | The instrument was returned and evaluated. Engineering observed a small amount of material to be missing from the proximal clevis seal. Engineering concluded that the missing material suggests that arcing occurred at the location of the proximal clevis seal as the area lines up with the small hole in the tip cover accessory removed during the procedure. Refer to medwatch mfr report # 2955842-2009-00082. The site returned a second tip cover accessory. Refer to medwatch mfr report # 2955842-2009-00089. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1459061 | 1334818 | 1272569 | 2955842200900088 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-01-26 | 2009-02-20 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-02-03 | Yes | No | Initial | 2009
40018005 | 2009-01-26 | Malfunction |  | It was reported while taking down the pt's bladder during a da vinci prostatectomy s procedure, the monopolar curved scissor (mcs) instrument with tip cover installed collided with another instrument of an unk model. The surgeon noticed that the monopolar curved scissor (mcs) instrument with tip cover was not energizing properly when cautery energy was applied. Reportedly, arcing took place during dissection of the seminal vesicle. When the surgical staff removed the instrument a hole was observed in the tip cover. The planned surgical procedure was completed with a replacement tip cover and without significant delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the tip cover has a. 055" x. 020" oblong hole in the silicone. The hole is located. 210" above the silicone to the sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. The ends of the hole appear to have mechanical tears. The tip cover also has various mechanical tears and/or scratches and two small (under. 010" in diameter) holes with evidence of arcing, approximately 90 degrees from the oblong hole. The location of the holes relative to the plastic sleeve interface is. 015" and. 135". Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by instrument collisions and/or rough handling. The site returned an mcs instrument and another tip cover accessory, both with evidence of arcing. Refer to medwatch mfr report #s 2955842-2009-00088 and 2955842-2009-00089. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1459277 | 1334819 | 1272570 | 2955842200900082 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2009-01-26 | 2009-02-20 | No | Yes | Health Professional | 40018005 | Yes | Yes | 2009-01-26 | Yes | Yes | Initial | 2009
42020704 |  | Malfunction |  | It was reported that during central processing the tenaculum forceps instrument had "fiberglass" peeling off the shaft. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube had a. 750 inch long section with material removed on one side of the tube. The damaged area is contained within a larger area that has the appearance of tube wear, but the surface is smooth. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if additional info is received. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1459284 | 1334820 | 1272571 | 2955842200900079 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-26 | 2009-02-20 | No | Yes | Health Professional | 42020704 | Yes | No | 2009-01-26 | Yes | No | Reuse | 2009
IS2000 A51P6 | 2009-01-21 | Other | Other | It was reported that during the beginning of a da vinci s hysterectomy surgical procedure, the right camera controller unit (ccu) was shutting itself off. The site inspected the camera head and noticed that the connectors appeared loose or "stripped," therefore, the cables would not seat properly. The customer exchanged the camera head and camera cable, however, when the replacement camera head was installed, a warning light was generated on the right ccu and only a greent tint was displayed on the right side image of the surgeon side console. The site then swapped cables around and was able to produce one good image on the right side of the console. The system was then placed in 2-d mode to have the viewer look the same, however, the surgeon decided not to continue with the scheduled case. The patient was under anesthesia for approximately 45 minutes when the surgical staff made the decision to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | Investigation conducted by the field service engineer concluded that the reported issues experienced by the customer were associated with the right camera control unit (ccu), camera heads and camera cables. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the right ccu, camera heads and camera cables. The right ccu, camera heads and camera cables were returned to isi for failure analysis evaluation. One camera head was found to have a loose connector due to the weight of the cable and over-torquing, thereby confirming the loose connector issue reported by the customer and the other camera head had a bent connector pin which most likely contributed to the "green tint" experienced by the customer. No issues were found with the ccu and camera cables. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1448944 | 1313223 | 1251387 | 2955842200900069 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-01-21 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-01-20 | Other | Other | It was reported that during a da vinci hysterectomy surgical procedure, the customer experienced a system error code #20013. With the assistance of an isi technical support engineer (tse), the site attempted to reposition the master tool manipulator (mtm), however, the fault re-occurred. Before further troubleshooting steps could be taken, the site made the decision to convert the procedure to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation, conducted by field service engineering, found that the system error code #20013 experienced by the customer was associated with the right master tool manipulator (mtmr). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #20013 system error code appeared when the da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state. " the mtmr was returned to isi for failure analysis investigation. Engineering evaluation found that the potentiometer had malfunctioned, thus generating the system errors experienced by the customer. The mtml's motor assembly was also replaced as an additional precaution. As of (b)(4), there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1448939 | 1313224 | 1251388 | 2955842200900065 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Service Personnel | IS1200 A43P9 | Yes | No | 2009-01-20 | Yes | No | Reuse | 2009
42004905 | 2009-01-15 | Malfunction |  | It was reported that during a da vinci s surgical procedure, pieces were observed falling off of the cadiere forceps instrument; however, nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed damage to the main tube where multiple deep scratches with material removed were found within an area extending approximately 1. 250 inches from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1466362 | 1350645 | 1288196 | 2955842200900063 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-02-13 | 2009-02-13 | No | Yes | Health Professional | 42004905 | Yes | No | 2009-01-20 | Yes | No | Reuse | 2009
42018404 | 2009-01-16 | Malfunction |  | It was reported that during a da vinci s mitral valve repair surgical procedure, the tip of the permanent cautery spatula instrument broke off and fell inside of the pt while the surgeon was cauterizing the opening of the pericardium. The broken tip was retrieved and discarded. The planned procedure was completed with a replacement instrument of the same type and without significant delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned with the proximal clevis completely detached from the main tube. Per the customer reported complaint, engineering evaluation observed that the proximal clevis and wrist are missing from the instrument. The main tube interface wall is also missing pieces. Engineering concluded that the failure mode experienced by the customer was due to a break at the proximal clevis to main tube interface and most likely caused by excessive side loading at the tip. Engineering determined that the wrist was possibly cut off in order to remove the instrument from the cannula accessory. Per the case notes, the broken instrument component was successfully retrieved from the pt. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1466369 | 1350647 | 1288198 | 2955842200900064 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-02-13 | 2009-02-13 | No | Yes | Health Professional | 42018404 | Yes | No | 2009-01-20 | Yes | No | Reuse | 2009
42017303 | 2008-11-14 | Malfunction | Required Intervention | It was reported that during a da vinci s cardiac surgical procedure, the blank cartridge used in conjunction with endopass delivery instrument fell into the pt while passing v-clips during the coronary anastomosis. A right thoracoscopy on the left side of the pt's chest was performed using the system and fluoro x-ray in order to search for the cartridge. The cartridge was retrieved and the planned procedure was completed with a 2 hour delay. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the metal shaft has a deformation at the hole feature. The metal is indented on the proximal side of the hole and bulging outwards on the distal side of the hole. Engineering determined that the deformation may have been due to excessive loading on the cartridge and possibly due to the cartridge not being properly installed on the instrument shaft. The instrument housing was removed and engineering discovered that the rear bearing was dislodged from the guide plate and the roll cables were derailed, allowing the shaft to have excess play in all directions. Engineering concluded that the bearing dislodgement led to derailment of the cables, thereby causing excess play in the shaft and most likely contributing to the cartridge falling off. No other damage was found. Per the case noted, the cartridge was successfully retrieved from the pt. | ENDOPASS DELIVERY INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1467663 | 1350649 | 1288200 | 2955842200900072 | 1 | NAY | Manufacturer | User facilityHealth Professional | Nurse | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Health Professional | 42017303 | Yes | Yes | 2009-01-22 | Yes | No | Reuse | 2009
ISI200 A43P9 | 2009-01-15 | Other | Other | It was reported that prior to starting the da vinci surgical procedure, the surgical staff noticed that the right master tool manipulator (mtmr) was broken. The pt had been under anesthesia when the surgeon decided to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | During the investigation conducted by the field service engineer, system error code #20003 was found in the error logs, indicating that the mtmr had malfunctioned. The system was repaired by replacing the affected mtmr. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). System error code #20003 appears when the da vinci safety system determines that the measured angular position of one or more robotic joints on the specified manipulator deviated from the commanded position by more than a specified tolerance. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering observed that the main link tube was missing both connection clips and engineering also found that an axis potentiometer was out of specification. Both issues could generate system error code #20003. As of february 5, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1486061 | 1371652 | 1308363 | 2955842200900075 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Health Professional | ISI200 A43P9 | Yes | No | 2009-01-15 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-01-14 | Other | Other | It was reported that prior to starting a da vinci s mitral valve repair surgical procedure, no video was displayed on the left monitor of the surgeon's console. The pt had been under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that the vision issue experienced by the customer was associated with the high resolution stereo viewer (hrsv). The system was repaired by replacing the affected hrsv. The hrsv is located on the surgeon console and provides the video image for the surgeon console operator. The hrsv was returned to isi for evaluation. Engineering discovered that the left side monitor of the hrsv had malfunctioned, thus causing the reported vision issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1486050 | 1371653 | 1308364 | 2955842200900073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-01-14 | Yes | No | Reuse | 2009
400110 | 2009-01-15 | Malfunction |  | It was reported that during a da vinci prostatectomy surgical procedure a piece of the precise bipolar forceps instrument fell into the pt. No additional info was provided. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was not returned for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post evaluation), and/or, if additional info is received. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1489094 | 1371963 | 1308674 | 2955842200900078 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2009-02-13 | 2009-02-13 | No | Yes | Health Professional | 400110 | No | No | 2008-01-15 | Device Not Retu | No | Unkown | 2009
IS2000 A51P7 | 2009-01-21 | Other | Other | It was reported that during a da vinci s cholecystectomy procedure, the left master tool manipulator (mtml) would not align with the instrument installed on the corresponding system arm. With the assistance of a isi technical support engineer, the customer reseated the sterile adapter, re-inserted the instrument, used another instrument, restarted the system, engaged a different system arm, and re-docked the system to the pt, however, the misalignment issue continued to occur. The pt was under anesthesia for approximately forty minutes when the surgeon decided to convert to a laparoscopic procedure to complete the planned surgical procedure. No pt harm was reported. | The investigation conducted by an isi field service engineer found the system to be performing in accordance to manufacturer specifications. System tests were performed and no errors were generated. Engagement and motion were tested in a wide variety of positions, with no unexpected results. Engineering concluded that the master tool manipulator alignment issue experienced by the customer was likely due to pt position at the pt at the pt side cart (psc). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). As of (b)(6) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2578426 | 2548242 | 2444935 | 2955842200900071 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2009-02-13 | 2009-02-13 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-01-21 | Yes | No | Reuse | 2009
| 2008-04-11 | Injury | HospitalizationOtherRequired Intervention | The patient was reevaluated 4 days after hysterectomy for abdominal pain and distension. She was found to have a perforated sigmoid colon. She was brought emergently to the operating room where she underwent the repair of perforation and an end-colostomy. The patient was admitted to the intensive care unit on ventilator and blood pressure support medications. She had an extensive stay in the intensive care unit where she suffered from bilateral pleural effusions and pneumonia. She was eventually able to be transferred to the general medical floor. |  | MONOPOLAR SCISSORS | MONOPOLAR SCISSORS | 1381311 | 1315479 | 1253608 | 1315479 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2009-02-04 | 2009-02-09 | Yes | No | Physician |  | No | No Answer Provided |  |  |  |  | 2009
| 2008-04-11 | Injury | HospitalizationOtherRequired Intervention | The patient was reevaluated 4 days after hysterectomy for abdominal pain and distension. She was found to have a perforated sigmoid colon. She was brought emergently to the operating room where she underwent the repair of perforation and an end-colostomy. The patient was admitted to the intensive care unit on ventilator and blood pressure support medications. She had an extensive stay in the intensive care unit where she suffered from bilateral pleural effusions and pneumonia. She was eventually able to be transferred to the general medical floor. |  | MONOPOLAR SCISSORS | MONOPOLAR SCISSORS | 1381311 | 1315479 | 1253608 | 1315479 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2009-02-04 | 2009-02-09 | Yes | No | Physician |  | No | No Answer Provided |  |  |  |  | 2009
40018005 | 2008-12-23 | Malfunction |  | It was reported that approximately 2 hours into a da vinci s prostatectomy procedure, while cauterizing the peritoneum, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. The planned surgical procedure was completed without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and mcs instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for evaluation a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1487051 | 1371267 | 1307978 | 2955842200900070 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 40018005 | No | No | 2009-01-14 | Device Not Retu | Yes | Unkown | 2009
42020504 |  | Malfunction |  | It was reported that the black tip cover of the fenestrated bipolar forceps instrument cracked and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not find any cracks on the instrument, however, engineering observed that one side of the distal end of main tube has a 1. 400" long section with material removed. The damaged area is located near the intersection with the proximal clevis, parallel to tube axis and has a rough surface finish. The damage could be what the customer referred to as the "cracked" black tip cover. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Deep scratches with material removed were also noticed around the same area. The wear pattern of the deep scratches suggests that the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care, avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1488078 | 1371269 | 1307980 | 2955842200900060 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42020504 | Yes | No | 2009-01-13 | Yes | No | Reuse | 2009
400180 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and mcs instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1487066 | 1371270 | 1307981 | 2955842200900068 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 400180 | No | No | 2009-01-14 | Device Not Retu | Yes | Unkown | 2009
42000605 |  | Malfunction |  | It was reported that the black tip cover of the large needle driver instrument cracked and nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not find any cracks on the instrument, however, engineering observed that one side of the distal end of main tube has a 1. 65" long section with material removed. The damaged area has a rough surface finish and contains some deep scratches. The damage could be what the customer referred to as the "cracked" black tip cover. Based on the location and appearance, the damage was most likely caused by a combination of instrument collisions and/or rough handling and a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1532861 | 1371271 | 1307982 | 2955842200900062 | 1 | GCJ | Manufacturer | OtherUser facility | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42000605 | Yes | No | 2009-01-13 | Yes | No | Reuse | 2009
40018005 | 2008-12-19 | Malfunction |  | It was reported that during a da vinci s nephrectomy procedure, the tip cover accessory developed a pin hole while installed on the monopolar curved scissors (mcs) instrument. The mcs instrument was removed and the tip cover was replaced. The procedure was completed and no adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 010" diameter hole burned through the silicone. The hole is located. 285" above the silicone to sleeve interface. The tip cover also has a. 060 inch long mechanical tear in the axial direction, adjacent to the hole. No other damage was found. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by instrument collisions and/or rough handling. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1532867 | 1371273 | 1307984 | 2955842200900057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-01-12 | Yes | Yes | Initial | 2009
42020504 | 2009-01-13 | Malfunction | Other | It was reported that prior to use, a small hole was observed at the tip of the fenestrated bipolar forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one small circular char mark at the grip base of one yaw pulley. This circular char mark is what the customer referred to as a "small hole at tip. " the charring may be due to prolonged cautery activation without tissue between the grips. The grip shank has scratches which possibly may have caused some of the yaw pulley glue and/or insulation to come off, creating a path for arcing. High generator settings may also contribute to damage. Electrical continuity passed. No other damage found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1911292 | 1371304 | 1308015 | 2955842200900077 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42020504 | Yes | Yes | 2009-01-16 | Yes | No | Reuse | 2009
42020906 |  | Malfunction |  | It was reported that the tips of the suturecut needle driver instrument would not align. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the scissors to cut cleanly through ethibond 2-0 suture. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly and is fully functional. The customer reported complaint could not be confirmed. Engineering also observed deep scratches on the main tube with material removed. The damage is located in a section extending approx 1. 825" from the intersection with the proximal clevis. Based on the location and appearance, the damage is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1512344 | 1371305 | 1308016 | 2955842200900059 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42020906 | Yes | Yes | 2009-01-13 | Yes | No | Reuse | 2009
400180 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The tip cover accessory and mcs instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1510769 | 1371312 | 1308023 | 2955842200900067 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 400180 | No | No | 2009-01-14 | Device Not Retu | Yes | Unkown | 2009
42009307 |  | Malfunction |  | It was reported that the black tip cover of the endowrist prograsp instrument cracked and nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not find any cracks on the instrument, however, engineering observed a few deep scratches with material removed on the distal end of the main tube. The damage could be what the customer referred to as the "cracked" black tip cover. The scratches are short in length and not aligned with the tube axis, suggesting that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1510776 | 1371314 | 1308025 | 2955842200900061 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42009307 | Yes | No | 2009-01-13 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that the end of the maryland bipolar forceps instrument was pushed in. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the luer plate to be dislodged from the chassis and pushed into the back end. The chassis feature, which retains the luer plate is broken at the inner wall. The chassis likely broke from application of a non-axial force at the flush port. The breakage is likely to have occurred due to mishandling and/or misuse during the cleaning process. Electrical continuity passed. Engineering also observed the distal end of the main tube to have a 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted, if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1487067 | 1371341 | 1308052 | 2955842200900058 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2009-01-13 | Yes | No | Reuse | 2009
400180 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The tip cover accessory and mcs instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If the accessory or instrument are returned for evaluation, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1487058 | 1371342 | 1308053 | 2955842200900066 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 400180 | No | No | 2009-01-14 | Device Not Retu | Yes | Unkown | 2009
42009307 |  | Malfunction |  | It was reported that the tips of the endowrist prograsp instrument no longer work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the instrument moved intuitively with full range of motion and in all directions. The grips open and close properly. Engineering was unable to replicate the customer reported problem. Engineering observed the distal end of the main tube has various deep scratches with material removed. The scratched are short in length and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1487043 | 1371343 | 1308054 | 2955842200900056 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-02-09 | 2009-02-09 | No | Yes | Health Professional | 42009307 | Yes | No | 2009-01-12 | Yes | No | Reuse | 2009
IS2000 A51P7 | 2009-01-15 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the surgeon could not move the endoscopic camera manipulator (ecm) arm. The surgeon decided to convert to tradition open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering found system error code #23007 in the system error logs, and was associated with the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code # 23007 appears when the da vinci s safety system determines, on startup, that one or more robotic joints on the manipulator did not make the prescribed test motion to within the specified tolerance. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". As of february 3, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1487089 | 1371600 | 1308311 | 2955842200900074 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2009-02-09 | 2009-02-09 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-01-15 | Yes | No | Reuse | 2009
IS2000 A51P6 | 2008-01-15 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced unrecoverable system error code #23002. With the assistance of an isi technical support engineer (tse), the site exercised the grip on the left master tool manipulator (mtml), however, the error persisted. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. No pt harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23002 was associated with the left master tool manipulator. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23002 appears when the da vinci s safety system determines that the position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer) were beyond mechanical limits. Upon determining this condition, the safety systems put davinci in a "recoverable safe state". The mtml was returned to isi for failure analysis investigation. Engineering found system error code #23002 in the system error log, however, during their evaluation of the mtml, they could not duplicate the system error code. A sensor and circuit board were replaced as a precaution for the reported complaint experienced by the customer. As of 02/06/2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2642321 | 2605036 | 2501729 | 2955842200900101 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-02-09 | 2009-02-09 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-01-15 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2009-01-06 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, system error code #20105 occurred. An isi technical support engineer directed the surgical staff to restart the system in an attempt to resolve the issue, however, system error code #20105 re-appeared during start-up. The pt had been under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule it for a later date. No pt harm, adverse outcome or injury was reported. | The investigation conducted by the isi field service engineer concluded that system error code #20105 was associated with one of the remote interface adapter printed circuit assembly (ria pca) boards. An ria pca is assigned to each system arm and performs numerous functions related to the function, control and hardware fault monitoring for its assigned arm. The system was repaired by replacing the affected ria pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #20105 appears when an auxiliary voltage error is detected during self-test. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The ria pca was returned to isi for failure analysis eval, however, engineering was unable to reproduce the customer reported failure mode. As of february 3, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1455561 | 1321294 | 1259365 | 2955842200900046 | 1 | NAY | Manufacturer | User facility |  | Initial | 2009-02-04 | 2009-02-04 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2009-01-06 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument was not working. The planned procedure was completed with no pt harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on an in-house system for testing. Recognition and engagement passed and the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly and electrical continuity passed multiple tests while moving or repositioning the grips. Engineering found the instrument to be fully functional and could not confirm the customer reported failure mode. Results and conclusions - engineering evaluation also observed that the distal end of the main tube has sections with material removed, located within 3. 5" from the intersection with the proximal clevis. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440524 | 1302609 | 1241580 | 2955842200900044 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42017206 | Yes | No | 2009-01-06 | Yes | No | Reuse | 2009
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument would not work. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed. The instrument is fully functional. Engineering also observed that the distal end of main tube has a 2. 876" long section with material removed on one side of the tube near the intersection with the proximal clevis. There are also deep scratches within an area extending approximately 2" from the same intersection where material was removed from the damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damaged were most likely caused by a combination of cannula accessory and instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440676 | 1302610 | 1241581 | 2955842200900049 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42009307 | Yes | No | 2009-01-07 | Yes | No | Reuse | 2009
42022702 | 2008-12-04 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the orange part of the pk dissecting forceps instrument cautery piece cracked, however, nothing fell into the pt. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not observe any cracks on the orange connector at the backend. The connector is properly seated and does not exhibit obvious damage. Engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a. 160" x. 060" piece on the opposite side of the conductor break, allowing the grip to move within the pulley and have a larger range of yaw motion. The added range of grip motion likely created excess tension on the wire, causing it to break. Per the cause notes, no instrument components fell inside of the pt during the surgical procedure. Engineering also observed a couple of deep scratches with light material removed on the distal end of the main tube. The scratches are not aligned with the tube axis. The location and appearance of the scratches suggests that they may have been caused by instrument collisions and/or rough handling. No other damage was found. Conclusions- the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440683 | 1302611 | 1241582 | 2955842200900054 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42022702 | Yes | No | 2009-01-09 | Yes | No | Reuse | 2009
42020705 |  | Malfunction |  | It was reported that the blue and gray part on the housing of the tenaculum forceps instrument came apart, because the clamp that holds them together will not hold. No additional info was provided. No pt harm, adverse outcome or injury was reported. | Results and conclusions- this instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the housing is broken. Three of four housing snap tabs and two adjacent housing pins are broken. One rear snap is completely broken off and two front snaps have the bottom tab feature broken. As a result, the housing can be lifted up on one side. The instrument may have been mishandled. Results and conclusions - engineering also observed the distal end of the main tube to have various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis. Based on the location and appearance, the damage is most likely due to instrument collisions and/or rough handling. No other damage found. Conclusions- the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440569 | 1302612 | 1241583 | 2955842200900055 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42020705 | Yes | No | 2009-01-09 | Yes | No | Reuse | 2009
42019006 |  | Malfunction |  | It was reported that the insulation on the endowrist cobra grasper instrument is splintering. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found deep scratches in an area extending approximately. 500" from the intersection with the proximal clevis where material was removed from the main tube. Engineering also noticed tube damage in different sections within 3" from the same interface where material was removed and fibers are exposed. Based on the location and appearance, the damage was most likely caused by instrument collisions combined with rough handling. Engineering also conducted performance tests and found that the grips would not close properly due to the large teeth not meshing properly. The grips were inspected under magnification and one grip hub was found to have a crack above the swaged section that is causing an offset at the tips. Crack may be due to a void in the grip or caused by mishandling or misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440589 | 1302613 | 1241584 | 2955842200900048 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42019006 | Yes | No | 2009-01-07 | Yes | No | Reuse | 2009
12019006 | 2009-12-16 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, parts broke off of the cobra grasper instrument. No harm to the pt occurred and the surgery was slightly delayed while the instrument was replaced. No additional info was provided. No adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the instrument and could not locate any broken or missing components. Engineering observed the distal end of the main tube has various scratch marks with light material removed. The scratches are small in length and not aligned with the tube axis, with one large scratch mark parallel to the tube axis. Based on the location ans appearance, the damage was most likely caused by instrument collisions and or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440643 | 1302614 | 1241585 | 2955842200900043 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 12019006 | Yes | No | 2009-01-06 | Yes | No | Reuse | 2009
42017908 | 2008-12-23 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the discs on the back of the monopolar curved scissors instrument were hard to turn. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned for evaluated. Per the customer reported complaint engineering manually rotated the instrument input discs and found that they moved with ease. Engineering observed that the scissor blades do not cut through their entire length due to one blade edge having an indention halfway between the midpoint and tip that creates a mechanical stop for the other blade and closing. Engineering concluded that the damage to the blade was likely caused by mishandling and or misuse. Pitting of the blades was also observed, likely caused by chemicals caused in the cleaning process. Two char marks on the proximal clevis were found that suggest an arcing event occurred. A tip cover accessory was not returned with the instrument for evaluation, therefore, the root cause of the arcing cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post evaluation) and/or if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440608 | 1302615 | 1241586 | 2955842200900045 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42017908 | Yes | No | 2009-01-06 | Yes | No | Reuse | 2009
40019004 | 2008-12-11 | Malfunction |  | It was reported that prior to starting a da vinci surgical procedure, the cobra grasper instrument was not recognized by the system. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaints, engineering placed the instrument on a system and found it was successfully recognized, however, the grips could not close properly due to the large teeth not meshing. Magnification showed that one grip hub has a hair line crack above the swaged section, causing the offset at the tips. Also observed were two housing snaps on one side to be broken off and the housing is cracked at the back luer surface. Engineering concluded that the grip hub crack and broken housing were likely caused by mishandling and or misuse. Engineering also observed the distal end of the main tube has a couple of scratch marks with light material removed. The scratches are small in length and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440622 | 1302616 | 1241587 | 2955842200900050 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 40019004 | Yes | No | 2009-01-07 | Yes | No | Reuse | 2009
42017908 | 2008-12-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the monopolar curved scissors instrument were defective. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to be fully functional. The scissors cut cleanly through. 006" latex and the blades are not damaged electrical continuity passed. Engineering observed the main tube is damaged. There are multiple deep scratches that can be seen within an area extending approximately 3" where material was removed from the main tube. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440669 | 1302617 | 1241588 | 2955842200900052 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2009-01-08 | Yes | No | Reuse | 2009
42009307 | 2008-12-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the endowrist prograsp forceps instrument was defective. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument is fully functional. Engineering also observed the main tube is damaged. There are multiple deep scratches within an area extending approximately 2" from the intersection with the proximal clevis where material was removed from the main tube. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWIRST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440633 | 1302618 | 1241589 | 2955842200900051 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2009-01-08 | Yes | No | Reuse | 2009
42022702 |  | Malfunction |  | It was reported that prior to starting a da vinci s hysterectomy procedure the pk dissecting forceps instrument was observed to have frayed wires. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip open cable to have fraying between the tip and the distal idlers. The grips open and close without issues. The proximal clevis, outer distal pulley, and the grip, located on the same side as the damaged cable, present damage and scratch marks indicative of instrument collisions that may have also caused damage to the cable. Engineering also observed a 2. 250" section where material was removed form the main tube. The damage area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. The instrument also has a deep scratch extending approximately 1. 625" from the intersection with the proximal clevis, with material removed from the main tube. Based on the location and appearance, the damage is most likely due to instrument collisions and/or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1439097 | 1302619 | 1241590 | 2955842200900053 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-09 | 2009-02-02 | No | Yes | Health Professional | 42022702 | Device Returned To Manufacturer | No | 2009-01-08 | Yes | No | Reuse | 2009
ISI200 A43P9 | 2009-01-05 | Other | Other | It was reported that during a da vinci prostatectomy surgical procedure, the cautery cable connecting the surgeon side console and the electrosurgical unit (esu) pulled apart at the connector. An isi technical support engineer indicated that the esu footswitch could be used to continue the procedure, however, the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | Evaluation of the cautery cable was not performed since the recommended esus and corresponding electrocautery cables to be used with the system are not manufactured by isi and the proper handling and maintenance of the esus and cables are the responsibility of the customer. As of january 28, 2009, there have been no reported recurrence of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1441490 | 1302873 | 1241844 | 2955842200900040 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-01-30 | 2009-02-02 | Yes | Yes | Health Professional | ISI200 A43P9 | No | No | 2009-01-05 | No | No | Reuse | 2009
IS2000 A51P7 | 2009-01-06 | Other | Other | It was reported that prior to starting a da vinci s hysterectomy surgical procedure, the customer experienced multiple system error code #23008. The site called an isi technical support engineer to assist in trouble shooting the issue and was advised that system error code #23008 may reoccur as a nonrecoverable error. The patient had been under anesthesia for approximately 20 minutes when the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #23008 was associated with the left master tool manipulator. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering found system error code #23008 in the system error log, however, the evaluation of the mtml could not duplicate the system error code. A potentiometer was replaced as a precaution for the reported complaint experienced by the customer. As of january 29, 2009, there have been no reported recurrence of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1441496 | 1302876 | 1241847 | 2955842200900047 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2009-01-29 | 2009-02-02 | Yes | Yes | Health Professional | IS2000 A51P7 | Yes | No | 2009-01-06 | Yes | No | Reuse | 2009
42018404 | 2008-10-21 | Malfunction |  | On 1/5/09, isi received maude event report from the fda: event description: "during robotic hysterectomy, when changing out the unipolar paddle to the needle driver on the right, the paddle became stuck at the end of the trocar. The surgeon was eventually able to remove the trocar and instrument. Upon removal, a 4mm x 5mm piece of sterile plastic appeared to be missing from the instrument. A thorough inspection was conducted but the piece was not found. The surgeon is unsure if the piece was missing at the beginning of the case. ". | The instrument had not been returned to isi for eval, therefore, the root cause of the customer reported problem cannot be determined. A follow-up mdr will be submitted if the instrument is returned (upon completion of eval) or if additional info is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1285680 | 1323618 | 1177273 | 2955842200900039 | 1 | GEI | Manufacturer | Other | Invalid Data | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 42018404 | No | No Answer Provided | 2009-01-05 | Device Not Retu | No | Unkown | 2009
40016906 | 2008-12-16 | Other |  | On (b)(6) 2009, isi received user facility medwatch #(b)(4) from the fda: event description: "during a laparoscopic robotic procedure, half of a harmonic scalpel tip broke off at the hinge. The piece that broke was approx 15 mm long and 22 mm thick. The physician performed an eval but could not observe the tip. Due to the location of this surgery (peri-aortic region) the physician decided to leave the device in pt at that time. The device was visible upon x-ray. It is uncertain if the retained piece of device will be removed at a later time or left in. The pt is to f/u with her surgeon". | The harmonic curved shears insert accessory was not returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info becomes available a f/u mdr will be submitted. There are no components in the harmonic curved shears inserts that meet the definition of "medical sharps" per the customer reported complaint, it has been concluded that the "scalpel" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. | HARMONIC CURVED SHEARS INSERT | ELECTROSURGICAL ACCESSORY | 2646197 | 2613690 | 2510386 | 2955842200900041 | 1 | NAY | Manufacturer | Other | NOT APPLICABLE | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 40016906 | No | No | 2009-01-05 | Device Not Retu | No | Unkown | 2009
40016906 | 2008-12-23 | Other | Required Intervention | On (b)(6) 2009, isi received user facility medwatch # (b)(4) from the fda: event description: "during a robotic hysterectomy being performed a tip of the harmonic insert dislodged in the abdomen. Unable to extract tip laparoscopically therefore a mini incision was performed and the tip was retrieved. The pt required a higher level of surgical care as a result of the event. This is the second harmonic tip that has broken during surgery within one week. Immediate review by the mfr will be anticipated". | The insert accessory was returned and evaluated. Engineering found the clamp arm detached and the hinge features on one side of the insert, that mate with the clamp arm pins, are bent backwards. This deformation is consistent with the hyperextension of the clamp arm. Engineering also observed scratch marks near the hinge which is indicative of collisions or rough handling. Engineering concluded that the damage is likely due to mishandling and/or misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ELECTROSURGICAL ACCESSORY | 2646137 | 2613692 | 2510388 | 2955842200900042 | 1 | NAY | Manufacturer | Other | NOT APPLICABLE | Initial | 2009-01-30 | 2009-02-02 | No | Yes | Health Professional | 40016906 | Yes | No | 2009-01-05 | Yes | No | Unkown | 2009
IS1200 A43P9 | 2008-12-29 | Other | Other | It was reported that 1. 5 hours into a da vinci prostatectomy procedure, the left eye image in the high resolution stereo viewer (hrsv) was flickering. With the assistance of an intuitive surgical technical support representative, the site first changed the optical view to 2-d, however, the image continued to flicker on and off. The site then raised and lowered the hrsv, however, the image quality did not stabilize. The site then shut down the surgeon side console and reseated all the cables, however, the problem was not resolved when the system was restarted. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure after an hour of troubleshooting measures were taken. No patient harm was reported. No patient harm, adverse outcome or injury was reported. | Results/conclusion - an investigation conducted by the field service engineer concluded that the vision issues experienced by the customer were associated with the high resolution stereo viewer (hrsv) and the printed circuit assembly (pca) power supply. The high resolution stereo viewer (hrsv) is located on the surgeon console and provides the video image for the surgeon console operator. The pca power supply supplies power to the da vinci's electronics systems. The system was repaired by replacing the affected hrsv and pca power supply. As of january 23, 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1483308 | 1366035 | 1302979 | 2955842200900032 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-01-23 | 2009-01-23 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-12-29 | Yes | No | Reuse | 2009
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument had a broken cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one pitch cable broken, resulting in imprecise grasping and preventing the instrument from rotating. Engineering concluded that cable wear on the proximal clevis combined with high grasping force likely contributed to the cable breakage. Engineering also observed the distal end of main tube has a 1. 375 inch long section with material removed on one side, near the intersection with the proximal clevis. The damaged area is parallel to the tube axis and has a rough surface finish. Deep scratches were found in a section extending approximately. 500 inches from the same intersection where material is removed from the main tube. Based on the location and appearance, the damages were most likely caused by a combination of a cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1483758 | 1366235 | 1303168 | 2955842200900030 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-12-29 | Yes | No | Reuse | 2009
4217206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the grip of the maryland bipolar forceps instrument would not close. The planned surgical procedure was completed with another instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip is bent, causing side to side misalignment of grips. There is a. 039" offset at the tips. The bent grip has separation of the yaw pulley and yaw pulley cover at the glue joint, indicating overloading at the tip. Electrical continuity passed. Engineering also observed that the distal end of main tube has a 3. 75" long section, extending from the interface with the proximal clevis, with material removed on one side of the tube, and an adjacent 3. 5" section has material removed around the tube was also observed. The damaged areas are parallel to tube axis and have a rough surface finish. The main tube also has deep scratches within a 2" section near the same interface where material was removed from one side of the tube. Based on the location and appearance, the damage was most likely caused by a combination of cannula accessory and instrument collision. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1483738 | 1366236 | 1303169 | 2955842200900036 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 4217206 | Yes | No | 2008-12-31 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that after a da vinci s prostatectomy procedure, the plug of the maryland bipolar forceps instrument was dislocated and broke when the bipolar cord was removed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the plug riser insert was removed from back end as well as one of the two conductor wire pins. The components returned with instrument. The remaining pin and conductor wires are sticking out of chassis. Engineering concluded that the insert and the damaged pin were likely pulled out when the bipolar cord was removed. Electrical continuity passed. Engineering also observed that the distal end of the main tube has multiple areas within a 6. 6" section, from the intersection with the proximal clevis, where material was removed. The damaged areas are parallel to tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1483747 | 1366237 | 1303170 | 2955842200900035 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-31 | Yes | No | Initial | 2009
42009307 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of the endowrist prograsp forceps instrument fell off, but did not fall into the pt. No additional information was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the luer plate is dislodged from the chassis and both the plate and the flush tube have been pulled out of the back end. The chassis feature which retains the luer plate is broken at the outer wall. The chassis most likely broke from non-axial force at the flush port. Engineering concluded that the chassis breakage is likely occurred due to mishandling/misuse during cleaning process. Engineering also observed the distal end of main tube has multiple areas within 6" from the intersection with the proximal clevis, with material removed. The damaged areas are parallel to tube axis and have a rough surface finish. The main tube also has multiple deep scratches within 3" from the same intersection where material was removed. Based on the location and appearance, the damage was most likely caused by a combination of a cannula accessory and by instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1486119 | 1366257 | 1303190 | 2955842200900031 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-12-29 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that the casing of the maryland bipolar forceps instrument broke. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that three of the four housing snap tabs and two adjacent housing pins were broken. Engineering concluded that the instrument may have been mishandled or the housing material may have become brittle as a result of the site's cleaning and sterilization process. Engineering also observed various deep scratches with material removed on the distal end of the main tube. The scratches are all under. 250" long and not aligned with the tube axis. The location and appearance of the scratches suggests that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2868529 | 1366290 | 1303223 | 2955842200900034 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-30 | Yes | No | Reuse | 2009
40020504 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the wire was pulled out of the fenestrated bipolar forceps instrument and nothing fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a. 140" x. 060" piece on the opposite side of the conductor break, allowing the grip to move within the pulley and have a larger range of yaw motion. The cause of the yaw pulley breakage cannot be determined, however, it was concluded that the added range of grip motion likely created excess tension on the wire and caused it to break. Per the case notes, no instrument components fell into the pt during the surgical procedure. Engineering also observed various deep scratches with material removed on the distal end of the main tube. The scratches are not aligned with the tube axis. The location and appearance of the scratches suggests that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1486190 | 1366291 | 1303224 | 2955842200900033 | 1 | GEI | Manufacturer | User facilityOther | Nurse | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 40020504 | Yes | Yes | 2008-12-30 | Yes | No | Reuse | 2009
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the "wrist joint" of the monopolar curved scissors instrument dislocated when trying to remove the instrument from the cannula accessory. The planned surgical procedure was completed with another instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys, as a result, the tube extension can be rotated 360 degrees. There are no pieces missing. Engineering concluded the damage was likely caused by excessive side loading. Electrical continuity passed. Engineering also observed that the distal end of the main tube has a 3" long section with material removed from around the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1482469 | 1366292 | 1303225 | 2955842200900037 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-31 | Yes | No | Reuse | 2009
42009307 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the rotation of the wrist of the endowrist prograsp forceps instrument was limited. The planned surgical procedure was completed with another instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests but could not confirm the reported complaint. Engineering also observed that the distal end of the main tube has a 3. 2" long section with material removed from different parts of the tube. The damaged areas are parallel to tube axis and have a rough surface finish. The main tube also has deep scratches within a 1" section extending from the interface with the proximal clevis where material was removed from one side of the tube. Based on the location and appearance, the damages were most likely caused by a combination of a cannula accessory and instrument collision. No other damage found. Additional investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1486179 | 1366293 | 1303226 | 2955842200900038 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-23 | 2009-01-23 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-12-31 | Yes | No | Reuse | 2009
40018005 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, electrical leakage occurred when using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The surgical staff also observed that the tip cover was burnt. No pieces fell into the patient and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument and two tip cover accessories were returned and evaluated. Per the customer reported complaint, one tip cover has a. 035" x. 115" hole at the silicone to sleeve interface. The silicone and plastic sleeve exhibit localized melting/charring, indicating that an arcing event occurred. At one end of the hole there is a small mechanical tear in the silicone, which may have been caused by instrument collisions, thereby, creating a path for the arcing to occur. The damaged area on the tip cover matches closely with one of the two char marks found on the proximal clevis of the mcs instrument. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength possibly weakened by damage from instrument collisions. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The additional tip cover accessory returned from the site also shows evidence of arcing. It cannot be determined which tip cover was used during the reported event, therefore, mfr report #2955842-2009-00028 was submitted for the initial report of arcing from the site and includes the detailed failure analysis results of the other tip cover accessory returned. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INTRUMENT ACCE | 1482697 | 1365511 | 1302484 | 2955842200900029 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2009-01-19 | 2009-01-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2009-01-14 | Yes | Yes | Initial | 2009
40018005 | 2008-12-10 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, electrical leakage occurred when using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed. The surgical staff also observed that the tip cover was burnt. No pieces fell into the patient and the planned procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument and two tip cover accessories were returned and evaluated. Per the customer reported complaint, one tip cover was returned with the silicone detached from the plastic sleeve. When the pieces are put together, the tip cover has a. 135" x. 270" inch hole at the silicone to sleeve exhibit localized melting/charring, indicating that an arcing event occurred. Scratch marks, which were likely caused by instrument collisions, were observed on the plastic sleeve directly below the charred section. The damaged area on the tip cover matches directly with one of the two char marks found on the proximal clevis of the mcs instrument. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength possibly weakened by damage from instrument collisions. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The additional tip cover accessory returned from the site also shows evidence of arcing. It cannot be determined which tip cover was used during the reported event, therefore, mfr report #2955842-2009-00029 was submitted for the second occurrence of arcing, which was discovered during failure analysis of the additional tip cover accessory returned from the site. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1482689 | 1365513 | 1302486 | 2955842200900028 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2009-01-19 | 2009-01-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-12-17 | Yes | Yes | Initial | 2009
42017206 |  | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff observed that the prongs were bent on the housing of the maryland bipolar forceps instrument. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the top plug riser pin is pushed into the back end and the bottom pin is slightly bent. The chassis feature that retains the bipolar insert and pins is broken. Engineering determined that the chassis feature most likely broke due to mishandling. Engineering evaluation also observed that one side of the distal end of the main tube has a 2" long section with light material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1433626 | 1290218 | 1230370 | 2955842200900026 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2008-12-23 | Yes | No | Reuse | 2009
IS1200 A43P9 | 2008-12-22 | Other | Other | It was reported that approximately 20 minutes into a da vinci sacrocolpopexy procedure, the customer experienced a non recoverable #20009 system error code. With the assistance of an isi technical support engineer (tse), the customer restarted the system, checked the cable connection, exercised the affected psm axis and pressed fault override, however, the #20009 system error code recurred. The pt had been under anesthesia for approximately one hour and 40 minutes when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #20009 experienced by the customer was associated with a pt side manipulator (psm). The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system error code appeared when the primary sensor in the specified manipulator measured an unexpected change in the angular position, resulting in system error code #20009. Upon determining this condition, the safety systems put the da vinci in a "recoverable safe state". The psm was returned to isi for failure analysis investigation and engineering found the psm axis-1 cable was broken, and the bearing and retaining ring was disengaged from the axis-1 second stage drum assembly, causing the customer reported problem. The system was repaired by replacing all axis-1 cables and the retaining ring to the axis-1 second stage drum assembly, as a precautionary measure. As of (b) (6) 2009, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1434704 | 1292206 | 1232257 | 2955842200900023 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-01-16 | 2009-01-16 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-12-22 | Yes | No | Reuse | 2009
42017908 | 2009-11-18 | Malfunction |  | It was reported that during a da vinci s hysterectomy surgical procedure, sparking was observed during dissection in the pelvis and it gave the patient "norvstimulation". No cautery went to the patient colon and the planned procedure was completed with a replacement instrument. No patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1456542 | 1326097 | 1264076 | 2955842200900021 | 1 | GEI | Manufacturer | ForeignUser facility | Physician | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42017908 | No | Yes | 2008-12-03 | Device Not Retu | No | Unkown | 2009
42014204 |  | Malfunction |  | It was reported that the shaft of the 5mm monopolar cautery instrument had shavings falling off. It was reported that the shavings did not fall into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the insulation on the 5 mm main tube to be damaged in a couple of areas located at approximately 6. 6" and 6. 76" from the intersection with the snake wrist. The insulation has small cuts that expose the bare metal tube and material is missing. The damaged tube insulation may be the result of instrument collisions or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1456532 | 1326098 | 1264077 | 2955842200900022 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42014204 | Yes | No | 2008-12-22 | Yes | No | Reuse | 2009
42004904 | 2008-10-30 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, the cadiere forceps instrument did not work. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument is fully functional. Engineering also observed the distal end of the main tube to have a 1. 75" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1456524 | 1326099 | 1264078 | 2955842200900024 | 1 | NAY | Manufacturer | DistributorOther | Other | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42004904 | Yes | No | 2008-12-23 | Yes | No | Reuse | 2009
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument had a broken wire. It was also reported that no pieces have fallen into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to not have any broken or frayed cables at the wrist. The grips open and close properly and the wrist moves fine in pitch and yaw directions when inputs are manually rotated. Engineering also observed that the distal end of the main tube has a 1. 60" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. There is also light wearing above the midpoint of the tube, but this is likely from rough handling or the cleaning process. No other damage found. The endowrist instruments instructions for use (ifu) specially states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1456517 | 1326100 | 1264079 | 2955842200900025 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-12-23 | Yes | No | Reuse | 2009
42009307 |  | Malfunction |  | It was reported that during a da vinci s procedure, the endowrist prograsp forceps instrument was not being recognized by the system. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the system successfully recognized the instrument. Engineering was unable to replicate recognition issue. One of the grip cables is derailed from the force amp pulley, which causes non-intuitive yaw motion. Engineering also observed the distal end of the main tube has a couple of deep scratches where material was removed. The scratches are not aligned with the tube axis. Based on the location and appearance, the damage is most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1456681 | 1326101 | 1264080 | 2955842200900027 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-01-16 | 2009-01-16 | No | Yes | Health Professional | 42009307 | Yes | Yes | 2008-12-24 | Yes | No | Reuse | 2009
IS2000 A51P6 | 2008-11-03 | Injury | Required InterventionHospitalization | It was reported that post a da vinci s hysterectomy procedure, the pt experienced unk complications. Exploratory surgery was performed shortly after and it was found that the pt required a vaginal cuff repair, which was performed the same day. No additional pt harm or adverse outcome was reported. | On december 22, 2008, additional information regarding the event was provided by the site external system manger. It was reported that the exploratory surgery was performed six days prior, and the pt's vaginal cuff was repaired from below. It was noted that the pt's vaginal cuff appeared necrotic, possibly from excessive cautery during the procedure. Per the external system manager the surgeon has since undergone additional proctoring. Based on the information provided, it was determined that the da vinci s surgical system did not malfunction in a way that would cause nor contribute to the pt injury. | ADA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1347345 | 1285724 | 1226594 | 2955842200900020 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-01-09 | 2009-01-09 | Yes | No | Health Professional | IS2000 A51P6 | No | No | 2008-12-12 | Device Not Retu | No | Reuse | 2009
IS1200 A43P9 | 2008-12-19 | Other |  | It was reported that prior to starting a da vinci prostatectomy procedure, the surgeon was unable to move the endoscopic camera manipulator (ecm) arm from the surgeon's console, however, manual position adjustments could still be made using the ecm clutch buttons. With the assistance of an isi technical support engineer, the customer tested the camera clutch foot switch, checked for damaged systems cables, and restarted the system. The system started up and homed normally, however, the ecm continued to be unresponsive. At this time the surgeon decided to abort the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer could not duplicate the customer reported ecm control problem. The ecm is the camera arm located on the patient side chart which provides the sterile interface for the 3d endoscope. As a precaution, the system was repaired by replacing the two printed circuit assembly boards related to the control of the camera arm. As of january 8, 2008, there have been no reported recurrence of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433866 | 1286773 | 1227504 | 2955842200900019 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2009-01-09 | 2009-01-09 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-12-19 | Yes | No | Reuse | 2009
IS2000 A51P6 | 2008-12-17 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a system error code # 23. The site restarted the system twice with no change. With the assistance of an isi technical support engineer (tse), the site attempted to access the error logs, power down the system and reset the breaker. The site also attempted to emergency power off (epo) the console and pt side cart (psc) as well as swap the surgeon console and psc connection cable ends. The system then faulted at power up with system error code # 903. The site powered down 3 additional times, attempted to epo and reset the breaker each time without success. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. | The investigation conducted by field service engineering concluded that the system error codes #23 and # 903 experienced by the customer were associated with a pt side manipulator (psm). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code # 903 is an indication that the error log is full and system error code # 23 is reported by software to denote that the hardware wheel "wdog" has tripped on one of the digital communication links in the system. This means that the system cannot reliably communicate over the digital link and therefore cannot continue normal operation. The psm was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, a printed circuit assembly board related to the transfer of communication signals and an internal cable harness were replaced. As on jan 9, 2009, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433887 | 1286774 | 1227505 | 2955842200900012 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2009-01-09 | 2009-01-09 | Yes | Yes | Health Professional | IS2000 A51P6 | Device Returned To Manufacturer | No | 2008-12-17 | Yes | No | Reuse | 2009
40018005 | 2008-12-15 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found holes burned through the silicone with localized melting, indicative of arcing. A large oblong shaped hole, approximately. 060 inches by. 040 inches in size and located. 215 inches above the silicone to sleeve interface, has a. 100 inch tear bisecting one edge with a small arced hole at one end of the tear. A. 010 inch diameter hole was observed on the opposite side of the oblong hole, located. 225 inches above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occurred. Engineering concluded that tearing may have been caused by instrument collisions or a cannula if the instrument wrist was bent during removal and re-insertion of the mcs instrument. The site also returned the mcs instrument used during the procedure. One proximal clevis ear has a section with a dark discoloration; however, it cannot be confirmed that this is a char mark. With the tip cover installed on the instrument, the discolored area is in the same location as the tip cover silicone damage. The instrument does not have any burrs or sharp edges that may have caused damage to the tip cover. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1453647 | 1316972 | 1255098 | 2955842200900005 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-12-15 | Yes | Yes | Initial | 2009
| 2009-12-10 | Malfunction |  | It was reported that during a da vinci s surgical procedure a piece of an instrument broke off and fell into the patient. No additional information was provided. The planned procedure was completed and not patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to obtain additional information from the customer. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. |  | ENDOSCOPIC INSTRUMENT | 1453637 | 1316973 | 1255099 | 2955842200900001 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional |  | No | No | 2008-12-11 | Device Not Retu | No | Unkown | 2009
| 2009-12-10 | Malfunction |  | It was reported that during a da vinci s surgical procedure a piece of an instrument broke off and fell into the patient. No additional information was provided. The planned procedure was completed and not patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to obtain additional information from the customer. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. |  | ENDOSCOPIC INSTRUMENT | 1453632 | 1316974 | 1255100 | 2955842200900002 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional |  | No | No | 2008-12-11 | Device Not Retu | No | Unkown | 2009
| 2009-12-10 | Malfunction |  | It was reported that during a da vinci s surgical procedure, a piece of an instrument broke off and fell into the patient. No additional information was provided. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Multiple attempts have been made to obtain additional information from the customer. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. |  | ENDOSCOPIC INSTRUMENT | 1453628 | 1316975 | 1255101 | 2955842200900003 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional |  | No | No | 2008-12-11 | Device Not Retu | No | Unkown | 2009
42011005 |  | Malfunction |  | It was reported that the precise bipolar forceps instrument has a wire sticking out and was not identified during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is fraying and is derailed at the distal idler pulley. The cable derailment is most likely due to the cable losing contact with pulley during wrist articulation. The fleet angle between proximal and distal pulleys may have also contributed to derailment. The fraying may be the result of additional friction against the distal pulley's shaft caused by the derailment. Electrical continuity passed. Engineering also observed the distal end of main tube has deep scratches and a 3. 250" long section with material removed on one side of the tube. The damaged areas are parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damages were most likely caused by a combination of a cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453616 | 1316976 | 1255102 | 2955842200900018 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-12-19 | Yes | No | Reuse | 2009
40022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument will not close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side-to-side misalignment of the grips. The bent grip shows separation at the glue joint of the yaw pulley and yaw pulley cover, indicating overloading at the tip. Engineering evaluation also observed a 3. 5" long section with light material removed on the distal end of the main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Multiple deep scratches with material removed were also noticed within an area extending approx 1. 700" from the interface with proximal clevis. Based on the location and appearance, the scratches were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1454646 | 1316977 | 1255103 | 2955842200900014 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 40022702 | Yes | No | 2008-12-18 | Yes | No | Reuse | 2009
42022702 |  | Malfunction |  | It was reported that the wires came out of the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | Result and conclusion: - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable is broken at the distal clevis hub. It appears that the cable most likely broke under tensile loading and may have incurred some damage from instrument collisions and/or rough handling prior to breakage. Engineering also observe various deep scratches with material removed on one side of the distal end of the main tube. The scratches are concentrated in a 1. 5" long section, are all under. 250" long and not aligned with the tube axis. The location and appearance of the scratches suggests that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endo wrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453710 | 1316980 | 1255106 | 2955842200900010 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-12-17 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the system did not recognize the maryland bipolar forceps instrument. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | Result and conclusion: the instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on an in-house test system and the instrument was successfully recognized, with zero uses remaining. When the instrument was removed from the system, engineering observed a pogo pin which sticks during actuation and does not return to its original position. It was concluded that the sticking pogo pin may have caused the recognition issue experience by the customer and improper cleaning techniques may have contributed to the sticking of the pogo pin. Engineering evaluation also observed that one side of the distal end of the main tube has a 2. 1" long section with material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453733 | 1316982 | 1255108 | 2955842200900007 | 1 | NAY | Manufacturer | DistributorOther | NOT APPLICABLE | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-16 | Yes | No | Reuse | 2009
42014304 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the 5mm maryland dissector instrument had a broken shaft and no parts fell into the patient. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | Result and conclusion - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument shaft is not broken, but has an approximate 1" long section of missing insulation all around the tube. The damaged area is located approximately 2" above the snake wrist. The location and appearance of the damage suggests that it may have been caused by instrument collisions and/or rough handling. No other damage was found. Per the complaint notes, no instrument fragments fell into the patient during the surgical procedure. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1453718 | 1316984 | 1255110 | 2955842200900004 | 1 | NAY | Manufacturer | ForeignOther | NOT APPLICABLE | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-12-17 | Yes | No | Reuse | 2009
42022702 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument "cord" broke. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint that the instrument "cord" was broken, engineering found a broken conductor wire. The yaw pulley is also broken and a 290 inch x. 090 inch piece is missing. The distal clevis on the same side as the broken yaw pulley has multiple scratches indicative of instrument collisions. Electrical continuity failed. Engineering also observed deep scratches on the main tube where material was removed. The damage is located within a 2" section extending from the interface with the proximal clevis. Based on the location and appearance, the damage is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1454632 | 1316985 | 1255111 | 2955842200900013 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-12-18 | Yes | No | Reuse | 2009
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument has a wire sticking out and was not identified during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that one grip close cable is broken near the proximal pulleys and proximal clevis hole allowing the pulley to spin freely. The proximal clevis cable hole exhibits wear on the edge. The other cables at the wrist are not damaged. The cable wear on the pulleys and clevis combined with high grasping force likely contributed to the breakage. Engineering also observed deep scratches on the main tube where material was removed. The damage is located in a section extending approximately 2. 250" from intersection with the proximal clevis. Based on the location and appearance, the damage is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1454466 | 1316987 | 1255113 | 2955842200900017 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-12-19 | Yes | No | Reuse | 2009
42020906 |  | Malfunction |  | It was reported that the suturecut needle driver instrument would not cut. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the scissors do not cut cleanly, however,the blades are not damaged. Engineering concluded that the failed shearing contact between blades is due to decreased cable tension and/or (b) (6) force, contributing to poor cutting performance. Engineering also observed deep scratches on the main tube with material removed. The damage is located in a section extending approximately 1. 2" from the intersection with the proximal clevis. Based on the location and appearance, the damage is most likely due to instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1454455 | 1316988 | 1255114 | 2955842200900009 | 1 | GCJ | Manufacturer | User facilityOther | Unknown | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42020906 | Yes | No | 2008-12-16 | Yes | No | Reuse | 2009
42000605 |  | Malfunction |  | It was reported that prior to the start of case, the large needle driver instrument was no recognized by the da vinci s surgical system. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, recognition testing was performed and the instrument was successfully recognized by the system. The pogo pins do not stick and are not contaminated. Verification of the instrument programming chip passed. Engineering also observed that the distal end of the instrument main tube has a 1. 25" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1448652 | 1316989 | 1255115 | 2955842200900015 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-12-18 | Yes | No | Reuse | 2009
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument has a bent shaft. The customer is not aware if this was noticed during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Result and conclusion - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube not to be bent. Engineering also conducted performance tests and found that the scissors won't cut through. 006" latex at the tips, however, the blades are not damaged. Engineering also observed the distal end of the main tube has a 2. 5" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453775 | 1316990 | 1255116 | 2955842200900011 | 1 | GEI | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-17 | Yes | No | Reuse | 2009
42022702 |  | Malfunction |  | It was reported that the tip of the pk dissecting forceps instrument would not open. It was noted that there was no delay in the procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted which found the instrument's grips to open and close properly. Engineering also observed the distal end of the main tube to have a 1. 250" and a 1. 800" section with material removed from the tube. The damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damge was most likely caused by a cannula accessory. No other damage found. Electrical continuity passed. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1448672 | 1316991 | 1255117 | 2955842200900006 | 1 | NAY | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-12-16 | Yes | No | Reuse | 2009
42000605 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the shaft of the large needle drive instrument was splintering. It was also reported that nothing had fallen into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the main tube to have light material removal all around the outer diameter, spread out through the entire length of the tube. The tube damage appears to be repeating patters of marks perpendicular to the tube axis. The unusual wear pattern and amount of surface area damage suggests that the damage may be caused by rough handling or chemicals used in the cleaning process. No other damage found. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1448626 | 1316992 | 1255118 | 2955842200900016 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-12-19 | Yes | No | Reuse | 2009
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgeon was removing the maryland bipolar forceps instrument to exchange for another instrument when the bipolar cable connector broke. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the top plug riser is pushed into the back end and the bottom pin is bent. The chassis feature that retains the bipolar insert and pins are broken. Engineering has concluded that the damage was most likely caused by the instrument being mishandled, resulting in the pins bending and breaking the chassis feature. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering also observed that the distal end of main tube has a 2. 1" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1454436 | 1316993 | 1255119 | 2955842200900008 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2009-01-09 | 2009-01-09 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-16 | Yes | No | Reuse | 2009
400180 | 2008-12-08 | Other | Other | It was reported that approximately 2 hours into a da vinci s hysterectomy procedure, during the pre-aortic node dissection, arcing was seen from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The mcs instrument was immediately removed, and the tip cover was replaced. A small burn was found on the vessel, however, no significant damage was reported and no additional procedure was required. The planned surgical procedure was completed without significant delay. No adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1431381 | 1278401 | 1219972 | 2955842200801477 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | 400180 | No | No | 2008-12-08 | Device Not Retu | Yes | Unkown | 2008
IS1000 A51P4 | 2008-12-01 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced system error code #252. With the assistance of an isi field service engineer, the surgical staff restarted the system three times, however, system error code #252 continued to recur approximately 5-10 minutes after each restart. The surgeon decided to covert the procedure to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by the field service engineer concluded that system error code #252 was associated with the system control processor printed circuit assembly (scp pca) board. The scp pca provides real-time servo processing & control, system supervisor & fault monitoring, and external system communication control. The system was repaired by replacing the affected scp pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #252 appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the safety systems put da vinci s in a "recoverable safe state". The scp pca was returned to isi for failure analysis investigation. Engineering was able to reproduce system error code #252, thus confirming the reported failure mode experienced by the customer. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1431667 | 1278402 | 1219973 | 2955842200801446 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | IS1000 A51P4 | Yes | No | 2008-12-01 | Yes |  | Reuse | 2008
IS200 A51P6 | 2008-11-14 | Other | Other | It was reported that prior to starting a davinci s hysterectomy procedure, the patient side cart (psc) switched to battery power use. With the assistance of an isi technical support engineer, the customer power cycled the psc, reset the circuit breaker, verified the ac power outlet and checked the power cord, however, the psc continued to use power from the battery. The patient was under anesthesia and the port incisions had already been made when the surgeon decided to abort the planned surgical procedure. | The investigation conducted by field service engineering found that the power issue experienced by the customer was associated with the psc power supply switcher. The power supply switcher provides reserve power in the event the or loses power during a procedure. The system was repaired by replacing the affected power supply switcher. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VICNCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1431639 | 1278403 | 1219974 | 2955842200801476 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | IS200 A51P6 | Yes | No | 2008-11-14 | Yes | No | Reuse | 2008
IS1200 A5 1P6 | 2008-12-10 | Other | Other | It was reported that at the start of a da vinci prostatectomy procedure, the vision side cart lost power. The site checked multiple outlets, verified power to the operating room and also verified that all electric cables were fully seated. The surgeon made a decision to delay the procedure in order for the isi field service engineer (fse) to arrive on site to repair the system. The procedure was delayed for approximately an additional hour while the vision side cart issue was being resolved. The planned surgical procedure was completed and no patient harm was reported. | The investigation conducted by an isi field service engineer found a blown fuse on the vision cart. The vision cart contains camera and image processing equipment. The system was repaired by replacing the blown fuse. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1431441 | 1278404 | 1219975 | 2955842200801483 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | IS1200 A5 1P6 | Yes | No | 2008-12-10 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-12-11 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure, the customer experienced an unrecoverable system error code #20013. With the assistance of an isi technician support engineer, the customer exercised the affected patient side manipulator (psm) arm axis, however, the system error code #20013 recurred. The patient was under anesthesia and the port incisions had already been made when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. | The investigation conducted by field service engineering, found system error code #20013 in the system error logs and determined that the error code was associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #20013 appears when the davinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system put da vinci in a "recoverable safe state". The psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1431668 | 1278405 | 1219976 | 2955842200701484 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2008-12-11 | Yes | No | Reuse | 2008
IS2000 A51 P6 | 2008-12-11 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy surgical procedure, the system illuminator light began to dim every time the shutter was opened. An isi representative, in attendance at the case, recommended replacing the illuminator bulb, however, the issue continued to occur after the replacement was made. The patient had been under anesthesia and the port incisions had been made when the surgeon decided to abort the planned procedure and reschedule it for a later date. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that the "dim" light issue experienced by the customer was associated with the system illuminator. The illuminator provides the light which is transported to the system endoscope and projected onto the surgical site. The system was repaired by replacing the affected illuminator. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1429927 | 1278406 | 1219977 | 2955842200801485 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-31 | 2008-12-31 | Yes | Yes | Health Professional | IS2000 A51 P6 | Yes | No | 2008-12-11 | Yes | No | Reuse | 2008
40016906 | 2008-12-08 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, a piece of the harmonic curved shears insert fell into the patient. The piece was retrieved and the insert was replaced with another insert to complete the planned procedure. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Engineering observed a. 350" x. 085" piece broken off of the curved blade. The blade is broken slightly above where it transitions from straight to curved. The cause of the breakage is unknown, however, mishandling/misuse may have contributed to the breakage. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ELECTROSURGICAL ACCESSORY | 1440278 | 1302897 | 1241868 | 2955842200801478 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 40016906 | Yes | Yes | 2008-12-09 | Yes | Yes | Initial | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the system was not recognizing the monopolar curved scissors instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the system successfully recognized instrument. Engineering was unable to replicate recognition issue, however, on the backend of the instrument, the housing has both rear snaps and adjacent locating pins completely broken off. This damage is most likely due to mishandling or the housing becoming brittle due to autoclaving. Engineering also observed the distal end of main tube has a 3. 5" long section with material removed on one side of the tube. Damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440228 | 1302898 | 1241869 | 2955842200801490 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-12 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that fiber glass of the endowrist prograsp instrument is splintering. It was noted that nothing fell into a patient during the uses of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the distal end of main tube has a couple of narrow streaks under one inch in length, with material removed. Two distinct areas located 180 degrees apart also have streaks of material removed. The damaged areas are parallel to tube axis, however, the length and width are smaller than typical cannula related tube abrasions. The spacing between the damaged areas is also uncommon. Engineering concluded that based on the location and appearance, the damage was most likely caused by a burr or localized sharp edge on the cannula accessory inner radius. Rough handling may have also contributed to the main tube damage. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440306 | 1302899 | 1241870 | 2955842200801489 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-12-12 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the system was not recognizing the maryland bipolar forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the system successfully recognized instrument. Engineering was unable to replicate recognition issue, however, on the backend of the instrument, both bipolar pins are bent, preventing proper installation of a bipolar cord. This damage was likely caused by mishandling. Engineering also observed the distal end of main tube has a 3. 6" long section with material removed all around the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1440219 | 1302900 | 1241871 | 2955842200801491 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-12 | Yes | No | Reuse | 2008
42014304 |  | Malfunction |  | It was reported that at the end of a da vinci s surgical procedure, when the instruments were being removed, the shaft coating of the 5mm maryland dissector instrument was coming off. Nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the tube insulation is damaged. The tube has a 1. 28" long section with missing insulation, located roughly 1. 28" above the snake wrist. The shape of damaged area is not typical for cannula related damage. It is possible that the tube insulation was initially damaged by instrument collisions and/or rough handling, which caused pieces of the insulation to stick out. Removing the instrument from the cannula may have further damaged the tube insulation, as the insulation that was sticking out may have caught on the cannula edge. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments of other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1440294 | 1302901 | 1241872 | 2955842200801482 | 1 | GCJ | Manufacturer | DistributorOther | Other | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-12-10 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors instrument broke and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is broken at the distal idlers and the pulley spins freely. Engineering believes the cable broke under tensile loading. The other cables at the wrist are not damaged. The cable wear on pulleys combined with high drive torques most likely contributed to breakage. Engineering also observed that the distal end of main tube has a section with material removal comprised of a combination of deep scratches and tube abrasions. Scratches are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by a combination of a cannula accessory and instrument collisions. No other damage found. The end wrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1441451 | 1302902 | 1241873 | 2955842200801480 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-10 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors instrument collided multiple times with another instrument. The wire broke and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument has frayed grip close and pitch up cables, however, there are no cables broken. Tip cover should protect cables when installed. Therefore, cables most likely have been damaged from rough handling during cleaning. Engineering also observed the distal end of main tube has a 4" long section with material removed all around the tube. The damaged area is parallel to tube axis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1441445 | 1302903 | 1241874 | 2955842200801481 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-10 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the pins on the back end of the maryland bipolar forceps instrument had been pushed in. This was not identified during a procedure. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the top plug riser pin to be pushed into the back end. The chassis feature that retains the bipolar insert and pins is broken. The instrument was likely mishandled, breaking the chassis feature. Engineering also observed the distal end of the main tube to have a 2. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1441433 | 1302904 | 1241875 | 2955842200801492 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-12 | Yes | No | Reuse | 2008
42017206 | 2008-11-11 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the tips of the maryland bipolar forceps instrument collided with the tip of another instrument. The surgeon was reported as being inexperienced. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip is bent, resulting in side to side misalignment of the grips. There is a. 025" offset at the tips. The bent grip has separation of the yaw pulley and pulley cover at the glue joint, which is indicative of overloading at the tip. Engineering also observed the distal end of main tube to have a 1. 1. " long section with light material removal on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. The wear pattern is consistent with cannula related tube abrasions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1441440 | 1302905 | 1241876 | 2955842200801479 | 1 | GEI | Manufacturer | OtherForeign | Other | Initial | 2008-12-31 | 2008-12-31 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-12-10 | Yes | No | Reuse | 2008
IS1200 | 2007-07-17 | Injury | Other | It was reported that at the end of a da vinci hysterectomy surgical procedure the surgical staff noticed a "dried or burn-like" area around the camera port incision, approx 2mm in width. The damaged skin had to be cuff off in order to close the incision. No visual damage nor signs of melting were observed on the cannula accessory used in conjunction with the camera. | The site indicated that the cannula accessory used in conjunction with the camera was not set aside, therefore, would not be returning for analysis. The root cause of the customer reported event cannot be determined. A follow-up mdr will be submitted if additional info is received. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1336685 | 1276423 | 1218034 | 2955842200801470 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2008-12-30 | 2008-12-30 | Yes | Yes | Health Professional | IS1200 | No | Yes | 2007-07-17 | Device Not Retu | No | Reuse | 2008
400180 | 2008-02-15 | Other | Other | It was reported that towards the end of a da vinci s myomectomy procedure, while the surgeon was exploring the abdominal cavity, a burn mark on the sigmoid colon was noticed. A general surgeon evaluated the burn and determined that no additional procedure was required to repair the burned tissue. The monopolar curved scissors instrument with the tip cover accessory installed was removed, and the tip cover was replaced. A burn hole was seen in the tip cover, however, the site discarded the accessory. The planned procedure was completed, with less than a 30 minutes delay. No patient harm or adverse outcome was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1431372 | 1278399 | 1219970 | 2955842200801474 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-12-30 | 2008-12-30 | Yes | Yes | Health Professional | 400180 | No | No | 2008-02-19 | Device Not Retu | Yes | Unkown | 2008
IS2000 A51P6 |  | Other | Other | It was reported that during a da vinci s gynecological procedure, while lifting the uterus, the pk dissecting forceps instrument broke. The instrument could not be removed through the cannula accessory due to the breakage, therefore, the patient side manipulator (psm) arm was moved to allow the removal of the cannula and instrument together. The customer experienced an unrecoverable system error at this time. The patient was kept under anesthesia while the site prepared their spare da vinci surgical system to finish the planned surgery. No patient harm was reported. | The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1430515 | 1278400 | 1219971 | 2955842200801472 | 1 | NAY | Manufacturer | User facilityHealth Professional | Physician | Initial | 2008-12-30 | 2008-12-30 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | Yes | 2008-08-24 | No | No | Reuse | 2008
420036 | 2007-08-16 | Malfunction |  | It was reported that during a da vinci s ileovesicostomy surgical procedure the tip of the debakey forceps instrument fell inside of the patient and was retrieved. A fragment of the instrument shaft also fell inside of the patient. No patient harm, adverse outcome or injury was reported and the planned surgical procedure was completed. | The instrument was not returned for evaluation, therefore, the root causes of the customer reported failure modes cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post evaluation) and/or if additional information is received. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1439471 | 1302788 | 1241759 | 2955842200801471 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-30 | 2008-12-30 | No | Yes | Health Professional | 420036 | No | No | 2007-08-16 | Device Not Retu | No | Unkown | 2008
400184 | 2008-03-24 | Malfunction |  | It was reported that during a da vinci surgical procedure, the permanent cautery spatula instrument broke and a piece fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1439488 | 1302803 | 1241774 | 2955842200801475 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-30 | 2008-12-30 | No | Yes | Health Professional | 400184 | No | No | 2008-03-24 | Device Not Retu | No | Unkown | 2008
|  | Malfunction |  | It was reported that shavings came off of the cannula accessory. No additional information was provided. No patient harm, outcome or injury was reported. | The cannula accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | CANNULA REGULAR IS2000 | ENDOSCOPIC ACCESSORY | 1439528 | 1302807 | 1241778 | 2955842200801473 | 1 | GCJ | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-30 | 2008-12-30 | No | Yes | Health Professional |  | No | No | 2007-09-12 | Device Not Retu | No | Unkown | 2008
IS2000 A51P6 | 2008-11-19 | Injury | Required InterventionHospitalization | It was reported that 8 days after a da vinci s hysterectomy procedure, the pt underwent an exploratory laparoscopy procedure where a bowel perforation was discovered. The perforation was repaired, however, the pt became septic due to a bilious fluid leakage. The pt was in critical care for 2 weeks and was then transferred to long term care. No add'l pt harm or adverse outcome was reported. | On december 23, 2008, add'l info regarding the event was provided by the assistant side surgeon present for the procedure. It was reported that the trocar made contact with the pt's small bowel at the beginning of the case. The console surgeon believed that extended cautery use may have contributed to the event. Based on the info provided, it was determined that the da vinci s surgical sys did not malfunction in a way that would cause nor contribute to the pt injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1332184 | 1273916 | 1215781 | 2955842200801448 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-24 | 2008-12-24 | Yes | No | Health Professional | IS2000 A51P6 | No | No | 2008-12-01 | Device Not Retu | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument expired prematurely and chipping at the base of the forceps was observed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering tested the instrument on an in-house system and found the instrument to be fully functional with 8 uses remaining. Engineering observed charring and melting of the yaw pulley cover at the base of the grips. Engineering also observed a deep scratch on the main tube where material has been removed. The damage is located approx. 150" from the interface with the proximal clevis. Based on the appearance, the damage may have been caused by instrument collisions. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429244 | 1277886 | 1219457 | 2955842200801452 | 1 | GEI | Manufacturer | User facilityOther | Physician | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42017206 | Yes | Yes | 2008-12-03 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s hysterectomy surgical procedure, a broken cable was observed on the monopolar curved scissors instrument. The planned surgical procedure was completed and no pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, no broken cables were found, however, engineering observed that the entire tube extension wall is circumferentially broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. No pieces are missing and it was concluded that the damage may have been due to excessive side loading. Engineering also observed that one side of the distal end of the main tube has a 1. 1" long section with material removed. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429242 | 1277887 | 1219458 | 2955842200801458 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument had burrs. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering inspected the metal wrist components for burrs and no burrs were found. Engineering observed that the distal end of the main lube has various scratch marks concentrated on one side of the tube. The scratches have removed a small amount of tube material. The scratches are small, narrow and not aligned with the tube axis, suggesting that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429231 | 1277888 | 1219459 | 2955842200801457 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument was not recognized by the system. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering tested the instrument on an in-house system and the instrument was successfully recognized. The pogo pins do not stick and are not contaminated. The programming chip also passed verification. The customer reported failure mode could not be confirmed. Engineering also observed various small scratch marks on the distal end of the main tube with light material removed. The scratches are under. 250" long and not aligned with the tube axis, suggesting that the damage may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429235 | 1277889 | 1219460 | 2955842200801464 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42018304 |  | Malfunction |  | It was reported that the tip of the shaft of the permanent cautery hook instrument is scratched and has gouges. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found a deep scratch parallel to tube axis that extends approximately. 654" from the intersection with the proximal clevis where material was removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. Engineering also observed that the luer plate is dislodged from chassis and is sticking out of the chassis. The chassis feature which retains luer plate is broken at the outer wall. The chassis most likely broke from an application of a non-axial force at the flush port. The breakage most likely occurred due to mishandling/misuse during cleaning the process. Electric continuity passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429225 | 1277890 | 1219461 | 2955842200801462 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tip of the shaft of the monopolar curved scissors instrument is damaged, and the scissors are dull and will not cut. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found that the scissors cut cleanly through. 006" of latex and the blades are not damaged. Electric continuity passed. Engineering also observed the distal end of main tube has two different sections approximately 2" and. 900", with material removed on the side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429222 | 1277891 | 1219462 | 2955842200801460 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the tip of the shaft of the monopolar curved scissors instrument has dings and scratches. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the distal end of the main tube has a couple of deep scratches with material removed. The scratches are located roughly 1. 5" above the tube extension, are under. 250" long, and are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429220 | 1277892 | 1219463 | 2955842200801468 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-12-05 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the shaft of the large needle driver instrument is damaged and has scratches and dings. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument has deep scratches on one side of the main tube with material removed. The damaged area is approximately. 345" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429214 | 1277894 | 1219465 | 2955842200801461 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the tips of the large needle driver instrument are out of alignment. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint, as the grips are not visibly misaligned, even under a microscope. The teeth of the carbide inserts mesh properly. Engineering observed the distal end of main tube has a 1. 15" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a canula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429207 | 1277896 | 1219467 | 2955842200801467 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-12-05 | Yes | No | Reuse | 2008
42017908 | 2008-11-18 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the monopolar curved scissors instrument broke and could not be removed from the trocar. The trocar was removed from the pt, the instrument was removed and replaced, and the planned surgical procedure was completed. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys and as a result, the tube extension can be rotated 360 degrees. Pieces from the main tube are missing. Engineering concluded that damage was likely caused by excessive side loading. Engineering also observed the distal end of main tube has a 3. 250 inch long section with material removed. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429200 | 1277899 | 1219470 | 2955842200801459 | 1 | GEI | Manufacturer | Health ProfessionalOther | Nurse | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42017908 | Yes | Yes | 2008-12-04 | Yes | No | Reuse | 2008
42018404 |  | Malfunction |  | It was reported that the permanent cautery spatula instrument was not recognized by the da vinci s system. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a system to check for recognition and the instrument was successfully recognized. Engineering observed that the distal end of the main tube has a couple of deep scratches with material removed. The scratches are all under. 250" long and are not aligned with the tube axis. The ceramic sleeve also has various hairline cracks. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429196 | 1277900 | 1219471 | 2955842200801454 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42018404 | Yes | No | 2008-12-03 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument would not close. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip did not respond intuitively when squeezing the masters to open and close. The grip open angle was larger than usual and one grip stayed in approximately the same position when masters were squeezed. Further inspection found that the yaw pulley is missing a. 200 inch by. 060 inch piece on the opposite side of the conductor wire, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that rough handling and or a collision may have contributed to the yaw pulley damage. Engineering also observed the distal end of main tube has a 2. 25 inch long section with light material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429191 | 1277901 | 1219472 | 2955842200801453 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-12-03 | Yes | No | Reuse | 2008
40018005 | 2008-12-02 | Malfunction |  | It was reported that approximately 15 minutes into a davinci s hysterectomy procedure, while the surgeon was cutting adhesions around the uterus to remove from the abdominal wall, the surgeon staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. The mcs instrument was immediately removed and the tip cover was replaced. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found two holes, approximately 180 degrees apart, burned through the silicone. Multiple holes indicate multiple arcing events occurred. The holes are oblong,. 025 inches by. 010 inches in size, and located. 330 inches and. 475 inches above the silicone to sleeve interface. Both holes appear to have internal scratches/tears adjacent to the hole that may have been a contributor to the arcing. Engineering also found various mechanical tears away from the holes. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by instrument collisions and/or rough handling. High generator settings may also have contributed to the arcing. The site also returned the monopolar curved scissors (mcs) instrument believed to have been used in the procedure. The instrument did not contain any char marks, burrs, frayed cables, or abnormally sharp edges that may have contributed to the silicone damage. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1429012 | 1277902 | 1219473 | 2955842200801456 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-12-03 | Yes | Yes | Initial | 2008
40018005 | 2008-12-01 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing occurred, hitting uterine tissue. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover has a. 035 inch diameter hole in the silicone. The hole is located. 190 inches above the silicone to sleeve interface and exhibits melting around the hole, indicative of arcing. The site also returned the 3 monopolar curved scissors instruments, one was believed to have been used during procedure. A char mark was found on one of the instruments, at the proximal clevis ear. With the tip cover installed on the instrument, the hole position is very close to the char mark. The hole position is near the section of maximum silicone stretching when the wrist is pitched. The instrument does not have any burrs or damage in the area of the char marks. Engineering concluded that insufficient silicone dielectric strength (due to damage or reduced silicone thickness during wrist articulation) may have contributed to arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1429003 | 1277903 | 1219474 | 2955842200801449 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-12-01 | Yes | Yes | Initial | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument had a loose cable. No patient harm, adverse outcome of injury was reported. | The instrument was evaluated. Per the customer reported complaint, engineering found a broken conductor wire at the yaw pulley hub, and the wire insulation and stands have been cut. Based on the appearance, the damage may have been caused by instrument collisions and/or rough handling. Engineering also observed various scratch marks towards the distal end of the main tube with veery light material removed. Based on the location and appearance, the damage may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instruction for use (ifu) specifically states: general precaution and warning - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1428925 | 1277904 | 1219475 | 2955842200801463 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | A tenaculum forceps instrument was returned without any information. No patient harm, adverse outcome or injury was reported. | The instrument was evaluated. The instrument was tested on an in-house system and recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions and the grips opened and closed properly. Engineering observed that the distal end of the main tube has a couple of deep scratches with material removed. The scratches are under. 250" in length and not aligned with the tube axis, suggesting that they may have been caused by instrument collisions and/or rough handling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429127 | 1279930 | 1221281 | 2955842200801447 | 1 | NAY | Manufacturer | User facilityOther | Physician | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42020704 | Yes | Yes | 2008-11-26 | Yes | No | Reuse | 2008
42020504 | 2008-11-20 | Malfunction |  | It was reported that prior to starting a da vinci surgical procedure, it was found that the fenestrated bipolar forceps instrument tips were not aligned. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip is bent, causing side to side misalignment of the grips. There is a. 080 inch offset at the tips. The bent grip is slightly twisted, suggesting overloading or misuse occurred. Engineering also observed the distal end of the main tube has a 1 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damages were most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429761 | 1279932 | 1221283 | 2955842200801469 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-12-24 | 2008-12-24 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-12-05 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-12-02 | Other |  | It was reported that prior to starting a da vinci surgical procedure, the customer experienced multiple non recoverable system error codes #208. The system was restarted, however, the error codes recurred upon start up. With the assistance of an isi technical support engineer, the customer shut down the system and reseated and checked the pins on the s3 and p3 cables. The system started up and homed normally. The customer called back and stated that the system error code #208 recurred. At this time the surgeon decided to abort the planned surgical procedure and reschedule for a later date. No patient harm was reported. | Conclusions - the investigation conducted by the field service engineer concluded that system error code #208 experienced by the customer was associated with a motor channel on the endoscopic camera manipulator (ecm) arm. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected multiple servo drive (msd) board on the ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error coe #208 appears when the da vinci safety system determines that the a fault occurred on one or more motor channels on a msd board. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1431313 | 1282026 | 1223257 | 2955842200801450 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-24 | 2008-12-24 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-12-02 | Yes | No | Reuse | 2008
IS2000 A5 1P6 | 2008-12-04 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer lost video in the right eye of the surgeon side console (ssc), and black lines appeared at top of the right 2d video which had a strobing effect. With the assistance of a isi technical support engineer, the customer power cycled the system, replaced the camera cable, re-seated the video cable, changed the camera controller unit (ccu) settings and reversed the console vision cable connections which restored the 3d video. The patient was under anesthesia for approximately an additional thirty minutes while the vision issue was being resolved. The planned surgical procedure was completed and no patient harm was reported. | The investigation conducted by an isi field service engineer found that the console vision cable was defective. The console vision cable connects the vision cart to the surgeon side console (ssc) and passes the video signals between the ssc and the vision cart. The system was repaired by replacing the console vision cable. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1434224 | 1282028 | 1223259 | 2955842200801465 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-24 | 2008-12-24 | Yes | Yes | Health Professional | IS2000 A5 1P6 | Yes | No | 2008-12-04 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-12-03 | Other | Other | It was reported that during a da vinci s gynecological procedure, the bipolar cautery would not work. The customer tried using an external foot switch, a different instrument, instrument cable and generator, however, the bipolar cautery remained nonfunctional. The surgeon decided to convert to traditional open surgical technique to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering was unable to replicate or confirm the bipolar cautery issue reported by the customer. Electrical tests were unable to detect any issues with the bipolar foot pedal switch or bipolar cautery cable. The electrosurgical unit (esu) is customer supplied equipment that is to be maintained by the hospital and is not the responsibility of intuitive surgical. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433052 | 1282029 | 1223260 | 2955842200801455 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-24 | 2008-12-24 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-12-03 | Yes | No | Reuse | 2008
400180 | 2008-11-10 | Malfunction |  | On december 2, 2008, isi received medwatch report from the customer regarding this event: "during coagulation of tissue intraoperatively, tip of instrument (monopolar curved robotic scissors) sent out an arc of sparks out of the side of instrument. Instrument replaced. No adverse tissue damage occurred. ". |  | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1377322 | 1311809 | 1249998 | 1311809 | 1 | GEI | User Facility |  | Nurse | Initial | 2008-11-21 | 2008-12-24 | No | Yes | Health Professional | 400180 | No | Yes |  |  |  |  | 2008
IS2000 A51 P6 | 2008-11-25 | Other | Other | It was reported that prior to starting a da vinci prostatectomy, the customer experienced multiple occurrences of system error code #23008 in conjunction with the led display on system arm #1 indicating an error. With the assistance of an isi technical support engineer, the site moved arm #3 in arm #1's place to minimize the fault. The site continued with their case, however, the site experienced a non recoverable system error #22001. The site then stowed arm #3 to resolve the error and tried to use arm #1 again. Arm #1 continued to display system error #23008 which could not be corrected. The patient had been under anesthesia and ports had been placed when the surgeon decided to abort the planned procedure. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering concluded that system error code #22001 and #23008 experienced by the customer were associated with a patient side manipulator (psm). The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system error codes #22001 and #23008 functioned as designed and there was no injury to the patient. The da vinci s safety system determined that there was a potentiometer limit error as well as a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a recoverable safe state. The psm was returned to isi for failure analysis investigation. Engineering evaluation determined that an axis potentiometer had malfunctioned, thus generating the system error code experienced by the customer. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1467273 | 1348951 | 1286550 | 2955842200801441 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-12-19 | 2008-12-23 | Yes | Yes | Health Professional | IS2000 A51 P6 | Device Returned To Manufacturer | No | 2008-11-25 | Yes | No | Reuse | 2008
IS1200 A43P1 | 2008-11-26 | Injury | Required InterventionHospitalization | After a da vinci hysterectomy procedure, during the cystoscopy, it was found that the patient's ureter had been damaged. The ureter was repaired via a ureteroneocystotomy procedure, and the patient was discharged two days later. No additional patient harm or adverse outcome was reported. | On december 04, 2008, additional information regarding the event was provided by the isi representative that was present for the procedure. It was reported that this was the surgeon's first da vinci procedure, and the injury was believed to have occurred during the benign hysterectomy, while transecting the patient's left ureter during dissection/transection of the ligaments and uterine arteries. Based on the information provided, it was determined that the da vinci surgical system did not malfunction in a way that would cause nor contribute to the patient injury. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1331318 | 1272893 | 1214798 | 2955842200801443 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-12-19 | 2008-12-22 | Yes | No | Health Professional | IS1200 A43P1 | No | No | 2008-11-26 | Device Not Retu | No | Reuse | 2008
40018005 |  | Malfunction |  | A monopolar curved scissors (mcs) instrument was returned with a tip cover accessory installed. No information was provided. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was evaluated and engineering found multiple holes burned through the silicone that exhibited melting, which is indicative of arcing. The largest hole is. 035 inches in diameter and is located. 230 inches above the silicone to sleeve interface. A mechanical tear adjacent to the hole connects to an oblong shaped arced section. There is also a. 015 inch diameter hole approximately 180 degrees from the large hole and. 280 inches above the silicone to sleeve interface. When the tip cover is installed on the mcs instrument, the hole locations are in areas where the silicone stretches over the proximal clevis ears when the instrument wrist is pitched. Multiple holes indicate multiple arcing events occurred. Engineering also observed various mechanical tears, including a large tear at the tip of the tip cover. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1467249 | 1349158 | 1286755 | 2955842200801445 | 1 | GEI | Manufacturer | Other | UNKNOWN | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-24 | Yes | Yes | Initial | 2008
40018005 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure arcing occurred through the monopolar curved scissors (mcs) tip cover accessory, which was installed on a mcs instrument. The surgical staff also discovered holes on the tip cover accessory. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed two holes which exhibit localized melting, indicating that arcing events had occurred. The holes are roughly 180 degrees apart and one hole is oblong, measuring. 015" x. 030", and the other is a. 020" diameter hole. The holes are located. 385" and. 270" above the silicone to sleeve interface. There is also a. 130" long tear in the axial direction at the distal tip of the silicone. The tear at the tip does not appear to be associated with either hole. One hole appears to have a tear line on one side. No other damage was found. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by damage or reduced silicone thickness during wrist articulation. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1467608 | 1349160 | 1286757 | 2955842200801440 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-24 | Yes | Yes | Initial | 2008
40018005 | 2008-11-24 | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure arcing occurred through the monopolar curved scissors (mcs) tip cover accessory, which was installed on a mcs instrument. The surgical staff also discovered holes on the tip cover accessory. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 165" long, jagged tear at the distal end. Close inspection of the tear shows some material removed along the tear on the outer surface of the silicone, which is not typical of a pure mechanical tear. The removed material appears to be from localized melting, suggesting that an arcing event occurred. Based on the appearance of the tear, the silicone in the affected area may have been damaged prior to the arcing event and repeated opening and closing of the scissor blades may have contributed to the elongation of the tear. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warning: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1471866 | 1349161 | 1286758 | 2955842200801437 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-24 | Yes | Yes | Initial | 2008
40018005 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, arcing occurred through the monopolar curved scissors (mcs) tip cover accessory, which was installed on a mcs instrument. The surgical staff also discovered holes on the tip cover accessory. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed various holes burnt through the silicone, suggesting that multiple arcing events occurred. The damage is contained within one-half of the silicone, between the two parting lines. The holes vary in round and oblong shapes and are located from. 030" to. 370" above the silicone to sleeve interface. Three distinct areas with signs of arcing are concentrated within. 100" of the silicone to sleeve interface. There are also a couple of mechanical tears in the general vicinity of the holes. The tears appear to have been caused by instrument collisions and/or rough handling. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by instrument collisions and/or rough handling. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1467243 | 1349162 | 1286759 | 2955842200801438 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-24 | Yes | Yes | Initial | 2008
40018005 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy surgical procedure, arcing occurred through the monopolar curved scissors (mcs) tip cover accessory, which was installed on a mcs instrument. The surgical staff also discovered holes on the tip cover accessory. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering observed a. 010" diameter hole in the silicone. The hole is located. 410" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, suggesting that an arcing event occurred. The tip cover also has various mechanical tears adjacent to and in the general vicinity of the hole. The tears appear to have been caused by instrument collisions and/or rough handling. Away from the hole, the plastic sleeve also exhibits scratches, which is indicative of collisions and/or rough handling. No other damage was found. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, with the dielectric strength weakened by instrument collisions and/or rough handling. High generator settings may also have contributed to the arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1467594 | 1349163 | 1286760 | 2955842200801439 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-24 | Yes | Yes | Initial | 2008
40018402 | 2008-11-26 | Malfunction |  | It was reported that during a da vinci thymectomy procedure, the surgeon collided two instruments together and the junction of the shaft and wrist of the permanent cautery spatula instrument broke. The connector of the instrument fell inside of the pt and was immediately retrieved. The incident did not significantly lengthen the procedure. No additional pt information was reported. | Results and conclusions - the instrument accessory was returned and evaluated. Per the customer reported complaint, engineering observed that one of the distal clevis ears is broken off at its base. A. 280" x. 215" piece of the clevis ear and the conductor cap are missing. The yaw pulley has the boss feature that mates with the groove on the ear broken off. The breakage of the clevis ear is consistent with overloading the tip with the wrist fully pitched. The corner base of the ear is a stress concentration area when tip is overloaded. The proximal clevis is cracked halfway between the proximal clevis ears on the side with the breakage, suggesting that overloading occurred with the wrist pitched. Per case notes an instrument collisions contributed to the breakage. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1470682 | 1349164 | 1286761 | 2955842200801442 | 1 | GEI | Manufacturer | User facility | Nurse | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40018402 | Yes | Yes | 2008-11-26 | Yes | No | Reuse | 2008
40016906 | 2008-12-01 | Malfunction |  | It was reported that during a da vinci total hysterectomy procedure, the tip of the harmonic curved shears insert broke off and fell inside of the patient. The piece was immediately retrieved and the planned surgical procedure was completed. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering observed that the insert was returned with a. 585" x. 085" piece of the curved blade detached. The blade is broken near where it transitions from straight to curved. The curved blade has a scratch mark on the inner surface that contacts the teflon teeth of the clamp arm, likely caused by an instrument collision or rough handling. The fracture surface of the curved blade has a section that exhibits striations, which are indicative of fatigue failure. The curved blade was most likely not able to withstand the repeated impact/bending loads imposed on it. Engineering also observed that the distal end of the shaft of the insert has scratches. Based on the location and appearance, the scratches were most likely caused by instrument collisions, or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ACCESSORY | 1468859 | 1349165 | 1286762 | 2955842200801444 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-12-19 | 2008-12-19 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-12-01 | Yes | Yes | Unkown | 2008
37053403 | 2008-11-14 | Malfunction |  | While pulling the camera attachment out of the arm the push-in tab broke off. I pushed both tabs in and removed the piece easily without resistance and one of the tabs fell onto the floor after it was apart from the arm. |  | INTUITIVE | CAMERA PUSHIN TAB | 1338010 | 1272533 | 1214462 | MW5009328 | 1 | GCJ | Voluntary |  | RISK MANAGER | Initial | 2008-12-09 | 2008-12-17 | No | Yes | Health Professional | 37053403 | Yes | No |  |  |  |  | 2008
IS2000 A51P6 | 2008-11-12 | Death | Death | It was reported that after a da vinci s mitral valve repair surgical procedure and while re-inflation of the pt's right lung, the surgical staff noticed a very low oxygen (o2) saturation. As a result of the low o2 saturation, the surgical staff decided to put the pt on bypass in order to scope the lungs. The surgical staff found that fluid had accumulated in the pt's right lung. Despite efforts, the pt was not able to regain function of the right lung and was transferred to the icu. The pt expired due to multi-organ failure, hypoxia, and acute pulmonary edema of the right lung. | No malfunction of the da vinci s system was alleged by the hosp and the info provided indicates that this event is related to the inherent risks associated with surgery and the pt's medical history. The hospital has continued to use the da vinci s system to perform surgery on pts. As of late 2008, no adverse events or system malfunctions have been experienced with the site's da vinci s system and as of the same day, no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1321561 | 1263424 | 1206134 | 2955842200801421 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-12-12 | 2008-12-12 | Yes | No | Health Professional | IS2000 A51P6 | No | No | 2008-11-12 | Device Not Retu | No | Reuse | 2008
IS2000 A51P6 | 2008-11-14 | Death | Death | It was reported that during a da vinci s surgical procedure to remove a pelvic mass, the console surgeon unintentionally clipped the pt's superior mesenteric artery. After the procedure was completed, the pt went into organ failure while on the floor and expired. No additional info has been provided. | No malfunction of the da vinci s system was alleged by the hosp and the info provided indicates that this event is related to the inherent risks associated with surgery. The hosp has continued to use the da vinci s system to perform surgery on pts. As of this report, no adverse events or system malfunctions have been experienced with the site's da vinci s system, and no similar instances of this event have been reported to isi. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1321551 | 1263425 | 1206135 | 2955842200801433 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-12-12 | 2008-12-12 | Yes | No | Health Professional | IS2000 A51P6 | No | No | 2008-11-20 | Device Not Retu | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the cable fillament on the pk dissecting forceps instrument was broken. It was also reported that nothing fell in the patient and that the case was successfully completed with no harm to the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a damaged conductor wire at the yaw pulley hub. Engineering also observed the distal end of the main tube to exhibit deep scratches with light material removal 90 degrees from the side with conductor wire damaged. Scratches are not aligned with the tube axis. Based on the location and appearance, the main tube and conductor wire damage were likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429718 | 1277052 | 1218622 | 2955842200801429 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-18 | 2008-12-12 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-11-18 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that after a da vinci s prostatectomy procedure, the monopolar curved scissors instrument had unknown black spots after it was cleaned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the surface of the scissor blades to have various black spots from pitting corrosion. Engineering concluded that the blade material was likely damaged due to elevated temperatures and/or cleaning agents. The endowrist instruments instructions for use (ifu) specifically states: caution: prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. Do not expose instruments to hydrogen peroxide, bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Engineering also observed the distal end of main tube has two sections, 2 inches apart, with light material removal on one side of the tube. The damaged areas are both approximately 1. 2 inches long and parallel to the tube axis with a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429725 | 1277053 | 1218623 | 2955842200801431 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-19 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the tips of the monopolar curved scissors instrument will not close. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the scissors do not fully close due to blade damage. As a result, the blades do not cut through their entire length, failing the cut test. One blade edge has an indentation near the midpoint that creates a mechanical stop for the other blade when closing. The damage to blade is most likely due to mishandling or misuse. Engineering also observed the distal end of main tube has a 1. 35" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429882 | 1277054 | 1218624 | 2955842200801430 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 40017908 | Device Returned To Manufacturer | No | 2008-11-19 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument wire was broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument does not have any broken or frayed cables at the wrist. The grips open and close properly and the wrist pitches up and down correctly when the imput discs are manually rotated. Engineering also observed the distal end of main tube has a couple of damaged areas with material removed. Some of the damaged area is parallel to tube axis but there are also some deep scratches. Based on the location and appearance, the damage was most likely caused by instrument collisions or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429595 | 1277055 | 1218625 | 2955842200801435 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-11-22 | Yes | No | Reuse | 2008
40016906 | 2008-11-19 | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure, a metallic fragment from the harmonic curved shears insert broke and fell into the pt. The surgical staff was unable to retrieve the fragment from the pt. The planned procedure was completed and no pt harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. Per the customer reported complaint, engineering confirmed that the clamp arm is missing from the insert. The hinge features that the clamp arm pins mate with are significantly bent. This damage suggest that instrument collisions and/or other misuse occurred during the surgical procedure, which likely caused the clamp arm to hyperextend and detach. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states; general precautions and warning; handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1429605 | 1277056 | 1218626 | 2955842200801432 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 40016906 | Yes | Yes | 2008-11-20 | Yes | Yes | Initial | 2008
42020704 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tenaculum forceps instrument failed. Nothing fell into the pt and the planned procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable broken at the distal idlers, allowing the pulley to spin freely. It appears that the cable broke under tensile loading. The other cables at the wrist are not damaged. Engineering also observed various deep scratches with material removed on the distal end of the main tube. The scratches are all under. 500" in length and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions and/or rough handling. No other damage found. The endowrist instruments instructions for the (ifu) specifically states: general precautions an warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429613 | 1277057 | 1218627 | 2955842200801428 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-11-17 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that a piece of the monopolar curved scissors instrument was broken. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the banana plug was removed. The lock washer and the nut used to fasten the connector were found inside of the housing. The condition was most likely the result of improper torque applied during assembly of the components. Electrical continuity from the ring lug to the instrument tip passed. Engineering also observed the instrument main tube has a 2. 25" long section with light material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429645 | 1277058 | 1218628 | 2955842200801434 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-11-21 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | Results: the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the scissors do not cut cleanly through. 006 inch latex. The latex gets pinched at the scissor tips. One blade edge has an indention near the tip that negatively affects cutting performance. Engineering concluded the blade damage was likely caused by mishandling and/or misuse. Conclusion: engineering also observed the distal end of the main tube has a 1 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damages were most likely cased by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429624 | 1277059 | 1218629 | 2955842200801436 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-12-12 | 2008-12-12 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-22 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not recognized by the da vinci surgical system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on a system to check for recognition and the system successfully recognized the instrument. Engineering also observed one of the scissor blades is melted and resolidified at the tip. The damaged blade is roughly. 060 inches shorter than the other blade and has dark char marks around it, indicating that an arcing event had occurred. Engineering concluded that the blade damage was likely caused by the blade making contact with another metal object during activation of cautery. High generator settings may have also contributed to the damage. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1435453 | 1296268 | 1235797 | 2955842200801427 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-12-05 | 2008-12-08 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-07-07 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the distal tip of the monopolar curved scissors instrument came off and nothing fell into the patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the tube extension is broken and missing a. 375" by. 130" piece at the proximal clevis interface. The clevis is dislodged from tube extension. One of the fracture planes of the tube extension is adjacent to an edge of one of the keys that mates with the proximal clevis. Engineering concluded that the tube extension most likely broke due to excessive side loading or rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering also observed the distal end of main tube has a. 900" long by. 100" wide section with light material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1435442 | 1296269 | 1235798 | 2955842200801425 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-05 | 2008-12-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-14 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, performance testing was conducted and the instrument was not able to cleanly cut through a piece of latex, with the latex getting pinched at the scissor tips. Engineering observed that the blade edges are slightly rounded at the tips which negatively affects cutting performance. Engineering also discovered that the distal end of the main tube has a. 130" x. 180" oblong hole burnt through it. The hole center is located roughly. 400" above the tube extension. Localized material removed around the hole indicates that an arcing event occurred. The instrument was disassembled and char marks were found on the tube extension and hypotubes in the same location as the hole. A crack in the distal end of the tube was also observed, starting at one of the slot features and running along the side of the hole. It was determined that the crack in the tube most likely created a path for arcing to occur, however, the source of the crack cannot be determined. High generator settings may have also contributed to the damage. Engineering also noticed two sections on the main tube, both approximately 1. 75" long and 180 degrees apart, with material removed. The damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1435482 | 1296270 | 1235799 | 2955842200801426 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-05 | 2008-12-08 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-15 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the jaws of the precise bipolar forceps instrument were misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one bent grip, causing side to side misalignment of the grips and a. 110" offset at the tips. The bent grip has a twist in it and there is separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Engineering also observed various scratches on the distal end of the main tube with light material removed. The scratches are under. 250" long and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions and/or rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1435435 | 1296271 | 1235800 | 2955842200801420 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-12-05 | 2008-12-08 | No | Yes | Health Professional | 42011005 | Yes | No | 2008-11-11 | Yes | No | Reuse | 2008
IS2000 A5 1P6 | 2008-11-13 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, unstable video image was experienced on the surgeon side console (ssc) during camera movement. While attempting to troubleshoot the issue with an isi technical support engineer the customer lost the right side video image. The pt was under anesthesia and the port incisions had already been made when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. | Conclusions- the investigation conducted by field service engineering found that the vision issues experienced by the customer were associated with the cable that passes video signals between the surgeon side console and the vision cart. The system was repaired by replacing the affected cable. As of (b) (4) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433832 | 1290925 | 1231036 | 2955842200801424 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-12-05 | 2008-12-05 | Yes | Yes | Health Professional | IS2000 A5 1P6 | Yes | No | 2008-11-13 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-11-13 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure, the customer experienced a system error code #23008. The site called an isi field service engineer for technical support and advised that a loose wire was observed on the right master tool manipulator (mtmr). It was reported that toward the end of the case the wire snapped. The surgeon decide to convert the procedure to traditional open surgical techniques to finish the planned procedure. No pt harm was reported. | Conclusions- the investigation conducted by the field service engineer concluded that system error code #23008 was associated with the right master tool manipulator which was found to have a broken cable. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the pt from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtmr. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put davinci in a "recoverable safe state". The mtmr was returned to isi for failure analysis investigation. Engineering evaluation of the mtmr and found that the upper axis-4 turret cables to be broken and an axis-3 outer drum cable to be loose, thus confirming the reported complaint experienced by the customer. As of late 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433815 | 1290926 | 1231037 | 2955842200801423 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2008-12-05 | 2008-12-05 | Yes | Yes | Health Professional | IS2000 A51P4 | Yes | No | 2008-11-13 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-11-07 | Other | Other | It was reported that approximately 1. 5 hours into a da vinci s prostatectomy procedure, double vision was experienced from the surgeon side console (ssc). With the assistance of an isi technical support engineer the customer removed the endoscope from the camera head, power cycled the high resolution stereo viewer (hrsv), the ssc, and the entire system, however, the double vision continued to occur. The pt was under anesthesia for approximately 2 hours when the surgeon made the decision to convert the procedure to traditional open surgical techniques. The planned procedure was completed and no pt harm was reported. | Results and conclusions - the investigation conducted by field service engineering found no problems with the vision system and the customer reported problem could not be duplicated. Engineering concluded that the problem experienced by the customer was likely due to a visual alignment/balance failure which was corrected when the system was completely turned off at the end of the case. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433822 | 1290927 | 1231038 | 2955842200801416 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | UNKNOWN | Initial | 2008-12-05 | 2008-12-05 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-11-07 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument bipolar connection broke off. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the bipolar connection is not broken off, however the top plug riser pin is pushed into the back end. The chassis feature that retains the bipolar insert and pins is broken. Engineering concluded that the instrument was most likely mishandled, breaking the chassis feature. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering also observed the distal end of the main tube has a 2. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1491671 | 1367631 | 1304463 | 2955842200801422 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-12-05 | 2008-12-05 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-11-13 | Yes | No | Reuse | 2008
40016906 | 2008-09-30 | Malfunction |  | It was reported that approximately ten minutes into a da vinci s thymectomy procedure, while the surgeon was dissecting the thymus, the harmonic curved shears (hcs) insert separated and fell into the patient. The hcs insert was retrieved and the planned surgical procedure was completed with a replacement instrument. The replacement harmonic curved shears (hcs) insert separated and fell into the patient. The second hcs insert was retrieved and the planned surgical procedure was completed with a replacement instrument. The procedure was lengthened slightly due to these events. No patient harm, adverse outcome or injury was reported. Medwatch mfr report # 2955842-2008-01351 was submitted to the fda for the first hcs insert malfunction which occurred during the surgical procedure. This medwatch report is for the malfunction that occurred with the replacement hcs insert. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering found the clamp arm and curved blade detached from the distal end. The broken curved blade piece is approximately. 870 inches long. The curved blade and shaft of the insert exhibit various scratch marks, likely caused by instrument collisions or rough handling. One of the side hinge features, that accepts the clamp arm pin, is bent. The fractured surface of the curved blade has a section that exhibits striations, indicative of fatigue failure. Based on the features of the damage, engineering concluded that the curved blade could not withstand the repeated impact and bending loads and failed, causing clamp arm and curved blade to detach. Per the case notes, the missing accessory component that fell inside of the patient during the surgical procedure, was successfully retrieved. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1453203 | 1315992 | 1169014 | 2955842200801419 | 1 | GEI | Manufacturer | OtherUser facility | Unknown | Initial | 2008-12-02 | 2008-12-02 | No | Yes | Health Professional | 40016906 | Yes | No | 2008-11-03 | Yes | Yes | Initial | 2008
IS2000 A51P4 | 2008-10-21 | Other | Other | It was reported that a hospital staff member was driving the patient side cart (psc) in reverse when her hand became caught between the cart handle and the wall. Her hand was red, however, was found to be without injury. No additional information was provided. | The patient side cart (psc) is the operative component of the da vinci s system and it's primary function is to support the instrument arms and camera arm. The psc has a drive function to allow for ease of movement and storage. When the power button on the psc handle is pressed the psc moves, then depressed the psc will discontinue to move. It was concluded that the user was not aware of the proper operation of the psc drive function, therefore, causing a hazardous situation. As of november 25, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1433462 | 1285924 | 1226795 | 2955842200801415 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-11-26 | 2008-11-26 | Yes | No | Health Professional | IS2000 A51P4 | No | Yes | 2008-11-07 | No | No | Reuse | 2008
IS1200 A43P9 | 2008-10-30 | Other | Other | It was reported that the surgeon was at the system console for approx 5 minutes when system error code #20008 occurred. The pt was under anesthesia and the port incisions had already been made when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. While the surgical staff was disconnecting the pt side cart (psc), a bent pin was found on one of the system cables connecting the psc to the surgeon's console. The bent pin was corrected and there were no recurrences of the issue during the time of the report. The site also indicated that servicing was not required. | The #20008 system error code appears when the da vinci safety system determines the differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". There are there system cables that connect the pt side cart to the surgeon's console. Each cable is assigned a pt side manipulator and passes video, audio, and data during system operation for that specific psm. The psm is an instrument arm located on the pt's side cart, that provides the sterile interface for the endowrist instruments. It was determined that the psm assigned to the damaged cable generated the system error when it was not able to process the data. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1434707 | 1285933 | 1226804 | 2955842200801405 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-26 | 2008-11-26 | Yes | Yes | Health Professional | IS1200 A43P9 | No | No | 2008-10-30 | No | No | Reuse | 2008
IS1200 A43P9 | 2008-11-06 | Other | Other | It was reported that prior to starting a da vinci hysterectomy surgical procedure, one of the system patient side manipulators would not move into docking positon. An isi technical support engineer attempted to troubleshoot the issue via telephone, however, the issue could not be resolved. The patient had been under anesthesia for 30 minutes when the surgeon decided to abort the planned procedure and reschedule it for a later date. | The investigation conducted by field service engineering found that an axis on the affected patient side manipulator (psm) was rotated 360 degrees out of range, causing the psm cable to wrap around the setup joint and limiting movement of the psm. The patient side manipulator (psm) is an instrument arm located on the patient's side cart that provides the sterile interface for the endowrist instruments. The system was repaired by placing the psm in its proper position and training was provided to the surgical staff regarding proper psm placement during case setup. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1434628 | 1285934 | 1226805 | 2955842200801414 | 1 | NAY | Manufacturer | User facility |  | Initial | 2008-11-26 | 2008-11-26 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | Yes | 2008-11-06 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that prior to a da vinci s surgical procedure, the system was not recognizing the endowrist prograsp forceps instrument and the planned surgical procedure was completed with another instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed. The instrument moved intuitively with a full range of motion and in all directions. The grips opened and closed properly. Engineering also observed deep scratches on the main tube where material was removed. The damage is parallel to the main tube axis and extends approximately. 5" from the interface with the proximal clevis. Based on the location and appearance, the damage was most likely cause by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1454797 | 1314464 | 1252595 | 2955842200801403 | 1 | NAY | Manufacturer | DistributorOther | Other | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-11-03 | Yes | No | Reuse | 2008
42014304 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the tip of the endowrist 5mm maryland dissector instrument broke and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument did not have a broken tip, but the snake wrist does appear to have an offset at the hinge feature that attaches to the shaft. The snake wrist and the shaft are not completely coaxial, most likely due to a bent/damaged wrist link. The instrument moves in all directions, the offset does not significantly affect range of motion. Engineering also observed that the tube insulation was damaged roughly 3" above the snake wrist. The insulation has a. 500" long section with missing material. A flap has been created due to the insulation tearing. There are other small gouge marks in the insulation in the vicinity of this damage. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1455931 | 1314465 | 1252596 | 2955842200801406 | 1 | NAY | Manufacturer | DistributorOther | Other | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-11-04 | Yes | No | Reuse | 2008
42017208 |  | Malfunction |  | It was reported that prior to a da vinci s surgical procedure, the bipolar maryland forceps instrument would not function. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one close grip cable is derailed at the distal idler pulley. The grips can still open and close, but movement may not be precise. The cable derailment is likely due to the cable losing contact with the pulley during wrist articulation. Engineering concluded that the fleet angel between proximal and distal pulley most likely contributed to the cable derailment. Engineering also observed that the distal end of the main tube has a few deep scratches with light material removed. Based on the location and appearance, the damage was most likely caused by instrument collisions, or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | BIPOLAR MARYLAND FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1478493 | 1314466 | 1252597 | 2955842200801412 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42017208 | Yes | No | 2008-11-05 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument would not open and close properly. No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the grips opened and closed properly and the instrument moved intuitively with full range of motion and in all directions. The pitch up cable was found slightly frayed at the distal clevis hub. The cable may have been nicked by another instrument or handling during the cleaning process may have contributed to the fraying. Engineering also observed that the distal end of the main tube has various deep scratches with material removed. The scratches are under. 300 inches in length and are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1449577 | 1314485 | 1252616 | 2955842200801408 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-11-05 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the maryland bipolar forceps instrument tips would not align. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip is bent causing side to side misalignment of the grips and a. 020 inch offset at the tips. The bent grip has separation of the yaw pulley and the pulley cover at the glue joint, indicating overloading at the tip. One grip is approximately. 020 inches longer than the other, causing improper teeth meshing. Engineering also observed the distal end of the main tube to have a. 25 inch long section with light material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1448727 | 1314487 | 1252618 | 2955842200801417 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-11-11 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the maryland bipolar forceps instrument tips were "locking up". No other information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly and system recognition and engagement passed. The instrument is fully functional, the tips "locking up" failure could not be replicated. Engineering also observed deep scratches on the main tube, located approximately 1 inch from the intersection with the proximal clevis, with material removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1448722 | 1314488 | 1252619 | 2955842200801409 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-11-04 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the scissors do not cut cleanly through. 006 inch latex. The latex gets snagged at the scissor tips. The blade edges are slightly rounded at the tips, negatively affecting cutting performance. Decreased cable tension and / or belleville spring force may also contribute to poor cutting. Engineering also found that the distal end of the main tube has an 8. 75 inch long section with material removed. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damge was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453872 | 1314489 | 1252620 | 2955842200801411 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-05 | Yes | No | Reuse | 2008
40018005 | 2008-11-05 | Malfunction |  | It was reported that one and a half hours into a da vinci total hysterectomy procedure and while retracting the patient's uterus, the surgical staff noticed arcing through the tip cover accessory installed on the monopolar curved scissor (mcs) instrument. Upon inspection, an isi rep noticed three holes in the tip cover accessory. The tip cover was replaced and the surgical procedure was completed without further issues and no patient harm was reported. It was also reported the esu settings were at 40 and that inter-operative instrument collisions had occurred. The patient was reported as recovering well with no complications. | The tip cover accesory was returned and evaluated. Per the customer reported complaint, engineering found four holes burned though the silicon with material removal and localized melting indicative of arcing. The holes are all under. 015" in size with varying shapes (round and oblong), and are located from. 190" to. 350" above the silicone to sleeve interface. Multiple holes indicate multiple arcing events occured. Engineering also found various mechanical tears in the general vicinity of the holes. It has been concluded that tearing may have been caused by instrument collisions. In addition, high generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1453771 | 1314490 | 1252621 | 2955842200801413 | 1 | GEI | Manufacturer | Company Representative | Other | Initial | 2008-12-01 | 2008-11-26 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-11-05 | Yes | Yes | Initial | 2008
400180 | 2008-10-13 | Malfunction |  | It was reported that less than an hour into a da vinci s prostatectomy procedure, a hole was observed in the tip cover accessory attached to the monopolar curved scissors (mcs) instrument. It was reported that the surgeon rarely applied cautery energy to the mcs instrument without also contacting tissue, and there was nothing unusual about the settings on the electrosurgical unit. No patient harm, adverse outcome or injury was reported. | The tip cover accessory an monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1453777 | 1314491 | 1252622 | 2955842200801418 | 1 | GEI | Manufacturer | Company Representative | UNKNOWN | Initial | 2008-11-07 | 2008-11-26 | No | Yes | Health Professional | 400180 | No | No | 2008-11-05 | Device Not Retu | Yes | Unkown | 2008
42017206 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the maryland bipolar forceps instrument had a broken cable and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the conductor wire is exposed at the grip hub where the wire insulation is removed. Electrical continuity passed. Engineering also observed multiple deep scratches where material was removed from different sections of the main tube within an area extending approx 4" from the intersection with the proximal clevis interface. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453863 | 1314492 | 1252623 | 2955842200801407 | 1 | NAY | Manufacturer | OtherDistributor | Other | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-11-04 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was out of alignment. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of the grips and a. 039" offset at the tips. It was determined that the grip damage was most likely caused by overloading at the tip. Engineering also observed multiple deep scratches on the main tube with material removed. The damage is located within an area of approx 2. 8" and extends from the intersection with the proximal clevis. Based on the location and appearance, the main tube damage was most likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1453853 | 1314493 | 1252624 | 2955842200801410 | 1 | GEI | Manufacturer | User facilityOther | NOT APPLICABLE | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-11-05 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure the large needle driver instrument had a broken cable and shaft material dropped, but there were no fallen parts. The planned procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering did not observe any broken cables. Instrument recognition and engagement passed and the grips opened and closed properly. The instrument moved intuitively with full range of motion in all directions. Engineering also observed that one side of the distal end of the main tube has a 1. 150" long section with material removed. The damage is parallel to the tube axis and has a rough surface finish. Based on the appearance, the damage was likely caused by user handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1453847 | 1314494 | 1252625 | 2955842200801404 | 1 | NAY | Manufacturer | DistributorOther | Other | Initial | 2008-11-26 | 2008-11-26 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-11-03 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the monopolar curved scissors instrument was found to have a broken wire and nothing fell into patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip cable is broken t the distal idlers allowing the pulley to spin freely. The cable broke under tensile loading. The cable wear on the pulleys combined with high drive torques most likely contributed to the breakage. Engineering also observed the distal end of the main tube has a 1. 2" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely by a cannula accessory. No other damge found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1429728 | 1277744 | 1219315 | 2955842200801400 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-01 | Yes | No | Reuse | 2008
42004905 |  | Malfunction |  | It was reported that the endowrist cadiere forceps instrument has a broken cable and nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one close grip cable is frayed near the distal idler. The condition may be the result of tensile loading combined with possible instrument collisions other cables at the wrist are not damaged. Engineering also observed multiple deep scratches on the main tube, located within an area extending approximately 1. 4" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1430442 | 1277745 | 1219316 | 2955842200801401 | 1 | GCJ | Manufacturer | DistributorOther | Other | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42004905 | Yes | No | 2008-11-01 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the fenestrated maryland bipolar instrument has a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip cable is frayed near the distal idler pulleys. The grips can still open and close. The cable is derailed from the pulley and wedged between both distal idlers. The derailment likely preceded cable fraying. Engineering concluded the cable fray is likely from the cable traveling over sharp pulley edges. Engineering also observed the distal end of main tube has various scratch marks, some of which have light material removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED MARYLAND BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1430449 | 1277746 | 1219317 | 2955842200801395 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-10-30 | Yes | No | Reuse | 2008
42018304 | 2008-10-27 | Malfunction |  | It was reported that upon completion of the fibroid dissection during a da vinci s assisted myomectomy surgical procedure, the permanent cautery hook instrument could not be removed from the system arm when the surgical staff attempted to exchange the instrument. Using the system camera, the surgeon checked the instrument tip and noticed that it was broken. The instrument and trocar were removed simultaneously and a replacement instrument and trocar of the same type were used to complete the planned surgical procedure. It is unknown if any broken instrument pieces fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument hook and ceramic sleeve are intact, however, the instrument is broken at the proximal clevis to main tube interface. The proximal clevis is dislodged from the main tube as a result of the breakage. The proximal clevis is intact, however, the main tube interface wall has collapsed and is missing pieces. It appears that the sections of the wall were cut off after the initial breakage. No other damage was found. Engineering evaluation determined that the failure mode experienced by the customer was a result of damage to the proximal clevis of the instrument. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1430475 | 1277747 | 1219318 | 2955842200801388 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42018304 | Yes | No | 2008-10-27 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument has cables coming out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is broken at the distal idlers and the pulley to spin freely. The cable broke under tensile loading. The cable wear on the pulleys combined with large fleet angle between proximal and distal pulleys and high grasping force likely contributed to breakage. Engineering also observed the distal end of main tube has various deep scratches with material removed. The scratches are all under. 250" long and not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1429208 | 1277763 | 1219334 | 2955842200801402 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42020704 | Yes | Yes | 2008-11-01 | Yes | No | Reuse | 2008
40018005 |  | Malfunction |  | A monopolar curved scissors instrument was returned with a tip cover accessory installed. No information was provided. No patient harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Engineering found various mechanical tears and holes burned through the silicone indicating that arcing events had occurred. The majority of the damage is contained within one half of the tip cover silicon, between the two parting lines. The holes are all under. 020 inches in size with varying shapes (oblong and round) and are located. 065 inches to. 325 inches above the silicone to sleeve interface. Engineering concluded that arcing is likely due to insufficient silicone dielectric strength, with dielectric strength weakened due to instrument collisions. High generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1430226 | 1277764 | 1219335 | 2955842200801399 | 1 | GEI | Manufacturer | Other | UNKNOWN | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 40018005 | Yes | No | 2008-10-31 | Yes | Yes | Initial | 2008
42000605 |  | Malfunction |  | It was reported that the tip of the large needle driver instrument broke off and nothing fell into pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported complaint could not be confirmed. Engineering could find no visible issues with the instrument having broken/missing components at the wrist. All of the wrist components are intact and undamaged. Engineering observed that the distal end of the main tube has deep scratches and a 2" long section with material removed. The tube abrasions are parallel to the longitudinal axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a combination of cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress than can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1430281 | 1277816 | 1219387 | 2955842200801374 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-10-28 | Yes | No | Reuse | 2008
40018005 | 2008-10-20 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing was observed. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover has a hole in the silicone. The hole is located. 156 inches above the silicone to sleeve interface and exhibits localized melting around the hole, indicative of arcing. A mechanical tear on the opposite side of the tip cover, located. 254 inches from the tip and visible under magnification, was observed, however, it does not appear to be related to the arcing event. The tear was likely caused by collisions with other instruments and or accessories. Engineering concluded that the mechanical tear along with stretching of the silicone over the edge of the mcs instrument proximal clevis ear while pitching the wrist, may have locally reduced the dielectric strength and increased the likelihood of arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress than can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1430288 | 1277834 | 1219405 | 2955842200801392 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-10-29 | Yes | Yes | Initial | 2008
40018005 | 2008-10-22 | Malfunction | Other | It was reported that during a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, arcing occurred, hitting uterine tissue. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover has an oval shaped hole in the silicone. The hole is located. 168 inches above the silicone to sleeve interface and exhibits melting around the hole, indicative of arcing. Also observed was a mechanical tear in the silicone, parallel to the hole, likely caused by collisions with other instruments and/or accessories. Engineering concluded that the mechanical tear along with stretching of silicone over the edge of the mcs instrument proximal clevis ear while pitching the wrist, may have locally reduced the dielectric strength and increased the likelihood of arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress than can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1430306 | 1277835 | 1219406 | 2955842200801393 | 1 | NAY | Manufacturer | User facility | RISK MANAGER | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-10-29 | Yes | Yes | Initial | 2008
40018005 | 2008-10-21 | Malfunction | Other | It was reported that during the beginning of a da vinci s hysterectomy procedure, while using the monopolar curved scissors (mcs) instrument with a tip cover accessory installed, a spark was observed coming through the tip cover and causing a burn on the pt's uterus. The tip cover accessory was immediately removed and replaced. The planned surgical procedure was completed with the replacement tip cover and without significant delay. No pt harm, adverse outcome or injury was reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover has an oval shaped hole in the silicone. The hole is located. 340 inches above the silicone to sleeve interface and exhibits melting around the hole, indicative of arcing. Mechanical and internal scratches to the right and left sides of the hole were observed, however, they do not appear to be related to the arcing event. Engineering concluded that the tip cover mechanical and internal tears along with stretching of the silicone over the edge of the mcs instrument proximal clevis ear while pitching the wrist, may have locally reduced the dielectric strength and increased the likelihood of arcing. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1430310 | 1277837 | 1219405 | 2955842200801394 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-10-29 | Yes | Yes | Initial | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was dull. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the scissors cut cleanly through. 006" of latex. The blades are not damaged. Electrical continuity passed. Engineering observed that the distal end of the main tube had a 1. 5" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1430270 | 1277839 | 1219410 | 2955842200801373 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-28 | Yes | No | Reuse | 2008
42009307 | 2008-10-10 | Malfunction |  | It was reported that during a da vinci s surgical procedure, part of the endowrist prograsp instrument shaft was splintering. Nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of main tube has a section containing various deep scratches with material removed. The scratches are narrow, under. 250 inches long, and are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1430382 | 1277840 | 1219411 | 2955842200801389 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-10-28 | Yes | No | Reuse | 2008
42020705 |  | Malfunction |  | It was reported that after a da vinci s surgical procedure, the "fiberglass" shaft of the tenaculum forceps instrument was found to be cracked. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the main tube has multiple deep scratches in different sections, with material removed, within an area extending approximately 4. 5" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1430395 | 1277842 | 1219413 | 2955842200801375 | 1 | GCJ | Manufacturer | OtherHealth Professional | Nurse | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42020705 | Yes | Yes | 2008-10-28 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-10-30 | Other | Other | It was reported that during a da vinci s prostatectomy surgical procedure the customer experienced frequent occurrences of system error code #23013. The procedure was delayed by approx 15 minutes while an isi technical support engineer attempted to troubleshoot the issue, however, the surgeon decided to convert to open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #23013 experienced by the customer was associated with the left a master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The #23013 system error code appears when the da vinci s safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The mtml was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, an axis potentiometer and motor encoder assembly were replaced. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1429708 | 1278063 | 1219634 | 2955842200801396 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-30 | Yes | No | Reuse | 2008
IS2000 A51 P46 | 2008-10-30 | Other | Other | During a da vinci s hysterectomy procedure, the surgeon experienced low "brightness" from the image on the surgeon side console (ssc). The site called isi technical support, however, the image began to improve as they were calling in. Approximately 30 minutes later, the site called back and stated the procedure had been converted to traditional open surgical techniques to complete the planned procedure, as the vision was too blurry to continue. No patient harm was reported. | The investigation conducted by isi field service engineering (fse) discovered that the issue experienced by the customer was associated with a defective endoscope. The endoscope is a fiber optic tubular accessory connected to the camera head. Images are acquired through separate optical channels of the endoscope and are directed to the camera head, which are then processed to be displayed on the surgeon side console (ssc) and assistant monitor. The system was repaired by replacing the affected endoscope. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1429762 | 1278064 | 1219634 | 2955842200801397 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | IS2000 A51 P46 | No | No | 2008-10-30 | Yes | No | Reuse | 2008
42004905 | 2008-09-10 | Malfunction |  | On november 6, 2008, isi rec'd medwatch uf/importer report # 4100120000-2008-8020 from the fda: "event desc: during surgical procedure, a splinter size piece of graphite material was noted in the pt's abdomen. It was apparent that the piece broke free from the robotic wristed instrument. This happened during the removal and insertion of the instrument from the robot arm. Dr was aware and indicated that it would not have any adverse effect on the pt. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no. Device usage problem: other". | The instrument had not been returned to isi for eval. A f/u mdr will be submitted if the instrument is returned (upon completion of eval) or if add'l info is rec'd. | ENDOWRIST CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1348737 | 1288023 | 1172114 | 2955842200801390 | 1 | GCJ | Manufacturer | Other | NOT APPLICABLE | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42004905 | No | No | 2008-11-06 | Device Not Retu | No | Unkown | 2008
42017908 | 2008-10-24 | No Answer Provided |  | It was reported that during a da vinci s prostatectomy procedure "shavings" from the tube of the monopolar curved scissors (mcs) instrument fell into the patient. The shavings were removed using irrigation and suction. The mcs instrument was observed rubbing against another instrument installed on a different system arm. Both system arms were readjusted and the mcs instrument was replaced. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned to isi for evaluation. A follow-up mdr will be submitted if the instrument is returned, upon completion of evaluation, or if additional information is received. | AAMONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1368819 | 1291065 | 1231176 | 2955842200801391 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-11-21 | 2008-11-21 | No | Yes | Health Professional | 42017908 | No | No | 2008-10-27 | Device Not Retu | No | Unkown | 2008
IS1200 A43P9 | 2008-10-30 | Other | Other | It was reported that the surgeon was at the system console for 5 mins when system error code #20013 occurred. The pt was under anesthesia and the port incisions had already been made when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. | The investigation conducted by field service engineering found system error code #21013 in the system error logs as well as system error code #20013 and determined that both error codes were associated with a pt side manipulator (psm). The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #20013 and #21013 appear when the davinci safety system determines a differential change in the angular position of one or more robotic joints on the specified manipulator, as measure by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), are out of specified tolerance for agreement. Upon determining these conditions, the safety system puts davinci in a "recoverable safe state. " the psm was returned to isi for failure analysis investigation. Engineering concluded that an axis potentiometer had malfunctioned, thus generating the system error experienced by the customer. As of (b) (6) 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1496824 | 1387539 | 1323945 | 2955842200701398 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-21 | 2008-11-21 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-10-30 | Yes | No | Reuse | 2008
|  | Malfunction |  | In 2008, a robotic assisted vaginal hysterectomy was performed using intuitive surgical instrumentation. During the course of the procedure, a harmonic curved shears insert- catalogue # 400169- broke off and was unable to be retrieved. Physical exam of the specimen was negative. Radiographs of the intra-abdominal cavity were negative. |  | INTUITIVE SURGICAL | HARMONIC CURVED SHEARS INSERT | 1302798 | 1245776 | 1191567 | MW5009066 | 1 | LFL | Voluntary |  | Nurse | Initial | 2008-11-20 | 2008-11-20 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2008
40018005 | 2008-09-22 | Malfunction | Other | In 2008, medwatch uf / importer report # 5201390000-2008-8001 was received: "event desc: during robotic surgical procedure for total laparoscopic hysterectomy, a few minutes into the procedure, an arc occurred and arc hit uterine tissue. No harm to patient as uterus was going to be removed. Tip cover replaced. Device being used was the monopolar shears davinci surgical system by intuitive. Tip cover on device noted to have brown spots and holes after arc of shears occurred. Machine at recommended settings and, tip cover applied to monopolar shears according to manufacturer's directions. Did this event involve an electrophysiology procedure or an attempted electrophysiology procedure? no. ". | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover silicone has a. 030 inch by. 050 inch oblong hole next to one of the parting lines, which exhibits localized melting indicative of arcing. The hole is approx. 200 inches above the silicone to sleeve interface. When placing the tip cover on a monopolar curved scissors instrument, the hole is in a location where the silicone gets thinner, as it stretched over the radius when the instrument wrist is pitched. Engineering concluded that the arcing was likely due to insufficient silicone dielectric strength, due to repetitive articulation of the instrument wrist. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1414805 | 1232772 | 1179803 | 2955842200801368 | 1 | GEI | Manufacturer | User facility | RISK MANAGER | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | 40018005 | Yes | No | 2008-10-22 | Yes | Yes | Initial | 2008
IS2000 A51P4 |  | Malfunction | Other | In 2008, intuitive surgical received notification dated early of the month of an event at hosp: a named pt "underwent approx one and on-half years ago, surgical intervention with the use of your davinci robotic system. This occurred at the hosp with a dr. Subsequent to the surgery, my client sustained severe pain. Through a series of misdiagnosis, dr had her undergo pain management. The original surgery was for a hysterectomy and some issues regarding adhesions. Thereafter, she went to another surgeon, who during the course of the surgery found certain pieces of instrumentation within her, which they removed. In processing a potential malpractice claim against the hosp and dr, they advised me, in their opinion, that there were not any pieces other than the rings from your equipment which was left in her. Apparently, the rings detached. They have indicated that was something they would not have been aware of during the course of the surgery. ". | The instrument and or accessory have not been returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1414065 | 1232773 | 1179804 | 2955842200801370 | 1 | NAY | Manufacturer | Other | ATTORNEY | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | IS2000 A51P4 | No | No | 2008-10-22 | Device Not Retu | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that after washing the monopolar curved scissors instrument, a bend at the assembly point was observed. The cause was unk but was not noticed during or post procedure. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the tube extension to be dislodged from the main tube. The entire tube extension wall is broken at the base of the orange-colored tube. As a result, the tube extension can be rotated 360 degrees. No pieces are missing. Engineering concluded that damage may be due to excessive side loading. Engineering also observed the distal end of the main tube to have a 6" long section with material removed around the outer diameter of the tube. The damaged area has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414787 | 1232774 | 1179805 | 2955842200801384 | 1 | NAY | Manufacturer | User facilityOther | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-25 | Yes | No | Reuse | 2008
42004905 | 2008-10-28 | Malfunction | Other | It was reported that during a da vinci s myomectomy procedure, while repair of the uterus after fibroid dissection, the cadiere forceps instrument rubbed against another cadiere forceps instrument causing fragments from the instrument shaft to fall into the surgical site. Fragments where removed and the damaged instrument was taken out of service. The instrument was in use approx 20 minutes when the event occurred. The procedure was completed using a replacement instrument of the same type. The incident lengthened the procedure an undetermined amount of time. No additional pt harm, adverse outcome or injury was reported. | The cadiere forceps instrument was not returned for eval. The root cause of the customer reported failure mode cannot be determined. If additional info becomes available, a follow-up mdr will be submitted. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC ACCESSORY | 1414052 | 1232775 | 1179806 | 2955842200801386 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | 42004905 | No | No | 2008-10-28 | Device Not Retu | Yes | Unkown | 2008
42018304 |  | Malfunction | Other | It was reported that during a da vinci s nephrectomy surgical procedure, the electronic power at the tip of the permanent cautery hook instrument did not work. The surgical staff observed the wrist of the instrument had broken. The surgeon cut the tip of the instrument and took the instrument out of the instrument arm. The surgeon observed two small black pieces in the patient body. The pieces were retrieved and the planned surgical procedure was completed with another instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the instrument's wrist to be detached below the proximal clevis and half of the main tube interface wall is missing. Engineering is unable to confirm some of the items mentioned in the complaint, as electric continuity is good from the banana plug to the cut end of the wire and from the cut end of the wire to the hook. There is no evidence of a broken cable prior to the wrist being cut off. The instrument may have become stuck and unable to be removed from cannula due to a break at the tube to proximal clevis interface. The proximal clevis is intact, but a section of the main tube interface wall most likely collapsed. The wrist was likely cut off to enable removal of the instrument from cannula. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414766 | 1232776 | 1179807 | 2955842200801387 | 1 | GEI | Manufacturer | ForeignDistributor | Other | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | 42018304 | Yes | No | 2008-10-22 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure the pk dissecting forceps instrument did not react to a movement. The planned surgical procedure was completed with another instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a. 150" by. 060" piece opposite the conductor break. The instrument has a broken section of yaw pulley that created a large range of motion in yaw for one grip. The added range of motion most likely created excess tension on the wire, causing it to break. The wire was most likely not securely attached to the grip, with poor strain relief, resulting in breaking during articulation of the wrist. Engineering also observed the distal end of the main tube to have a 1. 75" long section with material removed on one side of the tube. There are also a couple of deep scratches with light material removed. Based on the location and appearance, the damage was most likely caused by a combination of instrument collisions and a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELETROSURGICAL INSTRUMENT | 1414790 | 1232777 | 1179808 | 2955842200801383 | 1 | GEI | Manufacturer | OtherForeign | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-10-25 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument would not cauterize. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wires do not exhibit damage at the wrist. Engineering was unable to replicate the reported complaint of the instrument not cauterizing. Electrical continuity passed. Engineering eval also observed the distal end of the main tube to have various deep scratches with material removed. The scratches are all under. 250" in length and are not aligned with the tube axis. Based on the location and appearance, the damge was most likely caused by a cannula accessory. No other damage found. No other damge found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instrument or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414793 | 1232778 | 1179809 | 2955842200801376 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-10-28 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the cautery attachment will not stay in the maryland bipolar forceps instrument. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the top plug riser pin is pushed into the back end and the bottom pin is bent, preventing proper installation of a bipolar cord. The chassis feature that retains the bipolar insert and pins is broken. The instrument was most likely mishandled, breaking the chassis feature. Engineering also found that the distal end of the main tube has 3 inch long section with light material removed on the side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414849 | 1232780 | 1179811 | 2955842200801378 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-10-25 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument is splintering on the shaft. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of the main tube has a 5. 750 inch area, extended from the interface between the proximal clevis and the main tube, where material has been removed from different sections of the tube. The damaged area is parallel to the axis and has a rough surface finish. Based on the location and appearance, the damge was most likely caused by a cannula accessory. Deep scratches were also observed concentrated within 2 inches from the proximal clevis which may have been caused by instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions ans warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damge to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1414842 | 1232781 | 1179812 | 2955842200801380 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-11-06 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that he monopolar curved scissors instrument is splintering on the shaft. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of the main tube has a 5 inch long section, extended from the interface between the proximal clevis and the main tube, where material has been removed from different sections of the tube. The damaged area is parallel to the axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Deep scratches were also observed which may have been caused by instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. No other damage found. The endorwrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering also observed both scissor blades have indentations which negatively affect cutting performance. This damge is likely due to mishandling or misuse. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414854 | 1232782 | 1179813 | 2955842200801379 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-11-06 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp instrument jaws will not open. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip stuck in a neutral position. The grips would not dislodge from neutral position even by applying high torque manually at the instrument input discs. The instrument was disassembled and 2 hypotubes were found stuck to each other at the distal ends via dried bio debris. The hypotube sticking prevented the grip motion. Failure due to improper cleaning methods. Engineering also found that the distal end of the main tube has 1. 750 inch long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Deep scratches were found on the same section of tube, on the opposite side from the tube abrasions. Scratches were likely caused by instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1414864 | 1232783 | 1179814 | 2955842200801377 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-10-25 | Yes | No | Reuse | 2008
42017908 | 2008-10-14 | Malfunction |  | It was reported that the monopolar curved scissors instrument were observed to be dull. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance testing and found the scissors to cut cleanly through. 006" latex. The blades are not damaged. Engineering also observed the distal end of the tube extension to have a triangular shaped char mark with material removed due to burning. The char mark is. 170" long and. 185" wide (at triangle base). The charring suggests and arcing event occurred. When a tip cover is placed on the instrument, the char mark is completely covered by the ultem sleeve. Assuming a tip cover was installed during the procedure, it is unclear if energy would have escaped through the silicone part of the tip cover based on the location of the char mark. The tube extension may have had a crack which filled with conductive body fluids and created a path for arcing. High generator settings may have also contributed to the arcing. Electrical continuity passed. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL ISNTRUMENT | 1414775 | 1232784 | 1179815 | 2955842200801372 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-28 | 2008-11-14 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-28 | Yes | No | Initial | 2008
42020704 |  | Malfunction |  | It was reported that the cable of the tenaculum forceps instrument is broken. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable is broken near the proximal pulleys and the proximal clevis cable hole. The pulley spins freely. The proximal clevis cable hole exhibits slight wear on one edge. The other cables at the wrist are not damaged. The cable wear on pulleys/clevis combined with high grasping force most likely contributed to the breakage. Engineering also found that the distal end of the main tube has a 2" long section with material removed on one side of the tube. The damage area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damages were most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1414764 | 1232785 | 1179816 | 2955842200801382 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-10-25 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument would not work. No additional info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the top plug riser pin to be pushed into the back end, making it difficult for bipolar cautery to be functional. The chassis feature that retains the bipolar insert and pins is broken. The instrument was most likely mishandled, breaking the chassis feature. Electrical continuity passed. Engineering also found that the main tube has a large section with material removed all around the tube from just below midpoint to the distal end. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory, however, the size of the damaged area is longer than usual, suggesting possible additional tube damage occurred during the cleaning process. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1414750 | 1232786 | 1179817 | 2955842200801366 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-14 | 2008-11-14 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-10-21 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-10-24 | Other | Other | It was reported that during a da vinci s myomectomy procedure, while the camera was being moved during the dissection of the fibroid, the endoscopic camera manipulator (ecm) cannula mount broke. The customer attempted to use sterile adhesives to hold the ecm cannula mount back together, however, it was not effective. The pt was under anesthesia for an add'l 40 mins while a replacement ecm cannula mount was obtained from another site. The planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The investigation conducted by an isi rep confirmed that the endoscopic camera manipulator (ecm) cannula mount was broken. The ecm cannula mount holds the cannula securely in place while the camera is delivered through the cannula. The cannula provides the port through the body wall of the pt, through which the endoscope is inserted. The system was repaired by replacing the affected ecm cannula camera mount. As of 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1421749 | 1235692 | 1182483 | 2955842200801381 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-24 | Yes | No | Reuse | 2008
40018005 | 2008-10-21 | Other | Other | It was reported that during a da vinci total hysterectomy procedure and while the physician was performing a lymph node dissection, a hole was observed in the pt's aorta. The physician had been using the back of the monopolar curved scissor (mcs) instrument with a tip cover accessory installed to push against the aorta while dissecting the lymph nodes from the high left of the vessel at the time of the event and a small hole was noticed on the tip cover accessory. The surgical staff converted to an open procedure to repair the damaged vessel with 8 stitches. The pt underwent a blood transfusion of 2 units of blood. The pt was reported as recovering well with no complications. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found a hole located. 185" above the silicone to sleeve interface. The silicone exhibits localized melting around the hole, indicative of arcing. Engineering also observed a mechanical tear on one side of the hole and a larger tear in the silicone on the side opposite the hole. The tears are possibly caused by an instrument collision. The site returned an mcs instrument with a char mark at the proximal clevis ear. With the tip cover installed on the instrument, the hole position is very close to the char mark. Engineering concluded that the arcing was likely due to reduced dielectrical strength of the silicone caused by the mechanical tear and stretching of silicone over the edge of the proximal clevis while pitching the instrument's wrist. Instrument collision and high generator settings may have also contributed to arcing. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSSORY | ELECTROSURGICAL INSTRUMENT ACC | 1420801 | 1235693 | 1182484 | 2955842200801371 | 1 | NAY | Manufacturer | Company Representative | Physician | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | 40018005 | Yes | Yes | 2008-10-23 | Yes | Yes | Initial | 2008
IS2000 A51P6 | 2008-10-24 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, water was found to be dripping out of the camera head when moved. The picture was said to be blurry on both the touchscreen monitor and the physician side console. The pt was under anesthesia for 1 hr and 15 mins and port incisions had been created when the surgical staff made the decision to abort the procedure. No pt harm, adverse outcome or injury was reported. | An investigation conducted by the field svc engineer concluded the vision issues experienced by the customer were associated with the camera head. The camera head produces three dimensional images to the da vinci s vision system through right and left optical paths. The images are acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The system was repaired by replacing the affected camera head. As of (b)(4), 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1901848 | 1235694 | 1226795 | 2955842200801385 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-24 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-10-22 | Other | Other | It was reported that approx one hour into a da vinci prostatectomy procedure, the customer experienced multiple #20013 system error codes. With the assistance of an isi technical support engineer (tse) and with the error codes indicating which system pt side manipulator (psm) was faulting, the tse had the customer remove the sterile adapter from the psm, exercise psm axis-7 and press fault override, however, the #20008 and #20013 system error codes recurred. The surgeon decided to convert the procedure to traditional open surgical techniques to finish the planned procedure. No pt harm was reported. | The investigation conducted by a field svc engineer found a cable broken on the pt side manipulator (psm), causing an axis to spin freely. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system error codes appeared when the redundant sensors in the specified manipulator measured different values for the angular position (resulting in system error code #20008) or the rate of change in angular position (resulting in system error code #20013). Upon determining this condition, the safety system put the da vinci in a "recoverable safe state. " engineering confirmed the psm axis-7 cable was derailed, causing the customer reported problem. The system was repaired by replacing all axis-7 cables as a precautionary measure. As of (b)(4), 2008, there have been no reported recurrences of the issue at this hosp. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1901853 | 1235695 | 1182486 | 2955842200801367 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-14 | 2008-11-14 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-10-22 | Yes | No | Reuse | 2008
420184 | 2008-10-21 | Malfunction |  | During robotic hysterectomy, when changing out the unipolar paddle to the needle driver on the right, the paddle became stuck at the end of the trocar. Surgeon was eventually able to remove the trocar and instrument. Upon removal, a 4mmx5mm piece of sterile plastic appeared to be missing from the instrument. A thorough inspection was conducted but the piece was not found. The surgeon is unsure if the piece was missing at the beginning of the case. |  | INTUITIVE CAUTERY SPATULA | CAUTERY SPATULA | 1285680 | 1230183 | 1177273 | MW5008985 | 1 | GEI | Voluntary |  | Nurse | Initial | 2008-10-31 | 2008-11-13 | No | Yes | Health Professional | 420184 | Yes | Yes |  |  |  |  | 2008
IS2000 A51P6 | 2008-10-14 | Other | Other | It was reported during a da vinci s rectopexy procedure, the surgeon could not close the grips of an instrument on one of the patient side manipulator (psm) arms. Intermittent limitation to axis movements on one of the other patient side manipulator (psm) arms was also noted. The surgical procedure was delayed for approximately one hour due to this event and while the site converted to laparoscopic surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that the instrument grip issue and axis limitations reported by the customer were associated with two separate patient side manipulator (psm) arms. The psm is an instrument arm located on the patient side cart, that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the two affected psms. The two psms were returned to isi for failure analysis investigation. Engineering was able to replicate both the instrument grip issue on one psm and the axis limitation on the second psm. Evaluation of both psms found the wrist drive shafts spinning without stopping at the hard stop point and, the alloy screws used to hold the pulley inside the drive shaft brackets were found broken, causing the slippage. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1429771 | 1277513 | 1219084 | 2955842200801365 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2008-11-13 | 2008-11-13 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-14 | Yes | No | Reuse | 2008
4201726 | 2008-10-14 | Injury | Disability | This event has occurred twice. This device is marketed with a guarantee for ten uses before it is no longer able to be used for service. The unit was only used twice and sustained damage. There is a chip in the material at the base of the forceps jaws. No pt injury reported. The damage was detected by the spd staff during assessment for reprocessing. |  | DAVINCI | FORCEPS | 1283196 | 1227009 | 1174180 | MW5008943 | 1 | GCJ | Voluntary |  | RISK MANAGER | Initial | 2008-11-11 | 2008-11-11 | No | Yes | Health Professional | 4201726 | Yes | Yes |  |  |  |  | 2008
| 2008-11-03 | Malfunction |  | During robotic prostatectomy the robotic needle holder tip broke and the tip was not able to hold the needle. Some of the internal cables/wires were visible and still attached to the needle holder. There were 2 more uses left on the needle holder. |  | ENDOWRIST | NEEDLE HOLDER LARGE ROBOTIC | 1284902 | 1233446 | 1180415 | 1233446 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2008-11-11 | 2008-11-11 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2008
IS2000 A51P6 | 2008-10-17 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, the right eye in the high resolution stereo viewer (hrsv) was displaying horizontal lines. With the assistance of an intuitive surgical technical support representative, the site reset the surgeon side console, reseated the vision cable, and cycled power, however, the problem was not resolved. The pt was under anesthesia and port incisions had been created when the surgical staff made the decision to abort the procedure. No pt harm, adverse outcome or injury was reported. | Result and conclusions- an investigation conducted by the field service engineer concluded the vision issues experienced by the customer were associated with the high resolution stereo viewer (hrsv). The high resolution stereo viewer (hrsv) is located on the surgeon console and provides the video image for the surgeon console operator. The system was repaired by replacing the affected hrsv. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1412313 | 1230292 | 1177381 | 2955842200801359 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-17 | Yes | No | Reuse | 2008
IS1200 A4 3P9 | 2008-10-17 | Other | Other | It was reported that during a da vinci hysterectomy procedure one of the pt's side manipulator (psm) arms would not respond and the fault over-ride light was illuminated. With the assistance of an isi technical support engineer the system error logs were retrieved and system error #21003 was found. The customer tried to engage an instrument however, the engagement issue continued to occur. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure, no pt harm was reported. | Conclusions- the investigation conducted by field service engineering was unable to confirm the unresponsive psm issue reported by the customer. System error code #21003 is associated with a pt side manipulator (psm) arm. The psm is an instrument arm located on the pt side cart, that provides the sterile interface for the endowrist instruments. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. The #21003 system error code appears when the da vinci safety system detects that a system arm was "bumped" or restricted from moving properly during startup. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state". As of (b) (6), 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1412326 | 1230293 | 1177382 | 2955842200801360 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-07 | 2008-11-07 | Yes | Yes | Health Professional | IS1200 A4 3P9 | Yes | No | 2008-10-17 | Yes | No | Reuse | 2008
40016906 | 2008-09-22 | Malfunction | Other | It was reported that during a da vinci sacrocolpopexy with hysterectomy procedure, the cutting blade of the harmonic curved shears instrument fell in the pt. The medical staff attempted to locate the piece but was unsuccessful. The procedure was completed using a replacement instrument of the same type. After the procedure, the pt underwent a cat scan which resulted in locating the piece in the pt's abdomen. The incident did lengthened procedure an undetermined amount of time. No additional pt harm, adverse outcome or injury was reported. | The insert was returned and evaluated. Per the customer reported complaint, engineering confirmed that the clamp arm is missing from the insert. The hinge features that clamp arm pins mate with are not significantly bent. The distal end of the shaft exhibits scratch marks and one side of the curved blade has two gouge marks. The teflon teeth on the clamp arm make contact with the curved blade surface that contains the gouges. This damage suggests instrument collisions or other misuse occurred during a procedure, which may have caused the clamp arm to hyperextend and detach. No other damage found. There are no components in the harmonic curved shears instruments or inserts that meet the definition of "medical sharps". Per the customer reported complaint, it has been concluded that the "blade" referenced by the customer is actually the clamp arm of the harmonic curved shears insert. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1414895 | 1230424 | 1177515 | 2955842200801364 | 1 | LFL | Manufacturer | User facility | Unknown | Initial | 2008-11-07 | 2008-11-07 | Yes | Yes | Health Professional | 40016906 | Yes | No | 2008-09-23 | Yes | Yes | Initial | 2008
42020504 |  | Malfunction |  | It was reported that the peg of the cord attached to the fenestrated bipolar forceps instrument could not be pulled out. This was not noticed intraoperatively. No additional information provided. No pt harm, adverse outcome or injury was reported. | Results/conclusions: the instrument was retuned and evaluated. Per the reported complaint, engineering confirmed that the top plug riser pin is pushed into the back end and the bottom pin is bent. The chassis feature that retains bipolar insert and pins is broken. Instrument was likely mishandled, breaking the chassis feature. Engineering also observed the distal end of the main tube to have a 1. 2" long section with light material removal on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. There are also a few light scratches below this section, likely caused by rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1902068 | 1230425 | 1177516 | 2955842200801363 | 1 | NAY | Manufacturer | User facility | Unknown | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-10-20 | Yes | No | Reuse | 2008
400180 | 2008-10-13 | Malfunction | Other | It was reported that during a da vinci s surgical procedure, "a pin hole developed" in the tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The pt experienced a vessel injury. No other info was provided. | Conclusions- the tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined, if additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1414008 | 1230427 | 1177518 | 2955842200801353 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | Yes | Yes | Health Professional | 400180 | No | No | 2008-10-13 | Device Not Retu | Yes | Unkown | 2008
42017908 | 2008-10-14 | Malfunction |  | It was reported that near the end of a da vinci s hysterectomy procedure, after the uterus was removed, the surgeon notice arcing from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed with the original mcs instrument and tip cover. No pt harm adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1413999 | 1230428 | 1177519 | 2955842200801354 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42017908 | No | No | 2008-10-15 | Device Not Retu | No | Unkown | 2008
42009307 |  | Malfunction |  | It was reported that the shaft of the endowrist prograsp forceps instrument is chewed up. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering visually confirmed that the distal end of the main tube has a number of scratch marks all around the tube with varying degrees of material removal. Scratches are all under. 250" in length and are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instrument's instructions for use (ifu) specifically states: general precautions and warnings: handle instrument with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1414663 | 1230429 | 1177520 | 2955842200801355 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-10-16 | Yes | No | Reuse | 2008
42011005 |  | Malfunction |  | It was reported that the tips of the precise bipolar forceps instrument are bent. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip is bent, causing side to side misalignment of the grips. There is a. 020" offset at the tips. The bent grip has a slight separation of the yaw pulley and the conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed the distal end of the main tube has various scratch marks all around the tube. A few scratches are deep enough to have removed the tube material. The scratches are under. 250" long and are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1415022 | 1230431 | 1177522 | 2955842200801356 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42011005 | Device Returned To Manufacturer | No | 2008-10-16 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument were not working properly. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the tube extension is dislodged from the main tube. The entire tube extension wall is broken at the base of the axial keys. As a result, the tube extension can be rotated 360 degrees. There are no pieces missing. The damage may be due to excessive side loading. Electrical continuity passed. Engineering also found that the distal end of the main tube has 2" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish, located at approximately 3. 25" from intersection with the tube extension. Based on the location and appearance, the damages were most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MNOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1413951 | 1230437 | 1177528 | 2955842200801362 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-18 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not working properly, no additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument was returned with the luer plate dislodged from the chassis. The chassis feature that retains the luer plate is not broken, but the housing has the bottom part of both front snap tabs broken and a small crack on the side of the housing. Housing likely came off due to mishandling causing luer plate to dislodge. Engineering also observed the distal end of the main tube has two sections parallel to the tube axis with material removed. One section is roughly. 700" long, located directly above the tube extension. The other section is higher up the tube, roughly 1. 2" long with a lighter wear pattern. Based on the location and appearance, the damages were most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1413941 | 1230438 | 1177529 | 2955842200801361 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-18 | Yes | No | Reuse | 2008
400180 | 2008-10-16 | Malfunction |  | It was reported that during a da vinci sacrocolpopexy with hysterectomy procedure, the surgeon noticed arcing from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. No other info was provided. No pt harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional info is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 1413930 | 1230439 | 1177530 | 2955842200801357 | 1 | GEI | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-11-07 | 2008-11-07 | No | Yes | Health Professional | 400180 | No | No | 2008-10-16 | Device Not Retu | Yes | Unkown | 2008
| 2008-10-22 | Malfunction |  | Arc/spark occurred during robotic assisted surgery through scissor guard / accessory tip cover causing burn to uterus which was being removed. No harm then to patient. Settings for shears at 38 coag which is recommended level. Tip cover inspected prior to application using gage provided by company. Also application of tip cover performed according to recommendations. ======================manufacturer response for cover, esu shears, robotic device======================this is 9th event. Conference call ensued. Scheduling second conference call and future site visit by company enginneering to observe surgical procedure in future. |  | DAVINCI SURGICAL SYSTEM | COVER ESU SHEARS ROBOTIC DEVIC | 1272861 | 1224425 | 1171916 | 1224425 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2008-11-05 | 2008-11-05 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
| 2008-10-20 | Malfunction |  | During robotic assisted surgical procedure for laparoscopic hysterectomy with shears set at recommended level of 38 coag, arcing occurred causing burn to uterine tissue. No harm to patient since uterus being removed. Tip cover inspected before procedure and placed on shears according to company recommendations. ======================manufacturer response forcover, esu shears, robotic device======================conference call ensued. Will be following up with conference call with company and hospital. Robotic tech and company agreed to send engineering person out to observe future robotic assissed surgery. Date being scheduled. |  | DAVINCI SURGICAL SYSTEM | COVER ESU SHEARS ROBOTIC DEVIC | 1272857 | 1224423 | 1171914 | 1224423 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2008-11-04 | 2008-11-04 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
| 2008-10-21 | Malfunction |  | Arc/spark occurred during use. Actually occurred a few minutes into case. Burn to uterus resulted. No harm to patient since uterus being removed. Tip cover then noted to have holes. Removed and replaced. Machine settings at 38 coag per manufacturer recommendations. Also tip cover inspected prior to application using gage unit and applied as recommended. ======================manufacturer response for cover, esu shears, robotic device======================this is 8th event. Conference call ensued. Second one scheduled to talk with robotic tech. Also scheduling engineering company staff person to come an observe future surgery. |  | DAVINCI SURGICAL SYSTEM | COVER ESU SHEARS ROBOTIC DEVIC | 1272859 | 1224424 | 1171915 | 1224424 | 1 | GEI | User Facility |  | RISK MANAGER | Initial | 2008-11-04 | 2008-11-04 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
40016906 | 2008-09-30 | Malfunction | Other | It was reported that approximately ten minutes into a da vinci s thymectomy procedure, while the surgeon was dissecting the thymus, the harmonic curved shears (hcs) insert separated and fell into the patient. The hcs insert was retrieved and the planned surgical procedure was completed with a replacement instrument. The replacement harmonic curved shears (hcs) insert separated and fell into the patient. The second hcs insert was retrieved and the planned surgical procedure was completed with a replacement instrument. The procedure was lengthened slightly due to these events. No patient harm, adverse outcome or injury was reported. | The instrument and insert accessory were not returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted, if the instrument and/or insert accessory are returned (post engineering evaluation) or if additional information is received. | HARMONIC CURVED SHEARS INSERT | ENDOSCOPIC ACCESSORY | 1411601 | 1221483 | 1169014 | 2955842200801351 | 1 | GEI | Manufacturer | User facility | Unknown | Initial | 2008-10-31 | 2008-11-03 | Yes | Yes | Health Professional | 40016906 | No | No | 2008-10-10 | Device Not Retu | Yes | Unkown | 2008
42017908 |  | Malfunction |  | It was reported that the during a da vinci surgical procedure, the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument scissors did not cut cleanly through. 006" of latex. The latex gets snagged at scissor tips. The blade edges are slightly rounded at the tips, negatively affecting cut performance. The decreased cable tension and/or belleville spring force may also contribute to poor cutting. Electrical continuity passed. Engineering also observed the main tube to have a 2" long section just below the midpoint with light material removal on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1411592 | 1221484 | 1169015 | 2955842200801347 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2008-11-03 | 2008-11-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-08 | Yes | No | Reuse | 2008
42017908 | 2008-09-04 | Malfunction |  | On october 6, 2008, medwatch uf / importer report was received: while setting up for the case, the protective cover was placed properly on the monopolar scissors, and inspected prior to the case. Scissors were mounted on robot arm, robot signaled instrument was ok and ready to use. Scissors were in use when a charred area on the protective cover between the clear and colored area on the sleeve was noticed. The scissors were taken out of service immediately. During cleaning process, the protective sleeve was thrown away (it's a disposal part the set-up). This is a reusable instrument that has 10 lives. It had 2 lives remaining out of the 10. | The tip cover accessory has been discarded at the site, and will not be returned for evaluation, however, the monopolar curved scissors instrument used during the procedure was returned for analysis. Per the customer reported complaint, the instrument was visually inspected at the distal end for signs of arcing. A. 040 inch by. 060 inch char mark was found below one ear, on the lower portion of the proximal clevis. The char mark is likely an area where energy escaped through the tip cover. The char mark is located near or just slightly above the silicone to sleeve interface on the tip cover accessory. The instrument does not have any burrs or damage in the area of the char marks that would have contributed to arcing. Engineering concluded that arcing may have occurred due to high generator settings or damage to the tip cover silicone. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1411585 | 1221485 | 1169016 | 2955842200801346 | 1 | GEI | Manufacturer | Company Representative | RISK MANAGER | Initial | 2008-09-22 | 2008-11-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-06 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-10-06 | Other | Other | It was reported that prior to the starting of a da vinci s prostatectomy procedure, the site experienced the left and right video to be out of synchronization. With the help of an isi technical support engineer, the site attempted to troubleshoot the problem by powering down the system, reseating vision cables, and redoing the black and white balance without success. The patient was under anesthesia and port incisions had been created before the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | An investigation by the field service engineer confirmed the customer reported problem was associated with the camera control unit (ccu) and the digitizer. The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The digitizer converts the analog video images recorded by the camera into a digital format so that it can be transmitted over the fiber-optic cable and combined with other images. The system was repaired by replacing the affected ccu and digitizer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1409892 | 1221625 | 1169156 | 2955842200801345 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-11-03 | 2008-11-03 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-10-06 | Yes | No | Reuse | 2008
IS2000 A5 1P6 | 2008-10-09 | Other | Other | It was reported that prior to starting a da vinci s surgical procedure, the site experienced a blue and purple tint in the left eye of the high resolution stereo viewer (hrsv). The site attempted to troubleshoot the problem without success. The patient was under anesthesia and port incisions had been created before the site decided to complete the procedure using a back up da vinci surgical system. The surgical delay was approximately 45 minutes, while the surgical staff prepared the second da vinci system for surgery. No patient harm, adverse outcome injury was reported. | An investigation by the field service engineer confirmed with the customer reported problem was associated with the digitizer. The digitizer converts the analog video images recorded by the camera into a digital format, so that it can be transmitted over the fiber-optic cable and combined with other images. The system was repaired by replacing the affected digitizer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1409187 | 1221626 | 1169157 | 2955842200801350 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial |  | 2008-11-03 | Yes | Yes | Health Professional | IS2000 A5 1P6 | Yes | Yes | 2008-10-09 | Yes | No | Reuse | 2008
IS1200 A4 3P9 | 2008-10-14 | Other | Other | It was reported that during the beginning of a da vinci prostatectomy procedure and while performing the initial incision, insufflation was lost and the cannula came out of the patient's "ballooned" abdomen. The site advised that the initial port hole incisions were created slightly higher than normal. The site also reported the right and left patient side manipulator's (psm) seized during this time. It was reported that the patient was under anesthesia and, port incisions had been created before the surgical staff made the decision to abort the procedure. The procedure was successfully completed a few days later using the same port incisions. No patient harm, adverse outcome or injury was reported. | The investigation conducted the field service engineer consisted of testing the system, psm, and reviewing system error logs. Engineering was unable to confirm the customer reported problem. The patient side manipulator (psm) is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. As indicated in the event description, the procedure was successfully completed a few days post this event using a da vinci system and as of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1409193 | 1221630 | 1169161 | 2955842200801352 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-11-03 | 2008-11-03 | Yes | No | Health Professional | IS1200 A4 3P9 | Yes | No | 2008-10-06 | Yes | No | Reuse | 2008
42020705 |  | Malfunction |  | It was reported that the jaws of the tenaculum forceps instrument were having trouble opening and closing. No additional information was provided. No pt harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering was unable to replicate the reported problem. Engineering also observed the distal end of the main tube has a deep scratch with material removed on one side of the tube. The scratch is not aligned with the tube axis. There are also other small scratch marks on either side of the deep scratch. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1407806 | 1220095 | 1167646 | 2955842200801348 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-31 | 2008-10-31 | No | Yes | Health Professional | 42020705 | Yes | No | 2008-10-08 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument has a broken wire. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering visually confirmed that the instrument has a broken grip cable. One grip open cable is broken at the distal idlers. The cable broke under tensile loading and may have been kinked or damaged prior to breaking. The other cables at the wrist are not damaged. Engineering also observed the distal end of the main tube has a few deep scratches with light material removed. The scratch marks are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1408414 | 1220096 | 1167647 | 2955842200801349 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-31 | 2008-10-31 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-10-09 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that at the end of a da vinci s surgical procedure, the maryland bipolar forceps instrument was not working properly and the tip was bent. Nothing was seen falling into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grips are not bent and the grip teeth properly mesh. One of the conductor wires is damaged near the yaw pulley exit. The insulation has been cut, exposing bare wire. There is a char mark on the conductor cap near the damaged wire, likely caused by energy escaping through the bare wire and arcing to the cap. Engineering also observed various deep scratches with material removed on one side of the main tube. The scratches are not aligned with the tube axis. Based on the location and appearance, the damages are likely caused by instrument collisions/rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1408407 | 1220097 | 1167648 | 2955842200801344 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-31 | 2008-10-31 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-10-06 | Yes | No | Reuse | 2008
420049 VER05 | 2008-09-10 | Malfunction |  | During surgical procedure a splinter size piece of graphite material was noted in the patient's abdomen. It was apparent that the piece broke free from the robotic wristed instrument. This happened during the removal and insertion of the instrument from the robot arm. Doctor was aware and indicated that it would not have any adverse effect on the patient. Intuitive surgical has since provided the hospital with an inspection procedure for our 8mm cannulae, which involves a visual inspection and the use of a "gage pin", also provided by intuitive surgical. The procedure is recommended to be performed prior to sterilization and prior to surgery to detect damaged cannulae. |  | DA VINCI ENDOWRIST CADIERE FORCEPS | FORCEPS DISSECTING ROBOTIC SUR | 1273301 | 1224623 | 1172114 | 1224623 | 1 | GCJ | User Facility |  | RISK MANAGER | Initial | 2008-10-28 | 2008-10-28 | No | Yes | Physician | 420049 VER05 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not grasping. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no visible issues with the instrument grips/tip. Instrument performance could not be tested, as the instrument had 0 remaining uses. Engineering also observed a 2. 5" long section with material removed around the distal end of the main tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Deep scratches were also observed in the same section of the main tube where material was removed, which may have been caused by instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1406233 | 1215486 | 1163457 | 2955842200801343 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-10-07 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not grasping. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one bent grip, causing side to side misalignment of the grips. The is a. 050" offset at the tips. The bent grips has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Engineering also observed a 5. 5" long section with material removed around the distal end of the main tube, extending from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Deep scratches were also observed which may have been cause by instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405319 | 1215487 | 1163458 | 2955842200801342 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-10-07 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the wire broke off the pk dissecting forceps instrument. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a. 160" x. 060" piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of yaw motion. The cause of the yaw pulley breakage cannot be determined. The added range of grip motion likely created excess tension in the wire, causing it to break. There are no indications from the site that the missing instrument component fell inside of a pt during a surgical procedure. Engineering also observed material removed on the distal end of the main tube. Based on the location and appearance, the damage was most likely caused by a cannula accessory and instrument collisions and/or rough handling. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405296 | 1215488 | 1163459 | 2955842200801339 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-10-04 | Yes | No | Reuse | 2008
40017908 | 2008-08-28 | Malfunction | Other | It was reported that during a da vinci prostatectomy surgical procedure, the monopolar curved scissors instrument would stick, did not work and a broken piece fell inside of the pt. The broken piece was retrieved and the planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed damage on one blade edge where a section of material was removed below the midpoint. The size of the missing section is ~. 050" long. The missing section created a corner which acts as a mechanical stop for the other blade and the scissors cannot cut through the entire blade length. No other damage was found. Per the case notes, the instrument fragment was successfully retrieved from the pt. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405284 | 1215491 | 1163462 | 2955842200801337 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2008-10-24 | 2008-10-24 | Yes | Yes | Health Professional | 40017908 | Yes | Yes | 2008-10-02 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was defective. No additional info was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one grip close cable is broken at the scissor grip. The grip hub has a peak across the cable grove. The cable wear against the peak likely contributed to the breakage. Engineering also observed the distal end of the main tube has a 5" long section with material removed on one side of the tube. The damage area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405265 | 1215492 | 1163463 | 2955842200801336 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-10-01 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was defective. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that both yaw pulleys exhibited charring and localized melting at the grip base. The charring may be due to prolonged cautery activation without tissue between the grips. High generator settings may also contributed to the damage. Electrical continuity passed. Engineering also observed the distal end of the main tube has a 6" long section just below the midpoint with material removed on one side of the tube, extending from the intersection with the proximal clevis. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405239 | 1215495 | 1163466 | 2955842200801334 | 1 | GEI | Manufacturer | DistributorOther | Other | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-09-30 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument were dull. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument scissors did not cut through the entire blade length because the scissors don't fully close. One blade edge has a small burr near the tip that acts as a mechanical stop for the other blade. The blade damage may be caused by mishandling. Engineering also observed the main tube has a 2. 5" long section just below the midpoint with material removed on one side. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405229 | 1215497 | 1163468 | 2955842200801332 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-09-29 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was not recognized by the system. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering place the instrument on a test system and instrument recognition failed. The test error logs indicate the one or more of the tool sensors were not being detected. Engineering also observed a 3" long section with material removed around the distal end of the main tube, extending from the interface between the proximal clevis and main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Deep scratches were also observed within the same area, which may have been caused by instrument collisions. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. No other damage found. The endowirst instruments instructions for use (ifu) specifically states: general precautions and warnings- handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405218 | 1215499 | 1163470 | 2955842200801328 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-09-26 | Yes | No | Reuse | 2008
40017208 |  | Malfunction |  | It was reported that the tip of the maryland bipolar forceps instrument is broken and that no pieces fell into a pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to have a severely bent grip. The tips are offset by roughly. 200 inches. The bent grip has a separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. The severity of the bend suggests mishandling. Engineering also observed the distal end of the main tube to have various deep scratches with light material removed. The scratches are roughly. 250 inches and. 400 inches long and are not aligned with the tube axis and may be due to instrument collisions and / or rough handling. No other damage found. The endowirst instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1405209 | 1215500 | 1163471 | 2955842200801340 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-24 | 2008-10-24 | No | Yes | Health Professional | 40017208 | Yes | No | 2008-10-04 | Yes | No | Initial | 2008
ISI200 A43P9 | 2008-09-29 | Other | Other | It was reported that prior to a da vinci prostatectomy procedure, blurry vision was experienced from the surgeon side console (ssc). With the assistance of an isi technical support engineer, the customer tested two endoscopes and reseated the camera head adapter, however, the blurry vision continued to occur. The patient was under anesthesia and port incisions had been created before the surgical staff made the decision to abort the procedure. No patient harm, adverse outcome or injury was reported. | An investigation conducted by the field service engineer found that the vision system was out of tolerance due to a high lumen output on the camera controller unit (ccu). The ccu calibrates and adjusts the brightness levels of the video image from the two high magnification camera heads. The ccu then feeds the image through the video synchronizer to the surgeon's console and assistant monitor. The system was repaired by resetting the ccus and replacing the camera cable. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMET CONTROL S | 1404356 | 1216004 | 1163935 | 2955842200801333 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-10-24 | 2008-10-24 | Yes | Yes | Health Professional | ISI200 A43P9 | Yes | Yes | 2008-09-29 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-09-30 | Other | Other | It was reported that after a procedure, a hospital staff member was driving the patient side cart (psc) away from the surgical field when his hand became caught between the cart handle and the wall. The site was unable to get the cart drive into neutral to move the cart away and free the staff member's hand. No additional information was provided. | The investigation conducted by isi field service engineering (fse) consisted of testing the psc and reviewing the system logbook. The patient side cart is the operative component of the da vinci s system and it's primary function is to support the instrument arms and camera arm. Engineering found the psc drive function operating normally and was unable to duplicate the customer's reported problem. The fse verified proper drive operation and verified that the psc discontinued to move when the power button was depressed. Engineering also verified that the psc discontinued to drive when the users hand was removed from the handle. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1404055 | 1216005 | 1163936 | 2955842200801335 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-10-24 | 2008-10-24 | Yes | No | Health Professional | IS2000 A51P4 | Yes | Yes | 2008-09-30 | Yes | No | Reuse | 2008
| 2008-08-08 | Malfunction |  | Patient undergoing robotic assisted total laparoscopic hysterectomy. The harmonic curved shears were inserted into the trocar and prior to the machine being hooked up or used, the surgeon noticed that part of the tip was missing. This was a disposable item that had not been used yet in the case. The device was not broken when removed from packaging. The surgeon was able to find the rest of the tip and removed it from the patient. There was no harm to the patient. |  | DA VINCI HARMONIC CURVED SHEARS | INSTRUMENT ROBOTIC SURGICAL HA | 1274326 | 1225043 | 1172554 | 1225043 | 1 | LFL | User Facility |  | Nurse | Initial | 2008-10-21 | 2008-10-24 | No | Yes | Physician |  | Yes | Yes |  |  |  |  | 2008
40018005 | 2008-10-13 | Malfunction |  | During robotic surgical procedure for total laparoscopic hysterectomy, approximately one hour into surgery, device arc occurred and arc hit uterine tissue. Hysterectomy in progress and no harm to patient. Tip cover accessory being used with the monopolar curved shears da vinci surgical system by intuitive. Arc occurred during use resulting in tip cover accessory on device noted to have brown spots and holes after arc of shears occurred. Machine at recommended settings and tip cover applied to monopolar shears according to manufacturer's direction. |  | DA VINCI | ESU SHEARS COVER ROBOTIC | 1261917 | 1215263 | 1163254 | 1215263 | 1 | NEY | User Facility |  | RISK MANAGER | Initial | 2008-10-21 | 2008-10-21 | No | Yes | Physician | 40018005 | No | No Answer Provided |  |  |  |  | 2008
40018005 | 2008-09-22 | Malfunction |  | During a robotic surgical procedure for total laparoscopic hysterectomy, a few minutes into the procedure, an arc occurred and hit uterine tissue. No harm to patient as uterus was going to be removed. Tip cover replaced. Device being used was the monopolar shears davinci surgical system by intuitive. Tip cover on device noted to have brown spots and holes after arc of shears occurred. Machine at recommended settings and tip cover applied to monopolar shears according to manufacturer's directions. |  | DA VINCI | ESU SHEARS COVER ROBOTIC | 1261919 | 1215264 | 1163255 | 1215264 | 1 | NEY | User Facility |  | RISK MANAGER | Initial | 2008-10-21 | 2008-10-21 | No | Yes | Physician | 40018005 | Yes | No Answer Provided |  |  |  |  | 2008
IS2000 A51P6 | 2008-09-26 | Other | Other | It was reported that during a da vinci s total hysterectomy surgical procedure, the customer experienced system error code #25589 when starting the system and the led on the system camera arm was red. An isi technical support engineer (tse) attempted to troubleshoot the issue and had the customer power cycle and power off the system, however, the system would not boot up normally despite numerous attempts. The tse then had the customer attempt to deactivate the system camera arm, however, system error code #45049 appeared during every attempt. The pt was under anesthesia for approximately 40 mins when the surgeon decided to convert to laparoscopic techniques to complete the planned procedure. No pt harm was reported. | Conclusions - the investigation conducted by field service engineering concluded that system error codes #25515 and #45049 experienced by the customer were associated with the remote arm controller (rac). The rac consists of give printed circuit assembly (pca) boards which operate together to provide control of the system arms. The system was repaired by replacing the affected rac. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #25589 indicates that the rac brakes failed the brake voltage test, which is performed during power up. System error code #45049 indicates a communication error related to the error codes viewed by the customer. Upon determining these conditions, the safety systems put da vinci s in a "recoverable safe state". The rac was returned to isi for failure analysis investigation. Engineering found three pins with cold solder, thus affecting the electrical connection and generating system error code #25589 experienced by the customer. Engineering concluded that system error code #45049 was displayed as a result of the flood of error messages caused by system error code #25589. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1389144 | 1204827 | 1153778 | 2955842200801330 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-20 | 2008-10-20 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-09-26 | Yes | No | Reuse | 2008
ISI200 A43 P9 | 2008-09-26 | Other | Other | It was reported that during a da vinci hysterectomy procedure, the customer experienced multiple recurrences of system error code #20008. The site called for support and an isi technical support engineer (tse) viewed the system error logs and found system error code #21008. Based on the system error logs, it was determined that the right master tool manipulator (mtmr) was faulting. While the tse was assisting the surgical staff in exercising the mtmr, the surgical staff decided to convert the procedure to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error codes #20008 and #21008 were experienced by the customer and, were associated with the endoscopic camera manipulator (ecm) and mtmr. The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The master tool manipulator (mtm) refers to the master controllers, which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeons left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected ecm and mtmr. The system error codes #20008 and #21008 functioned as designed and there was no injury to the patient. The da vinci safety system determined a differential change in the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining these conditions, the safety systems put da vinci in a "recoverable safe state. " the ecm and mtmr were returned to isi for failure analysis investigation. Engineering evaluation of the ecm found an axis motor encoder assembly had malfunctioned. Engineering also evaluated the mtmr and found an axis potentiometer had malfunctioned. These malfunctions were found to have generated the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1392509 | 1202907 | 1152156 | 2955842200801331 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | Yes | Yes | Health Professional | ISI200 A43 P9 | Yes | No | 2008-09-26 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-05-22 | Other | Other | It was reported that three hours into a da vinci s prostatectomy procedure, the surgical staff experienced multiple occurrences of system error code #23017. The system error recurred after several system restarts. The patient was kept under anesthesia approximately an add'l 30 minutes while troubleshooting the reported problem and, preparing the site's spare da vinci surgical system to finish the planned surgery. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code #23017 experienced by the customer, was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the affected ecm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. The system error code appears when the da vinci s safety system determines that the motor did not respond as expected. This occurs when the measured voltage does not match the internal simulation of the motor circuit. This can be caused by an encoder failure, amplifier failure, or motor and or motor lead problems. Sporadic faults may be caused by aging brushes inside the motor. Upon determining these conditions, the safety system put da vinci s in a "recoverable safe state. " the ecm was returned to isi for failure analysis investigation, however, the fault experienced by the customer could not be duplicated. Based on the system error logs, the axis 4 potentiometer and motor assembly were replaced. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1389297 | 1202909 | 1152158 | 2955842200801325 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-09-23 | Yes | No | Reuse | 2008
40017908 |  | Malfunction |  | It was reported that during a da vinci surgical procedure the monopolar curved scissors instrument was dull. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the instrument scissors did not cut cleanly through. 006" of latex. The latex gets snagged at the scissors tips, but the blades are not damaged. The blade edges are slightly rounded at the tips, negatively affecting cut performance. The decreased cable tension and/or belleville spring force may also be contributing to the poor cutting performance. Electrical continuity passed. Engineering also observed the distal end of the main tube had a. 935" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1380053 | 1203499 | 1152705 | 2955842200801326 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 40017908 | Yes | No | 2008-09-24 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the spring came loose on the pk dissecting forceps instrument. It was noted that nothing fell into a pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley shows no signs of arcing, however, there is a. 150 inch by. 060 inch piece opposite the conductor break that is missing allowing the grip to move within the pulley and have a larger range of motion in yaw. The added range of motion of the grip likely created excess tension on the wire, causing it to break. Engineering also observed the distal end of the main tube to have a. 800 inch long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. A section 90 degrees from the tube abrasion, just above the proximal clevis, various deep scratches were found. The scratches are not aligned with the tube axis, suggesting they were caused by instrument collisions or rough handling. Electrical continuity failed. No other damage was found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1380061 | 1203500 | 1152706 | 2955842200801323 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-22 | Yes | No | Reuse | 2008
42014304 |  | Malfunction |  | It was reported that the 5mm maryland dissector instrument "coating" had peeled and nothing fell into pt. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that there was damage to the main tube insulation located 4. 125" from the intersection between the main tube and the instrument wrist. A. 200" by. 120" piece was removed, and is missing from the tube installation. Based on the location and the appearance, the damage to the tube insulation was most likely caused by instrument collisions. No other damage found. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a pt during a surgical procedure. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1380067 | 1203501 | 1152707 | 2955842200801322 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 42014304 | Yes | No | 2008-09-22 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument does not work. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering discovered a broken conductor wire at the yaw pulley exit. The yaw pulley is missing a. 150" x. 060" piece opposite of the conductor break, allowing the grip to move within the pulley and have a larger range of motion. It cannot be determined how the yaw pulley broke, however, the added range of grip motion likely created excess tension on the wire and caused it to break. Engineering evaluation also observed a 1. 75" long section with light material removed on the distal end of the main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. Adjacent to the main tube damaged section are deep scratches with material removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1380075 | 1203502 | 1152708 | 2955842200801324 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-09-23 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that during a da vinci a total hysterectomy surgical procedure, the monopolar curved scissors instrument was dull. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted tests and found the scissors did not cut cleanly through. 006" of latex with the latex getting snagged at the scissors tips. The blades are not damaged and the blade edges are slightly rounded at the tips, negatively affecting cut performance. Engineering also observed the distal end of the main tube to have a couple of sections, roughly 90 degrees a part, with material removed on one side of the tube. The damaged areas are 1" and 1. 5" long, parallel to tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1381999 | 1203503 | 1152709 | 2955842200801327 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-09-25 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that a piece had broken off of the endowrist prograsp forceps instrument. The broken piece did not fall into the pt. The broken pieces were discarded at the customer site. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to be missing one release lever. The housing is intact but one pin adjacent to a front snap tab is cracked. Engineering concluded that the housing likely popped off during mishandling, causing the lever to fall out. The housing was then put back on without the lever. Engineering's evaluation also observed the distal end of the main tube to have various deep scratches with material removed. Scratches are not aligned with the tube axis and may be due to instrument collisions and/or rough handling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1380086 | 1203504 | 1152710 | 2955842200801329 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-17 | 2008-10-17 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-09-26 | Yes | No | Initial | 2008
400180 | 2008-09-08 | Malfunction |  | It was reported that during a da vinci s gynecological procedure, the surgical staff observed arcing from the monopolar curved scissors (mcs) instrument with a mcs tip cover accessory installed. The site also reported a small hole in the tip cover accessory. The site discarded the msc instrument and the tip cover accessory. The planned surgical procedure was completed. No pt harm, adverse outcome or injury was reported. | Conclusions: the tip cover accessory and monopolar curved scissors (mcs) instrument have been discarded at this site and will not be returned for eval. The root cause of the customer reported failure mode cannot be determined. If add'l info is received, a f/u mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ELECTROSURGICAL INSTRUMENT ACC | 2157610 | 1198625 | 1148416 | 2955842200801317 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 400180 | No | No | 2008-09-17 | Device Not Retu | Yes | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, the endowrist prograsp forceps instrument had fiberglass shavings coming off the shaft that could be seen through the surgeon's console. Nothing was seen falling into the patient. No patient harm, adverse outcome, or injury was reported. | Result, conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end of the main tube had some deep scratch marks with material removed. The majority of scratches are not aligned with the tube axis. Based on the location and appearance, the damage was most likely caused by instrument collisions/rough handling. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1378033 | 1198626 | 1148417 | 2955842200801321 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-09-19 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, internal cable of the pk dissecting forceps instrument was beginning to fray and nothing was seen falling into the patient. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed no frayed cables at the wrist. Engineering conducted performance tests and found the instrument moved intuitively with a full range of motion in all directions. Results/methods: the distal end of the main tube has a couple of sections with material removed on one side of the tube. The damaged area is 2. 25" long and. 75" and 180 degrees apart, parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1378959 | 1198627 | 1148418 | 2955842200801320 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-19 | Yes | No | Reuse | 2008
40009305 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument is defective. No additional information was provided. No patient hard, adverse outcome or injury was reported. | Results & conclusions -the instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to be broken at the proximal clevis to main tube interface, causing the proximal clevis to be dislodged from the main tube. The main tube has missing pieces. The breakage is most likely due to overloading at the tip. Per e-mail, there are no indications from the site that the broken instrument components fell inside of a patient during a surgical procedure. Results & conclusion - engineering also observed that deep scratches in an area of about. 370"x. 240" of the main tube, where material was removed. Scratches are located at approximately. 400" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions/rough handling. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1378076 | 1198628 | 1148419 | 2955842200801318 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 40009305 | Yes | No | 2008-09-18 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was defective. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found the instrument to pass both recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly and is fully functional. Results & conclusions: engineering also observed the distal end of the main tube to have a. 225" long deep scratch with material removed. The scratch is not aligned with the tube axis and may have been caused by an instrument collision. No other damage found. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1378012 | 1198643 | 1148434 | 2955842200801311 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-15 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-09-16 | Other | Other | It was reported that during a da vinci s hysterectomy surgical procedure, when the surgeon was ready to use the camera, very little light was coming out of the light guide cable. It was also noticed that the tip of the light guide cable was "burnt up" after calibration had been completed. The patient ports had already been placed when the surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the low light output issue experienced by the customer was associated with the illuminator lamp module located within the illuminator. The illuminator lamp module provides light, which is then delivered to the endoscope via a fiber optic, bifurcated light guide cable and projected onto the surgical site. The system was repaired by replacing the illuminator lamp module and damaged light guide cable. The illuminator lamp module was returned to isi for failure analysis investigation. Engineering tested the lamp module with an undamaged light guide cable and observed that the proximal tip of the light guide cable melted within 1 minute of installation, with the lamp module set to maximum output. Engineering determined that the lamp module contained a faculty component which caused it to produce excess heat, thus burning the tip of the light guide cable and leading to the low light output experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1390587 | 1198983 | 1148774 | 2955842200801316 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2008-10-10 | 2008-10-10 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | Yes | 2008-09-16 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | It was reported that the bipolar plug in on the maryland bipolar forceps instrument broke. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the top plug riser pin is pushed into the back end and the bottom pin is bent. The chassis feature that keeps the pins in place is broken, most likely due to mishandling. Electrical continuity passed. Engineering also observed the distal end of the main tube to have a 2" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2157385 | 2123203 | 2020179 | 2955842200801319 | 1 | NAY | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-10 | 2008-10-10 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-09-19 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument was not cutting. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the scissors do not cut cleanly through. 006 inches of latex, where the latex gets snagged at the scissor tips. One blade has an indentation at approximately. 018 inches from the tip, however, both blade edges are slightly rounded at the tips negatively affecting cutting performance. Electrical continuity passed. Engineering also observed the distal end of main tube has a. 006 inch long section with material removed on one side of the tube. The damaged area has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367603 | 1186463 | 1136954 | 2955842200801314 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-09-15 | Yes | No | Reuse | 2008
42017805 |  | Malfunction |  | It was reported that the endowrist 8mm curved scissors instrument is dull. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance cut tests and found the scissors cut cleanly through. 006" of latex and the blades are not damaged. Engineering also observed multiple deep scratches that are parallel to the main tube, extending from the intersection with the proximal clevis, where material was removed. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. | ENDOWRIST 8MM CURVED SCISSORS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367458 | 1186464 | 1136955 | 2955842200801305 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42017805 | Yes | No | 2008-09-13 | Yes | No | Reuse | 2008
42020705 |  | Malfunction |  | It was reported that prior to starting a da vinci s procedure, the surgical staff found the tenaculum forceps instrument's shaft was splintering. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the distal end of the main tube to have a 1. 5 inch long section with material removed. The damaged area is parallel to the tube axis and has a rough surface finish with scratch marks within the same area. Based on the location and appearance, the damage was most likely caused by a combination of instrument collisions and a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367618 | 1186465 | 1136956 | 2955842200801313 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42020705 | Yes | No | 2008-09-12 | Yes | No | Reuse | 2008
42000604 |  | Malfunction |  | It was reported that the large needle driver instrument tip is damaged. No additional info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found no trouble with the instrument tip. Engineering observed the distal end of the main tube was a 3" long section with material removed on one side of the tube. The damaged area is parallel to the tube axis and has a rough surface finish. There are also some deep scratches in an approximately 1. 125" section of the main tube where material is removed. Based on the location and appearance, the damage was most likely caused by a combination of instrument collisions and a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional info is received. The endowrist instruments instructions for the (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1370197 | 1186466 | 1136957 | 2955842200801306 | 1 | GCJ | Manufacturer | OtherUser facility | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42000604 | Yes | No | 2008-09-13 | Yes | No | Reuse | 2008
42020501 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument was not working properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned, and evaluated. Per the customer reported complaint, testing was performed by engineering, and system recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed the distal end of the main tube to have various deep scratch marks on one side of the tube. A small amount of the tube material is removed due to the scratch marks. The scratches are short, and not aligned with the tube axis, suggesting they may have been caused by instrument collisions. The proximal clevis is missing a. 250" x. 130" section below the ear, but the clevis is not. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a patient during a surgical procedure. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367504 | 1186467 | 1136958 | 2955842200801301 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42020501 | Yes | No | 2008-09-12 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that prior to starting a da vinci s procedure, the surgical staff found the large needle driver instrument shaft was splintering. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the main tube to have a 1. 75 inch long section, approximately 2. 5 inches below the midpoint, with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367478 | 1186468 | 1136959 | 2955842200801312 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-09-12 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the endowrist prograsp forceps instrument has a broken cable. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable derailed at the distal idler pulley. The grips can still open and close, however, the movement is not precise. Engineering also observed multiple deep scratches on areas extending approximately 3. 25 inches from the main tube to proximal clevis intersection. The scratches range in size from. 037 inches by. 212 inches,. 130 inches by. 175 inches. 177 inches by. 240 inches and. 148 inches to. 219 inches. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367306 | 1186469 | 1136960 | 2955842200801284 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-09-11 | Yes | No | Reuse | 2008
42018304 | 2008-09-11 | Malfunction | Other | It was reported that approximately one hour into a da vinci prostatectomy procedure, arcing occurred with the permanent cautery hook instrument, which nicked a vessel and created minor controllable bleeding. The planned surgical procedure was completed with a replacement instrument of the same type and without significant delay. No additional patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one conductor wire had a section with insulation missing at one of the proximal pulleys. The pulley and top of one of the proximal clevis ears exhibit charring, indicating an arcing event occurred. Engineering has concluded that the wire insulation may have been damaged due to mishandling or misuse of the instrument, causing a path for energy to escape. High generator settings may have also contributed to arcing. Electrical continuity passed. No other damage found. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367300 | 1186470 | 1136961 | 2955842200801303 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2008-10-03 | 2008-10-03 | Yes | Yes | Health Professional | 42018304 | Yes | No | 2008-09-12 | Yes | No | Unkown | 2008
42020705 |  | Malfunction |  | It was reported that the tenaculum forceps instrument wires are broken. No additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one pitch cable is broken near the proximal clevis cable hole, but the hole does not exhibit wear. The other cables at the wrist are not damaged. Engineering concluded that the high grasping force combined with instrument collisions, most likely contributed to the breakage. Engineering also observed deep scratches in areas of about. 350" by. 350" and. 350" respectively, where material of the main tube was removed. Scratches are located on the clevis end of the main tube, one extending from the intersection with the proximal clevis, and the other at approximately 1. 125" from the same intersection. The damaged areas were most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367288 | 1186471 | 1136962 | 2955842200801309 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42020705 | Yes | No | 2008-09-13 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument has a "bent prong. " no additional information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the bottom plug riser pin is bent and the top pin is pushed into the backend of the instrument. The chassis feature that keeps the pins in place is broken, most likely due to mishandling. Engineering also observed the distal end of the main tube has a 3" long section with material removed all around the tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted, if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367281 | 1186472 | 1136963 | 2955842200801290 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-09-10 | Yes | No | Reuse | 2008
42017908 | 2008-09-02 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon saw arcing from the monopolar curved scissors (msc) instrument with a mcs tip cover accessory installed. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned for evaluation. Per the customer reported complaint, engineering found the proximal clevis exhibiting a series of char marks from. 110 inches to. 130 inches above the interface with the tube extension. The char marks are located near where the silicone to sleeve interface of the tip cover accessory would mate with the instrument. The instrument does not have any burrs or damage in the area of the char marks that would have contributed to arcing. The tip cover accessory was discarded at the site and will not be returned for evaluation. Engineering concluded that the char marks on the instrument are likely areas where energy escaped through the tip cover. Arcing was likely due to the tip cover having insufficient silicone dielectric strength caused by damage and or high generator settings. Engineering also observed the distal end of the main tube to have multiple 2 inch long sections, 90 degrees apart, with material removed on one side of the tube. The damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367274 | 1186473 | 1136964 | 2955842200801297 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-09-03 | Yes | No | Reuse | 2008
42017206 |  | Malfunction |  | The customer returned a maryland bipolar forceps instrument with no information provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one conductor wire broken at the yaw pulley exit. The yaw pulley is missing a. 210 inch by. 060 inch piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Added range of motion on the grip likely created excess tension on the wire, causing it to break. Engineering also observed the distal end of the main tube to have a 1. 5 inch long section with material removed on the side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367266 | 1186474 | 1136965 | 2955842200801307 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42017206 | Yes | No | 2008-09-13 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that during a da vinci s total hysterectomy surgical procedure, the surgical staff observed that the shaft of the fenestrated bipolar forceps instrument was a little bit dented. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed deep scratches with material removed in a. 500" x. 350" area on the main tube. The scratches are located on the clevis end of the main tube, approximately. 300" from the intersection with the proximal clevis. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instrument instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367261 | 1186475 | 1136966 | 2955842200801308 | 1 | GEI | Manufacturer | Company RepresentativeOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-09-13 | Yes | No | Reuse | 2008
42009307 |  | Malfunction |  | It was reported that the system was not reading the remaining uses left on the endowrist prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument had zero uses remaining, indicating that the instrument had been used for the maximum number of ten surgical procedures which it was originally programmed for. The system functioned as designed and did not "read" the remaining uses left on the instrument, thus preventing the instrument from being used for an additional surgical procedure. Engineering evaluation also observed multiple deep scratches with material removed, which extend approximately 2" from the intersection with the proximal clevis. The scratches range in size from. 090" x. 087",. 115" x. 317",. 066" x. 510" and. 125" x. 590". Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ENDOWRIST PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1367228 | 1186476 | 1136967 | 2955842200801285 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42009307 | Yes | No | 2008-09-11 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument do not meet. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument tips were properly aligned and no signs of damage/deformation of the tips were found. The customer reported failure mode cannot be confirmed. Engineering also observed a 4" long section on one side of the distal end of the main tube with material removed, which extends from the intersection with the proximal clevis. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1367220 | 1186477 | 1136968 | 2955842200801310 | 1 | GEI | Manufacturer | OtherUser facility | Other | Initial | 2008-10-03 | 2008-10-03 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-10 | Yes | No | Reuse | 2008
IS2000 A51P4 | 2008-09-11 | Other | Other | It was reported that during a da vinci s surgical procedure, the customer experienced system error code #201. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. No pt harm was reported. No additional information provided. | The investigation conducted by field service engineering concluded that system error code #201 experienced by the customer was associated with a multiple servo driver (msd) pca board. The system contains two multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom for each system arm. The system was repaired by replacing the affected msd pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #201 occurs when a motor is not responding to the voltage applied. Upon determining this condition, the safety systems put da vinci in a "recoverable safe state". The msd pca was returned to isi for failure analysis investigation. Engineering was unable to duplicate system error code #201 during testing, but did see error code #205 on the same axis which also relates to motor drive circuit. Engineering concluded that a motor controller chip within the board had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1401132 | 1211565 | 1159755 | 2955842200801300 | 1 | NAY | Manufacturer | ForeignUser facility | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | Yes | Yes | Health Professional | IS2000 A51P4 | Device Returned To Manufacturer | No | 2008-09-12 | Yes | No | Reuse | 2008
IS1200 A43P9 | 2008-09-17 | Other | Other | It was reported that prior to starting a da vinci myomectomy procedure, the camera head focus connector broke off, resulting in the site being unable to focus the camera. The pt was under anesthesia before the surgical staff made the decision to abort the procedure. No pt harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering confirmed that the camera head focus connector is broken. The camera head has three cables connected to it, two to the cameras (one for the left camera and one for the right camera) and one for the motor controlled focus. The system was repaired by replacing the affected camera head. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1399142 | 1211566 | 1159756 | 2955842200801315 | 1 | NAY | Manufacturer | Company Representative | NOT APPLICABLE | Initial | 2008-10-03 | 2008-10-03 | Yes | Yes | Health Professional | IS1200 A43P9 | Yes | No | 2008-09-17 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-09-12 | Other | Other | It was reported that during a da vinci s total hysterectomy surgical procedure, the surgeon experienced a pink hue on the left image viewed through the surgeon's console. An isi technical support engineer (tse) was called to assist the surgical staff in resolving the issue and the isi tse determined that the image discoloration was due to a faulty camera cable, however, the site did not have a backup camera cable available for replacement. The pt was under anesthesia and the port incisions have been made when the surgical staff made the decision to abort the planned procedure. No pt harm, adverse outcome or injury was reported. | Per the case notes, the investigation conducted by the isi tse discovered that the issue experienced by the customer was associated with a faulty camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by providing the site with two new camera cables. The da vinci s surgical system user's manual explicitly states that, "environmental or equipment failures may cause the da vinci s system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. " as of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1401118 | 1211567 | 1159757 | 2955842200801304 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | Yes | Yes | Health Professional | IS2000 A51 P4 | Yes | No | 2008-09-12 | Yes | No | Reuse | 2008
IS1200 A43 P9 | 2008-09-12 | Other | Other | It was reported that prior to starting a da vinci prostatectomy surgical procedure, the customer experienced non-recoverable system error code #212. An isi technical support engineer (tse) attempted to troubleshoot the issue via telephone and requested that the site shut down the system and examine the surgeon's console and pt side cart connectors for recessed/bent pins. A recessed pin was found and repaired and the system was restarted, however, non-recoverable system error code #212 later recurred. The pt was under anesthesia when the surgical staff made the decision to complete the planned procedure with their spare da vinci surgical system. The procedure was delayed for approx 45 mins and no pt harm, adverse outcome or injury was reported. | The investigation conducted by isi field service engineering concluded that system error code #212 experienced by the customer was associated with a multiple servo driver (msd) pca board. The system contains five multiple servo driver (msd) pca boards which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected msd pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. System error code #212 is a voltage tracking error which occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put davinci in a "recoverable safe state". The msd pca was returned to isi for failure analysis investigation. Engineering evaluation discovered that two motor drive circuits within the board had malfunctioned, thus generating the system error code experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1401125 | 1211568 | 1159758 | 2955842200801302 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-10-03 | 2008-10-03 | Yes | Yes | Health Professional | IS1200 A43 P9 | Device Returned To Manufacturer | No | 2008-09-12 | Yes | No | Reuse | 2008
IS1200 A43P9 |  | Other | Other | It was reported that during a da vinci surgical procedure, the physician at the surgeon side console felt a "shock" at the master tool manipulators (mtm). This occurred when the hospital experienced a temporary power outage. No harm to the physician or patient was reported. | Per the customer reported complaint, the investigation conducted by the field service engineer alongside hospital's biomed engineer on august 8, 2008 concluded that the system passed electrical safety testing and all values performed to specification. The reported complaint could not be verified at this time. A second evaluation of the system was performed on september 19, 2008 by an isi senior hardware engineer in which the system and the operating room environment were evaluated. The system was found to be within specification with no physical or electro-mechanical defects that would contribute to the reported complaint. During the evaluation of the operating room environment, it was discovered that their standard apparel (non-disposable type) is not antistatic and the floor coating appeared to be commercial wax. The surgeon side console chair was not present but was reported to be a standard office type with a fabric seat and back. Experimentation with an ohm-stat tester confirmed that the existing practices do not provide sufficient resistivity to ground to effectively mitigate electrostatic discharge (esd) events. The site has been advised on minimizing the occurrence of esd events by wearing apparel (scrubs, gowns, footwear, etc. ) and using chairs that are antistatic rated. The floor material should ideally be conductive tile or treated with antistatic coating to reduce the possibility of electro-static buildup and esd events. As of september 30, 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1369062 | 1188402 | 1138853 | 2955842200801286 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-09-30 | 2008-10-02 | Yes | No | Health Professional | IS1200 A43P9 | Yes | No | 2008-08-07 | Yes | No | Reuse | 2008
420179 | 2008-09-04 | Malfunction |  | Protective cover was placed properly on monopolar scissors and inspected prior to use. Scissors were mounted on robot arm, robot signaled instrument was ok and ready to use. While using at one point, the scissors would not close. Instrument was removed from pt and robot was cleaned. After cleaning scissors would not open. Scissors were taken out of service immediately. When the scissors were handed to staff, the wheel on the bottom right hand side did not appear to be working correctly. This is a reposable instrument that has 10 uses and this was the first out of 10. |  | INTUITIVE SURGICAL | MONOPOLAR CURVED SCISSORS | 1227318 | 1186183 | 1136694 | MW5008510 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2008-09-22 | 2008-10-01 | No | Yes | Health Professional | 420179 | Yes | No |  |  |  |  | 2008
42017204 |  | Malfunction |  | It was reported that the maryland bipolar forceps instrument was not working properly. No patient harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was broken at the proximal clevis to main tube interface. The clevis is dislodged from the main tube as a result of the breakage and the proximal clevis and main tube are missing pieces. Per the case notes, there are no indications from the site that the missing instrument components fell inside of a patient during a surgical procedure. Engineering also observed tube abrasions on the distal end of the main tube which extend approximately 5. 75" and is located 1" from the intersection between the proximal clevis and main tube. The damaged area is parallel to the tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MARLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1370798 | 1191162 | 1141453 | 2955842200801294 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42017204 | Yes | No | 2008-09-04 | Yes | No | Reuse | 2008
42020504 |  | Malfunction |  | It was reported that the fenestrated bipolar instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that recognition and engagement passed. Results/conclusions - engineering also observed that the distal end of main tube has a 3. 75" long section with material removed on one side of the tube, and deep scratches in an area of approximately 1. 125" on the opposite side and approximately. 340" from the intersection with the proximal clevis. The damaged areas are parallel to the tube axis and have a rough surface finish. Based on the location and appearance, the damaged areas were most likely caused by a combination of cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Conclusion - the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1370838 | 1191166 | 1141457 | 2955842200801296 | 1 | GEI | Manufacturer | User facilityOther | Unknown | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42020504 | Yes | No | 2008-09-04 | Yes | No | Reuse | 2008
42020805 |  | Malfunction |  | It was reported that the thoracic grasper instrument was not working properly. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Recognition and engagement passed. Results/conclusions - engineering also observed that the distal end of main tube has a 2. 380" long section, extending from the intersection between the proximal clevis and the main tube toward the housing, with material removed on different parts of the tube. The damaged area has a rough surface finish, combined with deep scratches within the same area. Based on the location and appearance, the damage was most likely caused by a combination of cannula accessory and instrument collisions. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. Conclusions: the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage ot the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1370857 | 1191167 | 1141458 | 2955842200801292 | 1 | GCJ | Manufacturer | User facilityOther | Unknown | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42020805 | Yes | No | 2008-09-03 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the pk dissecting forceps instrument was not working properly. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that the instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Recognition and engagement passed. Engineering also observed the distal end of main tube has a 1. 85" long section, extending from the intersection between the proximal clevis and the main tube toward the housing, with material removed on one side of the tube. The damaged area has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Add'l investigation is underway regarding the cannula accessories. A f/u mdr will be submitted if add'l info is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1369509 | 1191239 | 1141530 | 2955842200801291 | 1 | GCJ | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-03 | Yes | No | Reuse | 2008
40018302 | 2008-09-05 | Malfunction |  | It was reported that 1. 5 hrs into a da vinci thoracic surgical procedure, the small plastic cover that covers the permanent cautery hook instrument tip broke off inside the pt. The broken piece was retrieved and the planned surgical procedure was completed without significant delay. No pt harm, adverse outcome or injury was reported. | The customer reported that they would not be returning the instrument for eval, therefore, the root cause of the customer reported failure mode cannot be determined. A f/u mdr will be submitted if add'l info is received. Per the case notes, the broken piece was successfully retrieved from the pt. | PERMANENT CAUTERY HOOK INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1369500 | 1191240 | 1141531 | 2955842200801298 | 1 | GEI | Manufacturer | User facility | UNKNOWN | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 40018302 | No | No | 2008-09-05 | Device Not Retu | No | Unkown | 2008
42020503 |  | Malfunction |  | It was reported that during a da vinci s surgical procedure, a cable broke on the fenestrated bipolar forceps instrument. No pieces fell into the pt. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, testing was performed by engineering and system recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly and no broken cables were observed. The instrument was found to be fully functional. Engineering also observed multiple deep scratches with material removed located within an approx 2. 250" area from the interface between the proximal clevis and the main tube. The scratches range in size from. 025" x. 020",. 500" x. 070" and. 730" x. 090". Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1369363 | 1191241 | 1141532 | 2955842200801293 | 1 | GEI | Manufacturer | User facilityOther | UNKNOWN | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42020503 | Yes | No | 2008-09-03 | Yes | No | Reuse | 2008
42022702 |  | Malfunction |  | It was reported that the tips of the pk dissecting forceps instrument do not meet. No add'l info was provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering conducted performance tests and found that one grip is bent, causing side to side misalignment of grips with a. 045" offset at the tips. The bent grip has separation of the yaw pulley and conductor cap at the glue joint, indicating overloading at the tip. Electrical continuity passed. Engineering also observed a deep scratch that is approx. 240" x. 150" in size located near a second, narrower, scratch of approx. 550" long. The scratches are near the intersection of the proximal clevis where material of the main tube was removed. Based on the location and appearance, the damage was most likely caused by instrument collisions. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1369073 | 1191242 | 1141533 | 2955842200801295 | 1 | GEI | Manufacturer | User facilityOther | Other | Initial | 2008-09-26 | 2008-09-26 | No | Yes | Health Professional | 42022702 | Yes | No | 2008-09-04 | Yes | No | Reuse | 2008
IS2000 A51 P4 | 2008-09-05 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon experienced a discoloration of the image viewed through the physician's console and could not continue with the case. The surgeon decided to convert to traditional open surgical techniques to finish the planned surgery. No patient harm was reported. | The investigation conducted by isi field service engineering (fse) discovered that the issue experienced by the customer was associated with a defective camera cable. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cable. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1380003 | 1200184 | 1149775 | 2955842200801299 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2008-09-26 | 2008-09-26 | Yes | Yes | Health Professional | IS2000 A51 P4 | No | No | 2008-09-05 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tip of the tenaculum forceps instrument had a broken cable. No additional information has been provided. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found one grip close cable to be broken near the proximal pulleys, and proximal clevis cable hole resulting in the pulley to freely spin. The proximal clevis cable hole exhibits wear on the edge. Other cables at the wrist are not damaged. The cable wear on the pulleys/clevis combined with high grasping force likely contributed to the breakage. Engineering also observed the distal end of the main tube to have deep scratches on an area approx. 500" x. 300" where material of the main tube was removed. The scratches are located on the clevis end of the the main tube, at approx 1. 200" from the proximal clevis and likely caused by instrument collisions. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instrument. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | ATENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1371198 | 1191505 | 1141795 | 2955842200801280 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-09-26 | 2008-09-25 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
42000605 |  | Malfunction |  | It was reported that the large needle driver instrument is splintering. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned, and evaluated. Per the customer reported complaint, engineering found that the main tube has a 1" long section just below the midpoint with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1346446 | 1181970 | 1132700 | 2955842200801278 | 1 | GCJ | Manufacturer | User facilityOther | Unknown | Initial | 2008-09-22 | 2008-09-22 | No | Yes | Health Professional | 42000605 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
42020704 |  | Malfunction |  | It was reported that the tenaculum forceps instrument has "fiberglass" coming off and nothing fell into patient. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed that the distal end of main tube has a 1" long section with light material removed all around the tube. There is an area above this section with the appearance of light abrasions, however, the tube surface is smooth to the touch. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 1346440 | 1181971 | 1132701 | 2955842200801288 | 1 | GCJ | Manufacturer | User facilityOther | Other | Initial | 2008-09-22 | 2008-09-22 | No | Yes | Health Professional | 42020704 | Yes | No | 2008-08-27 | Yes | No | Reuse | 2008
42017908 |  | Malfunction |  | It was reported that the monopolar curved scissors instrument is splintering. No additional information was provided. No patient harm, adverse outcome or injury was reported. | Results/conclusions - the instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the distal end of main tube has a 1. 25" long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. Based on the location and appearance, the damage was most likely caused by a cannula accessory. No other damage found. Additional investigation is underway regarding the cannula accessories. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 1346432 | 1181973 | 1132703 | 2955842200801279 | 1 | NAY | Manufacturer | User facilityOther | Unknown | Initial | 2008-09-22 | 2008-09-22 | No | Yes | Health Professional | 42017908 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008
IS2000 A51P6 | 2008-08-25 | Other | Other | It was reported that while inserting instruments into the patient during the beginning of a da vinci s prostatectomy surgical procedure, the customer experienced system error codes #23008 and #212. The system was restarted, however, the error codes recurred when attempting to home the system. With the assistance of an isis technical support engineer, the customer shut down the system and exercised the left master tool manipulator. The system was restarted and the system error codes reappeared during homing. The surgeon decided to abort the planned surgical procedure and no patient harm was reported. | The investigation conducted by the field service engineer concluded that system error codes #23008 and #212 were associated with the left master tool manipulator and/or the input output distribution board (iodg pca). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and to his right (mtmr). The purpose of the iodg pca is to interface all external input/output to the internal system electronics. The system was repaired by replacing the affected mtml and iodg pca. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code #23008 appears when the da vinci s safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system put davinci in a "recoverable safe state". System error code #212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Since the motor's velocity is a component of the expected voltage calculation, this error can be caused by a faulty encoder. The mtml and iodg pca were returned to isi for failure analysis investigation. Engineering evaluation found no issues with the iodg pca, however, during evaluation of the mtml, engineering found that a motor encoder had malfunctioned, thus generating the system errors experienced by the customer. As of 2008, there have been no reported recurrences of the issue at this hospital. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 1366019 | 1182394 | 1133105 | 2955842200801282 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2008-09-22 | 2008-09-22 | Yes | Yes | Health Professional | IS2000 A51P6 | Yes | No | 2008-08-25 | Yes | No | Reuse | 2008