42014604 | 2012-12-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the scissor tips on the curved scissors instrument would not open. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was placed on an in-house system and driven. Recognition and engagement passed. The snake wrist was yawed over to one side instead of straight when first installed. Housing was removed to find input gears for pitch and yaw have skipped some teeth. Skipped teeth change the orientation of the snake wrist. Blades would not open as well, as they are stuck. Additional observation not reported was damage to the main tube. Insulation was found with several gouge marks. Marks were short in length and not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation removed, and material is missing. The main tube also exhibits discoloration (white striations). Evidence not conclusive, but damage may be a result of improper cleaning. The instrument has 6 uses left. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The cleaning and sterilization user manual specifically states: prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | CURVED SCISSORS | ENDOSCOPIC INSTRUMENT | 2921637 | 2895886 | 2792401 | 2955842201300051 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42014604 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-10 | Malfunction |  | It was reported that wires on the large suturecut needle driver instrument were exposed at the distal end during central processing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The instrument has 3 uses left. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921660 | 2895888 | 2792403 | 2955842201300053 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-03 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
420015 | 2013-01-03 | Malfunction |  | Approximately halfway through the procedure the c-arm drape began sticking making it not respond correctly to movement - no patient harm but potential - appeared to be lot related. |  | DA VINCI S | SYSTEMSURGICAL | 2926022 | 2900760 | 2797276 | 2900760 | 1 | NAY | User Facility |  |  | Initial | 2013-01-03 | 2013-01-03 | No | Yes | Physician | 420015 | Yes | No Answer Provided |  |  |  |  | 2013
42009308 | 2012-11-27 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the cable at the tip of the prograsp forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. A visual inspection shows all drive cables at wrist are intact. Engineering evaluation also found that one grip close cable is derailed from the distal idler pulley. The cable is exposed on outside of pulley, which may contribute to wrist motion not being smooth. The grips were unable to open/close properly when inputs were manually rotated. The grips remained stuck in open position. The housing was removed to find the flush tube dislodged from the luer plate. The flush tube is pushed inside the housing and the flange feature that helps it stay retained in the luer plate has become detached. The flush tube inner surface exhibits above average contamination. The instrument was disassembled to find excessive biodebris stuck to the distal end of the hypotubes. The biodebris prevents relative motion between hypotubes, causing the grips to stick. Engineering concluded that the damage is likely due to improper/insufficient cleaning. No other damage was found. The reprocessing instructions specifically states: cleaning, disinfection, and sterilization general information and warning: - read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | PROGRASPH FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2919342 | 2893581 | 2780658 | 2955842201300001 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Service Personnel | 42009308 | Device Returned To Manufacturer | No | 2012-12-03 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-05 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the frayed cables on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | (b)(4). The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable at distal idler location. There was no damage found on pulley or clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2919381 | 2893620 | 2790156 | 2955842201300002 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-02 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42019007 | 2012-11-29 | Malfunction |  | It was reported that during prior to starting a da vinci si procedure, it was noted that the head near port 1 was cracked. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a piece of the luer plate broken off. One section of the luer plate outer lip is broken. The chassis is cracked adjacent to the broken luer plate section. Engineering concluded that the damage is likely due to mishandling/misuse during cleaning process. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are. 070 -. 160 in length and not aligned with the tube axis. Engineering concluded that the damage is likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute reportable event; however; the damage to the instrument's main tube, could likely cause or contribute to an adverse event if the malfunction were to recur. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2919413 | 2893622 | 2790158 | 2955842201300003 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-04 | 2013-01-02 | No | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2012-12-04 | Other | Other | It was reported that prior to starting a da vinci si surgical procedure, the right eye image through the high resolution stereo viewer (hrsv) on the surgeon side cart (ssc) was dark. With the assistance of an isi technical support engineer, the site checked all of the system cables and the site found no issues. The tse instructed the site to verify if the camera control unit (ccu) was on. The site confirmed that the ccu was on. The tse asked the site if the surgeon was able to visualize icons through the hrsv. The site indicated no and that the image through the right eye of hrsv was black. Unable to resolve the issue, the surgeon made the decision to abort the planned surgical procedure. No other information was provided. No patient harm, adverse outcome or injury was reported. | The investigation conducted by field service engineering was unable to replicate the vision issue experienced by the customer, as functional testing of the system found that the system powered on with no issue noted and the 3d vision through the left and right eye of the high resolution stereo viewer was good. The fse power cycled the system and kept the system power on for 1 hour and no issues were noted. However, review of the site system log found intermittent occurrences of system error code 45312. The fse concluded that the system error code was likely associated with the camera head, the camera cable, the video input output processor board (vio), and/or the system video processor printed assembly (svp pca) board and were all replaced as a precaution. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The s5vp pca is the video processor for the system which creates the display for the hrsv. System error code 45312 appears when the da vinci si system detects a loss of one or both surgeon's video inputs. Upon determining this condition, the safety system puts the da vinci si in the non recoverable safe state. Isi has made several attempts to verify additional information concerning information concerning the reported incident; however, no additional information was provided. If additional information becomes available, a follow-up medwatch report will be submitted to the fda. As of (b)(6) 2012, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2919471 | 2893698 | 2790232 | 2955842201300004 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-04 | 2013-01-02 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-11-29 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the mega suture cut needle driver instrument 'would not cut suture, nothing fell in patient or harm' no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument scissors did not cut cleanly through a 0-silk suture. The suture got smashed between the blades and it required multiple attempts to be able to cut completely through. The blades were found to be damaged with indentations. Engineering concluded that damage was likely due to mishandling. Additional finding not reported by the site was main insulation tube scratches. The distal end of the main insulation tube exhibited various scratch marks with light material removal. Engineering concluded that damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2919482 | 2893712 | 2790246 | 2955842201300005 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-02 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-12-05 | Injury | Other | On (b)(6) 2012, md surgeon dr. (b)(6) reported, during a tough case with extreme angulation, a break in the mcs tip cover was seen. This resulted in a slight burn to the uterus. The tip cover was discarded after the case. It was a four hour case for node debunking/removal on a high bmi ((b)(6)) patient with a (b)(6) size uterus. Setup was a maryland in arm 2, mcs (monopolar curved scissors) instrument in arm 1, and prograsp in arm 3 on same side as arm most of the time surgeon was deep in the pelvis, performing fine dissection with few large movements. Surgeon was thoroughly articulating the mcs instrument and regularly used it to retract the uterus. Davinci robot system settings were at 30/30 and the mcs tip cover arced to the uterus about three hours into the case. The surgical staff replaced the mcs tip cover (same instrument) and had no further problems for the remaining hour of the case. | Intuitive surgical contacted the initial reporter of this complaint to request more information concerning the reported event; however, as of the date of this report no response has been received. The mcs tip cover was discarded and it was not returned for evaluation, the cause of the damage to the tip cover prior to the reported arcing event is unknown. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2919484 | 2893714 | 2790248 | 2955842201300006 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-12-05 | 2013-01-02 | Yes | Yes | Health Professional | 40018012 | No | Yes | 2012-12-05 | No | Yes | Unkown | 2013
42022703 | 2012-12-06 | Malfunction |  | It was reported that during cleaning and sterilization, frayed wires were found in the sterile processing on the pk dissecting forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2919520 | 2893749 | 2790283 | 2955842201300007 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-02 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-04 | Malfunction |  | It was reported that during reprocessing of the maryland bipolar forceps instrument the wires on the instrument were frayed. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the pitch cable frayed at the distal idler. There was no damage found on pulley and clevis. The instrument also passed electrical continuity testing. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2919531 | 2893761 | 2790295 | 2955842201300008 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-04 | 2013-01-02 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-04 | Malfunction |  | It was reported that during reprocessing of the large needle driver instrument, the wire on the instrument was observed to be frayed. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at distal idler is frayed. No damage found on pulley and clevis. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal, suggesting the may have been caused by instrument collisions or rough handling. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2919547 | 2893777 | 2790311 | 2955842201300009 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-04 | 2013-01-02 | No | Yes | Health Professional | 42000606 |  | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42000306 | 2012-12-03 | Malfunction |  | It was reported prior to starting a da vinci si procedure the small clip applier instrument had a broken tip. No patient harm, adverse outcome or injury was reported. | Instrument was returned and evaluated. Engineering observed both clip appliers are broken. The broken pieces were not returned with the instrument. Engineering concluded the damage is likely due to misuse/user error. Additionally, a frayed pitch was observed, though not reported by the site. The frayed pitch cable was located at the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2919561 | 2893790 | 2790324 | 2955842201201488 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2012-12-03 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-12-03 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si sacropolpexy surgery that the tip cover accessory was too large and slipped off into the patient. | Manufacturer Narrative: --/-- Intuitive surgical contacted the initial reporter of this complaint. It was indicated that the tip cover accessory was retrieved and there was no harm or injury to the patient. The procedure was completed as planned and the patient has recovered as expected without any post-surgical complications. The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The accessory was returned and evaluated. Engineering was unable to confirm the reported complaint that the tip cover was too large. Visual inspection showed no scratches, tears, or signs of arcing on the tip cover. Compared with another mcs tip cover -12, tip cover showed no abnormalities in size. The tip cover was installed on a monopolar curved scissors instrument and it fit snugly. The instrument was driven on the system and the tip cover did not fall off. Engineering was unable to make this tip cover slip off the monopolar curved scissors when installed correctly. No physical damage was found. The instruments and accessories user manual specifically states: warning: failure to install the tip cover accessory properly may result in: improper scissor opening, tip cover accessory falling off, electrical arcs and alternate site burns. | 8 MM MCS TIP COVER ACCESSORY | Initial | 2919654 | 2893871 | 2790387 | 2955842201300010 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-15 | 2013-01-02 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | Yes |  | 2013
42027501 | 2012-12-03 | Malfunction |  | It was reported prior to starting a da vinci si the harmonic ace curved shears had a broken tip. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The harmonic ace curved shears instrument which is used in conjunction with the harmonic ace curved shears insert was returned for evaluation: the ace instrument does not have a tip, as ace insert acts as tip when installed into instrument backend. Shaft tip on instrument is intact. Additional observations not reported by site is a bent shaft. Shaft is bent roughly 2. 9 above the distal end. Tube no longer rolls along a straight line due to bending. Shaft is also dented below the bend line. An ace insert can still be installed, but there is high friction during installation. Evidence not conclusive, but damage is likely due to mishandling. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE CURVED SHEARS | ENDOSCOPIC ELECTROSURGICAL INS | 2919655 | 2893891 | 2790407 | 2955842201300011 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Health Professional | 42027501 | No | No | 2012-12-03 | No | No | Unkown | 2013
40022702 | 2012-10-22 | Malfunction |  | It was reported during a da vinci standard prostatectomy surgery the pk dissecting forceps instrument was found defective. No further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2919645 | 2893892 | 2790408 | 2955842201300015 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Health Professional | 40022702 | Device Returned To Manufacturer | No | 2012-12-03 | Device Not Retu | No | Reuse | 2013
400272 | 2012-12-03 | Malfunction |  | It was reported prior to starting a da vinci si procedure the harmonic ace curved shears instrument had a broken tip. No patient harm, adverse outcome or injury was reported. | The insert accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The harmonic ace curved shears instrument, which is used with the harmonic ace curved shears insert was returned and evaluated. Instrument does not have a tip, as the ace insert acts as the tip when it is installed through backend. Visual inspection shows no damage to shaft or backend. An ace insert was installed without issues and the instrument was placed on an is3000 system and driven. Instrument moved intuitively with full range of motion in all directions. Clamp arm opened and closed properly. Functional performance testing did not find any issues. Instrument is fully functional. No physical damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2919659 | 2893893 | 2790409 | 2955842201300013 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Health Professional | 400272 | No | No | 2012-12-03 | No | No | Unkown | 2013
42020505 | 2012-10-30 | Malfunction |  | It was reported during a da vinci si procedure that the fenestrated bipolar forceps instrument was damaged. No other information was provided. No patient harm, adverse outcome or injury was reported. | Instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The electrical continuity passed. Additionally scratches on the maintube were observed, though not reported by the site. The distal end of the main tube has various scratch marks with light material removal. Engineering concluded this may have likely been caused by instrument collisions or rough handling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2919656 | 2893912 | 2779354 | 2955842201300014 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-03 | 2013-01-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-03 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-12-03 | Malfunction |  | It was reported during a da vinci si hysterectomy surgery that the mega needle driver instrument cable broke. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | Instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2919646 | 2893914 | 2790430 | 2955842201300017 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-04 | 2013-01-02 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013